## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
201209_20-196_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.6073 | -69.7271 |  |  |  | 2.3394 | 19.372
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.397 | 5.0405
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.58788 | 7.641
4 |  | EV | -10.3332 | -75.5061 |  |  |  | 0.20732 | 2.6419
5 |  | UNDE1-A-5 | -6.0614 | -56.5082 |  | 443.8227 |  | 0.1935 | 1.4517 | -14.4995 | -2.3605
6 |  | UNDE1-C-4 | -5.8025 | -58.498 |  | 4.8207 |  | 0.069081 | 0.23959
7 |  | UNDE1-E-4 | -4.9823 | -51.1432 |  | 85.6177 |  | 0.050498 | 0.49294 |  | 2.118
8 |  | CPER1-C-2 | -6.6641 | -84.3535 |  | 164.7827 |  | 0.048836 | 2.3954 |  | 0.4502
9 |  | UNDE1-E-1 | -5.7419 | -55.4786 |  | 49.3695 |  | 0.051933 | 2.3059 |  | 0.84895
10 |  | UNDE1-A-2 | -6.0383 | -56.9254 |  | 380.4985 |  | 0.032833 | 0.15508 | -12.3622 | -1.6268
11 |  | UNDE1-C-3 | -8.2753 | -76.0304 |  | 373.7757 |  | 0.10899 | 1.351 | -11.8513 | -2.883
12 |  | UNDE1-B-3 | -6.1009 | -58.7203 |  | 619.0312 |  | 0.085936 | 1.4805 | -21.2443 | -4.88
13 |  | EV | -10.0265 | -73.0186 |  | 3.6227 |  | 0.14033 | 0.92637
14 |  | UNDE1-E-3 | -5.2494 | -45.9874 |  | 60.8102 |  | 0.18219 | 1.9508 |  | 2.0152
15 |  | UNDE1-B-1 | -4.6535 | -49.0929 | -1.2543 | 757.3195 |  | 0.0085829 | 0.4322 | -27.4905 | -6.53
16 |  | UNDE1-B-4 | -6.4962 | -61.3802 |  | 415.9735 |  | 0.11183 | 1.0914 | -14.5732 | -2.2215
17 |  | UNDE1-D-1 | -6.0999 | -55.5965 |  | 136.5832 |  | 0.098899 | 0.63471 |  | 1.8162
18 |  | UNDE1-C-5 | -8.801 | -73.9253 |  | 407.857 |  | 0.12936 | 1.3906 | -12.4482 | -4.3493
19 |  | UNDE1-D-3 | -4.3563 | -49.875 | -0.8075 | 11.3227 |  | 0.19126 | 1.9169 |  | 4.647
20 |  | UNDE1-D-5 | -6.6104 | -61.5202 |  | 105.9197 |  | 0.049615 | 0.70359 |  | 1.5485
21 |  | UNDE1-A-1 | -4.6664 | -59.6502 |  | 206.198 |  | 0.1812 | 0.54851 |  | -0.21955
22 |  | EV | -9.9969 | -72.8876 |  | 2.2272 |  | 0.24401 | 0.9531
23 |  | UNDE1-D-2 | -5.9281 | -55.3517 |  | 60.9947 |  | 0.14839 | 1.4038 |  | 2.6737
24 |  | UNDE1-D-4 | -5.8977 | -56.6289 |  | 3.1202 |  | 0.059718 | 0.12945
25 |  | UNDE1-A-3 | -6.1681 | -57.0144 |  | 406.561 |  | 0.01772 | 0.1458 | -13.0955 | -2.628
26 |  | UNDE1-E-5 | -5.8792 | -55.1781 |  | 77.9922 |  | 0.039682 | 0.33351 |  | 0.72695
27 | * | UNDE1-B-5 | -4.7421 | -38.6511 |  |  |  | 2.111 | 25.2773 | -9.4992
28 | * | UNDE1-B-2 | -3.996 | -38.1811 | -0.8385 |  |  | 1.2607 | 24.7055 | -12.9678
29 | * | NIWO1-A-4 | -0.9227 | -64.3351 | -4.83 |  |  | 2.241 | 56.5326 | -17.5288
30 | * | WOOD1-C-2 | -4.8916 | -42.221 |  |  |  | 2.2811 | 29.6226
31 |  | EV | -9.991 | -72.8262 |  | 13.6922 | 10.2329 | 0.30676 | 2.3687 |  | 9.9502

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV |  |  | -2.8896 | -17.8792 |  |  | 2020/12/09 11:48:09
1 |  | EV | -7.967 | -67.819 | -10.9637 | -82.2318 | 0.566 | 1342.445 | 2020/12/09 11:58:08
1 |  | EV | -7.406 | -61.388 | -10.3951 | -76.1295 | 0.498 | 1314.397 | 2020/12/09 12:07:05
1 |  | EV | -7.279 | -60.115 | -10.2664 | -74.9216 | 0.5 | 1310.128 | 2020/12/09 12:16:03
1 |  | EV | -7.282 | -59.725 | -10.2695 | -74.5515 | 0.382 | 1307.85 | 2020/12/09 12:24:59
1 |  | EV | -7.338 | -59.748 | -10.3262 | -74.5733 | 0.304 | 1305.628 | 2020/12/09 12:33:57
1 |  | EV | -7.278 | -59.511 | -10.2654 | -74.3485 | 0.472 | 1305.895 | 2020/12/09 12:42:54
1 |  | EV | -7.286 | -59.483 | -10.2735 | -74.3219 | 0.568 | 1304.974 | 2020/12/09 12:51:51
1 |  | EV | -7.253 | -59.374 | -10.2401 | -74.2185 | 0.44 | 1304.092 | 2020/12/09 13:00:49
1 |  | EV | -7.197 | -59.244 | -10.1833 | -74.0951 | 0.628 | 1303.75 | 2020/12/09 13:09:45
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.6073 | -69.7271 |  |  |  | 2.3394 | 19.372
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.699 | 24.56 | 0.85919 | 5.4255 | 0.607 | 1303.901 | 2020/12/09 13:19:34
2 |  | PZ | 4.542 | 34.887 | 1.7135 | 15.2247 | 0.281 | 1301.604 | 2020/12/09 13:28:31
2 |  | PZ | 4.741 | 37.894 | 1.9152 | 18.078 | 0.263 | 1302.971 | 2020/12/09 13:37:27
2 |  | PZ | 4.87 | 39.109 | 2.0459 | 19.2309 | 0.363 | 1302.081 | 2020/12/09 13:46:24
2 |  | PZ | 4.899 | 39.792 | 2.0753 | 19.879 | 0.287 | 1302.831 | 2020/12/09 13:55:20
2 |  | PZ | 4.936 | 40.247 | 2.1128 | 20.3107 | 0.595 | 1302.841 | 2020/12/09 14:04:17
2 |  | PZ | 4.939 | 40.254 | 2.1159 | 20.3174 | 0.476 | 1301.194 | 2020/12/09 14:13:13
2 |  | PZ | 4.964 | 40.74 | 2.1412 | 20.7785 | 0.319 | 1302.748 | 2020/12/09 14:22:10
2 |  | PZ | 4.914 | 40.591 | 2.0905 | 20.6371 | 0.488 | 1302.458 | 2020/12/09 14:31:05
2 |  | PZ | 5.052 | 41.098 | 2.2304 | 21.1182 | 0.322 | 1301.51 | 2020/12/09 14:40:03
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.397 | 5.0405
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -11.209 | -86.131 | -14.2493 | -99.6079 | 0.425 | 1301.884 | 2020/12/09 14:49:53
3 |  | UT2 | -12.443 | -102.398 | -15.4999 | -115.0434 | 0.675 | 1302.138 | 2020/12/09 14:58:50
3 |  | UT2 | -12.821 | -106.808 | -15.883 | -119.228 | 0.368 | 1301.474 | 2020/12/09 15:07:47
3 |  | UT2 | -12.903 | -108.37 | -15.9661 | -120.7102 | 0.35 | 1302.175 | 2020/12/09 15:16:44
3 |  | UT2 | -13.02 | -109.515 | -16.0846 | -121.7967 | 0.364 | 1301.69 | 2020/12/09 15:25:40
3 |  | UT2 | -13.019 | -109.782 | -16.0836 | -122.05 | 0.493 | 1301.871 | 2020/12/09 15:34:38
3 |  | UT2 | -13.041 | -110.09 | -16.1059 | -122.3423 | 0.3 | 1301.587 | 2020/12/09 15:43:35
3 |  | UT2 | -13.069 | -110.35 | -16.1343 | -122.589 | 0.557 | 1301.866 | 2020/12/09 15:52:31
3 |  | UT2 | -13.128 | -111.08 | -16.1941 | -123.2817 | 0.54 | 1301.123 | 2020/12/09 16:01:28
3 |  | UT2 | -13.133 | -111.153 | -16.1992 | -123.3509 | 0.303 | 1301.579 | 2020/12/09 16:10:26
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.58788 | 7.641
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -7.915 | -68.019 | -10.911 | -82.4216 | 0.545 | 1301.22 | 2020/12/09 16:20:15
4 |  | EV | -7.328 | -61.412 | -10.3161 | -76.1523 | 0.43 | 1302.988 | 2020/12/09 16:29:11
4 |  | EV | -7.352 | -60.399 | -10.3404 | -75.1911 | 0.401 | 1301.633 | 2020/12/09 16:38:07
4 |  | EV | -7.354 | -60.349 | -10.3424 | -75.1436 | 0.756 | 1301.326 | 2020/12/09 16:47:04
4 |  | EV | -7.194 | -59.595 | -10.1803 | -74.4282 | 0.505 | 1301.72 | 2020/12/09 16:56:01
4 |  | EV | -7.272 | -59.482 | -10.2593 | -74.3209 | 0.101 | 1301.685 | 2020/12/09 17:04:57
4 |  | EV | -7.304 | -59.819 | -10.2918 | -74.6407 | 0.462 | 1300.743 | 2020/12/09 17:13:55
4 |  | EV | -7.261 | -59.636 | -10.2482 | -74.4671 | 0.498 | 1300.881 | 2020/12/09 17:22:51
4 |  | EV | -7.245 | -59.345 | -10.232 | -74.1909 | 0.677 | 1300.855 | 2020/12/09 17:31:48
4 |  | EV | -7.224 | -59.254 | -10.2107 | -74.1046 | 0.421 | 1301.065 | 2020/12/09 17:40:45
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3332 | -75.5061 |  |  |  | 0.20732 | 2.6419
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | UNDE1-A-5 | -3.408 | -42.792 | -6.3434 | -58.484 | -0.337 | 1741.756 | 2020/12/09 17:50:34
5 |  | UNDE1-A-5 | -3.097 | -40.613 | -6.0282 | -56.4164 | -0.126 | 1743.619 | 2020/12/09 17:59:32
5 |  | UNDE1-A-5 | -2.964 | -39.66 | -5.8934 | -55.5121 | -0.343 | 1748.481 | 2020/12/09 18:08:28
5 |  | UNDE1-A-5 | -3.05 | -39.774 | -5.9806 | -55.6203 | -0.172 | 1749.74 | 2020/12/09 18:17:26
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | UNDE1-A-5 | -6.0614 | -56.5082 |  | 443.8227 |  | 0.1935 | 1.4517 | -14.4995 | -2.3605
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | UNDE1-C-4 | -2.96 | -42.877 | -5.8894 | -58.5647 | 0.029 | 1309.239 | 2020/12/09 18:27:17
6 |  | UNDE1-C-4 | -2.798 | -42.477 | -5.7252 | -58.1851 | -0.118 | 1307.289 | 2020/12/09 18:36:13
6 |  | UNDE1-C-4 | -2.846 | -42.825 | -5.7738 | -58.5153 | 0.261 | 1306.545 | 2020/12/09 18:45:10
6 |  | UNDE1-C-4 | -2.893 | -43.048 | -5.8215 | -58.7269 | 0.146 | 1304.515 | 2020/12/09 18:54:08
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | UNDE1-C-4 | -5.8025 | -58.498 |  | 4.8207 |  | 0.069081 | 0.23959
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | UNDE1-E-4 | -2.045 | -35.654 | -4.9621 | -51.7108 | -0.571 | 1388.731 | 2020/12/09 19:03:57
7 |  | UNDE1-E-4 | -2.14 | -35.265 | -5.0583 | -51.3417 | -0.366 | 1387.193 | 2020/12/09 19:12:53
7 |  | UNDE1-E-4 | -2.045 | -34.677 | -4.9621 | -50.7838 | -0.215 | 1387.431 | 2020/12/09 19:21:50
7 |  | UNDE1-E-4 | -2.03 | -34.627 | -4.9469 | -50.7363 | -0.236 | 1387.421 | 2020/12/09 19:30:46
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | UNDE1-E-4 | -4.9823 | -51.1432 |  | 85.6177 |  | 0.050498 | 0.49294 |  | 2.118
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | CPER1-C-2 | -3.693 | -66.576 | -6.6322 | -81.0524 | -0.27 | 1470.296 | 2020/12/09 19:40:34
8 |  | CPER1-C-2 | -3.681 | -70.406 | -6.6201 | -84.6866 | 0.014 | 1465.428 | 2020/12/09 19:49:30
8 |  | CPER1-C-2 | -3.735 | -71.401 | -6.6748 | -85.6307 | -0.169 | 1466.407 | 2020/12/09 19:58:27
8 |  | CPER1-C-2 | -3.789 | -71.837 | -6.7295 | -86.0445 | -0.151 | 1465.305 | 2020/12/09 20:07:24
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | CPER1-C-2 | -6.6641 | -84.3535 |  | 164.7827 |  | 0.048836 | 2.3954 |  | 0.4502
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | UNDE1-E-1 | -2.882 | -42.984 | -5.8103 | -58.6662 | -0.242 | 1352.52 | 2020/12/09 20:17:12
9 |  | UNDE1-E-1 | -2.804 | -39.216 | -5.7313 | -55.0908 | -0.315 | 1351.563 | 2020/12/09 20:26:09
9 |  | UNDE1-E-1 | -2.816 | -38.421 | -5.7434 | -54.3364 | -0.328 | 1350.37 | 2020/12/09 20:35:05
9 |  | UNDE1-E-1 | -2.756 | -37.878 | -5.6826 | -53.8212 | -0.213 | 1351.33 | 2020/12/09 20:44:01
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | UNDE1-E-1 | -5.7419 | -55.4786 |  | 49.3695 |  | 0.051933 | 2.3059 |  | 0.84895
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | UNDE1-A-2 | -3.142 | -41.081 | -6.0738 | -56.8605 | 0.09 | 1687.042 | 2020/12/09 20:53:48
10 |  | UNDE1-A-2 | -3.121 | -41.353 | -6.0525 | -57.1185 | -0.186 | 1681.285 | 2020/12/09 21:02:44
10 |  | UNDE1-A-2 | -3.065 | -40.99 | -5.9958 | -56.7741 | -0.151 | 1676.859 | 2020/12/09 21:11:41
10 |  | UNDE1-A-2 | -3.1 | -41.174 | -6.0312 | -56.9487 | -0.027 | 1685.113 | 2020/12/09 21:20:37
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | UNDE1-A-2 | -6.0383 | -56.9254 |  | 380.4985 |  | 0.032833 | 0.15508 | -12.3622 | -1.6268
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | UNDE1-C-3 | -5.195 | -59.413 | -8.1544 | -74.2555 | -0.291 | 1672.264 | 2020/12/09 21:30:25
11 |  | UNDE1-C-3 | -5.283 | -61.221 | -8.2436 | -75.9711 | -0.325 | 1680.741 | 2020/12/09 21:39:22
11 |  | UNDE1-C-3 | -5.457 | -62.479 | -8.4199 | -77.1648 | -0.08 | 1671.38 | 2020/12/09 21:48:20
11 |  | UNDE1-C-3 | -5.322 | -62.021 | -8.2831 | -76.7302 | -0.139 | 1679.023 | 2020/12/09 21:57:17
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | UNDE1-C-3 | -8.2753 | -76.0304 |  | 373.7757 |  | 0.10899 | 1.351 | -11.8513 | -2.883
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | UNDE1-B-3 | -3.296 | -45.186 | -6.2299 | -60.7556 | -0.196 | 1920.384 | 2020/12/09 22:07:05
12 |  | UNDE1-B-3 | -3.141 | -42.834 | -6.0728 | -58.5239 | -0.641 | 1922.08 | 2020/12/09 22:16:02
12 |  | UNDE1-B-3 | -3.13 | -42.236 | -6.0616 | -57.9564 | -0.631 | 1922.577 | 2020/12/09 22:24:58
12 |  | UNDE1-B-3 | -3.108 | -41.908 | -6.0393 | -57.6452 | -0.856 | 1919.389 | 2020/12/09 22:33:55
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | UNDE1-B-3 | -6.1009 | -58.7203 |  | 619.0312 |  | 0.085936 | 1.4805 | -21.2443 | -4.88
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -6.845 | -56.79 | -9.8266 | -71.7665 | 0.374 | 1309.732 | 2020/12/09 22:43:44
13 |  | EV | -7.044 | -58.149 | -10.0283 | -73.0561 | 0.383 | 1305.811 | 2020/12/09 22:52:40
13 |  | EV | -7.117 | -58.67 | -10.1023 | -73.5504 | 0.32 | 1304.138 | 2020/12/09 23:01:36
13 |  | EV | -7.163 | -58.829 | -10.1489 | -73.7013 | 0.354 | 1303.115 | 2020/12/09 23:10:33
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.0265 | -73.0186 |  | 3.6227 |  | 0.14033 | 0.92637
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | UNDE1-E-3 | -2.587 | -32.474 | -5.5113 | -48.6934 | 0.077 | 1364.514 | 2020/12/09 23:20:23
14 |  | UNDE1-E-3 | -2.31 | -29.227 | -5.2306 | -45.6123 | -0.113 | 1362.881 | 2020/12/09 23:29:19
14 |  | UNDE1-E-3 | -2.251 | -28.629 | -5.1708 | -45.0449 | -0.144 | 1361.776 | 2020/12/09 23:38:15
14 |  | UNDE1-E-3 | -2.166 | -28.159 | -5.0847 | -44.5989 | 0.019 | 1362.375 | 2020/12/09 23:47:12
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | UNDE1-E-3 | -5.2494 | -45.9874 |  | 60.8102 |  | 0.18219 | 1.9508 |  | 2.0152
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | UNDE1-B-1 | -1.73 | -32.264 | -4.6428 | -48.4941 | -1.181 | 2052.994 | 2020/12/09 23:56:58
15 |  | UNDE1-B-1 | -1.741 | -33.123 | -4.654 | -49.3092 | -1.349 | 2058.633 | 2020/12/10 00:05:55
15 |  | UNDE1-B-1 | -1.751 | -33.217 | -4.6641 | -49.3984 | -1.343 | 2057.439 | 2020/12/10 00:14:52
15 |  | UNDE1-B-1 | -1.74 | -32.976 | -4.653 | -49.1697 | -1.144 | 2068.517 | 2020/12/10 00:23:49
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | UNDE1-B-1 | -4.6535 | -49.0929 | -1.2543 | 757.3195 |  | 0.0085829 | 0.4322 | -27.4905 | -6.53
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | UNDE1-B-4 | -3.399 | -44.244 | -6.3343 | -59.8618 | -0.793 | 1721.44 | 2020/12/10 00:33:36
16 |  | UNDE1-B-4 | -3.632 | -46.055 | -6.5704 | -61.5802 | -0.611 | 1717.725 | 2020/12/10 00:42:32
16 |  | UNDE1-B-4 | -3.64 | -46.577 | -6.5785 | -62.0755 | -0.757 | 1718.462 | 2020/12/10 00:51:27
16 |  | UNDE1-B-4 | -3.564 | -46.501 | -6.5015 | -62.0034 | -0.6 | 1714.572 | 2020/12/10 01:00:23
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | UNDE1-B-4 | -6.4962 | -61.3802 |  | 415.9735 |  | 0.11183 | 1.0914 | -14.5732 | -2.2215
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | UNDE1-D-1 | -3.314 | -40.692 | -6.2481 | -56.4913 | -0.055 | 1444.604 | 2020/12/10 01:10:10
17 |  | UNDE1-D-1 | -3.123 | -39.452 | -6.0546 | -55.3147 | -0.415 | 1437.953 | 2020/12/10 01:19:06
17 |  | UNDE1-D-1 | -3.137 | -39.532 | -6.0687 | -55.3906 | -0.42 | 1435.762 | 2020/12/10 01:28:02
17 |  | UNDE1-D-1 | -3.097 | -39.32 | -6.0282 | -55.1895 | -0.292 | 1436.319 | 2020/12/10 01:36:59
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | UNDE1-D-1 | -6.0999 | -55.5965 |  | 136.5832 |  | 0.098899 | 0.63471 |  | 1.8162
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | UNDE1-C-5 | -5.643 | -57.06 | -8.6084 | -72.0227 | 0.014 | 1708.466 | 2020/12/10 01:46:47
18 |  | UNDE1-C-5 | -5.873 | -59.223 | -8.8415 | -74.0752 | -0.319 | 1711.317 | 2020/12/10 01:55:44
18 |  | UNDE1-C-5 | -5.883 | -59.829 | -8.8517 | -74.6502 | 0.061 | 1709.999 | 2020/12/10 02:04:40
18 |  | UNDE1-C-5 | -5.933 | -60.148 | -8.9023 | -74.9529 | -0.034 | 1709.951 | 2020/12/10 02:13:37
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | UNDE1-C-5 | -8.801 | -73.9253 |  | 407.857 |  | 0.12936 | 1.3906 | -12.4482 | -4.3493
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | UNDE1-D-3 | -1.733 | -36.467 | -4.6459 | -52.4823 | -0.971 | 1317.527 | 2020/12/10 02:23:25
19 |  | UNDE1-D-3 | -1.376 | -33.594 | -4.2841 | -49.7561 | -0.628 | 1311.189 | 2020/12/10 02:32:21
19 |  | UNDE1-D-3 | -1.342 | -32.504 | -4.2496 | -48.7218 | -0.672 | 1312.549 | 2020/12/10 02:41:17
19 |  | UNDE1-D-3 | -1.338 | -32.312 | -4.2455 | -48.5397 | -0.959 | 1312.331 | 2020/12/10 02:50:14
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | UNDE1-D-3 | -4.3563 | -49.875 | -0.8075 | 11.3227 |  | 0.19126 | 1.9169 |  | 4.647
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | UNDE1-D-5 | -3.602 | -44.996 | -6.54 | -60.5753 | -0.221 | 1407.967 | 2020/12/10 03:00:03
20 |  | UNDE1-D-5 | -3.67 | -46.035 | -6.6089 | -61.5612 | -0.297 | 1408.088 | 2020/12/10 03:08:59
20 |  | UNDE1-D-5 | -3.703 | -46.33 | -6.6424 | -61.8412 | -0.173 | 1407.788 | 2020/12/10 03:17:55
20 |  | UNDE1-D-5 | -3.711 | -46.606 | -6.6505 | -62.1031 | -0.502 | 1408.141 | 2020/12/10 03:26:50
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | UNDE1-D-5 | -6.6104 | -61.5202 |  | 105.9197 |  | 0.049615 | 0.70359 |  | 1.5485
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | UNDE1-A-1 | -2.025 | -44.843 | -4.9418 | -60.4302 | 0.05 | 1505.952 | 2020/12/10 03:36:39
21 |  | UNDE1-A-1 | -1.658 | -43.759 | -4.5699 | -59.4016 | 0.116 | 1509.046 | 2020/12/10 03:45:35
21 |  | UNDE1-A-1 | -1.665 | -43.768 | -4.5769 | -59.4101 | 0.173 | 1508.755 | 2020/12/10 03:54:31
21 |  | UNDE1-A-1 | -1.665 | -43.714 | -4.5769 | -59.3589 | 0.201 | 1509.344 | 2020/12/10 04:03:28
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | UNDE1-A-1 | -4.6664 | -59.6502 |  | 206.198 |  | 0.1812 | 0.54851 |  | -0.21955
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -6.659 | -56.596 | -9.6381 | -71.5824 | 0.281 | 1306.31 | 2020/12/10 04:13:14
22 |  | EV | -7.044 | -58.063 | -10.0283 | -72.9745 | 0.668 | 1304.573 | 2020/12/10 04:22:10
22 |  | EV | -7.165 | -58.598 | -10.1509 | -73.4821 | 0.657 | 1303.365 | 2020/12/10 04:31:07
22 |  | EV | -7.184 | -58.629 | -10.1702 | -73.5115 | 0.654 | 1302.966 | 2020/12/10 04:40:02
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -9.9969 | -72.8876 |  | 2.2272 |  | 0.24401 | 0.9531
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | UNDE1-D-2 | -3.219 | -41.533 | -6.1518 | -57.2893 | -0.344 | 1361.69 | 2020/12/10 04:49:49
23 |  | UNDE1-D-2 | -2.935 | -39.287 | -5.864 | -55.1581 | -0.287 | 1363.945 | 2020/12/10 04:58:45
23 |  | UNDE1-D-2 | -2.898 | -38.649 | -5.8265 | -54.5528 | 0.076 | 1363.349 | 2020/12/10 05:07:42
23 |  | UNDE1-D-2 | -2.941 | -38.495 | -5.8701 | -54.4066 | -0.114 | 1363.3 | 2020/12/10 05:16:38
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | UNDE1-D-2 | -5.9281 | -55.3517 |  | 60.9947 |  | 0.14839 | 1.4038 |  | 2.6737
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | UNDE1-D-4 | -3.055 | -41.023 | -5.9856 | -56.8054 | 0.067 | 1304.994 | 2020/12/10 05:26:25
24 |  | UNDE1-D-4 | -2.959 | -40.786 | -5.8883 | -56.5805 | -0.217 | 1305.331 | 2020/12/10 05:35:21
24 |  | UNDE1-D-4 | -2.923 | -40.725 | -5.8519 | -56.5226 | 0.269 | 1305.097 | 2020/12/10 05:44:17
24 |  | UNDE1-D-4 | -2.936 | -40.814 | -5.865 | -56.6071 | -0.167 | 1305.364 | 2020/12/10 05:53:13
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | UNDE1-D-4 | -5.8977 | -56.6289 |  | 3.1202 |  | 0.059718 | 0.12945
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | UNDE1-A-3 | -3.21 | -41.37 | -6.1427 | -57.1347 | -0.352 | 1694.661 | 2020/12/10 06:03:00
25 |  | UNDE1-A-3 | -3.246 | -41.256 | -6.1792 | -57.0265 | -0.337 | 1706.418 | 2020/12/10 06:11:56
25 |  | UNDE1-A-3 | -3.235 | -41.036 | -6.1681 | -56.8178 | -0.361 | 1715.877 | 2020/12/10 06:20:52
25 |  | UNDE1-A-3 | -3.249 | -41.311 | -6.1822 | -57.0787 | -0.472 | 1717.593 | 2020/12/10 06:29:47
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | UNDE1-A-3 | -6.1681 | -57.0144 |  | 406.561 |  | 0.01772 | 0.1458 | -13.0955 | -2.628
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | UNDE1-E-5 | -3.007 | -39.775 | -5.937 | -55.6212 | -0.22 | 1382.26 | 2020/12/10 06:39:35
26 |  | UNDE1-E-5 | -2.915 | -39.04 | -5.8438 | -54.9238 | 0.084 | 1382.634 | 2020/12/10 06:48:31
26 |  | UNDE1-E-5 | -2.941 | -39.099 | -5.8701 | -54.9798 | -0.53 | 1379.97 | 2020/12/10 06:57:27
26 |  | UNDE1-E-5 | -2.937 | -39.318 | -5.8661 | -55.1876 | -0.186 | 1375.41 | 2020/12/10 07:06:24
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | UNDE1-E-5 | -5.8792 | -55.1781 |  | 77.9922 |  | 0.039682 | 0.33351 |  | 0.72695
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | UNDE1-B-5 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 07:16:13
27 |  | UNDE1-B-5 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 07:25:09
27 |  | UNDE1-B-5 | -3.616 | -43.719 | -6.5542 | -59.3636 | -0.749 | 1855.239 | 2020/12/10 07:34:05
27 |  | UNDE1-B-5 | -3.696 | -43.844 | -6.6353 | -59.4822 | -0.564 | 1851.661 | 2020/12/10 07:43:01
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 | * | UNDE1-B-5 | -4.7421 | -38.6511 |  |  |  | 2.111 | 25.2773 | -9.4992
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | UNDE1-B-2 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 07:52:48
28 |  | UNDE1-B-2 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 08:01:45
28 |  | UNDE1-B-2 | -2.164 | -42.87 | -5.0827 | -58.558 | -1.669 | 1986.751 | 2020/12/10 08:10:41
28 |  | UNDE1-B-2 | -2.203 | -42.712 | -5.1222 | -58.4081 | -1.685 | 1990.111 | 2020/12/10 08:19:38
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 | * | UNDE1-B-2 | -3.996 | -38.1811 | -0.8385 |  |  | 1.2607 | 24.7055 | -12.9678
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | NIWO1-A-4 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 08:29:24
29 |  | NIWO1-A-4 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 08:38:21
29 |  | NIWO1-A-4 | 3.902 | -97.624 | 1.0649 | -110.5134 | -9.699 | 1564.273 | 2020/12/10 08:47:17
29 |  | NIWO1-A-4 | 3.861 | -98.209 | 1.0234 | -111.0685 | -9.621 | 1565.117 | 2020/12/10 08:56:12
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 | * | NIWO1-A-4 | -0.9227 | -64.3351 | -4.83 |  |  | 2.241 | 56.5326 | -17.5288
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | WOOD1-C-2 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 09:05:58
30 |  | WOOD1-C-2 |  |  | -2.8896 | -17.8792 |  |  | 2020/12/10 09:14:54
30 |  | WOOD1-C-2 | -3.957 | -51.621 | -6.8998 | -66.8617 | -0.34 | 1442.601 | 2020/12/10 09:23:50
30 |  | WOOD1-C-2 | -3.945 | -50.991 | -6.8876 | -66.2639 | -0.463 | 1442.329 | 2020/12/10 09:32:46
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 | * | WOOD1-C-2 | -4.8916 | -42.221 |  |  |  | 2.2811 | 29.6226
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -6.552 | -54.359 | -9.5297 | -69.4598 | 0.881 | 1352.645 | 2020/12/10 09:42:32
31 |  | EV | -7.096 | -58.914 | -10.081 | -73.782 | 0.343 | 1304.121 | 2020/12/10 09:51:28
31 |  | EV | -7.188 | -59.222 | -10.1742 | -74.0742 | 0.365 | 1303.252 | 2020/12/10 10:00:24
31 |  | EV | -7.193 | -59.132 | -10.1793 | -73.9888 | 0.51 | 1303.056 | 2020/12/10 10:09:20
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -9.991 | -72.8262 |  | 13.6922 | 10.2329 | 0.30676 | 2.3687 |  | 9.9502

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-14347 | 2020/12/09 11:48:09 | 1-01 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-14347 | 2020/12/09 11:58:08 | 1-01 | 2 | -7.967 | -67.819 | 18709 | -1 | 1 | EV | standard | H2O | 1607540595 | 0.336 | 0.782 | 104.023 | 0 | 0.005 | 2.451 | 1342.445 | 1.305 | 3.076 | 115.208 | 0.863 | 0.566 | 0 | 0 | 1 | 46.438 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-14347 | 2020/12/09 12:07:05 | 1-01 | 3 | -7.406 | -61.388 | 18502 | -1 | 1 | EV | standard | H2O | 1607541132 | 0.321 | 0.705 | 105.733 | -0.001 | 0.001 | 2.521 | 1314.397 | 1.871 | 3.001 | 114.455 | 0.864 | 0.498 | 0 | 0 | 1 | 46.062 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-14347 | 2020/12/09 12:16:03 | 1-01 | 4 | -7.279 | -60.115 | 18489 | 0 | 1 | EV | standard | H2O | 1607541670 | 0.299 | 0.756 | 106.768 | -0.001 | 0.001 | 2.513 | 1310.128 | 1.658 | 2.965 | 114.534 | 0.865 | 0.5 | 0 | 0 | 1 | 45.938 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-14347 | 2020/12/09 12:24:59 | 1-01 | 5 | -7.282 | -59.725 | 18510 | 0 | 1 | EV | standard | H2O | 1607542206 | 0.31 | 0.701 | 107.018 | 0 | 0 | 2.533 | 1307.85 | 2.378 | 2.985 | 113.842 | 0.865 | 0.382 | 0 | 0 | 1 | 45.938 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-14347 | 2020/12/09 12:33:57 | 1-01 | 6 | -7.338 | -59.748 | 18450 | 0 | 1 | EV | standard | H2O | 1607542744 | 0.325 | 0.762 | 106.386 | 0 | 0.001 | 2.523 | 1305.628 | 2.49 | 2.928 | 112.801 | 0.863 | 0.304 | 0 | 0 | 1 | 45.875 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-14347 | 2020/12/09 12:42:54 | 1-01 | 7 | -7.278 | -59.511 | 18584 | 0 | 1 | EV | standard | H2O | 1607543281 | 0.311 | 0.742 | 105.89 | -0.001 | 0.003 | 2.513 | 1305.895 | 2.562 | 3.006 | 114.131 | 0.862 | 0.472 | 0 | 0 | 1 | 45.875 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-14347 | 2020/12/09 12:51:51 | 1-01 | 8 | -7.286 | -59.483 | 18548 | 0 | 1 | EV | standard | H2O | 1607543818 | 0.342 | 0.793 | 106.917 | 0 | 0.003 | 2.536 | 1304.974 | 2.241 | 2.989 | 114.507 | 0.863 | 0.568 | 0 | 0 | 1 | 45.875 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-14347 | 2020/12/09 13:00:49 | 1-01 | 9 | -7.253 | -59.374 | 18451 | 0 | 1 | EV | standard | H2O | 1607544356 | 0.302 | 0.751 | 107.758 | 0 | 0.001 | 2.565 | 1304.092 | 2.176 | 2.965 | 113.703 | 0.865 | 0.44 | 0 | 0 | 1 | 45.875 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-14347 | 2020/12/09 13:09:45 | 1-01 | 10 | -7.197 | -59.244 | 18492 | 0 | 1 | EV | standard | H2O | 1607544892 | 0.297 | 0.774 | 108.81 | -0.001 | 0.004 | 2.579 | 1303.75 | 2.848 | 2.982 | 113.735 | 0.863 | 0.628 | 0 | 0 | 1 | 45.875 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-14348 | 2020/12/09 13:19:34 | 1-02 | 1 | 3.699 | 24.56 | 18638 | -1 | 1 | PZ | standard | H2O | 1607545480 | 0.299 | 0.929 | 106.192 | -0.001 | -0.011 | 2.52 | 1303.901 | 2.489 | 3.039 | 114.938 | 0.863 | 0.607 | 0 | 0 | 1 | 45.938 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-14348 | 2020/12/09 13:28:31 | 1-02 | 2 | 4.542 | 34.887 | 18217 | -1 | 1 | PZ | standard | H2O | 1607546017 | 0.315 | 0.944 | 104.827 | -0.002 | -0.011 | 2.482 | 1301.604 | 2.296 | 2.914 | 110.767 | 0.865 | 0.281 | 0 | 0 | 1 | 45.875 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-14348 | 2020/12/09 13:37:27 | 1-02 | 3 | 4.741 | 37.894 | 18423 | -1 | 1 | PZ | standard | H2O | 1607546553 | 0.306 | 0.85 | 106.729 | -0.001 | -0.006 | 2.514 | 1302.971 | 2.016 | 2.985 | 112.947 | 0.864 | 0.263 | 0 | 0 | 1 | 45.875 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-14348 | 2020/12/09 13:46:24 | 1-02 | 4 | 4.87 | 39.109 | 18294 | 0 | 1 | PZ | standard | H2O | 1607547090 | 0.281 | 0.843 | 105.383 | -0.001 | -0.005 | 2.515 | 1302.081 | 2.219 | 2.916 | 111.713 | 0.862 | 0.363 | 0 | 0 | 1 | 45.875 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-14348 | 2020/12/09 13:55:20 | 1-02 | 5 | 4.899 | 39.792 | 18444 | 0 | 1 | PZ | standard | H2O | 1607547626 | 0.346 | 0.789 | 107.589 | -0.001 | -0.003 | 2.536 | 1302.831 | 2.687 | 2.955 | 112.472 | 0.863 | 0.287 | 0 | 0 | 1 | 45.875 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-14348 | 2020/12/09 14:04:17 | 1-02 | 6 | 4.936 | 40.247 | 18458 | 0 | 1 | PZ | standard | H2O | 1607548163 | 0.312 | 0.798 | 107.406 | 0 | -0.005 | 2.532 | 1302.841 | 2.031 | 3.008 | 114.345 | 0.863 | 0.595 | 0 | 0 | 1 | 45.875 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-14348 | 2020/12/09 14:13:13 | 1-02 | 7 | 4.939 | 40.254 | 18199 | 0 | 1 | PZ | standard | H2O | 1607548700 | 0.307 | 0.81 | 106.362 | 0 | 0 | 2.524 | 1301.194 | 1.559 | 2.949 | 112.818 | 0.863 | 0.476 | 0 | 0 | 1 | 45.875 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-14348 | 2020/12/09 14:22:10 | 1-02 | 8 | 4.964 | 40.74 | 18430 | 0 | 1 | PZ | standard | H2O | 1607549236 | 0.322 | 0.746 | 104.992 | -0.001 | -0.004 | 2.488 | 1302.748 | 2.446 | 2.965 | 112.079 | 0.863 | 0.319 | 0 | 0 | 1 | 45.812 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-14348 | 2020/12/09 14:31:05 | 1-02 | 9 | 4.914 | 40.591 | 18398 | 0 | 1 | PZ | standard | H2O | 1607549772 | 0.342 | 0.761 | 111.487 | 0.001 | 0.004 | 2.626 | 1302.458 | 2.089 | 2.994 | 112.899 | 0.864 | 0.488 | 0 | 0 | 1 | 45.75 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-14348 | 2020/12/09 14:40:03 | 1-02 | 10 | 5.052 | 41.098 | 18286 | 0 | 1 | PZ | standard | H2O | 1607550309 | 0.313 | 0.769 | 107.284 | -0.001 | -0.003 | 2.562 | 1301.51 | 2.49 | 2.933 | 111.766 | 0.863 | 0.322 | 0 | 0 | 1 | 45.688 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-14349 | 2020/12/09 14:49:53 | 1-03 | 1 | -11.209 | -86.131 | 18359 | -1 | 1 | UT2 | standard | H2O | 1607550900 | 0.295 | 1.143 | 104.274 | 0.001 | 0.02 | 2.48 | 1301.884 | 2.596 | 2.962 | 112.561 | 0.861 | 0.425 | 0 | 0 | 1 | 45.562 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-14349 | 2020/12/09 14:58:50 | 1-03 | 2 | -12.443 | -102.398 | 18411 | -1 | 1 | UT2 | standard | H2O | 1607551436 | 0.325 | 0.876 | 105.719 | 0 | 0.014 | 2.51 | 1302.138 | 3.262 | 2.926 | 112.709 | 0.862 | 0.675 | 0 | 0 | 1 | 45.562 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-14349 | 2020/12/09 15:07:47 | 1-03 | 3 | -12.821 | -106.808 | 18277 | -1 | 1 | UT2 | standard | H2O | 1607551973 | 0.323 | 0.942 | 107.299 | 0.001 | 0.015 | 2.545 | 1301.474 | 3.013 | 2.883 | 110.437 | 0.862 | 0.368 | 0 | 0 | 1 | 45.5 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-14349 | 2020/12/09 15:16:44 | 1-03 | 4 | -12.903 | -108.37 | 18397 | 0 | 1 | UT2 | standard | H2O | 1607552510 | 0.28 | 0.878 | 107.517 | 0.001 | 0.01 | 2.552 | 1302.175 | 3.005 | 2.95 | 112.203 | 0.862 | 0.35 | 0 | 0 | 1 | 45.5 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-14349 | 2020/12/09 15:25:40 | 1-03 | 5 | -13.02 | -109.515 | 18352 | 0 | 1 | UT2 | standard | H2O | 1607553047 | 0.314 | 0.823 | 106.117 | 0.001 | 0.012 | 2.531 | 1301.69 | 3.249 | 2.936 | 111.741 | 0.863 | 0.364 | 0 | 0 | 1 | 45.5 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-14349 | 2020/12/09 15:34:38 | 1-03 | 6 | -13.019 | -109.782 | 18366 | 0 | 1 | UT2 | standard | H2O | 1607553585 | 0.287 | 0.782 | 104.011 | 0 | 0.007 | 2.472 | 1301.871 | 3.085 | 2.919 | 112.165 | 0.861 | 0.493 | 0 | 0 | 1 | 45.438 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-14349 | 2020/12/09 15:43:35 | 1-03 | 7 | -13.041 | -110.09 | 18342 | 0 | 1 | UT2 | standard | H2O | 1607554121 | 0.3 | 0.848 | 106.571 | 0 | 0.007 | 2.527 | 1301.587 | 3.8 | 2.889 | 110.201 | 0.863 | 0.3 | 0 | 0 | 1 | 45.438 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-14349 | 2020/12/09 15:52:31 | 1-03 | 8 | -13.069 | -110.35 | 18338 | 0 | 1 | UT2 | standard | H2O | 1607554657 | 0.314 | 0.76 | 106.436 | 0 | 0.002 | 2.528 | 1301.866 | 2.588 | 2.957 | 112.97 | 0.862 | 0.557 | 0 | 0 | 1 | 45.375 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-14349 | 2020/12/09 16:01:28 | 1-03 | 9 | -13.128 | -111.08 | 18296 | 0 | 1 | UT2 | standard | H2O | 1607555194 | 0.31 | 0.802 | 106.909 | 0.001 | 0.008 | 2.545 | 1301.123 | 3.847 | 2.929 | 110.843 | 0.865 | 0.54 | 0 | 0 | 1 | 45.438 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-14349 | 2020/12/09 16:10:26 | 1-03 | 10 | -13.133 | -111.153 | 18314 | 0 | 1 | UT2 | standard | H2O | 1607555733 | 0.294 | 0.772 | 104.548 | 0 | 0.006 | 2.48 | 1301.579 | 3.142 | 2.889 | 110.717 | 0.862 | 0.303 | 0 | 0 | 1 | 45.375 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-14350 | 2020/12/09 16:20:15 | 1-04 | 1 | -7.915 | -68.019 | 18274 | -1 | 1 | EV | standard | H2O | 1607556322 | 0.31 | 0.784 | 104.072 | 0 | 0.001 | 2.471 | 1301.22 | 2.836 | 2.955 | 112.043 | 0.863 | 0.545 | 0 | 0 | 1 | 45.312 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-14350 | 2020/12/09 16:29:11 | 1-04 | 2 | -7.328 | -61.412 | 18488 | -1 | 1 | EV | standard | H2O | 1607556857 | 0.308 | 0.766 | 106.198 | 0 | 0 | 2.52 | 1302.988 | 2.675 | 2.96 | 113.003 | 0.863 | 0.43 | 0 | 0 | 1 | 45.312 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-14350 | 2020/12/09 16:38:07 | 1-04 | 3 | -7.352 | -60.399 | 18318 | -1 | 1 | EV | standard | H2O | 1607557393 | 0.32 | 0.699 | 106.998 | -0.001 | 0.002 | 2.545 | 1301.633 | 2.506 | 2.921 | 112.781 | 0.866 | 0.401 | 0 | 0 | 1 | 45.312 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-14350 | 2020/12/09 16:47:04 | 1-04 | 4 | -7.354 | -60.349 | 18298 | 0 | 1 | EV | standard | H2O | 1607557930 | 0.323 | 0.736 | 106.976 | 0.001 | 0.006 | 2.543 | 1301.326 | 2.764 | 3.008 | 113.117 | 0.861 | 0.756 | 0 | 0 | 1 | 45.375 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-14350 | 2020/12/09 16:56:01 | 1-04 | 5 | -7.194 | -59.595 | 18317 | 0 | 1 | EV | standard | H2O | 1607558467 | 0.274 | 0.72 | 105.973 | 0 | 0.001 | 2.531 | 1301.72 | 2.534 | 2.921 | 112.178 | 0.862 | 0.505 | 0 | 0 | 1 | 45.438 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-14350 | 2020/12/09 17:04:57 | 1-04 | 6 | -7.272 | -59.482 | 18292 | 0 | 1 | EV | standard | H2O | 1607559004 | 0.305 | 0.789 | 108.17 | 0.001 | 0.004 | 2.57 | 1301.685 | 2.531 | 2.876 | 110.608 | 0.862 | 0.101 | 0 | 0 | 1 | 45.5 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-14350 | 2020/12/09 17:13:55 | 1-04 | 7 | -7.304 | -59.819 | 18240 | 0 | 1 | EV | standard | H2O | 1607559541 | 0.287 | 0.775 | 106.508 | 0 | 0.006 | 2.546 | 1300.743 | 2.889 | 2.917 | 112.179 | 0.862 | 0.462 | 0 | 0 | 1 | 45.625 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-14350 | 2020/12/09 17:22:51 | 1-04 | 8 | -7.261 | -59.636 | 18250 | 0 | 1 | EV | standard | H2O | 1607560078 | 0.311 | 0.877 | 106.211 | 0.001 | 0.006 | 2.511 | 1300.881 | 3.19 | 2.956 | 111.318 | 0.866 | 0.498 | 0 | 0 | 1 | 45.688 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-14350 | 2020/12/09 17:31:48 | 1-04 | 9 | -7.245 | -59.345 | 18255 | 0 | 1 | EV | standard | H2O | 1607560614 | 0.318 | 0.747 | 106.103 | 0 | 0.004 | 2.518 | 1300.855 | 2.912 | 2.978 | 112.204 | 0.862 | 0.677 | 0 | 0 | 1 | 45.688 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-14350 | 2020/12/09 17:40:45 | 1-04 | 10 | -7.224 | -59.254 | 18263 | 0 | 1 | EV | standard | H2O | 1607561151 | 0.331 | 0.838 | 105.635 | -0.001 | 0.003 | 2.506 | 1301.065 | 3.025 | 2.887 | 111.567 | 0.863 | 0.421 | 0 | 0 | 1 | 45.688 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-14351 | 2020/12/09 17:50:34 | 1-05 | 1 | -3.408 | -42.792 | 18238 | -1 | 1 | UNDE1-A-5 | 20-196 | H2O | 1607561741 | 0.363 | 1.017 | 105.73 | 0 | 0.003 | 2.517 | 1741.756 | -14.572 | 2.925 | 109.44 | 0.865 | -0.337 | 0 | 0 | 1 | 45.688 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-14351 | 2020/12/09 17:59:32 | 1-05 | 2 | -3.097 | -40.613 | 18206 | -1 | 1 | UNDE1-A-5 | 20-196 | H2O | 1607562278 | 0.356 | 0.932 | 104.263 | 0 | 0.005 | 2.49 | 1743.619 | -14.187 | 2.974 | 109.908 | 0.868 | -0.126 | 0 | 0 | 1 | 45.688 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-14351 | 2020/12/09 18:08:28 | 1-05 | 3 | -2.964 | -39.66 | 18309 | -1 | 1 | UNDE1-A-5 | 20-196 | H2O | 1607562815 | 0.363 | 0.992 | 106.508 | 0 | 0.004 | 2.528 | 1748.481 | -14.702 | 2.953 | 110.014 | 0.867 | -0.343 | 0 | 0 | 1 | 45.688 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-14351 | 2020/12/09 18:17:26 | 1-05 | 4 | -3.05 | -39.774 | 18319 | 0 | 1 | UNDE1-A-5 | 20-196 | H2O | 1607563353 | 0.374 | 0.926 | 108.936 | 0.001 | 0.003 | 2.577 | 1749.74 | -14.537 | 2.94 | 109.854 | 0.865 | -0.172 | 0 | 0 | 1 | 45.688 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-14352 | 2020/12/09 18:27:17 | 1-06 | 1 | -2.96 | -42.877 | 18156 | -1 | 1 | UNDE1-C-4 | 20-196 | H2O | 1607563943 | 0.335 | 0.776 | 106.526 | 0 | 0.007 | 2.53 | 1309.239 | 0.9 | 2.938 | 112.137 | 0.862 | 0.029 | 0 | 0 | 1 | 45.688 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-14352 | 2020/12/09 18:36:13 | 1-06 | 2 | -2.798 | -42.477 | 18423 | -1 | 1 | UNDE1-C-4 | 20-196 | H2O | 1607564480 | 0.297 | 0.842 | 106.333 | -0.001 | 0.002 | 2.538 | 1307.289 | 1.127 | 2.896 | 112.776 | 0.863 | -0.118 | 0 | 0 | 1 | 45.562 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-14352 | 2020/12/09 18:45:10 | 1-06 | 3 | -2.846 | -42.825 | 18515 | -1 | 1 | UNDE1-C-4 | 20-196 | H2O | 1607565016 | 0.341 | 0.777 | 107.297 | -0.001 | 0.003 | 2.543 | 1306.545 | 1.454 | 2.97 | 114.694 | 0.863 | 0.261 | 0 | 0 | 1 | 45.688 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-14352 | 2020/12/09 18:54:08 | 1-06 | 4 | -2.893 | -43.048 | 18334 | 0 | 1 | UNDE1-C-4 | 20-196 | H2O | 1607565554 | 0.299 | 0.762 | 103.962 | 0.001 | 0.004 | 2.471 | 1304.515 | 1.694 | 2.97 | 113.479 | 0.861 | 0.146 | 0 | 0 | 1 | 45.688 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-14353 | 2020/12/09 19:03:57 | 1-07 | 1 | -2.045 | -35.654 | 18632 | -1 | 1 | UNDE1-E-4 | 20-196 | H2O | 1607566144 | 0.322 | 0.845 | 106.218 | -0.001 | -0.001 | 2.503 | 1388.731 | -2.444 | 2.934 | 113.947 | 0.864 | -0.571 | 0 | 0 | 1 | 45.688 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-14353 | 2020/12/09 19:12:53 | 1-07 | 2 | -2.14 | -35.265 | 18446 | -1 | 1 | UNDE1-E-4 | 20-196 | H2O | 1607566680 | 0.307 | 0.921 | 105.881 | 0.002 | 0.008 | 2.526 | 1387.193 | -2.45 | 2.927 | 114.106 | 0.864 | -0.366 | 0 | 0 | 1 | 45.625 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-14353 | 2020/12/09 19:21:50 | 1-07 | 3 | -2.045 | -34.677 | 18464 | -1 | 1 | UNDE1-E-4 | 20-196 | H2O | 1607567216 | 0.323 | 0.811 | 105.4 | 0 | 0.002 | 2.504 | 1387.431 | -2.191 | 2.961 | 114.464 | 0.862 | -0.215 | 0 | 0 | 1 | 45.688 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-14353 | 2020/12/09 19:30:46 | 1-07 | 4 | -2.03 | -34.627 | 18440 | 0 | 1 | UNDE1-E-4 | 20-196 | H2O | 1607567752 | 0.34 | 0.772 | 105.479 | 0 | 0.002 | 2.487 | 1387.421 | -2.637 | 2.947 | 114.613 | 0.863 | -0.236 | 0 | 0 | 1 | 45.688 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-14354 | 2020/12/09 19:40:34 | 1-08 | 1 | -3.693 | -66.576 | 18631 | -1 | 1 | CPER1-C-2 | 20-196 | H2O | 1607568340 | 0.329 | 0.914 | 106.455 | 0.001 | 0.011 | 2.524 | 1470.296 | -5.143 | 2.938 | 113.517 | 0.863 | -0.27 | 0 | 0 | 1 | 45.688 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-14354 | 2020/12/09 19:49:30 | 1-08 | 2 | -3.681 | -70.406 | 18307 | -1 | 1 | CPER1-C-2 | 20-196 | H2O | 1607568877 | 0.339 | 0.823 | 103.396 | 0 | 0.004 | 2.465 | 1465.428 | -4.684 | 2.941 | 113.276 | 0.863 | 0.014 | 0 | 0 | 1 | 45.75 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-14354 | 2020/12/09 19:58:27 | 1-08 | 3 | -3.735 | -71.401 | 18339 | -1 | 1 | CPER1-C-2 | 20-196 | H2O | 1607569413 | 0.325 | 0.853 | 103.324 | 0.001 | 0.007 | 2.435 | 1466.407 | -4.659 | 2.926 | 111.937 | 0.863 | -0.169 | 0 | 0 | 1 | 45.75 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-14354 | 2020/12/09 20:07:24 | 1-08 | 4 | -3.789 | -71.837 | 18275 | 0 | 1 | CPER1-C-2 | 20-196 | H2O | 1607569951 | 0.33 | 0.906 | 104.686 | 0 | 0.006 | 2.495 | 1465.305 | -4.572 | 2.952 | 111.729 | 0.864 | -0.151 | 0 | 0 | 1 | 45.688 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-14355 | 2020/12/09 20:17:12 | 1-09 | 1 | -2.882 | -42.984 | 18305 | -1 | 1 | UNDE1-E-1 | 20-196 | H2O | 1607570539 | 0.317 | 0.784 | 107 | -0.001 | 0 | 2.55 | 1352.52 | -1.403 | 2.912 | 114.004 | 0.864 | -0.242 | 0 | 0 | 1 | 45.688 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-14355 | 2020/12/09 20:26:09 | 1-09 | 2 | -2.804 | -39.216 | 18318 | -1 | 1 | UNDE1-E-1 | 20-196 | H2O | 1607571075 | 0.347 | 0.772 | 106.114 | -0.001 | 0.002 | 2.515 | 1351.563 | -1.051 | 2.959 | 113.287 | 0.863 | -0.315 | 0 | 0 | 1 | 45.75 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-14355 | 2020/12/09 20:35:05 | 1-09 | 3 | -2.816 | -38.421 | 18231 | -1 | 1 | UNDE1-E-1 | 20-196 | H2O | 1607571611 | 0.313 | 0.741 | 106.734 | 0 | 0.002 | 2.527 | 1350.37 | -0.418 | 2.834 | 111.621 | 0.864 | -0.328 | 0 | 0 | 1 | 45.75 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-14355 | 2020/12/09 20:44:01 | 1-09 | 4 | -2.756 | -37.878 | 18353 | 0 | 1 | UNDE1-E-1 | 20-196 | H2O | 1607572147 | 0.316 | 0.798 | 107.171 | 0 | 0.001 | 2.546 | 1351.33 | -0.625 | 2.961 | 113.142 | 0.861 | -0.213 | 0 | 0 | 1 | 45.688 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-14356 | 2020/12/09 20:53:48 | 1-10 | 1 | -3.142 | -41.081 | 18496 | -1 | 1 | UNDE1-A-2 | 20-196 | H2O | 1607572734 | 0.356 | 0.967 | 106.759 | 0.002 | 0.01 | 2.54 | 1687.042 | -12.216 | 2.993 | 111.963 | 0.866 | 0.09 | 0 | 0 | 1 | 45.75 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-14356 | 2020/12/09 21:02:44 | 1-10 | 2 | -3.121 | -41.353 | 18177 | -1 | 1 | UNDE1-A-2 | 20-196 | H2O | 1607573270 | 0.346 | 0.914 | 106.098 | -0.001 | 0.003 | 2.507 | 1681.285 | -12.395 | 2.919 | 109.716 | 0.866 | -0.186 | 0 | 0 | 1 | 45.75 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-14356 | 2020/12/09 21:11:41 | 1-10 | 3 | -3.065 | -40.99 | 18004 | -1 | 1 | UNDE1-A-2 | 20-196 | H2O | 1607573807 | 0.32 | 0.945 | 104.464 | 0 | 0.003 | 2.476 | 1676.859 | -12.347 | 2.877 | 108.999 | 0.864 | -0.151 | 0 | 0 | 1 | 45.688 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-14356 | 2020/12/09 21:20:37 | 1-10 | 4 | -3.1 | -41.174 | 18265 | 0 | 1 | UNDE1-A-2 | 20-196 | H2O | 1607574343 | 0.357 | 1.001 | 109.303 | 0 | 0.007 | 2.599 | 1685.113 | -12.491 | 2.978 | 111.473 | 0.866 | -0.027 | 0 | 0 | 1 | 45.75 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-14357 | 2020/12/09 21:30:25 | 1-11 | 1 | -5.195 | -59.413 | 18014 | -1 | 1 | UNDE1-C-3 | 20-196 | H2O | 1607574932 | 0.368 | 1.026 | 106.227 | 0 | 0.01 | 2.524 | 1672.264 | -11.117 | 2.855 | 106.955 | 0.866 | -0.291 | 0 | 0 | 1 | 45.75 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-14357 | 2020/12/09 21:39:22 | 1-11 | 2 | -5.283 | -61.221 | 18344 | -1 | 1 | UNDE1-C-3 | 20-196 | H2O | 1607575469 | 0.306 | 0.88 | 105.917 | 0 | 0.008 | 2.519 | 1680.741 | -12.263 | 2.92 | 110.449 | 0.865 | -0.325 | 0 | 0 | 1 | 45.75 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-14357 | 2020/12/09 21:48:20 | 1-11 | 3 | -5.457 | -62.479 | 17997 | -1 | 1 | UNDE1-C-3 | 20-196 | H2O | 1607576006 | 0.348 | 0.992 | 105.929 | 0.001 | 0.009 | 2.524 | 1671.38 | -11.906 | 2.884 | 108.806 | 0.864 | -0.08 | 0 | 0 | 1 | 45.75 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-14357 | 2020/12/09 21:57:17 | 1-11 | 4 | -5.322 | -62.021 | 18288 | 0 | 1 | UNDE1-C-3 | 20-196 | H2O | 1607576543 | 0.358 | 0.888 | 106.359 | 0 | 0.004 | 2.531 | 1679.023 | -12.119 | 2.94 | 110.916 | 0.865 | -0.139 | 0 | 0 | 1 | 45.75 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-14358 | 2020/12/09 22:07:05 | 1-12 | 1 | -3.296 | -45.186 | 18092 | -1 | 1 | UNDE1-B-3 | 20-196 | H2O | 1607577132 | 0.391 | 1.079 | 104.386 | 0.002 | 0.004 | 2.47 | 1920.384 | -21.062 | 2.917 | 108.718 | 0.87 | -0.196 | 0 | 0 | 1 | 45.688 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-14358 | 2020/12/09 22:16:02 | 1-12 | 2 | -3.141 | -42.834 | 18066 | -1 | 1 | UNDE1-B-3 | 20-196 | H2O | 1607577668 | 0.38 | 1.1 | 106.699 | 0 | 0.005 | 2.53 | 1922.08 | -21.44 | 2.94 | 107.619 | 0.868 | -0.641 | 0 | 0 | 1 | 45.688 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-14358 | 2020/12/09 22:24:58 | 1-12 | 3 | -3.13 | -42.236 | 18067 | -1 | 1 | UNDE1-B-3 | 20-196 | H2O | 1607578205 | 0.379 | 1.016 | 105.448 | 0 | 0.002 | 2.517 | 1922.577 | -21.591 | 2.939 | 107.821 | 0.867 | -0.631 | 0 | 0 | 1 | 45.75 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-14358 | 2020/12/09 22:33:55 | 1-12 | 4 | -3.108 | -41.908 | 17990 | 0 | 1 | UNDE1-B-3 | 20-196 | H2O | 1607578741 | 0.381 | 1.023 | 102.384 | -0.001 | 0 | 2.442 | 1919.389 | -20.884 | 2.873 | 104.98 | 0.868 | -0.856 | 0 | 0 | 1 | 45.688 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-14359 | 2020/12/09 22:43:44 | 1-13 | 1 | -6.845 | -56.79 | 17998 | -1 | 1 | EV | standard | H2O | 1607579331 | 0.323 | 0.882 | 106.199 | 0.001 | 0.011 | 2.519 | 1309.732 | 2.108 | 2.848 | 110.575 | 0.862 | 0.374 | 0 | 0 | 1 | 45.75 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-14359 | 2020/12/09 22:52:40 | 1-13 | 2 | -7.044 | -58.149 | 18287 | -1 | 1 | EV | standard | H2O | 1607579867 | 0.326 | 0.789 | 105.894 | -0.001 | 0.004 | 2.514 | 1305.811 | 2.299 | 2.927 | 111.895 | 0.862 | 0.383 | 0 | 0 | 1 | 45.75 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-14359 | 2020/12/09 23:01:36 | 1-13 | 3 | -7.117 | -58.67 | 18218 | -1 | 1 | EV | standard | H2O | 1607580403 | 0.303 | 0.777 | 106.427 | 0.001 | 0.006 | 2.541 | 1304.138 | 2.154 | 2.876 | 111.439 | 0.862 | 0.32 | 0 | 0 | 1 | 45.75 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-14359 | 2020/12/09 23:10:33 | 1-13 | 4 | -7.163 | -58.829 | 18221 | 0 | 1 | EV | standard | H2O | 1607580939 | 0.342 | 0.777 | 105.824 | -0.001 | 0.004 | 2.514 | 1303.115 | 2.738 | 2.879 | 110.995 | 0.861 | 0.354 | 0 | 0 | 1 | 45.812 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-14360 | 2020/12/09 23:20:23 | 1-14 | 1 | -2.587 | -32.474 | 18773 | -1 | 1 | UNDE1-E-3 | 20-196 | H2O | 1607581529 | 0.331 | 0.655 | 106.76 | 0 | -0.002 | 2.531 | 1364.514 | -0.854 | 3.04 | 116.055 | 0.863 | 0.077 | 0 | 0 | 1 | 45.812 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-14360 | 2020/12/09 23:29:19 | 1-14 | 2 | -2.31 | -29.227 | 18518 | -1 | 1 | UNDE1-E-3 | 20-196 | H2O | 1607582065 | 0.313 | 0.755 | 107.697 | -0.001 | -0.002 | 2.571 | 1362.881 | -0.784 | 3.005 | 113.498 | 0.862 | -0.113 | 0 | 0 | 1 | 45.812 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-14360 | 2020/12/09 23:38:15 | 1-14 | 3 | -2.251 | -28.629 | 18410 | -1 | 1 | UNDE1-E-3 | 20-196 | H2O | 1607582602 | 0.308 | 0.705 | 107.782 | -0.001 | 0.001 | 2.555 | 1361.776 | -0.835 | 2.972 | 113.081 | 0.863 | -0.144 | 0 | 0 | 1 | 45.812 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-14360 | 2020/12/09 23:47:12 | 1-14 | 4 | -2.166 | -28.159 | 18448 | 0 | 1 | UNDE1-E-3 | 20-196 | H2O | 1607583138 | 0.311 | 0.779 | 111.252 | -0.001 | 0 | 2.623 | 1362.375 | -1.306 | 2.963 | 114.085 | 0.863 | 0.019 | 0 | 0 | 1 | 45.812 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-14361 | 2020/12/09 23:56:58 | 1-15 | 1 | -1.73 | -32.264 | 17958 | -1 | 1 | UNDE1-B-1 | 20-196 | H2O | 1607583725 | 0.404 | 1.013 | 103.026 | 0 | 0.001 | 2.457 | 2052.994 | -26.928 | 2.848 | 105.509 | 0.868 | -1.181 | 0 | 0 | 1 | 45.812 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-14361 | 2020/12/10 00:05:55 | 1-15 | 2 | -1.741 | -33.123 | 17935 | -1 | 1 | UNDE1-B-1 | 20-196 | H2O | 1607584262 | 0.409 | 1.041 | 106.178 | 0.001 | 0.005 | 2.532 | 2058.633 | -27.563 | 2.846 | 104.533 | 0.868 | -1.349 | 0 | 0 | 1 | 45.75 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-14361 | 2020/12/10 00:14:52 | 1-15 | 3 | -1.751 | -33.217 | 17881 | -1 | 1 | UNDE1-B-1 | 20-196 | H2O | 1607584798 | 0.419 | 1.2 | 104.455 | -0.001 | 0.004 | 2.49 | 2057.439 | -27.693 | 2.849 | 105.65 | 0.869 | -1.343 | 0 | 0 | 1 | 45.75 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-14361 | 2020/12/10 00:23:49 | 1-15 | 4 | -1.74 | -32.976 | 18090 | 0 | 1 | UNDE1-B-1 | 20-196 | H2O | 1607585335 | 0.373 | 1.009 | 105.584 | 0 | 0.004 | 2.505 | 2068.517 | -27.778 | 2.908 | 106.846 | 0.872 | -1.144 | 0 | 0 | 1 | 45.812 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-14362 | 2020/12/10 00:33:36 | 1-16 | 1 | -3.399 | -44.244 | 18200 | -1 | 1 | UNDE1-B-4 | 20-196 | H2O | 1607585922 | 0.394 | 0.995 | 107.307 | 0 | 0.006 | 2.549 | 1721.44 | -14.706 | 2.879 | 109.819 | 0.866 | -0.793 | 0 | 0 | 1 | 45.75 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-14362 | 2020/12/10 00:42:32 | 1-16 | 2 | -3.632 | -46.055 | 18180 | -1 | 1 | UNDE1-B-4 | 20-196 | H2O | 1607586458 | 0.348 | 0.964 | 106.099 | 0 | 0.003 | 2.508 | 1717.725 | -14.566 | 2.906 | 109.723 | 0.866 | -0.611 | 0 | 0 | 1 | 45.75 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-14362 | 2020/12/10 00:51:27 | 1-16 | 3 | -3.64 | -46.577 | 18223 | -1 | 1 | UNDE1-B-4 | 20-196 | H2O | 1607586994 | 0.326 | 0.88 | 104.102 | 0.001 | 0.004 | 2.473 | 1718.462 | -14.68 | 2.925 | 110.251 | 0.866 | -0.757 | 0 | 0 | 1 | 45.75 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-14362 | 2020/12/10 01:00:23 | 1-16 | 4 | -3.564 | -46.501 | 18115 | 0 | 1 | UNDE1-B-4 | 20-196 | H2O | 1607587529 | 0.394 | 0.932 | 106.487 | -0.001 | 0.003 | 2.525 | 1714.572 | -14.341 | 2.928 | 109.979 | 0.868 | -0.6 | 0 | 0 | 1 | 45.812 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-14363 | 2020/12/10 01:10:10 | 1-17 | 1 | -3.314 | -40.692 | 18767 | -1 | 1 | UNDE1-D-1 | 20-196 | H2O | 1607588115 | 0.318 | 0.9 | 108.345 | 0 | 0.008 | 2.567 | 1444.604 | -4.043 | 3.017 | 116.201 | 0.864 | -0.055 | 0 | 0 | 1 | 45.812 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-14363 | 2020/12/10 01:19:06 | 1-17 | 2 | -3.123 | -39.452 | 18480 | -1 | 1 | UNDE1-D-1 | 20-196 | H2O | 1607588652 | 0.332 | 0.827 | 107.096 | -0.002 | 0.002 | 2.538 | 1437.953 | -4.218 | 2.975 | 113.599 | 0.864 | -0.415 | 0 | 0 | 1 | 45.812 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-14363 | 2020/12/10 01:28:02 | 1-17 | 3 | -3.137 | -39.532 | 18395 | -1 | 1 | UNDE1-D-1 | 20-196 | H2O | 1607589188 | 0.327 | 0.882 | 107.18 | -0.002 | 0.003 | 2.54 | 1435.762 | -3.883 | 2.931 | 112.344 | 0.864 | -0.42 | 0 | 0 | 1 | 45.812 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-14363 | 2020/12/10 01:36:59 | 1-17 | 4 | -3.097 | -39.32 | 18456 | 0 | 1 | UNDE1-D-1 | 20-196 | H2O | 1607589725 | 0.325 | 0.855 | 106.722 | -0.002 | -0.003 | 2.541 | 1436.319 | -3.973 | 2.964 | 113.779 | 0.864 | -0.292 | 0 | 0 | 1 | 45.812 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-14364 | 2020/12/10 01:46:47 | 1-18 | 1 | -5.643 | -57.06 | 18169 | -1 | 1 | UNDE1-C-5 | 20-196 | H2O | 1607590314 | 0.351 | 0.966 | 104.007 | -0.001 | 0.009 | 2.469 | 1708.466 | -12.323 | 2.951 | 108.515 | 0.866 | 0.014 | 0 | 0 | 1 | 45.812 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-14364 | 2020/12/10 01:55:44 | 1-18 | 2 | -5.873 | -59.223 | 18146 | -1 | 1 | UNDE1-C-5 | 20-196 | H2O | 1607590850 | 0.331 | 0.869 | 104.698 | -0.001 | 0.001 | 2.478 | 1711.317 | -12.785 | 2.893 | 107.349 | 0.865 | -0.319 | 0 | 0 | 1 | 45.875 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-14364 | 2020/12/10 02:04:40 | 1-18 | 3 | -5.883 | -59.829 | 18091 | -1 | 1 | UNDE1-C-5 | 20-196 | H2O | 1607591386 | 0.351 | 0.977 | 105.25 | 0 | 0.003 | 2.506 | 1709.999 | -12.139 | 2.897 | 107.658 | 0.865 | 0.061 | 0 | 0 | 1 | 45.812 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-14364 | 2020/12/10 02:13:37 | 1-18 | 4 | -5.933 | -60.148 | 18081 | 0 | 1 | UNDE1-C-5 | 20-196 | H2O | 1607591923 | 0.347 | 0.834 | 104.145 | -0.001 | 0 | 2.462 | 1709.951 | -12.546 | 2.846 | 107.739 | 0.866 | -0.034 | 0 | 0 | 1 | 45.875 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-14365 | 2020/12/10 02:23:25 | 1-19 | 1 | -1.733 | -36.467 | 18665 | -1 | 1 | UNDE1-D-3 | 20-196 | H2O | 1607592512 | 0.322 | 0.753 | 106.317 | -0.001 | 0.002 | 2.537 | 1317.527 | -1.524 | 2.976 | 117.375 | 0.863 | -0.971 | 0 | 0 | 1 | 45.875 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-14365 | 2020/12/10 02:32:21 | 1-19 | 2 | -1.376 | -33.594 | 18398 | -1 | 1 | UNDE1-D-3 | 20-196 | H2O | 1607593048 | 0.291 | 0.735 | 106.575 | -0.001 | 0.004 | 2.526 | 1311.189 | -1.134 | 2.96 | 115.666 | 0.863 | -0.628 | 0 | 0 | 1 | 45.875 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-14365 | 2020/12/10 02:41:17 | 1-19 | 3 | -1.342 | -32.504 | 18713 | -1 | 1 | UNDE1-D-3 | 20-196 | H2O | 1607593584 | 0.305 | 0.744 | 103.887 | -0.001 | 0 | 2.46 | 1312.549 | -0.746 | 3.02 | 117.121 | 0.864 | -0.672 | 0 | 0 | 1 | 45.875 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-14365 | 2020/12/10 02:50:14 | 1-19 | 4 | -1.338 | -32.312 | 18712 | 0 | 1 | UNDE1-D-3 | 20-196 | H2O | 1607594121 | 0.332 | 0.697 | 107.192 | -0.001 | 0 | 2.545 | 1312.331 | -1.318 | 2.973 | 117.084 | 0.862 | -0.959 | 0 | 0 | 1 | 45.875 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-14366 | 2020/12/10 03:00:03 | 1-20 | 1 | -3.602 | -44.996 | 18507 | -1 | 1 | UNDE1-D-5 | 20-196 | H2O | 1607594709 | 0.336 | 0.895 | 106.203 | -0.001 | 0.008 | 2.502 | 1407.967 | -2.93 | 2.991 | 114.915 | 0.864 | -0.221 | 0 | 0 | 1 | 45.875 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-14366 | 2020/12/10 03:08:59 | 1-20 | 2 | -3.67 | -46.035 | 18394 | -1 | 1 | UNDE1-D-5 | 20-196 | H2O | 1607595246 | 0.34 | 0.888 | 106.268 | 0 | 0.006 | 2.533 | 1408.088 | -3.141 | 2.991 | 114.333 | 0.863 | -0.297 | 0 | 0 | 1 | 45.875 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-14366 | 2020/12/10 03:17:55 | 1-20 | 3 | -3.703 | -46.33 | 18358 | -1 | 1 | UNDE1-D-5 | 20-196 | H2O | 1607595781 | 0.318 | 0.732 | 105.643 | -0.001 | 0 | 2.508 | 1407.788 | -3.361 | 2.989 | 114.163 | 0.862 | -0.173 | 0 | 0 | 1 | 45.875 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-14366 | 2020/12/10 03:26:50 | 1-20 | 4 | -3.711 | -46.606 | 18366 | 0 | 1 | UNDE1-D-5 | 20-196 | H2O | 1607596317 | 0.308 | 0.816 | 106.638 | 0 | 0.003 | 2.542 | 1408.141 | -2.63 | 2.902 | 111.441 | 0.864 | -0.502 | 0 | 0 | 1 | 45.875 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-14367 | 2020/12/10 03:36:39 | 1-21 | 1 | -2.025 | -44.843 | 18347 | -1 | 1 | UNDE1-A-1 | 20-196 | H2O | 1607596905 | 0.351 | 0.927 | 107.339 | 0.001 | 0.012 | 2.544 | 1505.952 | -4.752 | 2.95 | 110.765 | 0.863 | 0.05 | 0 | 0 | 1 | 45.875 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-14367 | 2020/12/10 03:45:35 | 1-21 | 2 | -1.658 | -43.759 | 18442 | -1 | 1 | UNDE1-A-1 | 20-196 | H2O | 1607597442 | 0.359 | 0.755 | 104.133 | -0.001 | 0 | 2.468 | 1509.046 | -5.29 | 2.956 | 111.808 | 0.863 | 0.116 | 0 | 0 | 1 | 45.812 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-14367 | 2020/12/10 03:54:31 | 1-21 | 3 | -1.665 | -43.768 | 18428 | -1 | 1 | UNDE1-A-1 | 20-196 | H2O | 1607597977 | 0.339 | 0.806 | 107.111 | -0.002 | 0.002 | 2.537 | 1508.755 | -5.208 | 3 | 112.381 | 0.864 | 0.173 | 0 | 0 | 1 | 45.875 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-14367 | 2020/12/10 04:03:28 | 1-21 | 4 | -1.665 | -43.714 | 18432 | 0 | 1 | UNDE1-A-1 | 20-196 | H2O | 1607598514 | 0.319 | 0.847 | 105.318 | 0 | 0.004 | 2.509 | 1509.344 | -5.713 | 2.966 | 112.826 | 0.864 | 0.201 | 0 | 0 | 1 | 45.938 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-14368 | 2020/12/10 04:13:14 | 1-22 | 1 | -6.659 | -56.596 | 18427 | -1 | 1 | EV | standard | H2O | 1607599100 | 0.323 | 0.789 | 105.552 | 0 | 0.002 | 2.518 | 1306.31 | 2.616 | 2.945 | 112.263 | 0.863 | 0.281 | 0 | 0 | 1 | 45.938 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-14368 | 2020/12/10 04:22:10 | 1-22 | 2 | -7.044 | -58.063 | 18523 | -1 | 1 | EV | standard | H2O | 1607599637 | 0.284 | 0.802 | 107.082 | 0 | 0.004 | 2.541 | 1304.573 | 2.938 | 3.022 | 114.116 | 0.863 | 0.668 | 0 | 0 | 1 | 45.938 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-14368 | 2020/12/10 04:31:07 | 1-22 | 3 | -7.165 | -58.598 | 18390 | -1 | 1 | EV | standard | H2O | 1607600173 | 0.307 | 0.802 | 108.124 | 0 | 0.006 | 2.567 | 1303.365 | 1.988 | 3.008 | 114.742 | 0.862 | 0.657 | 0 | 0 | 1 | 45.938 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-14368 | 2020/12/10 04:40:02 | 1-22 | 4 | -7.184 | -58.629 | 18396 | 0 | 1 | EV | standard | H2O | 1607600709 | 0.33 | 0.758 | 106.835 | 0 | 0.004 | 2.537 | 1302.966 | 2.985 | 2.984 | 113.085 | 0.862 | 0.654 | 0 | 0 | 1 | 45.875 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-14369 | 2020/12/10 04:49:49 | 1-23 | 1 | -3.219 | -41.533 | 18335 | -1 | 1 | UNDE1-D-2 | 20-196 | H2O | 1607601296 | 0.28 | 0.706 | 103.135 | -0.001 | -0.001 | 2.46 | 1361.69 | -1.374 | 2.934 | 113.367 | 0.863 | -0.344 | 0 | 0 | 1 | 45.875 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-14369 | 2020/12/10 04:58:45 | 1-23 | 2 | -2.935 | -39.287 | 18457 | -1 | 1 | UNDE1-D-2 | 20-196 | H2O | 1607601831 | 0.326 | 0.776 | 106.378 | -0.002 | 0 | 2.541 | 1363.945 | -1.493 | 2.96 | 114.553 | 0.862 | -0.287 | 0 | 0 | 1 | 45.938 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-14369 | 2020/12/10 05:07:42 | 1-23 | 3 | -2.898 | -38.649 | 18406 | -1 | 1 | UNDE1-D-2 | 20-196 | H2O | 1607602368 | 0.302 | 0.712 | 105.463 | -0.001 | -0.003 | 2.499 | 1363.349 | -1.736 | 2.973 | 115.711 | 0.864 | 0.076 | 0 | 0 | 1 | 45.875 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-14369 | 2020/12/10 05:16:38 | 1-23 | 4 | -2.941 | -38.495 | 18412 | 0 | 1 | UNDE1-D-2 | 20-196 | H2O | 1607602904 | 0.326 | 0.858 | 107.666 | -0.001 | 0.004 | 2.555 | 1363.3 | -1.888 | 2.98 | 115.722 | 0.863 | -0.114 | 0 | 0 | 1 | 45.875 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-14370 | 2020/12/10 05:26:25 | 1-24 | 1 | -3.055 | -41.023 | 18298 | -1 | 1 | UNDE1-D-4 | 20-196 | H2O | 1607603492 | 0.345 | 0.787 | 104.485 | -0.001 | 0.002 | 2.481 | 1304.994 | 0.825 | 2.986 | 113.192 | 0.862 | 0.067 | 0 | 0 | 1 | 45.875 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-14370 | 2020/12/10 05:35:21 | 1-24 | 2 | -2.959 | -40.786 | 18488 | -1 | 1 | UNDE1-D-4 | 20-196 | H2O | 1607604028 | 0.306 | 0.754 | 106.993 | 0 | 0.001 | 2.541 | 1305.331 | 1.999 | 2.93 | 112.889 | 0.862 | -0.217 | 0 | 0 | 1 | 45.875 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-14370 | 2020/12/10 05:44:17 | 1-24 | 3 | -2.923 | -40.725 | 18440 | -1 | 1 | UNDE1-D-4 | 20-196 | H2O | 1607604563 | 0.327 | 0.771 | 105.689 | -0.001 | 0.001 | 2.506 | 1305.097 | 0.854 | 2.953 | 115.013 | 0.863 | 0.269 | 0 | 0 | 1 | 45.875 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-14370 | 2020/12/10 05:53:13 | 1-24 | 4 | -2.936 | -40.814 | 18475 | 0 | 1 | UNDE1-D-4 | 20-196 | H2O | 1607605099 | 0.302 | 0.698 | 107.378 | 0.001 | -0.001 | 2.546 | 1305.364 | 1.236 | 2.95 | 113.476 | 0.863 | -0.167 | 0 | 0 | 1 | 45.938 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-14371 | 2020/12/10 06:03:00 | 1-25 | 1 | -3.21 | -41.37 | 17980 | -1 | 1 | UNDE1-A-3 | 20-196 | H2O | 1607605686 | 0.362 | 0.882 | 105 | -0.002 | 0.004 | 2.494 | 1694.661 | -12.375 | 2.855 | 106.627 | 0.866 | -0.352 | 0 | 0 | 1 | 45.938 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-14371 | 2020/12/10 06:11:56 | 1-25 | 2 | -3.246 | -41.256 | 18225 | -1 | 1 | UNDE1-A-3 | 20-196 | H2O | 1607606222 | 0.353 | 0.994 | 110.573 | 0 | 0.002 | 2.616 | 1706.418 | -13.341 | 2.909 | 109.984 | 0.865 | -0.337 | 0 | 0 | 1 | 45.875 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-14371 | 2020/12/10 06:20:52 | 1-25 | 3 | -3.235 | -41.036 | 18530 | -1 | 1 | UNDE1-A-3 | 20-196 | H2O | 1607606758 | 0.384 | 0.936 | 107.662 | 0 | 0.002 | 2.545 | 1715.877 | -13.362 | 2.966 | 111.014 | 0.866 | -0.361 | 0 | 0 | 1 | 45.812 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-14371 | 2020/12/10 06:29:47 | 1-25 | 4 | -3.249 | -41.311 | 18542 | 0 | 1 | UNDE1-A-3 | 20-196 | H2O | 1607607293 | 0.371 | 0.895 | 105.644 | 0 | 0.004 | 2.508 | 1717.593 | -13.304 | 2.949 | 110.521 | 0.866 | -0.472 | 0 | 0 | 1 | 45.875 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-14372 | 2020/12/10 06:39:35 | 1-26 | 1 | -3.007 | -39.775 | 18322 | -1 | 1 | UNDE1-E-5 | 20-196 | H2O | 1607607881 | 0.333 | 0.822 | 107.053 | 0.001 | 0.007 | 2.543 | 1382.26 | -2.018 | 2.93 | 112.882 | 0.862 | -0.22 | 0 | 0 | 1 | 45.812 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-14372 | 2020/12/10 06:48:31 | 1-26 | 2 | -2.915 | -39.04 | 18690 | -1 | 1 | UNDE1-E-5 | 20-196 | H2O | 1607608417 | 0.346 | 0.821 | 106.188 | -0.002 | 0 | 2.532 | 1382.634 | -1.401 | 3.067 | 115.709 | 0.864 | 0.084 | 0 | 0 | 1 | 45.875 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-14372 | 2020/12/10 06:57:27 | 1-26 | 3 | -2.941 | -39.099 | 18522 | -1 | 1 | UNDE1-E-5 | 20-196 | H2O | 1607608954 | 0.299 | 0.837 | 106.008 | 0 | 0.002 | 2.526 | 1379.97 | -0.9 | 2.913 | 111.823 | 0.864 | -0.53 | 0 | 0 | 1 | 45.812 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-14372 | 2020/12/10 07:06:24 | 1-26 | 4 | -2.937 | -39.318 | 18141 | 0 | 1 | UNDE1-E-5 | 20-196 | H2O | 1607609490 | 0.325 | 0.733 | 107.621 | 0 | 0 | 2.584 | 1375.41 | -1.413 | 2.879 | 111.152 | 0.863 | -0.186 | 0 | 0 | 1 | 45.812 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-14373 | 2020/12/10 07:16:13 | 1-27 | 1 | [5 empty] | UNDE1-B-5 | 20-196 | [18 empty] | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-14373 | 2020/12/10 07:25:09 | 1-27 | 2 | [5 empty] | UNDE1-B-5 | 20-196 | [18 empty] | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-14373 | 2020/12/10 07:34:05 | 1-27 | 3 | -3.616 | -43.719 | 18273 | -1 | 1 | UNDE1-B-5 | 20-196 | H2O | 1607611152 | 0.4 | 1.038 | 107.653 | -0.001 | -0.001 | 2.542 | 1855.239 | -19.303 | 2.862 | 108.35 | 0.871 | -0.749 | 0 | 0 | 1 | 45.812 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-14373 | 2020/12/10 07:43:01 | 1-27 | 4 | -3.696 | -43.844 | 18182 | 0 | 1 | UNDE1-B-5 | 20-196 | H2O | 1607611688 | 0.365 | 0.932 | 106.477 | 0.001 | 0 | 2.53 | 1851.661 | -18.694 | 2.922 | 108.518 | 0.866 | -0.564 | 0 | 0 | 1 | 45.812 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-14374 | 2020/12/10 07:52:48 | 1-28 | 1 | [5 empty] | UNDE1-B-2 | 20-196 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-14374 | 2020/12/10 08:01:45 | 1-28 | 2 | [5 empty] | UNDE1-B-2 | 20-196 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-14374 | 2020/12/10 08:10:41 | 1-28 | 3 | -2.164 | -42.87 | 18280 | -1 | 1 | UNDE1-B-2 | 20-196 | H2O | 1607613348 | 0.423 | 1.145 | 108.459 | -0.002 | 0.003 | 2.572 | 1986.751 | -25.954 | 2.921 | 109.312 | 0.868 | -1.669 | 0 | 0 | 1 | 45.75 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-14374 | 2020/12/10 08:19:38 | 1-28 | 4 | -2.203 | -42.712 | 18367 | 0 | 1 | UNDE1-B-2 | 20-196 | H2O | 1607613884 | 0.351 | 1.149 | 103.949 | -0.001 | -0.001 | 2.463 | 1990.111 | -25.917 | 2.943 | 109.568 | 0.869 | -1.685 | 0 | 0 | 1 | 45.75 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-14375 | 2020/12/10 08:29:24 | 1-29 | 1 | [5 empty] | NIWO1-A-4 | 20-196 | [18 empty] | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-14375 | 2020/12/10 08:38:21 | 1-29 | 2 | [5 empty] | NIWO1-A-4 | 20-196 | [18 empty] | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-14375 | 2020/12/10 08:47:17 | 1-29 | 3 | 3.902 | -97.624 | 18351 | -1 | 1 | NIWO1-A-4 | 20-196 | H2O | 1607615543 | 0.361 | 1.135 | 110.146 | -0.002 | 0.006 | 2.606 | 1564.273 | -34.776 | 5.146 | 132.111 | 0.865 | -9.699 | 0 | 0 | 1 | 45.875 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-14375 | 2020/12/10 08:56:12 | 1-29 | 4 | 3.861 | -98.209 | 18403 | 0 | 1 | NIWO1-A-4 | 20-196 | H2O | 1607616079 | 0.367 | 1.035 | 106.542 | -0.002 | 0.006 | 2.535 | 1565.117 | -35.339 | 5.082 | 133.606 | 0.865 | -9.621 | 0 | 0 | 1 | 45.875 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-14376 | 2020/12/10 09:05:58 | 1-30 | 1 | [5 empty] | WOOD1-C-2 | 20-196 | [18 empty] | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-14376 | 2020/12/10 09:14:54 | 1-30 | 2 | [5 empty] | WOOD1-C-2 | 20-196 | [18 empty] | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-14376 | 2020/12/10 09:23:50 | 1-30 | 3 | -3.957 | -51.621 | 18479 | -1 | 1 | WOOD1-C-2 | 20-196 | H2O | 1607617737 | 0.327 | 0.867 | 107.061 | -0.001 | -0.001 | 2.537 | 1442.601 | -4.999 | 3.017 | 115.061 | 0.863 | -0.34 | 0 | 0 | 1 | 45.875 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-14376 | 2020/12/10 09:32:46 | 1-30 | 4 | -3.945 | -50.991 | 18456 | 0 | 1 | WOOD1-C-2 | 20-196 | H2O | 1607618273 | 0.327 | 0.872 | 108.677 | -0.001 | -0.001 | 2.59 | 1442.329 | -4.756 | 2.932 | 113.678 | 0.864 | -0.463 | 0 | 0 | 1 | 45.875 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-14377 | 2020/12/10 09:42:32 | 1-31 | 1 | -6.552 | -54.359 | 24614 | -1 | 0 | EV | standard | H2O | 1607618858 | 0.342 | 0.688 | 120.545 | 0 | 0.003 | 2.803 | 1352.645 | 1.936 | 4.193 | 152.146 | 0.874 | 0.881 | 0 | 0 | 1 | 45.812 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-14377 | 2020/12/10 09:51:28 | 1-31 | 2 | -7.096 | -58.914 | 18376 | -1 | 1 | EV | standard | H2O | 1607619395 | 0.316 | 0.733 | 109.167 | -0.001 | 0.003 | 2.595 | 1304.121 | 2.462 | 2.914 | 112.285 | 0.861 | 0.343 | 0 | 0 | 1 | 45.875 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-14377 | 2020/12/10 10:00:24 | 1-31 | 3 | -7.188 | -59.222 | 18334 | -1 | 1 | EV | standard | H2O | 1607619931 | 0.303 | 0.777 | 108.198 | -0.001 | 0.003 | 2.563 | 1303.252 | 2.459 | 2.916 | 112.004 | 0.864 | 0.365 | 0 | 0 | 1 | 45.875 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-14377 | 2020/12/10 10:09:20 | 1-31 | 4 | -7.193 | -59.132 | 18391 | 0 | 1 | EV | standard | H2O | 1607620467 | 0.317 | 0.797 | 106.64 | -0.001 | 0.001 | 2.547 | 1303.056 | 3.158 | 2.962 | 112.024 | 0.863 | 0.51 | 0 | 0 | 1 | 45.875 | 1 | 31 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1302.0763
varray: [1303.901, 1301.604, 1302.971, 1302.081, 1302.831, 1302.841, 1301.194, 1302.748, 1302.458, 1301.51, 1301.884, 1302.138, 1301.474, 1302.175, 1301.69, 1301.871, 1301.587, 1301.866, 1301.123, 1301.579]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_shift
baseline_shift_STD | 0.697660307932
varray: [1303.901, 1301.604, 1302.971, 1302.081, 1302.831, 1302.841, 1301.194, 1302.748, 1302.458, 1301.51, 1301.884, 1302.138, 1301.474, 1302.175, 1301.69, 1301.871, 1301.587, 1301.866, 1301.123, 1301.579]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.4188
varray: [0.607, 0.281, 0.263, 0.363, 0.287, 0.595, 0.476, 0.319, 0.488, 0.322, 0.425, 0.675, 0.368, 0.35, 0.364, 0.493, 0.3, 0.557, 0.54, 0.303]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.125743096494
varray: [0.607, 0.281, 0.263, 0.363, 0.287, 0.595, 0.476, 0.319, 0.488, 0.322, 0.425, 0.675, 0.368, 0.35, 0.364, 0.493, 0.3, 0.557, 0.54, 0.303]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: ch4_ppm
residuals_MEAN | 2.9449
varray: [3.039, 2.914, 2.985, 2.916, 2.955, 3.008, 2.949, 2.965, 2.994, 2.933, 2.962, 2.926, 2.883, 2.95, 2.936, 2.919, 2.889, 2.957, 2.929, 2.889]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: residuals
residuals_STD | 0.0405331573173
varray: [3.039, 2.914, 2.985, 2.916, 2.955, 3.008, 2.949, 2.965, 2.994, 2.933, 2.962, 2.926, 2.883, 2.95, 2.936, 2.919, 2.889, 2.957, 2.929, 2.889]
inst: std_standards
operation: std_standards(residuals)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: residuals
slope_shift_MEAN | 2.69545
varray: [2.489, 2.296, 2.016, 2.219, 2.687, 2.031, 1.559, 2.446, 2.089, 2.49, 2.596, 3.262, 3.013, 3.005, 3.249, 3.085, 3.8, 2.588, 3.847, 3.142]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: slope_shift
slope_shift_STD | 0.599951092305
varray: [2.489, 2.296, 2.016, 2.219, 2.687, 2.031, 1.559, 2.446, 2.089, 2.49, 2.596, 3.262, 3.013, 3.005, 3.249, 3.085, 3.8, 2.588, 3.847, 3.142]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: slope_shift
baseline_curvature_MEAN | 112.16455
varray: [114.938, 110.767, 112.947, 111.713, 112.472, 114.345, 112.818, 112.079, 112.899, 111.766, 112.561, 112.709, 110.437, 112.203, 111.741, 112.165, 110.201, 112.97, 110.843, 110.717]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.22276289714
varray: [114.938, 110.767, 112.947, 111.713, 112.472, 114.345, 112.818, 112.079, 112.899, 111.766, 112.561, 112.709, 110.437, 112.203, 111.741, 112.165, 110.201, 112.97, 110.843, 110.717]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.9449
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1302.0763
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.4188
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 2.69545
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0607997359759
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 3.00569973598
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.0810663146345
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.02596631463
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.121599471952
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.06649947195
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.29449
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.377229289481
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.796029289481
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.348830153966
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1302.42513015
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1302.77396031
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 2.09298092379
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1304.16928092
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 12.5578855428
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1314.63418554
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.99975546152
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 5.69520546152
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 3.66828869143
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 115.832838691
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -6.74772934022
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197, 3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052, -11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133, -7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224, -3.408, -3.097, -2.964, -3.05, -2.96, -2.798, -2.846, -2.893, -2.045, -2.14, -2.045, -2.03, -3.693, -3.681, -3.735, -3.789, -2.882, -2.804, -2.816, -2.756, -3.142, -3.121, -3.065, -3.1, -5.195, -5.283, -5.457, -5.322, -3.296, -3.141, -3.13, -3.108, -6.845, -7.044, -7.117, -7.163, -2.587, -2.31, -2.251, -2.166, -1.73, -1.741, -1.751, -1.74, -3.399, -3.632, -3.64, -3.564, -3.314, -3.123, -3.137, -3.097, -5.643, -5.873, -5.883, -5.933, -1.733, -1.376, -1.342, -1.338, -3.602, -3.67, -3.703, -3.711, -2.025, -1.658, -1.665, -1.665, -6.659, -7.044, -7.165, -7.184, -3.219, -2.935, -2.898, -2.941, -3.055, -2.959, -2.923, -2.936, -3.21, -3.246, -3.235, -3.249, -3.007, -2.915, -2.941, -2.937, 0.0, 0.0, -3.616, -3.696, 0.0, 0.0, -2.164, -2.203, 0.0, 0.0, 3.902, 3.861, 0.0, 0.0, -3.957, -3.945, -6.552, -7.096, -7.188, -7.193]
fit_rsq: 1.0
fit_intercept: 2.85121024198
except: True
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.80697972973
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0, 5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0, 16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0, 21746.0, 22284.0, 22820.0, 23358.0, 23949.0, 24485.0, 25022.0, 25560.0, 26149.0, 26685.0, 27222.0, 27758.0, 28346.0, 28882.0, 29419.0, 29956.0, 30544.0, 31081.0, 31617.0, 32153.0, 32740.0, 33276.0, 33813.0, 34349.0, 34937.0, 35474.0, 36012.0, 36549.0, 37137.0, 37674.0, 38210.0, 38747.0, 39336.0, 39872.0, 40408.0, 40945.0, 41535.0, 42071.0, 42607.0, 43144.0, 43730.0, 44267.0, 44804.0, 45341.0, 45928.0, 46464.0, 46999.0, 47535.0, 48122.0, 48658.0, 49194.0, 49731.0, 50319.0, 50856.0, 51392.0, 51929.0, 52517.0, 53053.0, 53589.0, 54126.0, 54715.0, 55251.0, 55787.0, 56322.0, 56911.0, 57447.0, 57983.0, 58520.0, 59106.0, 59642.0, 60179.0, 60714.0, 61301.0, 61837.0, 62374.0, 62910.0, 63497.0, 64033.0, 64569.0, 65105.0, 65692.0, 66228.0, 66764.0, 67299.0, 67887.0, 68423.0, 68959.0, 69496.0, 70085.0, 70621.0, 71157.0, 71693.0, 72280.0, 72817.0, 73353.0, 73890.0, 74476.0, 75013.0, 75549.0, 76084.0, 76670.0, 77206.0, 77742.0, 78278.0, 78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -59.359783423
actual_array: [ 18.1 -119. ]
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244, 24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098, -86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153, -68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254, -42.792, -40.613, -39.66, -39.774, -42.877, -42.477, -42.825, -43.048, -35.654, -35.265, -34.677, -34.627, -66.576, -70.406, -71.401, -71.837, -42.984, -39.216, -38.421, -37.878, -41.081, -41.353, -40.99, -41.174, -59.413, -61.221, -62.479, -62.021, -45.186, -42.834, -42.236, -41.908, -56.79, -58.149, -58.67, -58.829, -32.474, -29.227, -28.629, -28.159, -32.264, -33.123, -33.217, -32.976, -44.244, -46.055, -46.577, -46.501, -40.692, -39.452, -39.532, -39.32, -57.06, -59.223, -59.829, -60.148, -36.467, -33.594, -32.504, -32.312, -44.996, -46.035, -46.33, -46.606, -44.843, -43.759, -43.768, -43.714, -56.596, -58.063, -58.598, -58.629, -41.533, -39.287, -38.649, -38.495, -41.023, -40.786, -40.725, -40.814, -41.37, -41.256, -41.036, -41.311, -39.775, -39.04, -39.099, -39.318, 0.0, 0.0, -43.719, -43.844, 0.0, 0.0, -42.87, -42.712, 0.0, 0.0, -97.624, -98.209, 0.0, 0.0, -51.621, -50.991, -54.359, -58.914, -59.222, -59.132]
fit_rsq: 1.0
fit_intercept: 18.8422423778
except: True
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.7149594595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0, 5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0, 10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0, 16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0, 21746.0, 22284.0, 22820.0, 23358.0, 23949.0, 24485.0, 25022.0, 25560.0, 26149.0, 26685.0, 27222.0, 27758.0, 28346.0, 28882.0, 29419.0, 29956.0, 30544.0, 31081.0, 31617.0, 32153.0, 32740.0, 33276.0, 33813.0, 34349.0, 34937.0, 35474.0, 36012.0, 36549.0, 37137.0, 37674.0, 38210.0, 38747.0, 39336.0, 39872.0, 40408.0, 40945.0, 41535.0, 42071.0, 42607.0, 43144.0, 43730.0, 44267.0, 44804.0, 45341.0, 45928.0, 46464.0, 46999.0, 47535.0, 48122.0, 48658.0, 49194.0, 49731.0, 50319.0, 50856.0, 51392.0, 51929.0, 52517.0, 53053.0, 53589.0, 54126.0, 54715.0, 55251.0, 55787.0, 56322.0, 56911.0, 57447.0, 57983.0, 58520.0, 59106.0, 59642.0, 60179.0, 60714.0, 61301.0, 61837.0, 62374.0, 62910.0, 63497.0, 64033.0, 64569.0, 65105.0, 65692.0, 66228.0, 66764.0, 67299.0, 67887.0, 68423.0, 68959.0, 69496.0, 70085.0, 70621.0, 71157.0, 71693.0, 72280.0, 72817.0, 73353.0, 73890.0, 74476.0, 75013.0, 75549.0, 76084.0, 76670.0, 77206.0, 77742.0, 78278.0, 78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6897
varray: [0.0, 3.076, 3.001, 2.965, 2.985, 2.928, 3.006, 2.989, 2.965, 2.982]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: residuals
SAMP_residuals_STD | 0.945824513909
varray: [0.0, 3.076, 3.001, 2.965, 2.985, 2.928, 3.006, 2.989, 2.965, 2.982]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4358
varray: [0.0, 0.566, 0.498, 0.5, 0.382, 0.304, 0.472, 0.568, 0.44, 0.628]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.179776280724
varray: [0.0, 0.566, 0.498, 0.5, 0.382, 0.304, 0.472, 0.568, 0.44, 0.628]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1179.9159
varray: [0.0, 1342.445, 1314.397, 1310.128, 1307.85, 1305.628, 1305.895, 1304.974, 1304.092, 1303.75]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 414.741299934
varray: [0.0, 1342.445, 1314.397, 1310.128, 1307.85, 1305.628, 1305.895, 1304.974, 1304.092, 1303.75]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.9529
varray: [0.0, 1.305, 1.871, 1.658, 2.378, 2.49, 2.562, 2.241, 2.176, 2.848]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.82336942566
varray: [0.0, 1.305, 1.871, 1.658, 2.378, 2.49, 2.562, 2.241, 2.176, 2.848]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: slope_shift
SAMP_d18O_STD | 2.33937437981
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 19.3719858332
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.6916
varray: [0.0, 115.208, 114.455, 114.534, 113.842, 112.801, 114.131, 114.507, 113.703, 113.735]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 36.0879289582
varray: [0.0, 115.208, 114.455, 114.534, 113.842, 112.801, 114.131, 114.507, 113.703, 113.735]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6897
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4358
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1179.9159
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.9529
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2552
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2552
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.66582906041
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.47295
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.6286
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
dD_mean | -54.6407
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
d18O_stdev | 2.33937437981
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
dD_stdev | 19.3719858332
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.6286
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.6407
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.33937437981
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
measured_dD_stdev | 19.3719858332
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.6072947725
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.967, -7.406, -7.279, -7.282, -7.338, -7.278, -7.286, -7.253, -7.197]
fit_rsq: 1.0
fit_intercept: 2.85121024198
except: True
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -6.6286
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.7270884362
actual_array: [ 18.1 -119. ]
varray: [0.0, -67.819, -61.388, -60.115, -59.725, -59.748, -59.511, -59.483, -59.374, -59.244]
fit_rsq: 1.0
fit_intercept: 18.8422423778
except: True
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -54.6407
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [1.0, 600.0, 1137.0, 1675.0, 2211.0, 2749.0, 3286.0, 3823.0, 4361.0, 4897.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3866
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -122.1604
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9658
varray: [3.039, 2.914, 2.985, 2.916, 2.955, 3.008, 2.949, 2.965, 2.994, 2.933]
inst: mean_all
operation: mean_all(residuals)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: residuals
SAMP_residuals_STD | 0.0407452792085
varray: [3.039, 2.914, 2.985, 2.916, 2.955, 3.008, 2.949, 2.965, 2.994, 2.933]
inst: std_all
operation: std_all(residuals)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4001
varray: [0.607, 0.281, 0.263, 0.363, 0.287, 0.595, 0.476, 0.319, 0.488, 0.322]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.130854838997
varray: [0.607, 0.281, 0.263, 0.363, 0.287, 0.595, 0.476, 0.319, 0.488, 0.322]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.4139
varray: [1303.901, 1301.604, 1302.971, 1302.081, 1302.831, 1302.841, 1301.194, 1302.748, 1302.458, 1301.51]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.820821600593
varray: [1303.901, 1301.604, 1302.971, 1302.081, 1302.831, 1302.841, 1301.194, 1302.748, 1302.458, 1301.51]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.2322
varray: [2.489, 2.296, 2.016, 2.219, 2.687, 2.031, 1.559, 2.446, 2.089, 2.49]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.324915989142
varray: [2.489, 2.296, 2.016, 2.219, 2.687, 2.031, 1.559, 2.446, 2.089, 2.49]
inst: std_all
operation: std_all(slope_shift)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: slope_shift
SAMP_d18O_STD | 0.396995158101
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.04046075495
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.6744
varray: [114.938, 110.767, 112.947, 111.713, 112.472, 114.345, 112.818, 112.079, 112.899, 111.766]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.240218995
varray: [114.938, 110.767, 112.947, 111.713, 112.472, 114.345, 112.818, 112.079, 112.899, 111.766]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9658
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4001
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.4139
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.2322
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.709701517878
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.50985
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.7556
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 37.9172
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.7556
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
dD_mean | 37.9172
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
d18O_stdev | 0.396995158101
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
dD_stdev | 5.04046075495
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.7556
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
measured_dD_mean | 37.9172
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.396995158101
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.04046075495
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.699, 4.542, 4.741, 4.87, 4.899, 4.936, 4.939, 4.964, 4.914, 5.052]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.7556
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [24.56, 34.887, 37.894, 39.109, 39.792, 40.247, 40.254, 40.74, 40.591, 41.098]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 37.9172
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [5486.0, 6023.0, 6559.0, 7096.0, 7632.0, 8169.0, 8705.0, 9242.0, 9777.0, 10315.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1105
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.3376
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.924
varray: [2.962, 2.926, 2.883, 2.95, 2.936, 2.919, 2.889, 2.957, 2.929, 2.889]
inst: mean_all
operation: mean_all(residuals)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: residuals
SAMP_residuals_STD | 0.0289405520949
varray: [2.962, 2.926, 2.883, 2.95, 2.936, 2.919, 2.889, 2.957, 2.929, 2.889]
inst: std_all
operation: std_all(residuals)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4375
varray: [0.425, 0.675, 0.368, 0.35, 0.364, 0.493, 0.3, 0.557, 0.54, 0.303]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.12441619223
varray: [0.425, 0.675, 0.368, 0.35, 0.364, 0.493, 0.3, 0.557, 0.54, 0.303]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.7387
varray: [1301.884, 1302.138, 1301.474, 1302.175, 1301.69, 1301.871, 1301.587, 1301.866, 1301.123, 1301.579]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.31704540509
varray: [1301.884, 1302.138, 1301.474, 1302.175, 1301.69, 1301.871, 1301.587, 1301.866, 1301.123, 1301.579]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 3.1587
varray: [2.596, 3.262, 3.013, 3.005, 3.249, 3.085, 3.8, 2.588, 3.847, 3.142]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.421207272808
varray: [2.596, 3.262, 3.013, 3.005, 3.249, 3.085, 3.8, 2.588, 3.847, 3.142]
inst: std_all
operation: std_all(slope_shift)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: slope_shift
SAMP_d18O_STD | 0.58788475817
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 7.6410270841
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.6547
varray: [112.561, 112.709, 110.437, 112.203, 111.741, 112.165, 110.201, 112.97, 110.843, 110.717]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.02010914993
varray: [112.561, 112.709, 110.437, 112.203, 111.741, 112.165, 110.201, 112.97, 110.843, 110.717]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.924
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.7387
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.1587
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.709701517878
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.50985
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.7786
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
dD_mean | -106.5677
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
d18O_stdev | 0.58788475817
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
dD_stdev | 7.6410270841
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.7786
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
measured_dD_mean | -106.5677
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.58788475817
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
measured_dD_stdev | 7.6410270841
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-11.209, -12.443, -12.821, -12.903, -13.02, -13.019, -13.041, -13.069, -13.128, -13.133]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -12.7786
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-86.131, -102.398, -106.808, -108.37, -109.515, -109.782, -110.09, -110.35, -111.08, -111.153]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -106.5677
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [10905.0, 11442.0, 11979.0, 12516.0, 13052.0, 13590.0, 14127.0, 14663.0, 15200.0, 15738.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1523
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.3376
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9379
varray: [2.955, 2.96, 2.921, 3.008, 2.921, 2.876, 2.917, 2.956, 2.978, 2.887]
inst: mean_all
operation: mean_all(residuals)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: residuals
SAMP_residuals_STD | 0.0409076195022
varray: [2.955, 2.96, 2.921, 3.008, 2.921, 2.876, 2.917, 2.956, 2.978, 2.887]
inst: std_all
operation: std_all(residuals)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4796
varray: [0.545, 0.43, 0.401, 0.756, 0.505, 0.101, 0.462, 0.498, 0.677, 0.421]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.175228232124
varray: [0.545, 0.43, 0.401, 0.756, 0.505, 0.101, 0.462, 0.498, 0.677, 0.421]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.4116
varray: [1301.22, 1302.988, 1301.633, 1301.326, 1301.72, 1301.685, 1300.743, 1300.881, 1300.855, 1301.065]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.658824576213
varray: [1301.22, 1302.988, 1301.633, 1301.326, 1301.72, 1301.685, 1300.743, 1300.881, 1300.855, 1301.065]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.7862
varray: [2.836, 2.675, 2.506, 2.764, 2.534, 2.531, 2.889, 3.19, 2.912, 3.025]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.228292113068
varray: [2.836, 2.675, 2.506, 2.764, 2.534, 2.531, 2.889, 3.19, 2.912, 3.025]
inst: std_all
operation: std_all(slope_shift)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: slope_shift
SAMP_d18O_STD | 0.207318032662
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.64194044848
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.0998
varray: [112.043, 113.003, 112.781, 113.117, 112.178, 110.608, 112.179, 111.318, 112.204, 111.567]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.778996904858
varray: [112.043, 113.003, 112.781, 113.117, 112.178, 110.608, 112.179, 111.318, 112.204, 111.567]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9379
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4796
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.4116
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.7862
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.007
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.007
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.237699072974
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.06475
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.3449
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
dD_mean | -60.731
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
d18O_stdev | 0.207318032662
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
dD_stdev | 2.64194044848
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.3449
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.731
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.207318032662
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.64194044848
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3332275781
actual_array: [ 1.93 -15.84]
varray: [-7.915, -7.328, -7.352, -7.354, -7.194, -7.272, -7.304, -7.261, -7.245, -7.224]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.3449
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.5061015373
actual_array: [ 18.1 -119. ]
varray: [-68.019, -61.412, -60.399, -60.349, -59.595, -59.482, -59.819, -59.636, -59.345, -59.254]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -60.731
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [16327.0, 16863.0, 17399.0, 17936.0, 18473.0, 19009.0, 19547.0, 20083.0, 20620.0, 21157.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1384
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.6647
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.948
varray: [2.925, 2.974, 2.953, 2.94]
inst: mean_all
operation: mean_all(residuals)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: residuals
SAMP_residuals_STD | 0.0207685659913
varray: [2.925, 2.974, 2.953, 2.94]
inst: std_all
operation: std_all(residuals)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.2445
varray: [-0.337, -0.126, -0.343, -0.172]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.111888337194
varray: [-0.337, -0.126, -0.343, -0.172]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1745.899
varray: [1741.756, 1743.619, 1748.481, 1749.74]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.82024667615
varray: [1741.756, 1743.619, 1748.481, 1749.74]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -14.4995
varray: [-14.572, -14.187, -14.702, -14.537]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.220094676598
varray: [-14.572, -14.187, -14.702, -14.537]
inst: std_all
operation: std_all(slope_shift)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: slope_shift
SAMP_d18O_STD | 0.193501722645
varray: [-3.408, -3.097, -2.964, -3.05]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.45174937575
varray: [-42.792, -40.613, -39.66, -39.774]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.804
varray: [109.44, 109.908, 110.014, 109.854]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.251602861669
varray: [109.44, 109.908, 110.014, 109.854]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.948
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2445
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1745.899
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.4995
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0031
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0031
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.105266732317
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.36055
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.408, -3.097, -2.964, -3.05]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-42.792, -40.613, -39.66, -39.774]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.12975
varray: [-3.408, -3.097, -2.964, -3.05]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
dD_mean | -40.70975
varray: [-42.792, -40.613, -39.66, -39.774]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
d18O_stdev | 0.193501722645
varray: [-3.408, -3.097, -2.964, -3.05]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
dD_stdev | 1.45174937575
varray: [-42.792, -40.613, -39.66, -39.774]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.12975
varray: [-3.408, -3.097, -2.964, -3.05]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
measured_dD_mean | -40.70975
varray: [-42.792, -40.613, -39.66, -39.774]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.193501722645
varray: [-3.408, -3.097, -2.964, -3.05]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.45174937575
varray: [-42.792, -40.613, -39.66, -39.774]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.06139222206
actual_array: [ 1.93 -15.84]
varray: [-3.408, -3.097, -2.964, -3.05]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.12975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.5081759755
actual_array: [ 18.1 -119. ]
varray: [-42.792, -40.613, -39.66, -39.774]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.70975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [21746.0, 22284.0, 22820.0, 23358.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1283
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 443.8227
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9435
varray: [2.938, 2.896, 2.97, 2.97]
inst: mean_all
operation: mean_all(residuals)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: residuals
SAMP_residuals_STD | 0.0350761077278
varray: [2.938, 2.896, 2.97, 2.97]
inst: std_all
operation: std_all(residuals)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0795
varray: [0.029, -0.118, 0.261, 0.146]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.162194327891
varray: [0.029, -0.118, 0.261, 0.146]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1306.897
varray: [1309.239, 1307.289, 1306.545, 1304.515]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.95247330327
varray: [1309.239, 1307.289, 1306.545, 1304.515]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.29375
varray: [0.9, 1.127, 1.454, 1.694]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.350582729181
varray: [0.9, 1.127, 1.454, 1.694]
inst: std_all
operation: std_all(slope_shift)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: slope_shift
SAMP_d18O_STD | 0.069081473638
varray: [-2.96, -2.798, -2.846, -2.893]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.239586275344
varray: [-42.877, -42.477, -42.825, -43.048]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.2715
varray: [112.137, 112.776, 114.694, 113.479]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09531928374
varray: [112.137, 112.776, 114.694, 113.479]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9435
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0795
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1306.897
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.29375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0014
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0014
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0475398145947
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.10695
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-2.96, -2.798, -2.846, -2.893]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-42.877, -42.477, -42.825, -43.048]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.87425
varray: [-2.96, -2.798, -2.846, -2.893]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
dD_mean | -42.80675
varray: [-42.877, -42.477, -42.825, -43.048]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
d18O_stdev | 0.069081473638
varray: [-2.96, -2.798, -2.846, -2.893]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
dD_stdev | 0.239586275344
varray: [-42.877, -42.477, -42.825, -43.048]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.87425
varray: [-2.96, -2.798, -2.846, -2.893]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.80675
varray: [-42.877, -42.477, -42.825, -43.048]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.069081473638
varray: [-2.96, -2.798, -2.846, -2.893]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.239586275344
varray: [-42.877, -42.477, -42.825, -43.048]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.80245625692
actual_array: [ 1.93 -15.84]
varray: [-2.96, -2.798, -2.846, -2.893]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.87425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.4979942887
actual_array: [ 18.1 -119. ]
varray: [-42.877, -42.477, -42.825, -43.048]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.80675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [23949.0, 24485.0, 25022.0, 25560.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1328
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.8207
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94225
varray: [2.934, 2.927, 2.961, 2.947]
inst: mean_all
operation: mean_all(residuals)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: residuals
SAMP_residuals_STD | 0.014997221965
varray: [2.934, 2.927, 2.961, 2.947]
inst: std_all
operation: std_all(residuals)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.347
varray: [-0.571, -0.366, -0.215, -0.236]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.163586878039
varray: [-0.571, -0.366, -0.215, -0.236]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1387.694
varray: [1388.731, 1387.193, 1387.431, 1387.421]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.700016190289
varray: [1388.731, 1387.193, 1387.431, 1387.421]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.4305
varray: [-2.444, -2.45, -2.191, -2.637]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.183089231433
varray: [-2.444, -2.45, -2.191, -2.637]
inst: std_all
operation: std_all(slope_shift)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0504975246918
varray: [-2.045, -2.14, -2.045, -2.03]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.49293905979
varray: [-35.654, -35.265, -34.677, -34.627]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.2825
varray: [113.947, 114.106, 114.464, 114.613]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.308698882408
varray: [113.947, 114.106, 114.464, 114.613]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.347
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1387.694
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.4305
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00265
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00265
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0899860776257
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.11795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-2.045, -2.14, -2.045, -2.03]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-35.654, -35.265, -34.677, -34.627]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.065
varray: [-2.045, -2.14, -2.045, -2.03]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
dD_mean | -35.05575
varray: [-35.654, -35.265, -34.677, -34.627]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0504975246918
varray: [-2.045, -2.14, -2.045, -2.03]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
dD_stdev | 0.49293905979
varray: [-35.654, -35.265, -34.677, -34.627]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.065
varray: [-2.045, -2.14, -2.045, -2.03]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.05575
varray: [-35.654, -35.265, -34.677, -34.627]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0504975246918
varray: [-2.045, -2.14, -2.045, -2.03]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.49293905979
varray: [-35.654, -35.265, -34.677, -34.627]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.9823234593
actual_array: [ 1.93 -15.84]
varray: [-2.045, -2.14, -2.045, -2.03]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.065
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -51.1431627457
actual_array: [ 18.1 -119. ]
varray: [-35.654, -35.265, -34.677, -34.627]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -35.05575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [26149.0, 26685.0, 27222.0, 27758.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.13405
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 85.6177
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.93925
varray: [2.938, 2.941, 2.926, 2.952]
inst: mean_all
operation: mean_all(residuals)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: residuals
SAMP_residuals_STD | 0.0106887791632
varray: [2.938, 2.941, 2.926, 2.952]
inst: std_all
operation: std_all(residuals)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.144
varray: [-0.27, 0.014, -0.169, -0.151]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117635028797
varray: [-0.27, 0.014, -0.169, -0.151]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1466.859
varray: [1470.296, 1465.428, 1466.407, 1465.305]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.34378255533
varray: [1470.296, 1465.428, 1466.407, 1465.305]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -4.7645
varray: [-5.143, -4.684, -4.659, -4.572]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.256858586256
varray: [-5.143, -4.684, -4.659, -4.572]
inst: std_all
operation: std_all(slope_shift)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0488364617883
varray: [-3.693, -3.681, -3.735, -3.789]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.39540407169
varray: [-66.576, -70.406, -71.401, -71.837]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.61475
varray: [113.517, 113.276, 111.937, 111.729]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.911995020089
varray: [113.517, 113.276, 111.937, 111.729]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.93925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.144
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1466.859
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 4.7645
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00565
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00565
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.1918571089
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.4502
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.693, -3.681, -3.735, -3.789]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-66.576, -70.406, -71.401, -71.837]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7245
varray: [-3.693, -3.681, -3.735, -3.789]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
dD_mean | -70.055
varray: [-66.576, -70.406, -71.401, -71.837]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0488364617883
varray: [-3.693, -3.681, -3.735, -3.789]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
dD_stdev | 2.39540407169
varray: [-66.576, -70.406, -71.401, -71.837]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7245
varray: [-3.693, -3.681, -3.735, -3.789]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.055
varray: [-66.576, -70.406, -71.401, -71.837]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0488364617883
varray: [-3.693, -3.681, -3.735, -3.789]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.39540407169
varray: [-66.576, -70.406, -71.401, -71.837]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.66414042272
actual_array: [ 1.93 -15.84]
varray: [-3.693, -3.681, -3.735, -3.789]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.7245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -84.3535340371
actual_array: [ 18.1 -119. ]
varray: [-66.576, -70.406, -71.401, -71.837]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -70.055
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [28346.0, 28882.0, 29419.0, 29956.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.13705
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 164.7827
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9165
varray: [2.912, 2.959, 2.834, 2.961]
inst: mean_all
operation: mean_all(residuals)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: residuals
SAMP_residuals_STD | 0.0594782873549
varray: [2.912, 2.959, 2.834, 2.961]
inst: std_all
operation: std_all(residuals)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.2745
varray: [-0.242, -0.315, -0.328, -0.213]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0558002389481
varray: [-0.242, -0.315, -0.328, -0.213]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1351.44575
varray: [1352.52, 1351.563, 1350.37, 1351.33]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.882873480932
varray: [1352.52, 1351.563, 1350.37, 1351.33]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.87425
varray: [-1.403, -1.051, -0.418, -0.625]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.440116177844
varray: [-1.403, -1.051, -0.418, -0.625]
inst: std_all
operation: std_all(slope_shift)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0519326486904
varray: [-2.882, -2.804, -2.816, -2.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.30591895998
varray: [-42.984, -39.216, -38.421, -37.878]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.0135
varray: [114.004, 113.287, 111.621, 113.142]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.00190834577
varray: [114.004, 113.287, 111.621, 113.142]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9165
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2745
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1351.44575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.87425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0284
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0284
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.964379096064
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.84895
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-2.882, -2.804, -2.816, -2.756]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-42.984, -39.216, -38.421, -37.878]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.8145
varray: [-2.882, -2.804, -2.816, -2.756]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
dD_mean | -39.62475
varray: [-42.984, -39.216, -38.421, -37.878]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0519326486904
varray: [-2.882, -2.804, -2.816, -2.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
dD_stdev | 2.30591895998
varray: [-42.984, -39.216, -38.421, -37.878]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.8145
varray: [-2.882, -2.804, -2.816, -2.756]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.62475
varray: [-42.984, -39.216, -38.421, -37.878]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0519326486904
varray: [-2.882, -2.804, -2.816, -2.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.30591895998
varray: [-42.984, -39.216, -38.421, -37.878]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.74190273865
actual_array: [ 1.93 -15.84]
varray: [-2.882, -2.804, -2.816, -2.756]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.8145
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.4786324038
actual_array: [ 18.1 -119. ]
varray: [-42.984, -39.216, -38.421, -37.878]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.62475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [30544.0, 31081.0, 31617.0, 32153.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1598
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 49.36945
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94175
varray: [2.993, 2.919, 2.877, 2.978]
inst: mean_all
operation: mean_all(residuals)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: residuals
SAMP_residuals_STD | 0.0536990068189
varray: [2.993, 2.919, 2.877, 2.978]
inst: std_all
operation: std_all(residuals)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0685
varray: [0.09, -0.186, -0.151, -0.027]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.125773606134
varray: [0.09, -0.186, -0.151, -0.027]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1682.57475
varray: [1687.042, 1681.285, 1676.859, 1685.113]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.49934397996
varray: [1687.042, 1681.285, 1676.859, 1685.113]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -12.36225
varray: [-12.216, -12.395, -12.347, -12.491]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.114412630422
varray: [-12.216, -12.395, -12.347, -12.491]
inst: std_all
operation: std_all(slope_shift)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0328329103188
varray: [-3.142, -3.121, -3.065, -3.1]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.155075250551
varray: [-41.081, -41.353, -40.99, -41.174]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.53775
varray: [111.963, 109.716, 108.999, 111.473]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.40819680798
varray: [111.963, 109.716, 108.999, 111.473]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0685
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1682.57475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 12.36225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00315
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00315
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.106964582838
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.6268
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.142, -3.121, -3.065, -3.1]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-41.081, -41.353, -40.99, -41.174]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.107
varray: [-3.142, -3.121, -3.065, -3.1]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
dD_mean | -41.1495
varray: [-41.081, -41.353, -40.99, -41.174]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0328329103188
varray: [-3.142, -3.121, -3.065, -3.1]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
dD_stdev | 0.155075250551
varray: [-41.081, -41.353, -40.99, -41.174]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.107
varray: [-3.142, -3.121, -3.065, -3.1]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.1495
varray: [-41.081, -41.353, -40.99, -41.174]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0328329103188
varray: [-3.142, -3.121, -3.065, -3.1]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.155075250551
varray: [-41.081, -41.353, -40.99, -41.174]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.03833627996
actual_array: [ 1.93 -15.84]
varray: [-3.142, -3.121, -3.065, -3.1]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.9254495106
actual_array: [ 18.1 -119. ]
varray: [-41.081, -41.353, -40.99, -41.174]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.1495
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [32740.0, 33276.0, 33813.0, 34349.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.13455
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 380.49845
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89975
varray: [2.855, 2.92, 2.884, 2.94]
inst: mean_all
operation: mean_all(residuals)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: residuals
SAMP_residuals_STD | 0.0377745505157
varray: [2.855, 2.92, 2.884, 2.94]
inst: std_all
operation: std_all(residuals)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.20875
varray: [-0.291, -0.325, -0.08, -0.139]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117927590778
varray: [-0.291, -0.325, -0.08, -0.139]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1675.852
varray: [1672.264, 1680.741, 1671.38, 1679.023]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.7198195587
varray: [1672.264, 1680.741, 1671.38, 1679.023]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -11.85125
varray: [-11.117, -12.263, -11.906, -12.119]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.510995352229
varray: [-11.117, -12.263, -11.906, -12.119]
inst: std_all
operation: std_all(slope_shift)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: slope_shift
SAMP_d18O_STD | 0.108987384591
varray: [-5.195, -5.283, -5.457, -5.322]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.35102467286
varray: [-59.413, -61.221, -62.479, -62.021]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.2815
varray: [106.955, 110.449, 108.806, 110.916]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7956752676
varray: [106.955, 110.449, 108.806, 110.916]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.20875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1675.852
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 11.85125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.53315902068
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.88305
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-5.195, -5.283, -5.457, -5.322]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-59.413, -61.221, -62.479, -62.021]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.31425
varray: [-5.195, -5.283, -5.457, -5.322]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
dD_mean | -61.2835
varray: [-59.413, -61.221, -62.479, -62.021]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
d18O_stdev | 0.108987384591
varray: [-5.195, -5.283, -5.457, -5.322]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
dD_stdev | 1.35102467286
varray: [-59.413, -61.221, -62.479, -62.021]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.31425
varray: [-5.195, -5.283, -5.457, -5.322]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.2835
varray: [-59.413, -61.221, -62.479, -62.021]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.108987384591
varray: [-5.195, -5.283, -5.457, -5.322]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.35102467286
varray: [-59.413, -61.221, -62.479, -62.021]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.27526938782
actual_array: [ 1.93 -15.84]
varray: [-5.195, -5.283, -5.457, -5.322]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.31425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.0303622039
actual_array: [ 18.1 -119. ]
varray: [-59.413, -61.221, -62.479, -62.021]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -61.2835
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [34937.0, 35474.0, 36012.0, 36549.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.17655
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 373.7757
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91725
varray: [2.917, 2.94, 2.939, 2.873]
inst: mean_all
operation: mean_all(residuals)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: residuals
SAMP_residuals_STD | 0.0313515018247
varray: [2.917, 2.94, 2.939, 2.873]
inst: std_all
operation: std_all(residuals)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.581
varray: [-0.196, -0.641, -0.631, -0.856]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.276857363998
varray: [-0.196, -0.641, -0.631, -0.856]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1921.1075
varray: [1920.384, 1922.08, 1922.577, 1919.389]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.48120232694
varray: [1920.384, 1922.08, 1922.577, 1919.389]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -21.24425
varray: [-21.062, -21.44, -21.591, -20.884]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.327388021569
varray: [-21.062, -21.44, -21.591, -20.884]
inst: std_all
operation: std_all(slope_shift)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0859355378564
varray: [-3.296, -3.141, -3.13, -3.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.48049406168
varray: [-45.186, -42.834, -42.236, -41.908]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.2845
varray: [108.718, 107.619, 107.821, 104.98]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.60886761419
varray: [108.718, 107.619, 107.821, 104.98]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.581
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1921.1075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 21.24425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02765
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02765
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.938911338246
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.88005
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.296, -3.141, -3.13, -3.108]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-45.186, -42.834, -42.236, -41.908]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.16875
varray: [-3.296, -3.141, -3.13, -3.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
dD_mean | -43.041
varray: [-45.186, -42.834, -42.236, -41.908]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0859355378564
varray: [-3.296, -3.141, -3.13, -3.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
dD_stdev | 1.48049406168
varray: [-45.186, -42.834, -42.236, -41.908]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.16875
varray: [-3.296, -3.141, -3.13, -3.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.041
varray: [-45.186, -42.834, -42.236, -41.908]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0859355378564
varray: [-3.296, -3.141, -3.13, -3.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.48049406168
varray: [-45.186, -42.834, -42.236, -41.908]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.10091669423
actual_array: [ 1.93 -15.84]
varray: [-3.296, -3.141, -3.13, -3.108]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.16875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.7202713225
actual_array: [ 18.1 -119. ]
varray: [-45.186, -42.834, -42.236, -41.908]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.041
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [37137.0, 37674.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.15905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 619.0312
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8825
varray: [2.848, 2.927, 2.876, 2.879]
inst: mean_all
operation: mean_all(residuals)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: residuals
SAMP_residuals_STD | 0.0327871926215
varray: [2.848, 2.927, 2.876, 2.879]
inst: std_all
operation: std_all(residuals)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.35775
varray: [0.374, 0.383, 0.32, 0.354]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0279329554469
varray: [0.374, 0.383, 0.32, 0.354]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1305.699
varray: [1309.732, 1305.811, 1304.138, 1303.115]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.90926164745
varray: [1309.732, 1305.811, 1304.138, 1303.115]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.32475
varray: [2.108, 2.299, 2.154, 2.738]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.287271503402
varray: [2.108, 2.299, 2.154, 2.738]
inst: std_all
operation: std_all(slope_shift)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: slope_shift
SAMP_d18O_STD | 0.140331452877
varray: [-6.845, -7.044, -7.117, -7.163]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.926369436744
varray: [-56.79, -58.149, -58.67, -58.829]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.226
varray: [110.575, 111.895, 111.439, 110.995]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.568651035346
varray: [110.575, 111.895, 111.439, 110.995]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.35775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1305.699
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.32475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0624
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0624
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.11891745051
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.93855
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-6.845, -7.044, -7.117, -7.163]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-56.79, -58.149, -58.67, -58.829]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.04225
varray: [-6.845, -7.044, -7.117, -7.163]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
dD_mean | -58.1095
varray: [-56.79, -58.149, -58.67, -58.829]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
d18O_stdev | 0.140331452877
varray: [-6.845, -7.044, -7.117, -7.163]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
dD_stdev | 0.926369436744
varray: [-56.79, -58.149, -58.67, -58.829]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.04225
varray: [-6.845, -7.044, -7.117, -7.163]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.1095
varray: [-56.79, -58.149, -58.67, -58.829]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.140331452877
varray: [-6.845, -7.044, -7.117, -7.163]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.926369436744
varray: [-56.79, -58.149, -58.67, -58.829]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0265075396
actual_array: [ 1.93 -15.84]
varray: [-6.845, -7.044, -7.117, -7.163]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.04225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.0185914237
actual_array: [ 18.1 -119. ]
varray: [-56.79, -58.149, -58.67, -58.829]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.1095
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [39336.0, 39872.0, 40408.0, 40945.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1938
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.6227
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.995
varray: [3.04, 3.005, 2.972, 2.963]
inst: mean_all
operation: mean_all(residuals)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: residuals
SAMP_residuals_STD | 0.0350142828
varray: [3.04, 3.005, 2.972, 2.963]
inst: std_all
operation: std_all(residuals)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.04025
varray: [0.077, -0.113, -0.144, 0.019]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105379868413
varray: [0.077, -0.113, -0.144, 0.019]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1362.8865
varray: [1364.514, 1362.881, 1361.776, 1362.375]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.17524876799
varray: [1364.514, 1362.881, 1361.776, 1362.375]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.94475
varray: [-0.854, -0.784, -0.835, -1.306]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.242640165677
varray: [-0.854, -0.784, -0.835, -1.306]
inst: std_all
operation: std_all(slope_shift)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: slope_shift
SAMP_d18O_STD | 0.182187632218
varray: [-2.587, -2.31, -2.251, -2.166]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.95075598594
varray: [-32.474, -29.227, -28.629, -28.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.17975
varray: [116.055, 113.498, 113.081, 114.085]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.31625412313
varray: [116.055, 113.498, 113.081, 114.085]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.995
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.04025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1362.8865
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.94475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0501
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0501
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.70124622228
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.0152
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-2.587, -2.31, -2.251, -2.166]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-32.474, -29.227, -28.629, -28.159]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.3285
varray: [-2.587, -2.31, -2.251, -2.166]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
dD_mean | -29.62225
varray: [-32.474, -29.227, -28.629, -28.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
d18O_stdev | 0.182187632218
varray: [-2.587, -2.31, -2.251, -2.166]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
dD_stdev | 1.95075598594
varray: [-32.474, -29.227, -28.629, -28.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.3285
varray: [-2.587, -2.31, -2.251, -2.166]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
measured_dD_mean | -29.62225
varray: [-32.474, -29.227, -28.629, -28.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.182187632218
varray: [-2.587, -2.31, -2.251, -2.166]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.95075598594
varray: [-32.474, -29.227, -28.629, -28.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.24936700847
actual_array: [ 1.93 -15.84]
varray: [-2.587, -2.31, -2.251, -2.166]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.3285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -45.9873793386
actual_array: [ 18.1 -119. ]
varray: [-32.474, -29.227, -28.629, -28.159]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -29.62225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [41535.0, 42071.0, 42607.0, 43144.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0813
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 60.8102
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86275
varray: [2.848, 2.846, 2.849, 2.908]
inst: mean_all
operation: mean_all(residuals)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: residuals
SAMP_residuals_STD | 0.0301924383469
varray: [2.848, 2.846, 2.849, 2.908]
inst: std_all
operation: std_all(residuals)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.25425
varray: [-1.181, -1.349, -1.343, -1.144]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107043215572
varray: [-1.181, -1.349, -1.343, -1.144]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2059.39575
varray: [2052.994, 2058.633, 2057.439, 2068.517]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.54701618424
varray: [2052.994, 2058.633, 2057.439, 2068.517]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -27.4905
varray: [-26.928, -27.563, -27.693, -27.778]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.385281282528
varray: [-26.928, -27.563, -27.693, -27.778]
inst: std_all
operation: std_all(slope_shift)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00858292879306
varray: [-1.73, -1.741, -1.751, -1.74]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.432199799476
varray: [-32.264, -33.123, -33.217, -32.976]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.6345
varray: [105.509, 104.533, 105.65, 106.846]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.948160499775
varray: [105.509, 104.533, 105.65, 106.846]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.25425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2059.39575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.4905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08215
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08215
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.7895684064
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.53005
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-1.73, -1.741, -1.751, -1.74]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-32.264, -33.123, -33.217, -32.976]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.7405
varray: [-1.73, -1.741, -1.751, -1.74]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
dD_mean | -32.895
varray: [-32.264, -33.123, -33.217, -32.976]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00858292879306
varray: [-1.73, -1.741, -1.751, -1.74]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
dD_stdev | 0.432199799476
varray: [-32.264, -33.123, -33.217, -32.976]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.7405
varray: [-1.73, -1.741, -1.751, -1.74]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.895
varray: [-32.264, -33.123, -33.217, -32.976]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00858292879306
varray: [-1.73, -1.741, -1.751, -1.74]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.432199799476
varray: [-32.264, -33.123, -33.217, -32.976]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.65345958185
actual_array: [ 1.93 -15.84]
varray: [-1.73, -1.741, -1.751, -1.74]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.7405
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.0928528171
actual_array: [ 18.1 -119. ]
varray: [-32.264, -33.123, -33.217, -32.976]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -32.895
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [43730.0, 44267.0, 44804.0, 45341.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.21355
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 757.31945
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9095
varray: [2.879, 2.906, 2.925, 2.928]
inst: mean_all
operation: mean_all(residuals)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: residuals
SAMP_residuals_STD | 0.0225462487641
varray: [2.879, 2.906, 2.925, 2.928]
inst: std_all
operation: std_all(residuals)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.69025
varray: [-0.793, -0.611, -0.757, -0.6]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0990601669021
varray: [-0.793, -0.611, -0.757, -0.6]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1718.04975
varray: [1721.44, 1717.725, 1718.462, 1714.572]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.82039907046
varray: [1721.44, 1717.725, 1718.462, 1714.572]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -14.57325
varray: [-14.706, -14.566, -14.68, -14.341]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.166343770548
varray: [-14.706, -14.566, -14.68, -14.341]
inst: std_all
operation: std_all(slope_shift)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111825384715
varray: [-3.399, -3.632, -3.64, -3.564]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.09139952202
varray: [-44.244, -46.055, -46.577, -46.501]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.943
varray: [109.819, 109.723, 110.251, 109.979]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.230893915035
varray: [109.819, 109.723, 110.251, 109.979]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9095
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.69025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1718.04975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.57325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0354
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0354
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.20207816904
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.22155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.399, -3.632, -3.64, -3.564]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-44.244, -46.055, -46.577, -46.501]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.55875
varray: [-3.399, -3.632, -3.64, -3.564]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
dD_mean | -45.84425
varray: [-44.244, -46.055, -46.577, -46.501]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111825384715
varray: [-3.399, -3.632, -3.64, -3.564]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
dD_stdev | 1.09139952202
varray: [-44.244, -46.055, -46.577, -46.501]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.55875
varray: [-3.399, -3.632, -3.64, -3.564]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.84425
varray: [-44.244, -46.055, -46.577, -46.501]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111825384715
varray: [-3.399, -3.632, -3.64, -3.564]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.09139952202
varray: [-44.244, -46.055, -46.577, -46.501]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.49616141598
actual_array: [ 1.93 -15.84]
varray: [-3.399, -3.632, -3.64, -3.564]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.55875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.3802418453
actual_array: [ 18.1 -119. ]
varray: [-44.244, -46.055, -46.577, -46.501]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -45.84425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [45928.0, 46464.0, 46999.0, 47535.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1668
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 415.97345
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97175
varray: [3.017, 2.975, 2.931, 2.964]
inst: mean_all
operation: mean_all(residuals)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: residuals
SAMP_residuals_STD | 0.0354906090865
varray: [3.017, 2.975, 2.931, 2.964]
inst: std_all
operation: std_all(residuals)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.2955
varray: [-0.055, -0.415, -0.42, -0.292]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.170912258191
varray: [-0.055, -0.415, -0.42, -0.292]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1438.6595
varray: [1444.604, 1437.953, 1435.762, 1436.319]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.07061293501
varray: [1444.604, 1437.953, 1435.762, 1436.319]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -4.02925
varray: [-4.043, -4.218, -3.883, -3.973]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.141855266616
varray: [-4.043, -4.218, -3.883, -3.973]
inst: std_all
operation: std_all(slope_shift)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0988985170094
varray: [-3.314, -3.123, -3.137, -3.097]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.634714634042
varray: [-40.692, -39.452, -39.532, -39.32]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.98075
varray: [116.201, 113.599, 112.344, 113.779]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.6119239798
varray: [116.201, 113.599, 112.344, 113.779]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2955
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1438.6595
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 4.02925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.911745729906
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8162
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.314, -3.123, -3.137, -3.097]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-40.692, -39.452, -39.532, -39.32]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.16775
varray: [-3.314, -3.123, -3.137, -3.097]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
dD_mean | -39.749
varray: [-40.692, -39.452, -39.532, -39.32]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0988985170094
varray: [-3.314, -3.123, -3.137, -3.097]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
dD_stdev | 0.634714634042
varray: [-40.692, -39.452, -39.532, -39.32]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.16775
varray: [-3.314, -3.123, -3.137, -3.097]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.749
varray: [-40.692, -39.452, -39.532, -39.32]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0988985170094
varray: [-3.314, -3.123, -3.137, -3.097]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.634714634042
varray: [-40.692, -39.452, -39.532, -39.32]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.09990324623
actual_array: [ 1.93 -15.84]
varray: [-3.314, -3.123, -3.137, -3.097]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.16775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.5965317483
actual_array: [ 18.1 -119. ]
varray: [-40.692, -39.452, -39.532, -39.32]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [48122.0, 48658.0, 49194.0, 49731.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.10455
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 136.5832
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89675
varray: [2.951, 2.893, 2.897, 2.846]
inst: mean_all
operation: mean_all(residuals)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: residuals
SAMP_residuals_STD | 0.0429447319237
varray: [2.951, 2.893, 2.897, 2.846]
inst: std_all
operation: std_all(residuals)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0695
varray: [0.014, -0.319, 0.061, -0.034]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.170795198996
varray: [0.014, -0.319, 0.061, -0.034]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1709.93325
varray: [1708.466, 1711.317, 1709.999, 1709.951]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.16507864541
varray: [1708.466, 1711.317, 1709.999, 1709.951]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -12.44825
varray: [-12.323, -12.785, -12.139, -12.546]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.279451098883
varray: [-12.323, -12.785, -12.139, -12.546]
inst: std_all
operation: std_all(slope_shift)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: slope_shift
SAMP_d18O_STD | 0.12935738608
varray: [-5.643, -5.873, -5.883, -5.933]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.39063223032
varray: [-57.06, -59.223, -59.829, -60.148]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.81525
varray: [108.515, 107.349, 107.658, 107.739]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.495842969094
varray: [108.515, 107.349, 107.658, 107.739]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0695
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1709.93325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 12.44825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04815
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04815
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.63503005195
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.3493
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-5.643, -5.873, -5.883, -5.933]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-57.06, -59.223, -59.829, -60.148]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.833
varray: [-5.643, -5.873, -5.883, -5.933]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
dD_mean | -59.065
varray: [-57.06, -59.223, -59.829, -60.148]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
d18O_stdev | 0.12935738608
varray: [-5.643, -5.873, -5.883, -5.933]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
dD_stdev | 1.39063223032
varray: [-57.06, -59.223, -59.829, -60.148]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.833
varray: [-5.643, -5.873, -5.883, -5.933]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.065
varray: [-57.06, -59.223, -59.829, -60.148]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.12935738608
varray: [-5.643, -5.873, -5.883, -5.933]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.39063223032
varray: [-57.06, -59.223, -59.829, -60.148]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.80099554014
actual_array: [ 1.93 -15.84]
varray: [-5.643, -5.873, -5.883, -5.933]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.833
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.9252539885
actual_array: [ 18.1 -119. ]
varray: [-57.06, -59.223, -59.829, -60.148]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.065
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [50319.0, 50856.0, 51392.0, 51929.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.17955
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 407.85695
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98225
varray: [2.976, 2.96, 3.02, 2.973]
inst: mean_all
operation: mean_all(residuals)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: residuals
SAMP_residuals_STD | 0.0261071509999
varray: [2.976, 2.96, 3.02, 2.973]
inst: std_all
operation: std_all(residuals)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8075
varray: [-0.971, -0.628, -0.672, -0.959]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.182815936577
varray: [-0.971, -0.628, -0.672, -0.959]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1313.399
varray: [1317.527, 1311.189, 1312.549, 1312.331]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.81588399382
varray: [1317.527, 1311.189, 1312.549, 1312.331]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.1805
varray: [-1.524, -1.134, -0.746, -1.318]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.330580802024
varray: [-1.524, -1.134, -0.746, -1.318]
inst: std_all
operation: std_all(slope_shift)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: slope_shift
SAMP_d18O_STD | 0.191261383104
varray: [-1.733, -1.376, -1.342, -1.338]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.91685530231
varray: [-36.467, -33.594, -32.504, -32.312]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 116.8115
varray: [117.375, 115.666, 117.121, 117.084]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.774542660758
varray: [117.375, 115.666, 117.121, 117.084]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1313.399
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.1805
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.03735
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03735
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.26829433937
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.64695
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-1.733, -1.376, -1.342, -1.338]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-36.467, -33.594, -32.504, -32.312]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.44725
varray: [-1.733, -1.376, -1.342, -1.338]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
dD_mean | -33.71925
varray: [-36.467, -33.594, -32.504, -32.312]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
d18O_stdev | 0.191261383104
varray: [-1.733, -1.376, -1.342, -1.338]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
dD_stdev | 1.91685530231
varray: [-36.467, -33.594, -32.504, -32.312]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.44725
varray: [-1.733, -1.376, -1.342, -1.338]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
measured_dD_mean | -33.71925
varray: [-36.467, -33.594, -32.504, -32.312]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.191261383104
varray: [-1.733, -1.376, -1.342, -1.338]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.91685530231
varray: [-36.467, -33.594, -32.504, -32.312]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.35626595453
actual_array: [ 1.93 -15.84]
varray: [-1.733, -1.376, -1.342, -1.338]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.44725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.8749738208
actual_array: [ 18.1 -119. ]
varray: [-36.467, -33.594, -32.504, -32.312]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -33.71925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [52517.0, 53053.0, 53589.0, 54126.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.09405
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 11.3227
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96825
varray: [2.991, 2.991, 2.989, 2.902]
inst: mean_all
operation: mean_all(residuals)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: residuals
SAMP_residuals_STD | 0.0441767284136
varray: [2.991, 2.991, 2.989, 2.902]
inst: std_all
operation: std_all(residuals)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.29825
varray: [-0.221, -0.297, -0.173, -0.502]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.145110015735
varray: [-0.221, -0.297, -0.173, -0.502]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1407.996
varray: [1407.967, 1408.088, 1407.788, 1408.141]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.156624817106
varray: [1407.967, 1408.088, 1407.788, 1408.141]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -3.0155
varray: [-2.93, -3.141, -3.361, -2.63]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.311470169786
varray: [-2.93, -3.141, -3.361, -2.63]
inst: std_all
operation: std_all(slope_shift)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0496151858473
varray: [-3.602, -3.67, -3.703, -3.711]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.70358717773
varray: [-44.996, -46.035, -46.33, -46.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.713
varray: [114.915, 114.333, 114.163, 111.441]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.54851412651
varray: [114.915, 114.333, 114.163, 111.441]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1407.996
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.0155
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02335
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02335
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.792896193419
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.54845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.602, -3.67, -3.703, -3.711]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-44.996, -46.035, -46.33, -46.606]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.6715
varray: [-3.602, -3.67, -3.703, -3.711]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
dD_mean | -45.99175
varray: [-44.996, -46.035, -46.33, -46.606]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0496151858473
varray: [-3.602, -3.67, -3.703, -3.711]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
dD_stdev | 0.70358717773
varray: [-44.996, -46.035, -46.33, -46.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.6715
varray: [-3.602, -3.67, -3.703, -3.711]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.99175
varray: [-44.996, -46.035, -46.33, -46.606]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0496151858473
varray: [-3.602, -3.67, -3.703, -3.711]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.70358717773
varray: [-44.996, -46.035, -46.33, -46.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.61042767848
actual_array: [ 1.93 -15.84]
varray: [-3.602, -3.67, -3.703, -3.711]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.6715
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.5202028378
actual_array: [ 18.1 -119. ]
varray: [-44.996, -46.035, -46.33, -46.606]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -45.99175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [54715.0, 55251.0, 55787.0, 56322.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.10805
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 105.9197
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.968
varray: [2.95, 2.956, 3.0, 2.966]
inst: mean_all
operation: mean_all(residuals)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: residuals
SAMP_residuals_STD | 0.0223308456326
varray: [2.95, 2.956, 3.0, 2.966]
inst: std_all
operation: std_all(residuals)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.135
varray: [0.05, 0.116, 0.173, 0.201]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0667982035687
varray: [0.05, 0.116, 0.173, 0.201]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1508.27425
varray: [1505.952, 1509.046, 1508.755, 1509.344]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.56673000758
varray: [1505.952, 1509.046, 1508.755, 1509.344]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -5.24075
varray: [-4.752, -5.29, -5.208, -5.713]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.39386746925
varray: [-4.752, -5.29, -5.208, -5.713]
inst: std_all
operation: std_all(slope_shift)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: slope_shift
SAMP_d18O_STD | 0.181196716306
varray: [-2.025, -1.658, -1.665, -1.665]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.548508887804
varray: [-44.843, -43.759, -43.768, -43.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.945
varray: [110.765, 111.808, 112.381, 112.826]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.89021083645
varray: [110.765, 111.808, 112.381, 112.826]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.968
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1508.27425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 5.24075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0231
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0231
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.784406940813
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.21955
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-2.025, -1.658, -1.665, -1.665]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-44.843, -43.759, -43.768, -43.714]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.75325
varray: [-2.025, -1.658, -1.665, -1.665]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
dD_mean | -44.021
varray: [-44.843, -43.759, -43.768, -43.714]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
d18O_stdev | 0.181196716306
varray: [-2.025, -1.658, -1.665, -1.665]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
dD_stdev | 0.548508887804
varray: [-44.843, -43.759, -43.768, -43.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.75325
varray: [-2.025, -1.658, -1.665, -1.665]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.021
varray: [-44.843, -43.759, -43.768, -43.714]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.181196716306
varray: [-2.025, -1.658, -1.665, -1.665]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.548508887804
varray: [-44.843, -43.759, -43.768, -43.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.6663810439
actual_array: [ 1.93 -15.84]
varray: [-2.025, -1.658, -1.665, -1.665]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.75325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.6501816453
actual_array: [ 18.1 -119. ]
varray: [-44.843, -43.759, -43.768, -43.714]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -44.021
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [56911.0, 57447.0, 57983.0, 58520.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1083
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 206.19795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98975
varray: [2.945, 3.022, 3.008, 2.984]
inst: mean_all
operation: mean_all(residuals)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: residuals
SAMP_residuals_STD | 0.0337083075814
varray: [2.945, 3.022, 3.008, 2.984]
inst: std_all
operation: std_all(residuals)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.565
varray: [0.281, 0.668, 0.657, 0.654]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.189428966458
varray: [0.281, 0.668, 0.657, 0.654]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.3035
varray: [1306.31, 1304.573, 1303.365, 1302.966]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.50203872564
varray: [1306.31, 1304.573, 1303.365, 1302.966]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.63175
varray: [2.616, 2.938, 1.988, 2.985]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.459433255943
varray: [2.616, 2.938, 1.988, 2.985]
inst: std_all
operation: std_all(slope_shift)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: slope_shift
SAMP_d18O_STD | 0.244009562654
varray: [-6.659, -7.044, -7.165, -7.184]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.953096882099
varray: [-56.596, -58.063, -58.598, -58.629]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.5515
varray: [112.263, 114.116, 114.742, 113.085]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09754346915
varray: [112.263, 114.116, 114.742, 113.085]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.3035
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.63175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.52297191755
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.38695
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-6.659, -7.044, -7.165, -7.184]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-56.596, -58.063, -58.598, -58.629]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.013
varray: [-6.659, -7.044, -7.165, -7.184]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
dD_mean | -57.9715
varray: [-56.596, -58.063, -58.598, -58.629]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
d18O_stdev | 0.244009562654
varray: [-6.659, -7.044, -7.165, -7.184]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
dD_stdev | 0.953096882099
varray: [-56.596, -58.063, -58.598, -58.629]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.013
varray: [-6.659, -7.044, -7.165, -7.184]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.9715
varray: [-56.596, -58.063, -58.598, -58.629]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.244009562654
varray: [-6.659, -7.044, -7.165, -7.184]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.953096882099
varray: [-56.596, -58.063, -58.598, -58.629]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.99686418542
actual_array: [ 1.93 -15.84]
varray: [-6.659, -7.044, -7.165, -7.184]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.013
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.8876448681
actual_array: [ 18.1 -119. ]
varray: [-56.596, -58.063, -58.598, -58.629]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -57.9715
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [59106.0, 59642.0, 60179.0, 60714.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.08655
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.2272
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96175
varray: [2.934, 2.96, 2.973, 2.98]
inst: mean_all
operation: mean_all(residuals)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: residuals
SAMP_residuals_STD | 0.0202710795634
varray: [2.934, 2.96, 2.973, 2.98]
inst: std_all
operation: std_all(residuals)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.16725
varray: [-0.344, -0.287, 0.076, -0.114]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.189373308573
varray: [-0.344, -0.287, 0.076, -0.114]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1363.071
varray: [1361.69, 1363.945, 1363.349, 1363.3]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.966223231626
varray: [1361.69, 1363.945, 1363.349, 1363.3]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.62275
varray: [-1.374, -1.493, -1.736, -1.888]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.232303501193
varray: [-1.374, -1.493, -1.736, -1.888]
inst: std_all
operation: std_all(slope_shift)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: slope_shift
SAMP_d18O_STD | 0.148389970461
varray: [-3.219, -2.935, -2.898, -2.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.40384709519
varray: [-41.533, -39.287, -38.649, -38.495]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.83825
varray: [113.367, 114.553, 115.711, 115.722]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.12378093209
varray: [113.367, 114.553, 115.711, 115.722]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.16725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1363.071
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.62275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.572175625658
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.6737
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.219, -2.935, -2.898, -2.941]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-41.533, -39.287, -38.649, -38.495]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.99825
varray: [-3.219, -2.935, -2.898, -2.941]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
dD_mean | -39.491
varray: [-41.533, -39.287, -38.649, -38.495]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
d18O_stdev | 0.148389970461
varray: [-3.219, -2.935, -2.898, -2.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
dD_stdev | 1.40384709519
varray: [-41.533, -39.287, -38.649, -38.495]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.99825
varray: [-3.219, -2.935, -2.898, -2.941]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.491
varray: [-41.533, -39.287, -38.649, -38.495]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.148389970461
varray: [-3.219, -2.935, -2.898, -2.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.40384709519
varray: [-41.533, -39.287, -38.649, -38.495]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.92812380947
actual_array: [ 1.93 -15.84]
varray: [-3.219, -2.935, -2.898, -2.941]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.99825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.3517186225
actual_array: [ 18.1 -119. ]
varray: [-41.533, -39.287, -38.649, -38.495]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.491
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [61301.0, 61837.0, 62374.0, 62910.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.11455
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 60.9947
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95475
varray: [2.986, 2.93, 2.953, 2.95]
inst: mean_all
operation: mean_all(residuals)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: residuals
SAMP_residuals_STD | 0.0232002155162
varray: [2.986, 2.93, 2.953, 2.95]
inst: std_all
operation: std_all(residuals)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.012
varray: [0.067, -0.217, 0.269, -0.167]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.224538044289
varray: [0.067, -0.217, 0.269, -0.167]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1305.1965
varray: [1304.994, 1305.331, 1305.097, 1305.364]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.179863837388
varray: [1304.994, 1305.331, 1305.097, 1305.364]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.2285
varray: [0.825, 1.999, 0.854, 1.236]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.546744608265
varray: [0.825, 1.999, 0.854, 1.236]
inst: std_all
operation: std_all(slope_shift)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0597180877122
varray: [-3.055, -2.959, -2.923, -2.936]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.129447544074
varray: [-41.023, -40.786, -40.725, -40.814]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.6425
varray: [113.192, 112.889, 115.013, 113.476]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.944581918099
varray: [113.192, 112.889, 115.013, 113.476]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.012
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1305.1965
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.2285
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.334476552684
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.47795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.055, -2.959, -2.923, -2.936]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-41.023, -40.786, -40.725, -40.814]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.96825
varray: [-3.055, -2.959, -2.923, -2.936]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
dD_mean | -40.837
varray: [-41.023, -40.786, -40.725, -40.814]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0597180877122
varray: [-3.055, -2.959, -2.923, -2.936]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
dD_stdev | 0.129447544074
varray: [-41.023, -40.786, -40.725, -40.814]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.96825
varray: [-3.055, -2.959, -2.923, -2.936]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
measured_dD_mean | -40.837
varray: [-41.023, -40.786, -40.725, -40.814]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0597180877122
varray: [-3.055, -2.959, -2.923, -2.936]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.129447544074
varray: [-41.023, -40.786, -40.725, -40.814]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.89772036934
actual_array: [ 1.93 -15.84]
varray: [-3.055, -2.959, -2.923, -2.936]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.96825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.6289219842
actual_array: [ 18.1 -119. ]
varray: [-41.023, -40.786, -40.725, -40.814]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.837
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [63497.0, 64033.0, 64569.0, 65105.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.12155
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.1202
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91975
varray: [2.855, 2.909, 2.966, 2.949]
inst: mean_all
operation: mean_all(residuals)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: residuals
SAMP_residuals_STD | 0.0493381191372
varray: [2.855, 2.909, 2.966, 2.949]
inst: std_all
operation: std_all(residuals)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.3805
varray: [-0.352, -0.337, -0.361, -0.472]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0617980582219
varray: [-0.352, -0.337, -0.361, -0.472]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1708.63725
varray: [1694.661, 1706.418, 1715.877, 1717.593]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 10.5337564169
varray: [1694.661, 1706.418, 1715.877, 1717.593]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -13.0955
varray: [-12.375, -13.341, -13.362, -13.304]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.480931388038
varray: [-12.375, -13.341, -13.362, -13.304]
inst: std_all
operation: std_all(slope_shift)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0177200451467
varray: [-3.21, -3.246, -3.235, -3.249]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.145797519412
varray: [-41.37, -41.256, -41.036, -41.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.5365
varray: [106.627, 109.984, 111.014, 110.521]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.98474960637
varray: [106.627, 109.984, 111.014, 110.521]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3805
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1708.63725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.0955
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.854018812184
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.62805
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.21, -3.246, -3.235, -3.249]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-41.37, -41.256, -41.036, -41.311]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.235
varray: [-3.21, -3.246, -3.235, -3.249]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
dD_mean | -41.24325
varray: [-41.37, -41.256, -41.036, -41.311]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0177200451467
varray: [-3.21, -3.246, -3.235, -3.249]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
dD_stdev | 0.145797519412
varray: [-41.37, -41.256, -41.036, -41.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.235
varray: [-3.21, -3.246, -3.235, -3.249]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.24325
varray: [-41.37, -41.256, -41.036, -41.311]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0177200451467
varray: [-3.21, -3.246, -3.235, -3.249]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.145797519412
varray: [-41.37, -41.256, -41.036, -41.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.16805762453
actual_array: [ 1.93 -15.84]
varray: [-3.21, -3.246, -3.235, -3.249]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.0144077686
actual_array: [ 18.1 -119. ]
varray: [-41.37, -41.256, -41.036, -41.311]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.24325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [65692.0, 66228.0, 66764.0, 67299.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.15655
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 406.56095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94725
varray: [2.93, 3.067, 2.913, 2.879]
inst: mean_all
operation: mean_all(residuals)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: residuals
SAMP_residuals_STD | 0.08260094834
varray: [2.93, 3.067, 2.913, 2.879]
inst: std_all
operation: std_all(residuals)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.213
varray: [-0.22, 0.084, -0.53, -0.186]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.251313880768
varray: [-0.22, 0.084, -0.53, -0.186]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1380.0685
varray: [1382.26, 1382.634, 1379.97, 1375.41]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.32143478033
varray: [1382.26, 1382.634, 1379.97, 1375.41]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.433
varray: [-2.018, -1.401, -0.9, -1.413]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.457434148266
varray: [-2.018, -1.401, -0.9, -1.413]
inst: std_all
operation: std_all(slope_shift)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0396820698385
varray: [-3.007, -2.915, -2.941, -2.937]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.333513617913
varray: [-39.775, -39.04, -39.099, -39.318]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.8915
varray: [112.882, 115.709, 111.823, 111.152]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.00880968071
varray: [112.882, 115.709, 111.823, 111.152]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.213
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1380.0685
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.433
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00235
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00235
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0797989744983
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.72695
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-3.007, -2.915, -2.941, -2.937]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-39.775, -39.04, -39.099, -39.318]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.95
varray: [-3.007, -2.915, -2.941, -2.937]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
dD_mean | -39.308
varray: [-39.775, -39.04, -39.099, -39.318]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0396820698385
varray: [-3.007, -2.915, -2.941, -2.937]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
dD_stdev | 0.333513617913
varray: [-39.775, -39.04, -39.099, -39.318]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.95
varray: [-3.007, -2.915, -2.941, -2.937]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.308
varray: [-39.775, -39.04, -39.099, -39.318]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0396820698385
varray: [-3.007, -2.915, -2.941, -2.937]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.333513617913
varray: [-39.775, -39.04, -39.099, -39.318]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.87922494325
actual_array: [ 1.93 -15.84]
varray: [-3.007, -2.915, -2.941, -2.937]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.95
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.178072103
actual_array: [ 18.1 -119. ]
varray: [-39.775, -39.04, -39.099, -39.318]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [67887.0, 68423.0, 68959.0, 69496.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.12905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 77.9922
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.446
varray: [0.0, 0.0, 2.862, 2.922]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: residuals
SAMP_residuals_STD | 1.66987664215
varray: [0.0, 0.0, 2.862, 2.922]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.32825
varray: [0.0, 0.0, -0.749, -0.564]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.386481888321
varray: [0.0, 0.0, -0.749, -0.564]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 926.725
varray: [0.0, 0.0, 1855.239, 1851.661]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1070.09085339
varray: [0.0, 0.0, 1855.239, 1851.661]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -9.49925
varray: [0.0, 0.0, -19.303, -18.694]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: slope_shift
SAMP_slope_shift_STD | 10.9716064267
varray: [0.0, 0.0, -19.303, -18.694]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: slope_shift
SAMP_d18O_STD | 2.11104523874
varray: [0.0, 0.0, -3.616, -3.696]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 25.2773123225
varray: [0.0, 0.0, -43.719, -43.844]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.217
varray: [0.0, 0.0, 108.35, 108.518]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 62.6044366585
varray: [0.0, 0.0, 108.35, 108.518]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.446
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.32825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 926.725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.49925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.4989
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.4989
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.8981629257
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.94755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [0.0, 0.0, -3.616, -3.696]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [0.0, 0.0, -43.719, -43.844]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.828
varray: [0.0, 0.0, -3.616, -3.696]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
dD_mean | -21.89075
varray: [0.0, 0.0, -43.719, -43.844]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
d18O_stdev | 2.11104523874
varray: [0.0, 0.0, -3.616, -3.696]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
dD_stdev | 25.2773123225
varray: [0.0, 0.0, -43.719, -43.844]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.828
varray: [0.0, 0.0, -3.616, -3.696]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.89075
varray: [0.0, 0.0, -43.719, -43.844]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.11104523874
varray: [0.0, 0.0, -3.616, -3.696]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
measured_dD_stdev | 25.2773123225
varray: [0.0, 0.0, -43.719, -43.844]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.74213628224
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -3.616, -3.696]
fit_rsq: 1.0
fit_intercept: 2.85121024198
except: True
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.828
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -38.6510511133
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -43.719, -43.844]
fit_rsq: 1.0
fit_intercept: 18.8422423778
except: True
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -21.89075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [70085.0, 70621.0, 71157.0, 71693.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.6303
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -375.3513
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.466
varray: [0.0, 0.0, 2.921, 2.943]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: residuals
SAMP_residuals_STD | 1.69281481563
varray: [0.0, 0.0, 2.921, 2.943]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8385
varray: [0.0, 0.0, -1.669, -1.685]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.968238434822
varray: [0.0, 0.0, -1.669, -1.685]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 994.2155
varray: [0.0, 0.0, 1986.751, 1990.111]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1148.02199261
varray: [0.0, 0.0, 1986.751, 1990.111]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -12.96775
varray: [0.0, 0.0, -25.954, -25.917]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: slope_shift
SAMP_slope_shift_STD | 14.9738755254
varray: [0.0, 0.0, -25.954, -25.917]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: slope_shift
SAMP_d18O_STD | 1.26074485259
varray: [0.0, 0.0, -2.164, -2.203]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 24.7054795744
varray: [0.0, 0.0, -42.87, -42.712]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.72
varray: [0.0, 0.0, 109.312, 109.568]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 63.1852998938
varray: [0.0, 0.0, 109.312, 109.568]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.466
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8385
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 994.2155
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 12.96775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.4789
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.4789
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.2190227172
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.44455
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [0.0, 0.0, -2.164, -2.203]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [0.0, 0.0, -42.87, -42.712]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.09175
varray: [0.0, 0.0, -2.164, -2.203]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
dD_mean | -21.3955
varray: [0.0, 0.0, -42.87, -42.712]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
d18O_stdev | 1.26074485259
varray: [0.0, 0.0, -2.164, -2.203]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
dD_stdev | 24.7054795744
varray: [0.0, 0.0, -42.87, -42.712]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.09175
varray: [0.0, 0.0, -2.164, -2.203]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.3955
varray: [0.0, 0.0, -42.87, -42.712]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.26074485259
varray: [0.0, 0.0, -2.164, -2.203]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
measured_dD_stdev | 24.7054795744
varray: [0.0, 0.0, -42.87, -42.712]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.99598518895
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -2.164, -2.203]
fit_rsq: 1.0
fit_intercept: 2.85121024198
except: True
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.09175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -38.1811142895
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -42.87, -42.712]
fit_rsq: 1.0
fit_intercept: 18.8422423778
except: True
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -21.3955
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [72280.0, 72817.0, 73353.0, 73890.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.6103
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -307.8608
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.557
varray: [0.0, 0.0, 5.146, 5.082]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: residuals
SAMP_residuals_STD | 2.9526848799
varray: [0.0, 0.0, 5.146, 5.082]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.83
varray: [0.0, 0.0, -9.699, -9.621]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 5.5772945054
varray: [0.0, 0.0, -9.699, -9.621]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 782.3475
varray: [0.0, 0.0, 1564.273, 1565.117]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 903.37714516
varray: [0.0, 0.0, 1564.273, 1565.117]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -17.52875
varray: [0.0, 0.0, -34.776, -35.339]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: slope_shift
SAMP_slope_shift_STD | 20.2417620342
varray: [0.0, 0.0, -34.776, -35.339]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: slope_shift
SAMP_d18O_STD | 2.2410475787
varray: [0.0, 0.0, 3.902, 3.861]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 56.5326221007
varray: [0.0, 0.0, -97.624, -98.209]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.42925
varray: [0.0, 0.0, 132.111, 133.606]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 76.7083188334
varray: [0.0, 0.0, 132.111, 133.606]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.557
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.83
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 782.3475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 17.52875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3879
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3879
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -13.1719243438
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -45.7353
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [0.0, 0.0, 3.902, 3.861]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [0.0, 0.0, -97.624, -98.209]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.94075
varray: [0.0, 0.0, 3.902, 3.861]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
dD_mean | -48.95825
varray: [0.0, 0.0, -97.624, -98.209]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
d18O_stdev | 2.2410475787
varray: [0.0, 0.0, 3.902, 3.861]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
dD_stdev | 56.5326221007
varray: [0.0, 0.0, -97.624, -98.209]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.94075
varray: [0.0, 0.0, 3.902, 3.861]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.95825
varray: [0.0, 0.0, -97.624, -98.209]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.2410475787
varray: [0.0, 0.0, 3.902, 3.861]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
measured_dD_stdev | 56.5326221007
varray: [0.0, 0.0, -97.624, -98.209]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.922704115386
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 3.902, 3.861]
fit_rsq: 1.0
fit_intercept: 2.85121024198
except: True
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 1.94075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.3350793405
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -97.624, -98.209]
fit_rsq: 1.0
fit_intercept: 18.8422423778
except: True
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -48.95825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [74476.0, 75013.0, 75549.0, 76084.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.5193
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -519.7288
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.48725
varray: [0.0, 0.0, 3.017, 2.932]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: residuals
SAMP_residuals_STD | 1.71767893294
varray: [0.0, 0.0, 3.017, 2.932]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.20075
varray: [0.0, 0.0, -0.34, -0.463]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.237182594921
varray: [0.0, 0.0, -0.34, -0.463]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 721.2325
varray: [0.0, 0.0, 1442.601, 1442.329]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 832.80756345
varray: [0.0, 0.0, 1442.601, 1442.329]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.43875
varray: [0.0, 0.0, -4.999, -4.756]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.81777280549
varray: [0.0, 0.0, -4.999, -4.756]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: slope_shift
SAMP_d18O_STD | 2.28111617416
varray: [0.0, 0.0, -3.957, -3.945]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 29.6226494764
varray: [0.0, 0.0, -51.621, -50.991]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.18475
varray: [0.0, 0.0, 115.061, 113.678]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 66.033675436
varray: [0.0, 0.0, 115.061, 113.678]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.48725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.20075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 721.2325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.43875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.45765
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.45765
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -49.4974362457
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -54.9798
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [0.0, 0.0, -3.957, -3.945]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [0.0, 0.0, -51.621, -50.991]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.9755
varray: [0.0, 0.0, -3.957, -3.945]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
dD_mean | -25.653
varray: [0.0, 0.0, -51.621, -50.991]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
d18O_stdev | 2.28111617416
varray: [0.0, 0.0, -3.957, -3.945]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
dD_stdev | 29.6226494764
varray: [0.0, 0.0, -51.621, -50.991]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.9755
varray: [0.0, 0.0, -3.957, -3.945]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.653
varray: [0.0, 0.0, -51.621, -50.991]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.28111617416
varray: [0.0, 0.0, -3.957, -3.945]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
measured_dD_stdev | 29.6226494764
varray: [0.0, 0.0, -51.621, -50.991]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.8916198629
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -3.957, -3.945]
fit_rsq: 1.0
fit_intercept: 2.85121024198
except: True
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.9755
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.2210053092
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -51.621, -50.991]
fit_rsq: 1.0
fit_intercept: 18.8422423778
except: True
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -25.653
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [76670.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.58905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -580.8438
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.24625
varray: [4.193, 2.914, 2.916, 2.962]
inst: mean_all
operation: mean_all(residuals)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: residuals
SAMP_residuals_STD | 0.63155594896
varray: [4.193, 2.914, 2.916, 2.962]
inst: std_all
operation: std_all(residuals)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.52475
varray: [0.881, 0.343, 0.365, 0.51]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.248786890062
varray: [0.881, 0.343, 0.365, 0.51]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1315.7685
varray: [1352.645, 1304.121, 1303.252, 1303.056]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 24.5886893849
varray: [1352.645, 1304.121, 1303.252, 1303.056]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.50375
varray: [1.936, 2.462, 2.459, 3.158]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.501374361131
varray: [1.936, 2.462, 2.459, 3.158]
inst: std_all
operation: std_all(slope_shift)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: slope_shift
SAMP_d18O_STD | 0.306758727124
varray: [-6.552, -7.096, -7.188, -7.193]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.36869885169
varray: [-54.359, -58.914, -59.222, -59.132]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 122.11475
varray: [152.146, 112.285, 112.004, 112.024]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 20.0212425751
varray: [152.146, 112.285, 112.004, 112.024]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.24625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.52475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1315.7685
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.50375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.30135
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.30135
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 10.2329450915
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 9.9502
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.7786
varray: [-6.552, -7.096, -7.188, -7.193]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.5677
varray: [-54.359, -58.914, -59.222, -59.132]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.00725
varray: [-6.552, -7.096, -7.188, -7.193]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
dD_mean | -57.90675
varray: [-54.359, -58.914, -59.222, -59.132]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
d18O_stdev | 0.306758727124
varray: [-6.552, -7.096, -7.188, -7.193]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
dD_stdev | 2.36869885169
varray: [-54.359, -58.914, -59.222, -59.132]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.00725
varray: [-6.552, -7.096, -7.188, -7.193]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.90675
varray: [-54.359, -58.914, -59.222, -59.132]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.306758727124
varray: [-6.552, -7.096, -7.188, -7.193]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.36869885169
varray: [-54.359, -58.914, -59.222, -59.132]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.99103685939
actual_array: [ 1.93 -15.84]
varray: [-6.552, -7.096, -7.188, -7.193]
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.00725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.8262043646
actual_array: [ 18.1 -119. ]
varray: [-54.359, -58.914, -59.222, -59.132]
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -57.90675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: [78864.0, 79400.0, 79936.0, 80472.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.83005
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 13.6922
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.967
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.819
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9636935817
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.967
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.2318202802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -67.819
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.406
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.388
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3951492512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.406
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1295210088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -61.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.279
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.115
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2664413546
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.279
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9215864772
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -60.115
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.282
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.725
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2694816986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5515201242
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.338
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3262347869
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.338
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5733445502
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.278
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.511
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2654279066
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.278
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3484580742
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.511
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.286
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.483
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2735354906
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.286
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3218892078
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.483
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.253
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.374
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2400917065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.253
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2184604066
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.374
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.244
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1833386182
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0951049556
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.244
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.699
operation: col(d(18_16)Mean)
inst: col
dD_actual | 24.56
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.859190838476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 3.699
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.42551207773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 24.56
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.542
operation: col(d(18_16)Mean)
inst: col
dD_actual | 34.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.71352750624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.542
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.2246793264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 34.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | 37.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.91520365913
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.0779857964
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 37.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.87
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.109
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04593845171
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.87
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.2308848191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 39.109
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.899
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.792
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07532844384
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.899
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.8789753808
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 39.792
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.936
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.247
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.11282602001
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.936
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3107194593
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 40.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.939
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.254
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.11586636402
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.939
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3173616759
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 40.254
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.74
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.14120256413
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.7785212849
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 40.74
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.591
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.09053016391
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 4.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6371369603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 40.591
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 5.052
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.098
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.7556
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.9172
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.098
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.23038598853
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 5.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.1182232192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 41.098
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.209
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.131
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.2492920122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -11.209
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.6078588835
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -86.131
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.443
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.398
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.4998868497
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -12.443
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.043421354
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -102.398
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | -106.808
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8829701954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -12.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.228017807
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -106.808
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.903
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.37
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9660729318
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -12.903
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.710180995
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -108.37
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.515
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0846463483
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -13.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.796657851
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -109.515
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.782
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0836329003
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -13.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.05001097
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -109.782
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.09
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1059287564
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -13.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.3422685
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -110.09
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.069
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.35
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1343053005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -13.069
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.588979402
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -110.35
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.128
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.08
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1940987328
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -13.128
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.281667704
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -111.08
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.133
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.153
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1991659728
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -13.133
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.350936534
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -111.153
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.915
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.019
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9109942855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.4215978971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -68.019
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.328
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3161003068
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.328
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1522943228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -61.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.352
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3404230589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.352
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1910706932
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -60.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.354
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.349
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3424499549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.354
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.143626289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -60.349
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.194
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1802982742
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.194
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4281646733
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.272
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.482
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2593472186
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.272
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3209403197
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.482
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.304
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.819
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2917775547
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.304
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6407156042
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.819
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.261
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2481992905
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.261
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4670690847
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2319841225
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1909426521
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.224
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.254
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2107017144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.224
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1045938364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.254
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.408
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.792
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.3433841293
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.408
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.4839981894
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.792
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.097
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.613
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.02820179991
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.097
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.4163710533
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.613
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.66
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.89341321532
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.5120807088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.66
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.05
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.774
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.9805697437
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.05
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.6202539504
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.774
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.96
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.8893594233
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.96
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5646536766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.798
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.477
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.72518084657
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.798
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.1850984428
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.477
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.825
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.77382635079
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5153114962
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.048
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.821458407
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.7269135391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.048
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.654
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.96205449921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.7108350423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -35.654
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.14
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.265
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.05833205963
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.14
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.3417175774
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -35.265
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.677
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.96205449921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.7837713837
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -34.677
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.03
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.627
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.94685277914
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.03
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.7363269795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -34.627
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.693
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.576
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.63221681058
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.693
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.0523523911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -66.576
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.681
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62005543452
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.681
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.6865937548
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -70.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.735
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.401
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.67478162676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.735
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.6307373989
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -71.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.789
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.837
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.72950781901
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.789
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.0444526037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -71.837
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.882
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.984
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.81031047895
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.882
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.6661847017
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.984
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.804
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.216
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.7312615346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.804
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.0907743993
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.216
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.816
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.421
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.74342291065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.816
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.3364083721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -38.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.756
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.878
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.68261603039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.756
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.8211621422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -37.878
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.07380696011
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.8604506769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.121
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.353
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.05252455202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.121
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.1185482358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.353
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.065
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.99577146377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.065
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.7741018612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.1
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.03124214392
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.1
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.9486972687
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.195
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.413
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.15441571329
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2554670419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.413
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.283
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.221
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.24359913769
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.283
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9710566987
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -61.221
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.457
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.479
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.41993909046
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.457
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.164757909
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -62.479
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.322
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.021
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.28312360986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.322
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.7301671663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -62.021
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.296
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.186
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.2298779528
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.296
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.7556362637
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -45.186
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.141
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.834
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.07279351211
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.141
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.523851489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.834
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.13
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.236
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.06164558406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.13
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.9564164144
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.236
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.908
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.03934972796
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.6451811227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.908
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.845
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.82660492067
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -6.845
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.7665335962
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -56.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.149
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0282810736
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0560725031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.149
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1022627779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5504431951
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.163
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1488813861
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.163
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7013164005
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.474
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.51134331763
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.6933709336
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -32.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.31
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.227
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.23061822039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.31
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.6123313232
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -29.227
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.251
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.629
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17082478813
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.251
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.0448962487
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -28.629
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.159
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.08468170775
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.598918849
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -28.159
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.264
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.6428183778
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.4941044358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -32.264
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.123
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.65396630585
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.3091993004
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -33.123
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.751
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.217
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.66410078589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.3983947804
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -33.217
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.74
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.976
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.65295285784
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.74
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.169712752
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -32.976
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.399
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.244
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.33426309726
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.399
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.8617836881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -44.244
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.055
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5703964823
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.5802200091
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -46.055
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.64
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.577
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.57850406634
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.64
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.0755395893
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -46.577
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.564
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.501
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.501482018
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.564
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.0034240948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -46.501
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.314
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.692
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.24812001688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.314
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.491333212
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.692
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.123
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.452
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.05455144803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.123
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.3147119872
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.452
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.137
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.532
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.06873972009
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.137
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.390623034
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.532
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.097
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.32
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.02820179991
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.097
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.1894587601
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.32
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.06
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.60844041929
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0227333791
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -57.06
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.873
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.223
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.84153346032
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.873
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0751783058
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.223
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.883
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.85166794037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.883
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.650204485
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.933
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.148
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.90234034059
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -5.933
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.952899784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -60.148
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.733
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.467
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.64585872181
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.733
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.482281055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -36.467
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.376
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.594
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.28405778422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.376
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.7561255882
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -33.594
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.342
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.504
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.24960055206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.342
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.7218375761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -32.504
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.338
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.312
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.24554676005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.338
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.5396510639
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -32.312
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.996
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.53999304217
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.5753475277
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -44.996
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.67
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.035
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.60890750647
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.67
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.5612422475
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -46.035
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.703
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.33
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.64235129062
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.703
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.8411642324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -46.33
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.711
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.606
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.65045887466
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.711
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.1030573437
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -46.606
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.025
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.843
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.94178553912
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.4301676507
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -44.843
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.658
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.759
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.56985012148
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.658
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.4015729671
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.759
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.665
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.57694425751
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.665
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.4101129599
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.665
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.57694425751
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -1.665
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.3588730033
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.659
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.596
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.63810359184
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -6.659
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5824493079
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -56.596
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.063
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0282810736
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9744681278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.063
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.165
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.598
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1509082821
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.165
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.482123253
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.598
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.184
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.629
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1701637942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.184
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5115387836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.629
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.219
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.533
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.15184245646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.219
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.289348091
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.533
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.935
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.287
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.86402322319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.1581454533
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.287
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.898
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.649
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.82652564702
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.898
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.5527548554
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -38.649
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.941
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.495
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.87010391121
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.941
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.4066260903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -38.495
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.055
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.023
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.98563698372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.055
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.805415168
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.023
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.959
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.786
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.88834597529
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.959
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.5805286919
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.786
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.725
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.85186184713
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.5226465188
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.936
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.814
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.86503667119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.936
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.6070975583
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -40.814
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.21
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.37
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.14272142442
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.21
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.1346793333
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.37
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.246
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.256
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.17920555258
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.246
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.0265060916
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.256
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.235
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.036
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.16805762453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.8177507131
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.036
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.249
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.311
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.18224589659
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.249
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.0786949363
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -41.311
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.007
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.775
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.93699147951
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.007
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.6212028385
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.915
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.8437542631
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9237700964
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.941
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.099
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.87010391121
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.941
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9797544934
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.099
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.937
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.318
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.8660501192
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.937
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.1875609839
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -39.318
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.616
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.719
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.55418131423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.616
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.3636174438
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.719
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.696
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.63525715459
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.696
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.4822284543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -43.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.164
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.87
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.08265481174
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.164
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.55801146
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.87
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.203
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.712
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.12217928391
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -2.203
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.4080871427
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -42.712
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -97.624
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.06492078338
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 3.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.513429639
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -97.624
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.861
operation: col(d(18_16)Mean)
inst: col
dD_actual | -98.209
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.0233694152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 3.861
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -111.068529168
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -98.209
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.88955333006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.8791792776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.957
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.89976708376
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.957
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.8617310875
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -51.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.991
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.8876057077
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -3.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.2639315942
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -50.991
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:17:06
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.552
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.359
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.52966465536
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -6.552
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4597866628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -54.359
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.096
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.914
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0809803698
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.096
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7819718877
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -58.914
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.188
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1742175862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.188
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0742294177
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.7556
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.9172
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.133
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.153
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.7786
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.5677
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.193
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.132
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1792848262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.85121024198
measure_array: [ 4.7556 -12.7786]
inst: fit_type1
actual_mean: -7.193
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.986730444569
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9888294901
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.8422423778
measure_array: [ 37.9172 -106.5677]
inst: fit_type1
actual_mean: -59.132
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.053865062
tarray: []
source_col: d(D_H)Mean