## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
210104_20-196_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.0832 | -73.5933 |  | 20.3733 |  | 0.17155 | 0.58389
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.40486 | 5.1397
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.55887 | 7.4784
4 |  | EV | -10.3813 | -75.7321 |  |  |  | 0.20241 | 2.5887
5 | * | CPER1-E-4 | -2.7533 | -17.3383 |  |  |  | 0 | 0
6 | * | CPER1-A-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
7 |  | WOOD1-D-1 | 4.9972 | -63.7524 | -7.1548 | 28.0919 | 42.5632 | 0.06661 | 1.0193 | -20.6375 | 11.5877
8 | * | NIWO1-B-4 | -3.7659 | -57.6501 |  |  |  | 1.1576 | 49.0694
9 | * | CPER1-D-2 | 1.2951 | -42.0713 | -2.629 |  |  | 4.6269 | 30.1061 | -14.9335
10 | * | UNDE1-A-4 | 0.19435 | -25.813 |  |  |  | 5.835 | 17.867
11 | * | CPER1-B-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
12 |  | CPER1-E-2 | 1.4767 | -64.1063 |  | 92.3401 |  | 0.022277 | 0.63956 |  | -14.4298
13 |  | EV | -9.8983 | -75.3626 |  |  |  | 0.57331 | 1.1763
14 | * | WOOD1-A-2 | -3.166 | -37.1915 |  |  |  | 0.4717 | 24.1656 | -9.695
15 | * | WOOD1-E-4 | -3.1165 | -39.0622 |  |  |  | 0.41566 | 26.4428
16 | * | WOOD1-A-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
17 |  | NIWO1-2030-2 | -4.1453 | -76.7429 |  |  |  | 0.13831 | 2.6566
18 | * | NIWO1-010-1 | -7.5339 | -69.8626 |  |  |  | 5.4638 | 63.9337
19 | * | CPER1-3040-3 | -5.7156 | -57.9733 |  |  |  | 3.3856 | 49.4616
20 | * | NIWO1-010-3 | -2.7533 | -17.3383 |  |  |  | 0 | 0
21 | * | CPER1-3040-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
22 |  | EV | -10.7213 | -81.096 |  |  |  | 0.18387 | 1.2127
23 | * | NIWO1-010-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
24 | * | CPER1-deep-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
25 | * | CPER1-1020-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
26 | * | CPER1-1020-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
27 | * | CPER-010-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
28 | * | WOOD1-E-3 | -2.7533 | -17.3383 |  |  |  | 0 | 0
29 | * | EV | -2.7533 | -17.3383 |  |  |  | 0 | 0

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -7.29 | -59.425 | -10.1186 | -73.7113 | 0.615 | 1320.074 | 2021/01/04 10:26:44
1 |  | EV | -7.256 | -59.427 | -10.0842 | -73.7132 | 0.314 | 1324.079 | 2021/01/04 10:35:40
1 |  | EV | -6.939 | -58.182 | -9.764 | -72.5321 | 0.517 | 1356.821 | 2021/01/04 10:44:37
1 |  | EV | -6.955 | -58.294 | -9.7801 | -72.6384 | 0.558 | 1348.801 | 2021/01/04 10:53:33
1 |  | EV | -7.251 | -59.278 | -10.0792 | -73.5718 | 0.488 | 1316.905 | 2021/01/04 11:02:30
1 |  | EV | -7.349 | -59.541 | -10.1782 | -73.8213 | 0.541 | 1312.11 | 2021/01/04 11:11:26
1 |  | EV | -7.4 | -59.786 | -10.2297 | -74.0537 | 0.595 | 1307.918 | 2021/01/04 11:20:23
1 |  | EV | -7.399 | -59.754 | -10.2287 | -74.0234 | 0.492 | 1308.25 | 2021/01/04 11:29:19
1 |  | EV | -7.394 | -59.747 | -10.2237 | -74.0167 | 0.576 | 1306.042 | 2021/01/04 11:38:16
1 |  | EV | -7.317 | -59.572 | -10.1459 | -73.8507 | 0.244 | 1307.014 | 2021/01/04 11:47:13
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.0832 | -73.5933 |  | 20.3733 |  | 0.17155 | 0.58389
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.556 | 23.639 | 0.83945 | 5.0866 | 0.177 | 1300.045 | 2021/01/04 11:57:00
2 |  | PZ | 4.437 | 34.658 | 1.7296 | 15.5396 | 0.16 | 1304.005 | 2021/01/04 12:05:57
2 |  | PZ | 4.636 | 37.253 | 1.9306 | 18.0013 | 0.246 | 1304.084 | 2021/01/04 12:14:53
2 |  | PZ | 4.689 | 38.378 | 1.9842 | 19.0686 | 0.273 | 1302.202 | 2021/01/04 12:23:50
2 |  | PZ | 4.763 | 39.257 | 2.0589 | 19.9024 | 0.362 | 1303.34 | 2021/01/04 12:32:46
2 |  | PZ | 4.771 | 39.53 | 2.067 | 20.1614 | 0.183 | 1302.899 | 2021/01/04 12:41:43
2 |  | PZ | 4.847 | 39.897 | 2.1438 | 20.5095 | 0.385 | 1302.27 | 2021/01/04 12:50:40
2 |  | PZ | 4.893 | 40.068 | 2.1903 | 20.6718 | 0.427 | 1301.857 | 2021/01/04 12:59:36
2 |  | PZ | 4.854 | 40.335 | 2.1509 | 20.925 | 0.375 | 1301.805 | 2021/01/04 13:08:32
2 |  | PZ | 4.908 | 40.555 | 2.2054 | 21.1338 | 0.651 | 1302.902 | 2021/01/04 13:17:29
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.40486 | 5.1397
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -11.444 | -87.202 | -14.3155 | -100.0616 | 0.332 | 1299.795 | 2021/01/04 13:27:17
3 |  | UT2 | -12.691 | -103.028 | -15.5754 | -115.0748 | 0.422 | 1298.075 | 2021/01/04 13:36:15
3 |  | UT2 | -12.962 | -107.254 | -15.8492 | -119.0838 | 0.387 | 1297.91 | 2021/01/04 13:45:11
3 |  | UT2 | -13.102 | -108.924 | -15.9906 | -120.668 | 0.334 | 1297.36 | 2021/01/04 13:54:08
3 |  | UT2 | -13.175 | -109.825 | -16.0644 | -121.5227 | -0.055 | 1297.594 | 2021/01/04 14:03:04
3 |  | UT2 | -13.229 | -110.647 | -16.119 | -122.3025 | 0.171 | 1297.317 | 2021/01/04 14:12:00
3 |  | UT2 | -13.239 | -110.982 | -16.1291 | -122.6203 | 0.512 | 1297.079 | 2021/01/04 14:20:57
3 |  | UT2 | -13.163 | -110.758 | -16.0523 | -122.4078 | 0.614 | 1303.175 | 2021/01/04 14:29:54
3 |  | UT2 | -13.281 | -111.451 | -16.1715 | -123.0652 | 0.642 | 1297.344 | 2021/01/04 14:38:50
3 |  | UT2 | -13.243 | -111.586 | -16.1331 | -123.1933 | 0.659 | 1297.504 | 2021/01/04 14:47:47
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.55887 | 7.4784
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -8.101 | -68.55 | -10.938 | -82.3676 | 0.179 | 1297.7 | 2021/01/04 14:57:36
4 |  | EV | -7.608 | -62.748 | -10.4399 | -76.8636 | 0.404 | 1297.534 | 2021/01/04 15:06:33
4 |  | EV | -7.561 | -61.511 | -10.3924 | -75.6901 | 0.391 | 1296.774 | 2021/01/04 15:15:30
4 |  | EV | -7.506 | -61.01 | -10.3368 | -75.2149 | 0.259 | 1296.984 | 2021/01/04 15:24:26
4 |  | EV | -7.452 | -60.405 | -10.2823 | -74.6409 | 0.471 | 1298 | 2021/01/04 15:33:23
4 |  | EV | -7.441 | -60.315 | -10.2712 | -74.5556 | 0.578 | 1297.638 | 2021/01/04 15:42:20
4 |  | EV | -7.463 | -60.22 | -10.2934 | -74.4654 | 0.219 | 1297.504 | 2021/01/04 15:51:15
4 |  | EV | -7.423 | -60.373 | -10.253 | -74.6106 | 0.425 | 1297.18 | 2021/01/04 16:00:11
4 |  | EV | -7.437 | -60.242 | -10.2671 | -74.4863 | 0.213 | 1297.833 | 2021/01/04 16:09:08
4 |  | EV | -7.508 | -60.178 | -10.3389 | -74.4256 | 0.346 | 1298.177 | 2021/01/04 16:18:05
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3813 | -75.7321 |  |  |  | 0.20241 | 2.5887
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | CPER1-E-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 16:27:54
5 |  | CPER1-E-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 16:37:52
5 |  | CPER1-E-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 16:47:51
5 |  | CPER1-E-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 16:57:50
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 | * | CPER1-E-4 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | CPER1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 17:08:41
6 |  | CPER1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 17:18:40
6 |  | CPER1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 17:28:39
6 |  | CPER1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 17:38:38
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 | * | CPER1-A-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | WOOD1-D-1 | 7.669 | -47.454 | 4.9949 | -62.3551 | -7.752 | 1336.075 | 2021/01/04 17:49:30
7 |  | WOOD1-D-1 | 7.578 | -49.152 | 4.903 | -63.9659 | -7.455 | 1332.1 | 2021/01/04 17:58:27
7 |  | WOOD1-D-1 | 7.724 | -49.308 | 5.0505 | -64.1139 | -6.862 | 1326.458 | 2021/01/04 18:07:23
7 |  | WOOD1-D-1 | 7.714 | -49.794 | 5.0404 | -64.5749 | -6.55 | 1319.447 | 2021/01/04 18:16:20
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | WOOD1-D-1 | 4.9972 | -63.7524 | -7.1548 | 28.0919 | 42.5632 | 0.06661 | 1.0193 | -20.6375 | 11.5877
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | NIWO1-B-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 18:26:07
8 |  | NIWO1-B-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 18:35:03
8 |  | NIWO1-B-4 | -1.97 | -84.553 | -4.7436 | -97.5487 | 0.06 | 1281.908 | 2021/01/04 18:43:59
8 |  | NIWO1-B-4 | -2.039 | -85.424 | -4.8134 | -98.375 | 0.061 | 1273.942 | 2021/01/04 18:52:56
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 | * | NIWO1-B-4 | -3.7659 | -57.6501 |  |  |  | 1.1576 | 49.0694
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | CPER1-D-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 19:02:43
9 |  | CPER1-D-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 19:11:39
9 |  | CPER1-D-2 | 8.033 | -51.877 | 5.3627 | -66.5509 | -5.334 | 1791.7 | 2021/01/04 19:20:36
9 |  | CPER1-D-2 | 7.995 | -52.411 | 5.3243 | -67.0575 | -5.182 | 1769.042 | 2021/01/04 19:29:33
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 | * | CPER1-D-2 | 1.2951 | -42.0713 | -2.629 |  |  | 4.6269 | 30.1061 | -14.9335
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | UNDE1-A-4 | 11.67 | -35.734 | 9.0373 | -51.237 | -0.33 | 1395.029 | 2021/01/04 19:39:21
10 |  | UNDE1-A-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 19:48:17
10 |  | UNDE1-A-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 19:58:16
10 |  | UNDE1-A-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 20:08:14
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 | * | UNDE1-A-4 | 0.19435 | -25.813 |  |  |  | 5.835 | 17.867
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | CPER1-B-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 20:19:04
11 |  | CPER1-B-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 20:29:03
11 |  | CPER1-B-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 20:39:02
11 |  | CPER1-B-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 20:49:00
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 | * | CPER1-B-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | CPER1-E-2 | 4.19 | -48.346 | 1.48 | -63.2013 | 0.225 | 1404.017 | 2021/01/04 20:59:51
12 |  | CPER1-E-2 | 4.169 | -49.544 | 1.4588 | -64.3378 | -0.061 | 1390.448 | 2021/01/04 21:08:48
12 |  | CPER1-E-2 | 4.171 | -49.707 | 1.4608 | -64.4924 | 0.254 | 1393.468 | 2021/01/04 21:17:45
12 |  | CPER1-E-2 | 4.217 | -49.603 | 1.5073 | -64.3937 | 0.05 | 1383.14 | 2021/01/04 21:26:41
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | CPER1-E-2 | 1.4767 | -64.1063 |  | 92.3401 |  | 0.022277 | 0.63956 |  | -14.4298
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -6.244 | -59.697 | -9.0618 | -73.9693 | 0.227 | 1282.981 | 2021/01/04 21:36:30
13 |  | EV | -7.15 | -60.991 | -9.9772 | -75.1968 | 0.304 | 1278.79 | 2021/01/04 21:45:26
13 |  | EV | -7.366 | -61.434 | -10.1954 | -75.6171 | 0.163 | 1280.407 | 2021/01/04 21:54:22
13 |  | EV | -7.528 | -62.541 | -10.3591 | -76.6672 | 0.267 | 1267.677 | 2021/01/04 22:03:18
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -9.8983 | -75.3626 |  |  |  | 0.57331 | 1.1763
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD1-A-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 22:13:08
14 |  | WOOD1-A-2 | -0.817 | -41.856 | -3.5787 | -57.0446 | -1.163 | 1802.319 | 2021/01/04 22:22:05
14 |  | WOOD1-A-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 22:31:00
14 |  | WOOD1-A-2 | -0.817 | -41.856 | -3.5787 | -57.0446 | -1.163 | 1802.319 | 2021/01/04 22:40:59
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 | * | WOOD1-A-2 | -3.166 | -37.1915 |  |  |  | 0.4717 | 24.1656 | -9.695
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD1-E-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 22:51:47
15 |  | WOOD1-E-4 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 23:00:43
15 |  | WOOD1-E-4 | -0.693 | -45.694 | -3.4534 | -60.6855 | -0.059 | 1271.803 | 2021/01/04 23:09:40
15 |  | WOOD1-E-4 | -0.745 | -45.906 | -3.506 | -60.8866 | 0.16 | 1265.829 | 2021/01/04 23:18:36
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 | * | WOOD1-E-4 | -3.1165 | -39.0622 |  |  |  | 0.41566 | 26.4428
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | WOOD1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 23:28:23
16 |  | WOOD1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 23:38:21
16 |  | WOOD1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 23:48:20
16 |  | WOOD1-A-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/04 23:58:18
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 | * | WOOD1-A-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | NIWO1-2030-2 | -1.285 | -59.05 | -4.0516 | -73.3555 | 0.416 | 1323.831 | 2021/01/05 00:09:07
17 |  | NIWO1-2030-2 | -1.248 | -62.231 | -4.0142 | -76.3732 | 0.183 | 1306.707 | 2021/01/05 00:18:03
17 |  | NIWO1-2030-2 | -1.429 | -64.139 | -4.197 | -78.1832 | 0.49 | 1287.208 | 2021/01/05 00:26:59
17 |  | NIWO1-2030-2 | -1.549 | -65.063 | -4.3183 | -79.0597 | 0.274 | 1282.988 | 2021/01/05 00:35:56
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | NIWO1-2030-2 | -4.1453 | -76.7429 |  |  |  | 0.13831 | 2.6566
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | NIWO1-010-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 00:45:43
18 |  | NIWO1-010-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 00:54:39
18 |  | NIWO1-010-1 | -9.426 | -110.536 | -12.2767 | -122.1972 | 0.037 | 1246.74 | 2021/01/05 01:03:35
18 |  | NIWO1-010-1 | -9.501 | -110.936 | -12.3524 | -122.5767 | 0.235 | 1251.069 | 2021/01/05 01:12:30
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 | * | NIWO1-010-1 | -7.5339 | -69.8626 |  |  |  | 5.4638 | 63.9337
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | CPER1-3040-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 01:22:18
19 |  | CPER1-3040-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 01:31:15
19 |  | CPER1-3040-3 | -5.864 | -85.67 | -8.6779 | -98.6083 | -0.156 | 1256.3 | 2021/01/05 01:41:13
19 |  | CPER1-3040-3 | -5.864 | -85.67 | -8.6779 | -98.6083 | -0.156 | 1256.3 | 2021/01/05 01:50:09
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 | * | CPER1-3040-3 | -5.7156 | -57.9733 |  |  |  | 3.3856 | 49.4616
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | NIWO1-010-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 02:00:59
20 |  | NIWO1-010-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 02:10:57
20 |  | NIWO1-010-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 02:20:55
20 |  | NIWO1-010-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 02:30:54
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 | * | NIWO1-010-3 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | CPER1-3040-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 02:41:45
21 |  | CPER1-3040-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 02:51:43
21 |  | CPER1-3040-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 03:01:42
21 |  | CPER1-3040-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 03:11:41
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 | * | CPER1-3040-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -7.749 | -67.837 | -10.5823 | -81.6912 | 0.404 | 1278.289 | 2021/01/05 03:22:27
22 |  | EV | -7.732 | -65.516 | -10.5652 | -79.4894 | 0.244 | 1264.667 | 2021/01/05 03:31:23
22 |  | EV | -7.943 | -67.203 | -10.7783 | -81.0898 | 0.065 | 1241.353 | 2021/01/05 03:40:19
22 |  | EV | -8.122 | -68.282 | -10.9592 | -82.1134 | 0.213 | 1225.225 | 2021/01/05 03:49:15
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.7213 | -81.096 |  |  |  | 0.18387 | 1.2127
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | NIWO1-010-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 03:59:00
23 |  | NIWO1-010-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 04:08:59
23 |  | NIWO1-010-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 04:18:57
23 |  | NIWO1-010-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 04:28:55
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 | * | NIWO1-010-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | CPER1-deep-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 04:39:44
24 |  | CPER1-deep-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 04:49:42
24 |  | CPER1-deep-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 04:59:40
24 |  | CPER1-deep-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 05:09:38
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 | * | CPER1-deep-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | CPER1-1020-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 05:20:26
25 |  | CPER1-1020-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 05:30:25
25 |  | CPER1-1020-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 05:40:23
25 |  | CPER1-1020-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 05:50:21
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 | * | CPER1-1020-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | CPER1-1020-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 06:01:10
26 |  | CPER1-1020-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 06:11:08
26 |  | CPER1-1020-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 06:21:06
26 |  | CPER1-1020-2 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 06:31:04
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 | * | CPER1-1020-2 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | CPER-010-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 06:41:53
27 |  | CPER-010-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 06:51:52
27 |  | CPER-010-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 07:01:50
27 |  | CPER-010-1 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 07:11:48
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 | * | CPER-010-1 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | WOOD1-E-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 07:22:39
28 |  | WOOD1-E-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 07:32:37
28 |  | WOOD1-E-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 07:42:34
28 |  | WOOD1-E-3 |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 07:52:33
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 | * | WOOD1-E-3 | -2.7533 | -17.3383 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | EV |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 08:03:19
29 |  | EV |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 08:13:17
29 |  | EV |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 08:23:15
29 |  | EV |  |  | -2.7533 | -17.3383 |  |  | 2021/01/05 08:33:13
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 | * | EV | -2.7533 | -17.3383 |  |  |  | 0 | 0

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-14438 | 2021/01/04 10:26:44 | 1-01 | 1 | -7.29 | -59.425 | 19368 | -1 | 1 | EV | standard | H2O | 1609781510 | 0.316 | 0.733 | 101.578 | -0.001 | 0 | 2.408 | 1320.074 | 2.355 | 3.09 | 118.898 | 0.868 | 0.615 | 0 | 0 | 1 | 45 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-14438 | 2021/01/04 10:35:40 | 1-01 | 2 | -7.256 | -59.427 | 18518 | -1 | 1 | EV | standard | H2O | 1609782047 | 0.302 | 0.665 | 97.168 | -0.001 | 0.001 | 2.296 | 1324.079 | 2.165 | 2.919 | 112.592 | 0.866 | 0.314 | 0 | 0 | 1 | 45 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-14438 | 2021/01/04 10:44:37 | 1-01 | 3 | -6.939 | -58.182 | 18920 | -1 | 1 | EV | standard | H2O | 1609782583 | 0.304 | 0.681 | 99.717 | 0 | 0.002 | 2.366 | 1356.821 | 0.579 | 3.064 | 115.598 | 0.873 | 0.517 | 0 | 0 | 1 | 45.062 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-14438 | 2021/01/04 10:53:33 | 1-01 | 4 | -6.955 | -58.294 | 18916 | 0 | 1 | EV | standard | H2O | 1609783120 | 0.343 | 0.777 | 102.608 | -0.001 | 0.001 | 2.431 | 1348.801 | 1.195 | 3.041 | 116.445 | 0.868 | 0.558 | 0 | 0 | 1 | 45.062 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-14438 | 2021/01/04 11:02:30 | 1-01 | 5 | -7.251 | -59.278 | 18532 | 0 | 1 | EV | standard | H2O | 1609783656 | 0.284 | 0.761 | 102.703 | 0 | 0.003 | 2.418 | 1316.905 | 2.578 | 2.942 | 112.917 | 0.864 | 0.488 | 0 | 0 | 1 | 45.062 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-14438 | 2021/01/04 11:11:26 | 1-01 | 6 | -7.349 | -59.541 | 18528 | 0 | 1 | EV | standard | H2O | 1609784193 | 0.274 | 0.711 | 101.479 | 0 | 0 | 2.397 | 1312.11 | 2.795 | 2.943 | 113.271 | 0.872 | 0.541 | 0 | 0 | 1 | 45.125 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-14438 | 2021/01/04 11:20:23 | 1-01 | 7 | -7.4 | -59.786 | 18210 | 0 | 1 | EV | standard | H2O | 1609784729 | 0.301 | 0.746 | 99.373 | -0.001 | 0 | 2.376 | 1307.918 | 2.423 | 2.925 | 111.969 | 0.872 | 0.595 | 0 | 0 | 1 | 45.188 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-14438 | 2021/01/04 11:29:19 | 1-01 | 8 | -7.399 | -59.754 | 18537 | 0 | 1 | EV | standard | H2O | 1609785266 | 0.295 | 0.649 | 99.716 | 0 | 0.003 | 2.369 | 1308.25 | 2.364 | 2.979 | 113.747 | 0.877 | 0.492 | 0 | 0 | 1 | 45.188 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-14438 | 2021/01/04 11:38:16 | 1-01 | 9 | -7.394 | -59.747 | 18387 | 0 | 1 | EV | standard | H2O | 1609785802 | 0.33 | 0.712 | 101.004 | -0.001 | 0.001 | 2.396 | 1306.042 | 2.572 | 2.941 | 113.294 | 0.868 | 0.576 | 0 | 0 | 1 | 45.125 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-14438 | 2021/01/04 11:47:13 | 1-01 | 10 | -7.317 | -59.572 | 18614 | 0 | 1 | EV | standard | H2O | 1609786339 | 0.308 | 0.766 | 102.801 | -0.001 | 0.005 | 2.439 | 1307.014 | 2.611 | 2.92 | 113.376 | 0.868 | 0.244 | 0 | 0 | 1 | 45.25 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-14439 | 2021/01/04 11:57:00 | 1-02 | 1 | 3.556 | 23.639 | 18102 | -1 | 1 | PZ | standard | H2O | 1609786926 | 0.319 | 0.861 | 99.407 | -0.002 | -0.011 | 2.368 | 1300.045 | 2.409 | 2.843 | 110.428 | 0.869 | 0.177 | 0 | 0 | 1 | 45.375 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-14439 | 2021/01/04 12:05:57 | 1-02 | 2 | 4.437 | 34.658 | 18407 | -1 | 1 | PZ | standard | H2O | 1609787463 | 0.314 | 0.868 | 100.007 | -0.002 | -0.01 | 2.385 | 1304.005 | 2.093 | 2.927 | 112.111 | 0.874 | 0.16 | 0 | 0 | 1 | 45.375 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-14439 | 2021/01/04 12:14:53 | 1-02 | 3 | 4.636 | 37.253 | 18405 | -1 | 1 | PZ | standard | H2O | 1609787999 | 0.308 | 0.836 | 103.695 | -0.001 | -0.004 | 2.457 | 1304.084 | 2.426 | 2.915 | 111.386 | 0.874 | 0.246 | 0 | 0 | 1 | 45.312 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-14439 | 2021/01/04 12:23:50 | 1-02 | 4 | 4.689 | 38.378 | 18198 | 0 | 1 | PZ | standard | H2O | 1609788536 | 0.344 | 0.79 | 102.923 | -0.001 | -0.004 | 2.455 | 1302.202 | 1.775 | 2.927 | 111.901 | 0.868 | 0.273 | 0 | 0 | 1 | 45.312 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-14439 | 2021/01/04 12:32:46 | 1-02 | 5 | 4.763 | 39.257 | 18377 | 0 | 1 | PZ | standard | H2O | 1609789072 | 0.322 | 0.709 | 100.588 | -0.002 | -0.004 | 2.38 | 1303.34 | 1.845 | 3.001 | 112.963 | 0.865 | 0.362 | 0 | 0 | 1 | 45.312 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-14439 | 2021/01/04 12:41:43 | 1-02 | 6 | 4.771 | 39.53 | 18339 | 0 | 1 | PZ | standard | H2O | 1609789609 | 0.288 | 0.722 | 98.293 | 0 | -0.005 | 2.35 | 1302.899 | 2.074 | 2.929 | 111.991 | 0.872 | 0.183 | 0 | 0 | 1 | 45.312 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-14439 | 2021/01/04 12:50:40 | 1-02 | 7 | 4.847 | 39.897 | 18304 | 0 | 1 | PZ | standard | H2O | 1609790146 | 0.304 | 0.719 | 100.152 | -0.001 | -0.001 | 2.381 | 1302.27 | 2.571 | 2.907 | 111.689 | 0.872 | 0.385 | 0 | 0 | 1 | 45.375 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-14439 | 2021/01/04 12:59:36 | 1-02 | 8 | 4.893 | 40.068 | 18269 | 0 | 1 | PZ | standard | H2O | 1609790682 | 0.287 | 0.753 | 100.597 | 0 | -0.004 | 2.387 | 1301.857 | 2.092 | 2.972 | 112.724 | 0.873 | 0.427 | 0 | 0 | 1 | 45.312 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-14439 | 2021/01/04 13:08:32 | 1-02 | 9 | 4.854 | 40.335 | 18270 | 0 | 1 | PZ | standard | H2O | 1609791218 | 0.321 | 0.762 | 101.953 | -0.001 | -0.003 | 2.407 | 1301.805 | 2.281 | 3 | 111.848 | 0.867 | 0.375 | 0 | 0 | 1 | 45.375 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-14439 | 2021/01/04 13:17:29 | 1-02 | 10 | 4.908 | 40.555 | 18447 | 0 | 1 | PZ | standard | H2O | 1609791755 | 0.291 | 0.706 | 101.613 | -0.001 | -0.002 | 2.42 | 1302.902 | 1.859 | 3.01 | 115.062 | 0.865 | 0.651 | 0 | 0 | 1 | 45.312 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-14440 | 2021/01/04 13:27:17 | 1-03 | 1 | -11.444 | -87.202 | 18148 | -1 | 1 | UT2 | standard | H2O | 1609792343 | 0.328 | 0.981 | 97.153 | 0.001 | 0.016 | 2.352 | 1299.795 | 3.361 | 2.873 | 110.119 | 0.865 | 0.332 | 0 | 0 | 1 | 45.375 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-14440 | 2021/01/04 13:36:15 | 1-03 | 2 | -12.691 | -103.028 | 17824 | -1 | 1 | UT2 | standard | H2O | 1609792881 | 0.333 | 0.982 | 100.011 | 0 | 0.014 | 2.37 | 1298.075 | 3.301 | 2.83 | 108.35 | 0.864 | 0.422 | 0 | 0 | 1 | 45.312 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-14440 | 2021/01/04 13:45:11 | 1-03 | 3 | -12.962 | -107.254 | 17781 | -1 | 1 | UT2 | standard | H2O | 1609793418 | 0.299 | 0.884 | 102.127 | 0 | 0.009 | 2.418 | 1297.91 | 2.639 | 2.812 | 108.434 | 0.865 | 0.387 | 0 | 0 | 1 | 45.375 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-14440 | 2021/01/04 13:54:08 | 1-03 | 4 | -13.102 | -108.924 | 17715 | 0 | 1 | UT2 | standard | H2O | 1609793955 | 0.293 | 0.731 | 97.134 | -0.001 | 0.005 | 2.31 | 1297.36 | 3.114 | 2.766 | 107.204 | 0.866 | 0.334 | 0 | 0 | 1 | 45.438 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-14440 | 2021/01/04 14:03:04 | 1-03 | 5 | -13.175 | -109.825 | 17724 | 0 | 1 | UT2 | standard | H2O | 1609794490 | 0.31 | 0.875 | 99.999 | -0.001 | 0.009 | 2.39 | 1297.594 | 2.96 | 2.694 | 105.991 | 0.861 | -0.055 | 0 | 0 | 1 | 45.312 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-14440 | 2021/01/04 14:12:00 | 1-03 | 6 | -13.229 | -110.647 | 17721 | 0 | 1 | UT2 | standard | H2O | 1609795027 | 0.314 | 0.869 | 99.363 | -0.001 | 0.009 | 2.369 | 1297.317 | 3.047 | 2.774 | 106.801 | 0.868 | 0.171 | 0 | 0 | 1 | 45.25 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-14440 | 2021/01/04 14:20:57 | 1-03 | 7 | -13.239 | -110.982 | 17721 | 0 | 1 | UT2 | standard | H2O | 1609795564 | 0.333 | 0.76 | 99.448 | -0.001 | 0.003 | 2.375 | 1297.079 | 2.877 | 2.79 | 108.166 | 0.868 | 0.512 | 0 | 0 | 1 | 45 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-14440 | 2021/01/04 14:29:54 | 1-03 | 8 | -13.163 | -110.758 | 18572 | 0 | 1 | UT2 | standard | H2O | 1609796100 | 0.283 | 0.789 | 101.147 | 0 | 0.003 | 2.401 | 1303.175 | 2.487 | 2.966 | 115.128 | 0.868 | 0.614 | 0 | 0 | 1 | 44.938 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-14440 | 2021/01/04 14:38:50 | 1-03 | 9 | -13.281 | -111.451 | 17785 | 0 | 1 | UT2 | standard | H2O | 1609796637 | 0.291 | 0.744 | 102.015 | 0 | 0.001 | 2.421 | 1297.344 | 2.752 | 2.809 | 110.237 | 0.867 | 0.642 | 0 | 0 | 1 | 44.875 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-14440 | 2021/01/04 14:47:47 | 1-03 | 10 | -13.243 | -111.586 | 17810 | 0 | 1 | UT2 | standard | H2O | 1609797173 | 0.303 | 0.677 | 100.301 | -0.001 | 0.005 | 2.391 | 1297.504 | 2.767 | 2.84 | 110.629 | 0.874 | 0.659 | 0 | 0 | 1 | 44.75 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-14441 | 2021/01/04 14:57:36 | 1-04 | 1 | -8.101 | -68.55 | 17864 | -1 | 1 | EV | standard | H2O | 1609797763 | 0.3 | 0.683 | 100.396 | -0.001 | 0 | 2.39 | 1297.7 | 3.518 | 2.812 | 107.639 | 0.87 | 0.179 | 0 | 0 | 1 | 44.688 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-14441 | 2021/01/04 15:06:33 | 1-04 | 2 | -7.608 | -62.748 | 17776 | -1 | 1 | EV | standard | H2O | 1609798299 | 0.295 | 0.772 | 98.615 | 0 | 0.001 | 2.355 | 1297.534 | 2.522 | 2.816 | 108.795 | 0.873 | 0.404 | 0 | 0 | 1 | 44.625 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-14441 | 2021/01/04 15:15:30 | 1-04 | 3 | -7.561 | -61.511 | 17675 | -1 | 1 | EV | standard | H2O | 1609798836 | 0.288 | 0.73 | 100.303 | -0.001 | 0.001 | 2.39 | 1296.774 | 2.602 | 2.792 | 107.964 | 0.874 | 0.391 | 0 | 0 | 1 | 44.625 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-14441 | 2021/01/04 15:24:26 | 1-04 | 4 | -7.506 | -61.01 | 17690 | 0 | 1 | EV | standard | H2O | 1609799373 | 0.326 | 0.704 | 99.867 | 0 | -0.002 | 2.379 | 1296.984 | 2.639 | 2.759 | 107.374 | 0.874 | 0.259 | 0 | 0 | 1 | 44.562 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-14441 | 2021/01/04 15:33:23 | 1-04 | 5 | -7.452 | -60.405 | 17866 | 0 | 1 | EV | standard | H2O | 1609799910 | 0.286 | 0.682 | 100.702 | 0 | 0 | 2.392 | 1298 | 2.853 | 2.867 | 109.355 | 0.867 | 0.471 | 0 | 0 | 1 | 44.688 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-14441 | 2021/01/04 15:42:20 | 1-04 | 6 | -7.441 | -60.315 | 17840 | 0 | 1 | EV | standard | H2O | 1609800446 | 0.317 | 0.753 | 100.618 | -0.001 | 0 | 2.404 | 1297.638 | 3.117 | 2.914 | 109.241 | 0.872 | 0.578 | 0 | 0 | 1 | 44.812 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-14441 | 2021/01/04 15:51:15 | 1-04 | 7 | -7.463 | -60.22 | 17773 | 0 | 1 | EV | standard | H2O | 1609800982 | 0.285 | 0.748 | 98.068 | 0 | 0.003 | 2.349 | 1297.504 | 2.885 | 2.782 | 107.423 | 0.87 | 0.219 | 0 | 0 | 1 | 44.875 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-14441 | 2021/01/04 16:00:11 | 1-04 | 8 | -7.423 | -60.373 | 17786 | 0 | 1 | EV | standard | H2O | 1609801518 | 0.294 | 0.695 | 101.293 | -0.001 | -0.001 | 2.409 | 1297.18 | 2.946 | 2.869 | 108.99 | 0.871 | 0.425 | 0 | 0 | 1 | 44.938 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-14441 | 2021/01/04 16:09:08 | 1-04 | 9 | -7.437 | -60.242 | 17834 | 0 | 1 | EV | standard | H2O | 1609802055 | 0.318 | 0.756 | 101.535 | -0.001 | -0.001 | 2.431 | 1297.833 | 2.939 | 2.776 | 107.656 | 0.865 | 0.213 | 0 | 0 | 1 | 44.938 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-14441 | 2021/01/04 16:18:05 | 1-04 | 10 | -7.508 | -60.178 | 17900 | 0 | 1 | EV | standard | H2O | 1609802591 | 0.337 | 0.767 | 100.542 | 0 | 0 | 2.386 | 1298.177 | 2.884 | 2.866 | 108.795 | 0.868 | 0.346 | 0 | 0 | 1 | 45 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-14442 | 2021/01/04 16:27:54 | 1-05 | 1 | [5 empty] | CPER1-E-4 | 20-196 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-14442 | 2021/01/04 16:37:52 | 1-05 | 2 | [5 empty] | CPER1-E-4 | 20-196 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-14442 | 2021/01/04 16:47:51 | 1-05 | 3 | [5 empty] | CPER1-E-4 | 20-196 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-14442 | 2021/01/04 16:57:50 | 1-05 | 4 | [5 empty] | CPER1-E-4 | 20-196 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-14443 | 2021/01/04 17:08:41 | 1-06 | 1 | [5 empty] | CPER1-A-1 | 20-196 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-14443 | 2021/01/04 17:18:40 | 1-06 | 2 | [5 empty] | CPER1-A-1 | 20-196 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-14443 | 2021/01/04 17:28:39 | 1-06 | 3 | [5 empty] | CPER1-A-1 | 20-196 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-14443 | 2021/01/04 17:38:38 | 1-06 | 4 | [5 empty] | CPER1-A-1 | 20-196 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-14444 | 2021/01/04 17:49:30 | 1-07 | 1 | 7.669 | -47.454 | 17798 | -1 | 1 | WOOD1-D-1 | 20-196 | H2O | 1609808077 | 0.361 | 0.936 | 84.232 | -0.003 | 0.003 | 2.006 | 1336.075 | -21.858 | 4.31 | 127.058 | 0.869 | -7.752 | 0 | 0 | 1 | 45.375 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-14444 | 2021/01/04 17:58:27 | 1-07 | 2 | 7.578 | -49.152 | 17427 | -1 | 1 | WOOD1-D-1 | 20-196 | H2O | 1609808613 | 0.294 | 0.987 | 95.289 | -0.001 | 0.006 | 2.273 | 1332.1 | -21.196 | 4.196 | 124.188 | 0.873 | -7.455 | 0 | 0 | 1 | 45.188 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-14444 | 2021/01/04 18:07:23 | 1-07 | 3 | 7.724 | -49.308 | 16849 | -1 | 1 | WOOD1-D-1 | 20-196 | H2O | 1609809150 | 0.305 | 0.894 | 93.505 | -0.002 | 0 | 2.238 | 1326.458 | -20.391 | 4.016 | 121.703 | 0.87 | -6.862 | 0 | 0 | 1 | 45 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-14444 | 2021/01/04 18:16:20 | 1-07 | 4 | 7.714 | -49.794 | 16102 | 0 | 1 | WOOD1-D-1 | 20-196 | H2O | 1609809687 | 0.325 | 0.968 | 93.226 | -0.002 | 0.001 | 2.232 | 1319.447 | -19.105 | 3.897 | 116.034 | 0.865 | -6.55 | 0 | 0 | 1 | 45.062 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-14445 | 2021/01/04 18:26:07 | 1-08 | 1 | [5 empty] | NIWO1-B-4 | 20-196 | [18 empty] | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-14445 | 2021/01/04 18:35:03 | 1-08 | 2 | [5 empty] | NIWO1-B-4 | 20-196 | [18 empty] | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-14445 | 2021/01/04 18:43:59 | 1-08 | 3 | -1.97 | -84.553 | 15181 | -1 | 1 | NIWO1-B-4 | 20-196 | H2O | 1609811346 | 0.363 | 0.842 | 91.718 | 0 | 0.005 | 2.206 | 1281.908 | 2.38 | 2.332 | 91.993 | 0.862 | 0.06 | 0 | 0 | 1 | 45.125 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-14445 | 2021/01/04 18:52:56 | 1-08 | 4 | -2.039 | -85.424 | 14069 | 0 | 1 | NIWO1-B-4 | 20-196 | H2O | 1609811882 | 0.308 | 0.817 | 86.828 | -0.001 | 0.006 | 2.123 | 1273.942 | 1.827 | 2.177 | 86.705 | 0.863 | 0.061 | 0 | 0 | 1 | 45.125 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-14446 | 2021/01/04 19:02:43 | 1-09 | 1 | [5 empty] | CPER1-D-2 | 20-196 | [18 empty] | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-14446 | 2021/01/04 19:11:39 | 1-09 | 2 | [5 empty] | CPER1-D-2 | 20-196 | [18 empty] | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-14446 | 2021/01/04 19:20:36 | 1-09 | 3 | 8.033 | -51.877 | 17034 | -1 | 1 | CPER1-D-2 | 20-196 | H2O | 1609813542 | 0.382 | 1.062 | 99.079 | -0.003 | 0.004 | 2.386 | 1791.7 | -30.614 | 3.373 | 112.917 | 0.87 | -5.334 | 0 | 0 | 1 | 45.125 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-14446 | 2021/01/04 19:29:33 | 1-09 | 4 | 7.995 | -52.411 | 16404 | 0 | 1 | CPER1-D-2 | 20-196 | H2O | 1609814080 | 0.392 | 1.067 | 98.541 | -0.002 | 0.004 | 2.353 | 1769.042 | -29.12 | 3.245 | 107.357 | 0.867 | -5.182 | 0 | 0 | 1 | 45.188 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-14447 | 2021/01/04 19:39:21 | 1-10 | 1 | 11.67 | -35.734 | 6698 | -1 | 1 | UNDE1-A-4 | 20-196 | H2O | 1609814717 | 0.379 | 1.363 | 18.394 | -0.006 | -0.02 | 1.75 | 1395.029 | -5.735 | 1.046 | 42.231 | 0.859 | -0.33 | 0 | 0 | 1 | 45.25 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-14447 | 2021/01/04 19:48:17 | 1-10 | 2 | [5 empty] | UNDE1-A-4 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-14447 | 2021/01/04 19:58:16 | 1-10 | 3 | [5 empty] | UNDE1-A-4 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-14447 | 2021/01/04 20:08:14 | 1-10 | 4 | [5 empty] | UNDE1-A-4 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-14448 | 2021/01/04 20:19:04 | 1-11 | 1 | [5 empty] | CPER1-B-1 | 20-196 | H2O | [17 empty] | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-14448 | 2021/01/04 20:29:03 | 1-11 | 2 | [5 empty] | CPER1-B-1 | 20-196 | H2O | [17 empty] | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-14448 | 2021/01/04 20:39:02 | 1-11 | 3 | [5 empty] | CPER1-B-1 | 20-196 | H2O | [17 empty] | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-14448 | 2021/01/04 20:49:00 | 1-11 | 4 | [5 empty] | CPER1-B-1 | 20-196 | H2O | [17 empty] | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-14449 | 2021/01/04 20:59:51 | 1-12 | 1 | 4.19 | -48.346 | 16348 | -1 | 1 | CPER1-E-2 | 20-196 | H2O | 1609819498 | 0.317 | 0.828 | 82.035 | -0.002 | 0.003 | 1.972 | 1404.017 | -2.143 | 2.598 | 99.927 | 0.871 | 0.225 | 0 | 0 | 1 | 45.562 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-14449 | 2021/01/04 21:08:48 | 1-12 | 2 | 4.169 | -49.544 | 15658 | -1 | 1 | CPER1-E-2 | 20-196 | H2O | 1609820035 | 0.339 | 0.835 | 88.638 | -0.001 | 0.001 | 2.135 | 1390.448 | -1.583 | 2.404 | 94.161 | 0.871 | -0.061 | 0 | 0 | 1 | 45.312 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-14449 | 2021/01/04 21:17:45 | 1-12 | 3 | 4.171 | -49.707 | 15975 | -1 | 1 | CPER1-E-2 | 20-196 | H2O | 1609820572 | 0.345 | 0.862 | 90.848 | 0 | 0.006 | 2.178 | 1393.468 | -1.455 | 2.527 | 97.574 | 0.869 | 0.254 | 0 | 0 | 1 | 45.188 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-14449 | 2021/01/04 21:26:41 | 1-12 | 4 | 4.217 | -49.603 | 15244 | 0 | 1 | CPER1-E-2 | 20-196 | H2O | 1609821108 | 0.326 | 0.877 | 87.723 | -0.001 | 0.001 | 2.11 | 1383.14 | -2.034 | 2.396 | 93.251 | 0.871 | 0.05 | 0 | 0 | 1 | 45.125 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-14450 | 2021/01/04 21:36:30 | 1-13 | 1 | -6.244 | -59.697 | 15450 | -1 | 1 | EV | standard | H2O | 1609821696 | 0.313 | 0.793 | 90.105 | 0.001 | 0.007 | 2.163 | 1282.981 | 2.582 | 2.405 | 93.994 | 0.868 | 0.227 | 0 | 0 | 1 | 45.188 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-14450 | 2021/01/04 21:45:26 | 1-13 | 2 | -7.15 | -60.991 | 14931 | -1 | 1 | EV | standard | H2O | 1609822233 | 0.312 | 0.852 | 88.531 | 0.001 | 0.006 | 2.14 | 1278.79 | 2.738 | 2.375 | 91.142 | 0.862 | 0.304 | 0 | 0 | 1 | 45.188 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-14450 | 2021/01/04 21:54:22 | 1-13 | 3 | -7.366 | -61.434 | 15178 | -1 | 1 | EV | standard | H2O | 1609822768 | 0.31 | 0.791 | 90.419 | -0.001 | 0.003 | 2.193 | 1280.407 | 2.618 | 2.37 | 91.911 | 0.865 | 0.163 | 0 | 0 | 1 | 45.25 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-14450 | 2021/01/04 22:03:18 | 1-13 | 4 | -7.528 | -62.541 | 13394 | 0 | 1 | EV | standard | H2O | 1609823306 | 0.283 | 0.897 | 84.616 | 0 | 0.003 | 2.072 | 1267.677 | 2.423 | 2.082 | 82.344 | 0.862 | 0.267 | 0 | 0 | 1 | 45.25 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-14451 | 2021/01/04 22:13:08 | 1-14 | 1 | [5 empty] | WOOD1-A-2 | 20-196 | [18 empty] | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-14451 | 2021/01/04 22:22:05 | 1-14 | 2 | -0.817 | -41.856 | 15319 | -1 | 1 | WOOD1-A-2 | 20-196 | H2O | 1609824431 | 0.374 | 1.007 | 93.545 | 0 | 0.004 | 2.249 | 1802.319 | -19.39 | 2.403 | 91.987 | 0.873 | -1.163 | 0 | 0 | 1 | 45.312 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-14451 | 2021/01/04 22:31:00 | 1-14 | 3 | [5 empty] | WOOD1-A-2 | 20-196 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-14451 | 2021/01/04 22:40:59 | 1-14 | 4 | -0.817 | -41.856 | 15319 | -1 | 1 | WOOD1-A-2 | 20-196 | H2O | 1609824431 | 0.374 | 1.007 | 93.545 | 0 | 0.004 | 2.249 | 1802.319 | -19.39 | 2.403 | 91.987 | 0.873 | -1.163 | 0 | 0 | 1 | 45.312 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-14452 | 2021/01/04 22:51:47 | 1-15 | 1 | [5 empty] | WOOD1-E-4 | 20-196 | [18 empty] | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-14452 | 2021/01/04 23:00:43 | 1-15 | 2 | [5 empty] | WOOD1-E-4 | 20-196 | [18 empty] | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-14452 | 2021/01/04 23:09:40 | 1-15 | 3 | -0.693 | -45.694 | 13455 | -1 | 1 | WOOD1-E-4 | 20-196 | H2O | 1609827287 | 0.33 | 0.888 | 85.504 | -0.001 | 0.001 | 2.081 | 1271.803 | 1.66 | 2.057 | 81.964 | 0.861 | -0.059 | 0 | 0 | 1 | 45.25 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-14452 | 2021/01/04 23:18:36 | 1-15 | 4 | -0.745 | -45.906 | 12705 | 0 | 1 | WOOD1-E-4 | 20-196 | H2O | 1609827824 | 0.343 | 0.897 | 82.572 | -0.003 | 0.001 | 2.036 | 1265.829 | 1.853 | 1.976 | 77.715 | 0.861 | 0.16 | 0 | 0 | 1 | 45.312 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-14453 | 2021/01/04 23:28:23 | 1-16 | 1 | [5 empty] | WOOD1-A-1 | 20-196 | H2O | [17 empty] | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-14453 | 2021/01/04 23:38:21 | 1-16 | 2 | [5 empty] | WOOD1-A-1 | 20-196 | H2O | [17 empty] | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-14453 | 2021/01/04 23:48:20 | 1-16 | 3 | [5 empty] | WOOD1-A-1 | 20-196 | H2O | [17 empty] | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-14453 | 2021/01/04 23:58:18 | 1-16 | 4 | [5 empty] | WOOD1-A-1 | 20-196 | H2O | [17 empty] | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-14454 | 2021/01/05 00:09:07 | 1-17 | 1 | -1.285 | -59.05 | 20827 | -1 | 1 | NIWO1-2030-2 | 20-196 | H2O | 1609830854 | 0.344 | 0.637 | 93.786 | -0.002 | 0.003 | 2.206 | 1323.831 | 2.392 | 3.344 | 126.598 | 0.87 | 0.416 | 0 | 0 | 1 | 45.625 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-14454 | 2021/01/05 00:18:03 | 1-17 | 2 | -1.248 | -62.231 | 18914 | -1 | 1 | NIWO1-2030-2 | 20-196 | H2O | 1609831390 | 0.326 | 0.773 | 100.201 | -0.001 | 0.005 | 2.381 | 1306.707 | 2.949 | 3.009 | 114.813 | 0.878 | 0.183 | 0 | 0 | 1 | 45.375 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-14454 | 2021/01/05 00:26:59 | 1-17 | 3 | -1.429 | -64.139 | 16268 | -1 | 1 | NIWO1-2030-2 | 20-196 | H2O | 1609831926 | 0.327 | 0.808 | 93.204 | -0.001 | 0.004 | 2.25 | 1287.208 | 2.331 | 2.573 | 100.822 | 0.869 | 0.49 | 0 | 0 | 1 | 45.312 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-14454 | 2021/01/05 00:35:56 | 1-17 | 4 | -1.549 | -65.063 | 15633 | 0 | 1 | NIWO1-2030-2 | 20-196 | H2O | 1609832462 | 0.315 | 0.763 | 91.394 | 0.001 | 0.004 | 2.201 | 1282.988 | 2.157 | 2.448 | 95.681 | 0.87 | 0.274 | 0 | 0 | 1 | 45.188 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-14455 | 2021/01/05 00:45:43 | 1-18 | 1 | [5 empty] | NIWO1-010-1 | 20-196 | [18 empty] | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-14455 | 2021/01/05 00:54:39 | 1-18 | 2 | [5 empty] | NIWO1-010-1 | 20-196 | [18 empty] | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-14455 | 2021/01/05 01:03:35 | 1-18 | 3 | -9.426 | -110.536 | 10331 | -1 | 1 | NIWO1-010-1 | 20-196 | H2O | 1609834123 | 0.305 | 1.101 | 73.446 | 0.002 | 0.011 | 1.861 | 1246.74 | 2.309 | 1.563 | 62.734 | 0.859 | 0.037 | 0 | 0 | 1 | 45.125 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-14455 | 2021/01/05 01:12:30 | 1-18 | 4 | -9.501 | -110.936 | 11004 | 0 | 1 | NIWO1-010-1 | 20-196 | H2O | 1609834657 | 0.345 | 1.086 | 76.778 | 0.001 | 0.011 | 1.924 | 1251.069 | 2.396 | 1.682 | 67.706 | 0.858 | 0.235 | 0 | 0 | 1 | 45.188 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-14456 | 2021/01/05 01:22:18 | 1-19 | 1 | [5 empty] | CPER1-3040-3 | 20-196 | [18 empty] | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-14456 | 2021/01/05 01:31:15 | 1-19 | 2 | [5 empty] | CPER1-3040-3 | 20-196 | H2O | [17 empty] | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-14456 | 2021/01/05 01:41:13 | 1-19 | 3 | -5.864 | -85.67 | 11739 | -1 | 1 | CPER1-3040-3 | 20-196 | H2O | 1609836380 | 0.292 | 0.86 | 71.827 | 0 | -0.002 | 1.795 | 1256.3 | 2.31 | 1.753 | 70.475 | 0.86 | -0.156 | 0 | 0 | 1 | 45.25 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-14456 | 2021/01/05 01:50:09 | 1-19 | 4 | -5.864 | -85.67 | 11739 | -1 | 1 | CPER1-3040-3 | 20-196 | H2O | 1609836380 | 0.292 | 0.86 | 71.827 | 0 | -0.002 | 1.795 | 1256.3 | 2.31 | 1.753 | 70.475 | 0.86 | -0.156 | 0 | 0 | 1 | 45.25 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-14457 | 2021/01/05 02:00:59 | 1-20 | 1 | [5 empty] | NIWO1-010-3 | 20-196 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-14457 | 2021/01/05 02:10:57 | 1-20 | 2 | [5 empty] | NIWO1-010-3 | 20-196 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-14457 | 2021/01/05 02:20:55 | 1-20 | 3 | [5 empty] | NIWO1-010-3 | 20-196 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-14457 | 2021/01/05 02:30:54 | 1-20 | 4 | [5 empty] | NIWO1-010-3 | 20-196 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-14458 | 2021/01/05 02:41:45 | 1-21 | 1 | [5 empty] | CPER1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-14458 | 2021/01/05 02:51:43 | 1-21 | 2 | [5 empty] | CPER1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-14458 | 2021/01/05 03:01:42 | 1-21 | 3 | [5 empty] | CPER1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-14458 | 2021/01/05 03:11:41 | 1-21 | 4 | [5 empty] | CPER1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-14459 | 2021/01/05 03:22:27 | 1-22 | 1 | -7.749 | -67.837 | 15055 | -1 | 1 | EV | standard | H2O | 1609842454 | 0.282 | 0.841 | 73.886 | -0.001 | -0.003 | 1.781 | 1278.289 | 2.887 | 2.326 | 91.75 | 0.864 | 0.404 | 0 | 0 | 1 | 45.688 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-14459 | 2021/01/05 03:31:23 | 1-22 | 2 | -7.732 | -65.516 | 13064 | -1 | 1 | EV | standard | H2O | 1609842990 | 0.291 | 0.88 | 78.127 | 0 | -0.003 | 1.923 | 1264.667 | 2.528 | 2 | 79.842 | 0.862 | 0.244 | 0 | 0 | 1 | 45.375 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-14459 | 2021/01/05 03:40:19 | 1-22 | 3 | -7.943 | -67.203 | 9541 | -1 | 1 | EV | standard | H2O | 1609843527 | 0.332 | 0.975 | 68.899 | 0 | 0.003 | 1.768 | 1241.353 | 2.059 | 1.455 | 58.427 | 0.859 | 0.065 | 0 | 0 | 1 | 45.312 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-14459 | 2021/01/05 03:49:15 | 1-22 | 4 | -8.122 | -68.282 | 7052 | 0 | 1 | EV | standard | H2O | 1609844070 | 0.382 | 1.369 | 53.872 | -0.002 | 0.007 | 1.589 | 1225.225 | 1.581 | 1.069 | 44.547 | 0.856 | 0.213 | 0 | 0 | 1 | 45.312 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-14460 | 2021/01/05 03:59:00 | 1-23 | 1 | [5 empty] | NIWO1-010-2 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-14460 | 2021/01/05 04:08:59 | 1-23 | 2 | [5 empty] | NIWO1-010-2 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-14460 | 2021/01/05 04:18:57 | 1-23 | 3 | [5 empty] | NIWO1-010-2 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-14460 | 2021/01/05 04:28:55 | 1-23 | 4 | [5 empty] | NIWO1-010-2 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-14461 | 2021/01/05 04:39:44 | 1-24 | 1 | [5 empty] | CPER1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-14461 | 2021/01/05 04:49:42 | 1-24 | 2 | [5 empty] | CPER1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-14461 | 2021/01/05 04:59:40 | 1-24 | 3 | [5 empty] | CPER1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-14461 | 2021/01/05 05:09:38 | 1-24 | 4 | [5 empty] | CPER1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-14462 | 2021/01/05 05:20:26 | 1-25 | 1 | [5 empty] | CPER1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-14462 | 2021/01/05 05:30:25 | 1-25 | 2 | [5 empty] | CPER1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-14462 | 2021/01/05 05:40:23 | 1-25 | 3 | [5 empty] | CPER1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-14462 | 2021/01/05 05:50:21 | 1-25 | 4 | [5 empty] | CPER1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-14463 | 2021/01/05 06:01:10 | 1-26 | 1 | [5 empty] | CPER1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-14463 | 2021/01/05 06:11:08 | 1-26 | 2 | [5 empty] | CPER1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-14463 | 2021/01/05 06:21:06 | 1-26 | 3 | [5 empty] | CPER1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-14463 | 2021/01/05 06:31:04 | 1-26 | 4 | [5 empty] | CPER1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-14464 | 2021/01/05 06:41:53 | 1-27 | 1 | [5 empty] | CPER-010-1 | 20-196 | H2O | [17 empty] | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-14464 | 2021/01/05 06:51:52 | 1-27 | 2 | [5 empty] | CPER-010-1 | 20-196 | H2O | [17 empty] | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-14464 | 2021/01/05 07:01:50 | 1-27 | 3 | [5 empty] | CPER-010-1 | 20-196 | H2O | [17 empty] | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-14464 | 2021/01/05 07:11:48 | 1-27 | 4 | [5 empty] | CPER-010-1 | 20-196 | H2O | [17 empty] | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-14465 | 2021/01/05 07:22:39 | 1-28 | 1 | [5 empty] | WOOD1-E-3 | 20-196 | H2O | [17 empty] | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-14465 | 2021/01/05 07:32:37 | 1-28 | 2 | [5 empty] | WOOD1-E-3 | 20-196 | H2O | [17 empty] | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-14465 | 2021/01/05 07:42:34 | 1-28 | 3 | [5 empty] | WOOD1-E-3 | 20-196 | H2O | [17 empty] | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-14465 | 2021/01/05 07:52:33 | 1-28 | 4 | [5 empty] | WOOD1-E-3 | 20-196 | H2O | [17 empty] | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-14466 | 2021/01/05 08:03:19 | 1-29 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-14466 | 2021/01/05 08:13:17 | 1-29 | 2 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-14466 | 2021/01/05 08:23:15 | 1-29 | 3 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-14466 | 2021/01/05 08:33:13 | 1-29 | 4 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1300.4281
varray: [1300.045, 1304.005, 1304.084, 1302.202, 1303.34, 1302.899, 1302.27, 1301.857, 1301.805, 1302.902, 1299.795, 1298.075, 1297.91, 1297.36, 1297.594, 1297.317, 1297.079, 1303.175, 1297.344, 1297.504]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_shift
baseline_shift_STD | 2.65285536138
varray: [1300.045, 1304.005, 1304.084, 1302.202, 1303.34, 1302.899, 1302.27, 1301.857, 1301.805, 1302.902, 1299.795, 1298.075, 1297.91, 1297.36, 1297.594, 1297.317, 1297.079, 1303.175, 1297.344, 1297.504]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.36285
varray: [0.177, 0.16, 0.246, 0.273, 0.362, 0.183, 0.385, 0.427, 0.375, 0.651, 0.332, 0.422, 0.387, 0.334, -0.055, 0.171, 0.512, 0.614, 0.642, 0.659]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.189981792756
varray: [0.177, 0.16, 0.246, 0.273, 0.362, 0.183, 0.385, 0.427, 0.375, 0.651, 0.332, 0.422, 0.387, 0.334, -0.055, 0.171, 0.512, 0.614, 0.642, 0.659]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: ch4_ppm
residuals_MEAN | 2.87925
varray: [2.843, 2.927, 2.915, 2.927, 3.001, 2.929, 2.907, 2.972, 3.0, 3.01, 2.873, 2.83, 2.812, 2.766, 2.694, 2.774, 2.79, 2.966, 2.809, 2.84]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: residuals
residuals_STD | 0.0896407302866
varray: [2.843, 2.927, 2.915, 2.927, 3.001, 2.929, 2.907, 2.972, 3.0, 3.01, 2.873, 2.83, 2.812, 2.766, 2.694, 2.774, 2.79, 2.966, 2.809, 2.84]
inst: std_standards
operation: std_standards(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: residuals
slope_shift_MEAN | 2.5365
varray: [2.409, 2.093, 2.426, 1.775, 1.845, 2.074, 2.571, 2.092, 2.281, 1.859, 3.361, 3.301, 2.639, 3.114, 2.96, 3.047, 2.877, 2.487, 2.752, 2.767]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: slope_shift
slope_shift_STD | 0.485926840931
varray: [2.409, 2.093, 2.426, 1.775, 1.845, 2.074, 2.571, 2.092, 2.281, 1.859, 3.361, 3.301, 2.639, 3.114, 2.96, 3.047, 2.877, 2.487, 2.752, 2.767]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: slope_shift
baseline_curvature_MEAN | 110.6581
varray: [110.428, 112.111, 111.386, 111.901, 112.963, 111.991, 111.689, 112.724, 111.848, 115.062, 110.119, 108.35, 108.434, 107.204, 105.991, 106.801, 108.166, 115.128, 110.237, 110.629]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_curvature
baseline_curvature_STD | 2.54425823954
varray: [110.428, 112.111, 111.386, 111.901, 112.963, 111.991, 111.689, 112.724, 111.848, 115.062, 110.119, 108.35, 108.434, 107.204, 105.991, 106.801, 108.166, 115.128, 110.237, 110.629]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.87925
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1300.4281
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.36285
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 2.5365
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.13446109543
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 3.01371109543
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.179281460573
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.05853146057
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.26892219086
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.14817219086
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.287925
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.569945378269
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.932795378269
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 1.32642768069
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1301.75452768
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1303.08095536
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 7.95856608415
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1308.38666608
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 47.7513965049
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1348.1794965
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.42963420465
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.96613420465
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 7.63277471862
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 118.290874719
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -4.62366905435
actual_array: [ 1.93 -15.84]
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317, 3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908, -11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243, -8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 7.669, 7.578, 7.724, 7.714, 0.0, 0.0, -1.97, -2.039, 0.0, 0.0, 8.033, 7.995, 11.67, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 4.19, 4.169, 4.171, 4.217, -6.244, -7.15, -7.366, -7.528, 0.0, -0.817, 0.0, -0.817, 0.0, 0.0, -0.693, -0.745, 0.0, 0.0, 0.0, 0.0, -1.285, -1.248, -1.429, -1.549, 0.0, 0.0, -9.426, -9.501, 0.0, 0.0, -5.864, -5.864, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -7.749, -7.732, -7.943, -8.122, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.85125714286
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0, 5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0, 16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0, 21671.0, 22269.0, 22868.0, 23467.0, 24118.0, 24717.0, 25316.0, 25915.0, 26567.0, 27104.0, 27640.0, 28177.0, 28764.0, 29300.0, 29836.0, 30373.0, 30960.0, 31496.0, 32033.0, 32570.0, 33158.0, 33694.0, 34293.0, 34891.0, 35541.0, 36140.0, 36739.0, 37337.0, 37988.0, 38525.0, 39062.0, 39598.0, 40187.0, 40723.0, 41259.0, 41795.0, 42385.0, 42922.0, 43457.0, 44056.0, 44704.0, 45240.0, 45777.0, 46313.0, 46900.0, 47498.0, 48097.0, 48695.0, 49344.0, 49880.0, 50416.0, 50953.0, 51540.0, 52076.0, 52612.0, 53147.0, 53735.0, 54272.0, 54870.0, 55406.0, 56056.0, 56654.0, 57252.0, 57851.0, 58502.0, 59100.0, 59699.0, 60298.0, 60944.0, 61480.0, 62016.0, 62552.0, 63137.0, 63736.0, 64334.0, 64932.0, 65581.0, 66179.0, 66777.0, 67375.0, 68023.0, 68622.0, 69220.0, 69818.0, 70467.0, 71065.0, 71663.0, 72261.0, 72910.0, 73509.0, 74107.0, 74705.0, 75356.0, 75954.0, 76551.0, 77150.0, 77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -44.0470262358
actual_array: [ 18.1 -119. ]
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572, 23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555, -87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586, -68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -47.454, -49.152, -49.308, -49.794, 0.0, 0.0, -84.553, -85.424, 0.0, 0.0, -51.877, -52.411, -35.734, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -48.346, -49.544, -49.707, -49.603, -59.697, -60.991, -61.434, -62.541, 0.0, -41.856, 0.0, -41.856, 0.0, 0.0, -45.694, -45.906, 0.0, 0.0, 0.0, 0.0, -59.05, -62.231, -64.139, -65.063, 0.0, 0.0, -110.536, -110.936, 0.0, 0.0, -85.67, -85.67, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -67.837, -65.516, -67.203, -68.282, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -28.1547142857
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0, 5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0, 10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0, 16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0, 21671.0, 22269.0, 22868.0, 23467.0, 24118.0, 24717.0, 25316.0, 25915.0, 26567.0, 27104.0, 27640.0, 28177.0, 28764.0, 29300.0, 29836.0, 30373.0, 30960.0, 31496.0, 32033.0, 32570.0, 33158.0, 33694.0, 34293.0, 34891.0, 35541.0, 36140.0, 36739.0, 37337.0, 37988.0, 38525.0, 39062.0, 39598.0, 40187.0, 40723.0, 41259.0, 41795.0, 42385.0, 42922.0, 43457.0, 44056.0, 44704.0, 45240.0, 45777.0, 46313.0, 46900.0, 47498.0, 48097.0, 48695.0, 49344.0, 49880.0, 50416.0, 50953.0, 51540.0, 52076.0, 52612.0, 53147.0, 53735.0, 54272.0, 54870.0, 55406.0, 56056.0, 56654.0, 57252.0, 57851.0, 58502.0, 59100.0, 59699.0, 60298.0, 60944.0, 61480.0, 62016.0, 62552.0, 63137.0, 63736.0, 64334.0, 64932.0, 65581.0, 66179.0, 66777.0, 67375.0, 68023.0, 68622.0, 69220.0, 69818.0, 70467.0, 71065.0, 71663.0, 72261.0, 72910.0, 73509.0, 74107.0, 74705.0, 75356.0, 75954.0, 76551.0, 77150.0, 77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9764
varray: [3.09, 2.919, 3.064, 3.041, 2.942, 2.943, 2.925, 2.979, 2.941, 2.92]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: residuals
SAMP_residuals_STD | 0.0644880695254
varray: [3.09, 2.919, 3.064, 3.041, 2.942, 2.943, 2.925, 2.979, 2.941, 2.92]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.494
varray: [0.615, 0.314, 0.517, 0.558, 0.488, 0.541, 0.595, 0.492, 0.576, 0.244]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.121719166755
varray: [0.615, 0.314, 0.517, 0.558, 0.488, 0.541, 0.595, 0.492, 0.576, 0.244]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1320.8014
varray: [1320.074, 1324.079, 1356.821, 1348.801, 1316.905, 1312.11, 1307.918, 1308.25, 1306.042, 1307.014]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 17.9964047286
varray: [1320.074, 1324.079, 1356.821, 1348.801, 1316.905, 1312.11, 1307.918, 1308.25, 1306.042, 1307.014]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.1637
varray: [2.355, 2.165, 0.579, 1.195, 2.578, 2.795, 2.423, 2.364, 2.572, 2.611]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.709505156195
varray: [2.355, 2.165, 0.579, 1.195, 2.578, 2.795, 2.423, 2.364, 2.572, 2.611]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_d18O_STD | 0.171548503021
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.58389310285
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.2107
varray: [118.898, 112.592, 115.598, 116.445, 112.917, 113.271, 111.969, 113.747, 113.294, 113.376]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.13024944027
varray: [118.898, 112.592, 115.598, 116.445, 112.917, 113.271, 111.969, 113.747, 113.294, 113.376]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9764
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.494
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1320.8014
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.1637
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.09715
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09715
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.37414257185
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.5526
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.255
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_mean | -59.3006
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
d18O_stdev | 0.171548503021
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_stdev | 0.58389310285
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.255
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.3006
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.171548503021
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.58389310285
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0832365266
actual_array: [ 1.93 -15.84]
varray: [-7.29, -7.256, -6.939, -6.955, -7.251, -7.349, -7.4, -7.399, -7.394, -7.317]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.255
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.5932562151
actual_array: [ 18.1 -119. ]
varray: [-59.425, -59.427, -58.182, -58.294, -59.278, -59.541, -59.786, -59.754, -59.747, -59.572]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.3006
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3756.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.4517
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 20.3733
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9431
varray: [2.843, 2.927, 2.915, 2.927, 3.001, 2.929, 2.907, 2.972, 3.0, 3.01]
inst: mean_all
operation: mean_all(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: residuals
SAMP_residuals_STD | 0.0524202039082
varray: [2.843, 2.927, 2.915, 2.927, 3.001, 2.929, 2.907, 2.972, 3.0, 3.01]
inst: std_all
operation: std_all(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3239
varray: [0.177, 0.16, 0.246, 0.273, 0.362, 0.183, 0.385, 0.427, 0.375, 0.651]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149953659509
varray: [0.177, 0.16, 0.246, 0.273, 0.362, 0.183, 0.385, 0.427, 0.375, 0.651]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.5409
varray: [1300.045, 1304.005, 1304.084, 1302.202, 1303.34, 1302.899, 1302.27, 1301.857, 1301.805, 1302.902]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.19484824234
varray: [1300.045, 1304.005, 1304.084, 1302.202, 1303.34, 1302.899, 1302.27, 1301.857, 1301.805, 1302.902]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.1425
varray: [2.409, 2.093, 2.426, 1.775, 1.845, 2.074, 2.571, 2.092, 2.281, 1.859]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.272400134606
varray: [2.409, 2.093, 2.426, 1.775, 1.845, 2.074, 2.571, 2.092, 2.281, 1.859]
inst: std_all
operation: std_all(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: slope_shift
SAMP_d18O_STD | 0.404864641303
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.13971594546
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.2103
varray: [110.428, 112.111, 111.386, 111.901, 112.963, 111.991, 111.689, 112.724, 111.848, 115.062]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.21907688665
varray: [110.428, 112.111, 111.386, 111.901, 112.963, 111.991, 111.689, 112.724, 111.848, 115.062]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9431
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3239
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.5409
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.1425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06385
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06385
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.21759138665
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.5522
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.6354
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 37.357
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.6354
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
dD_mean | 37.357
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
d18O_stdev | 0.404864641303
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
dD_stdev | 5.13971594546
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.6354
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
measured_dD_mean | 37.357
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.404864641303
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.13971594546
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.556, 4.437, 4.636, 4.689, 4.763, 4.771, 4.847, 4.893, 4.854, 4.908]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.6354
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [23.639, 34.658, 37.253, 38.378, 39.257, 39.53, 39.897, 40.068, 40.335, 40.555]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 37.357
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [5417.0, 5954.0, 6490.0, 7027.0, 7563.0, 8100.0, 8637.0, 9173.0, 9709.0, 10246.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.485
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.1128
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8154
varray: [2.873, 2.83, 2.812, 2.766, 2.694, 2.774, 2.79, 2.966, 2.809, 2.84]
inst: mean_all
operation: mean_all(residuals)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: residuals
SAMP_residuals_STD | 0.0718071491334
varray: [2.873, 2.83, 2.812, 2.766, 2.694, 2.774, 2.79, 2.966, 2.809, 2.84]
inst: std_all
operation: std_all(residuals)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4018
varray: [0.332, 0.422, 0.387, 0.334, -0.055, 0.171, 0.512, 0.614, 0.642, 0.659]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.224363692844
varray: [0.332, 0.422, 0.387, 0.334, -0.055, 0.171, 0.512, 0.614, 0.642, 0.659]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.3153
varray: [1299.795, 1298.075, 1297.91, 1297.36, 1297.594, 1297.317, 1297.079, 1303.175, 1297.344, 1297.504]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.87343274054
varray: [1299.795, 1298.075, 1297.91, 1297.36, 1297.594, 1297.317, 1297.079, 1303.175, 1297.344, 1297.504]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.9305
varray: [3.361, 3.301, 2.639, 3.114, 2.96, 3.047, 2.877, 2.487, 2.752, 2.767]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.281629484725
varray: [3.361, 3.301, 2.639, 3.114, 2.96, 3.047, 2.877, 2.487, 2.752, 2.767]
inst: std_all
operation: std_all(slope_shift)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: slope_shift
SAMP_d18O_STD | 0.558872565478
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 7.47835568751
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.1059
varray: [110.119, 108.35, 108.434, 107.204, 105.991, 106.801, 108.166, 115.128, 110.237, 110.629]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.6125746539
varray: [110.119, 108.35, 108.434, 107.204, 105.991, 106.801, 108.166, 115.128, 110.237, 110.629]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8154
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4018
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.3153
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.9305
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06385
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06385
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.21759138665
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.5522
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.9529
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
dD_mean | -107.1657
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
d18O_stdev | 0.558872565478
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
dD_stdev | 7.47835568751
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.9529
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.1657
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.558872565478
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
measured_dD_stdev | 7.47835568751
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-11.444, -12.691, -12.962, -13.102, -13.175, -13.229, -13.239, -13.163, -13.281, -13.243]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -12.9529
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-87.202, -103.028, -107.254, -108.924, -109.825, -110.647, -110.982, -110.758, -111.451, -111.586]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -107.1657
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [10834.0, 11372.0, 11908.0, 12445.0, 12981.0, 13517.0, 14054.0, 14591.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.6127
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.1128
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8253
varray: [2.812, 2.816, 2.792, 2.759, 2.867, 2.914, 2.782, 2.869, 2.776, 2.866]
inst: mean_all
operation: mean_all(residuals)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: residuals
SAMP_residuals_STD | 0.0508222392265
varray: [2.812, 2.816, 2.792, 2.759, 2.867, 2.914, 2.782, 2.869, 2.776, 2.866]
inst: std_all
operation: std_all(residuals)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3485
varray: [0.179, 0.404, 0.391, 0.259, 0.471, 0.578, 0.219, 0.425, 0.213, 0.346]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.12922525553
varray: [0.179, 0.404, 0.391, 0.259, 0.471, 0.578, 0.219, 0.425, 0.213, 0.346]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.5324
varray: [1297.7, 1297.534, 1296.774, 1296.984, 1298.0, 1297.638, 1297.504, 1297.18, 1297.833, 1298.177]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.44280274013
varray: [1297.7, 1297.534, 1296.774, 1296.984, 1298.0, 1297.638, 1297.504, 1297.18, 1297.833, 1298.177]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.8905
varray: [3.518, 2.522, 2.602, 2.639, 2.853, 3.117, 2.885, 2.946, 2.939, 2.884]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.285627282388
varray: [3.518, 2.522, 2.602, 2.639, 2.853, 3.117, 2.885, 2.946, 2.939, 2.884]
inst: std_all
operation: std_all(slope_shift)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: slope_shift
SAMP_d18O_STD | 0.202412669784
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.58868867705
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.3232
varray: [107.639, 108.795, 107.964, 107.374, 109.355, 109.241, 107.423, 108.99, 107.656, 108.795]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.785311799365
varray: [107.639, 108.795, 107.964, 107.374, 109.355, 109.241, 107.423, 108.99, 107.656, 108.795]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8253
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3485
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.5324
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.8905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05395
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05395
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.8737518451
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.3349
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.55
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
dD_mean | -61.5552
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
d18O_stdev | 0.202412669784
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
dD_stdev | 2.58868867705
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.55
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.5552
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.202412669784
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.58868867705
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3812840923
actual_array: [ 1.93 -15.84]
varray: [-8.101, -7.608, -7.561, -7.506, -7.452, -7.441, -7.463, -7.423, -7.437, -7.508]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.55
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.7320597387
actual_array: [ 18.1 -119. ]
varray: [-68.55, -62.748, -61.511, -61.01, -60.405, -60.315, -60.22, -60.373, -60.242, -60.178]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -61.5552
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [16253.0, 16790.0, 17327.0, 17863.0, 18400.0, 18937.0, 19472.0, 20008.0, 20545.0, 21082.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.6028
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.8957
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [21671.0, 22269.0, 22868.0, 23467.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [24118.0, 24717.0, 25316.0, 25915.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.10475
varray: [4.31, 4.196, 4.016, 3.897]
inst: mean_all
operation: mean_all(residuals)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: residuals
SAMP_residuals_STD | 0.183930013864
varray: [4.31, 4.196, 4.016, 3.897]
inst: std_all
operation: std_all(residuals)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -7.15475
varray: [-7.752, -7.455, -6.862, -6.55]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.547199765473
varray: [-7.752, -7.455, -6.862, -6.55]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1328.52
varray: [1336.075, 1332.1, 1326.458, 1319.447]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.22185567104
varray: [1336.075, 1332.1, 1326.458, 1319.447]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -20.6375
varray: [-21.858, -21.196, -20.391, -19.105]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.18474483891
varray: [-21.858, -21.196, -20.391, -19.105]
inst: std_all
operation: std_all(slope_shift)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0666101844065
varray: [7.669, 7.578, 7.724, 7.714]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.0193468497
varray: [-47.454, -49.152, -49.308, -49.794]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 122.24575
varray: [127.058, 124.188, 121.703, 116.034]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.68367735688
varray: [127.058, 124.188, 121.703, 116.034]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.10475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 7.15475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1328.52
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.6375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.2255
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.2255
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 42.5631674915
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 11.58765
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [7.669, 7.578, 7.724, 7.714]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-47.454, -49.152, -49.308, -49.794]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 7.67125
varray: [7.669, 7.578, 7.724, 7.714]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
dD_mean | -48.927
varray: [-47.454, -49.152, -49.308, -49.794]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0666101844065
varray: [7.669, 7.578, 7.724, 7.714]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
dD_stdev | 1.0193468497
varray: [-47.454, -49.152, -49.308, -49.794]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
measured_d18O_mean | 7.67125
varray: [7.669, 7.578, 7.724, 7.714]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.927
varray: [-47.454, -49.152, -49.308, -49.794]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0666101844065
varray: [7.669, 7.578, 7.724, 7.714]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.0193468497
varray: [-47.454, -49.152, -49.308, -49.794]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.99721255039
actual_array: [ 1.93 -15.84]
varray: [7.669, 7.578, 7.724, 7.714]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 7.67125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.7524453252
actual_array: [ 18.1 -119. ]
varray: [-47.454, -49.152, -49.308, -49.794]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -48.927
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [26567.0, 27104.0, 27640.0, 28177.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1296.32335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 28.0919
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.12725
varray: [0.0, 0.0, 2.332, 2.177]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: residuals
SAMP_residuals_STD | 1.30317340238
varray: [0.0, 0.0, 2.332, 2.177]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.03025
varray: [0.0, 0.0, 0.06, 0.061]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0349320769494
varray: [0.0, 0.0, 0.06, 0.061]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 638.9625
varray: [0.0, 0.0, 1281.908, 1273.942]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 737.817510003
varray: [0.0, 0.0, 1281.908, 1273.942]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.05175
varray: [0.0, 0.0, 2.38, 1.827]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.23526201674
varray: [0.0, 0.0, 2.38, 1.827]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: slope_shift
SAMP_d18O_STD | 1.15764138805
varray: [0.0, 0.0, -1.97, -2.039]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 49.0694217504
varray: [0.0, 0.0, -84.553, -85.424]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 44.6745
varray: [0.0, 0.0, 91.993, 86.705]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 51.6308217799
varray: [0.0, 0.0, 91.993, 86.705]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.12725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.03025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 638.9625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.05175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.752
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.752
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -60.8491794738
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.9836
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, -1.97, -2.039]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, -84.553, -85.424]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.00225
varray: [0.0, 0.0, -1.97, -2.039]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
dD_mean | -42.49425
varray: [0.0, 0.0, -84.553, -85.424]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
d18O_stdev | 1.15764138805
varray: [0.0, 0.0, -1.97, -2.039]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
dD_stdev | 49.0694217504
varray: [0.0, 0.0, -84.553, -85.424]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.00225
varray: [0.0, 0.0, -1.97, -2.039]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.49425
varray: [0.0, 0.0, -84.553, -85.424]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.15764138805
varray: [0.0, 0.0, -1.97, -2.039]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
measured_dD_stdev | 49.0694217504
varray: [0.0, 0.0, -84.553, -85.424]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.76589104689
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -1.97, -2.039]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.00225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.650081994
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -84.553, -85.424]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -42.49425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [28764.0, 29300.0, 29836.0, 30373.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.30085
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -661.4656
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.6545
varray: [0.0, 0.0, 3.373, 3.245]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: residuals
SAMP_residuals_STD | 1.91116657219
varray: [0.0, 0.0, 3.373, 3.245]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.629
varray: [0.0, 0.0, -5.334, -5.182]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 3.03634187798
varray: [0.0, 0.0, -5.334, -5.182]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 890.1855
varray: [0.0, 0.0, 1791.7, 1769.042]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1027.93929622
varray: [0.0, 0.0, 1791.7, 1769.042]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -14.9335
varray: [0.0, 0.0, -30.614, -29.12]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: slope_shift
SAMP_slope_shift_STD | 17.2545038275
varray: [0.0, 0.0, -30.614, -29.12]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: slope_shift
SAMP_d18O_STD | 4.62691106463
varray: [0.0, 0.0, 8.033, 7.995]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 30.1061417544
varray: [0.0, 0.0, -51.877, -52.411]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.0685
varray: [0.0, 0.0, 112.917, 107.357]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 63.628126823
varray: [0.0, 0.0, 112.917, 107.357]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.6545
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.629
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 890.1855
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.9335
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.22475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.22475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -42.5371190414
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -55.5896
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 8.033, 7.995]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, -51.877, -52.411]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.007
varray: [0.0, 0.0, 8.033, 7.995]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
dD_mean | -26.072
varray: [0.0, 0.0, -51.877, -52.411]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
d18O_stdev | 4.62691106463
varray: [0.0, 0.0, 8.033, 7.995]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
dD_stdev | 30.1061417544
varray: [0.0, 0.0, -51.877, -52.411]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.007
varray: [0.0, 0.0, 8.033, 7.995]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.072
varray: [0.0, 0.0, -51.877, -52.411]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.62691106463
varray: [0.0, 0.0, 8.033, 7.995]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
measured_dD_stdev | 30.1061417544
varray: [0.0, 0.0, -51.877, -52.411]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.2951081685
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 8.033, 7.995]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.007
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.0712803594
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -51.877, -52.411]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -26.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [30960.0, 31496.0, 32033.0, 32570.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.7736
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -410.2426
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.2615
varray: [1.046, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: residuals
SAMP_residuals_STD | 0.523
varray: [1.046, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0825
varray: [-0.33, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.165
varray: [-0.33, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 348.75725
varray: [1395.029, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 697.5145
varray: [1395.029, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.43375
varray: [-5.735, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.8675
varray: [-5.735, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: slope_shift
SAMP_d18O_STD | 5.835
varray: [11.67, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 17.867
varray: [-35.734, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 10.55775
varray: [42.231, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 21.1155
varray: [42.231, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.2615
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 348.75725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.43375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.61775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.61775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -90.9177737258
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -100.10035
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [11.67, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-35.734, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.9175
varray: [11.67, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
dD_mean | -8.9335
varray: [-35.734, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
d18O_stdev | 5.835
varray: [11.67, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
dD_stdev | 17.867
varray: [-35.734, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.9175
varray: [11.67, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
measured_dD_mean | -8.9335
varray: [-35.734, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.835
varray: [11.67, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
measured_dD_stdev | 17.867
varray: [-35.734, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.194352836829
actual_array: [ 1.93 -15.84]
varray: [11.67, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 2.9175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.8130153948
actual_array: [ 18.1 -119. ]
varray: [-35.734, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -8.9335
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [33158.0, 33694.0, 34293.0, 34891.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.1666
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -951.67085
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [35541.0, 36140.0, 36739.0, 37337.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.48125
varray: [2.598, 2.404, 2.527, 2.396]
inst: mean_all
operation: mean_all(residuals)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: residuals
SAMP_residuals_STD | 0.0982492578428
varray: [2.598, 2.404, 2.527, 2.396]
inst: std_all
operation: std_all(residuals)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.117
varray: [0.225, -0.061, 0.254, 0.05]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149003355667
varray: [0.225, -0.061, 0.254, 0.05]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1392.76825
varray: [1404.017, 1390.448, 1393.468, 1383.14]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.66238582897
varray: [1404.017, 1390.448, 1393.468, 1383.14]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.80375
varray: [-2.143, -1.583, -1.455, -2.034]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.335888250266
varray: [-2.143, -1.583, -1.455, -2.034]
inst: std_all
operation: std_all(slope_shift)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0222766694099
varray: [4.19, 4.169, 4.171, 4.217]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.639559744408
varray: [-48.346, -49.544, -49.707, -49.603]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 96.22825
varray: [99.927, 94.161, 97.574, 93.251]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.08919270738
varray: [99.927, 94.161, 97.574, 93.251]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.48125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.117
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1392.76825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.80375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.398
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.398
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -13.8230441955
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -14.42985
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [4.19, 4.169, 4.171, 4.217]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-48.346, -49.544, -49.707, -49.603]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.18675
varray: [4.19, 4.169, 4.171, 4.217]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
dD_mean | -49.3
varray: [-48.346, -49.544, -49.707, -49.603]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0222766694099
varray: [4.19, 4.169, 4.171, 4.217]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
dD_stdev | 0.639559744408
varray: [-48.346, -49.544, -49.707, -49.603]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.18675
varray: [4.19, 4.169, 4.171, 4.217]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.3
varray: [-48.346, -49.544, -49.707, -49.603]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0222766694099
varray: [4.19, 4.169, 4.171, 4.217]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.639559744408
varray: [-48.346, -49.544, -49.707, -49.603]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.47671511744
actual_array: [ 1.93 -15.84]
varray: [4.19, 4.169, 4.171, 4.217]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.18675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.1062880087
actual_array: [ 18.1 -119. ]
varray: [-48.346, -49.544, -49.707, -49.603]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.3
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [37988.0, 38525.0, 39062.0, 39598.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.94685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 92.34015
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.308
varray: [2.405, 2.375, 2.37, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: residuals
SAMP_residuals_STD | 0.151457364738
varray: [2.405, 2.375, 2.37, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.24025
varray: [0.227, 0.304, 0.163, 0.267]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0603400088388
varray: [0.227, 0.304, 0.163, 0.267]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1277.46375
varray: [1282.981, 1278.79, 1280.407, 1267.677]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.7488809628
varray: [1282.981, 1278.79, 1280.407, 1267.677]
inst: std_all
operation: std_all(baseline_shift)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.59025
varray: [2.582, 2.738, 2.618, 2.423]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.129924016256
varray: [2.582, 2.738, 2.618, 2.423]
inst: std_all
operation: std_all(slope_shift)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: slope_shift
SAMP_d18O_STD | 0.573306200908
varray: [-6.244, -7.15, -7.366, -7.528]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.17629853779
varray: [-59.697, -60.991, -61.434, -62.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 89.84775
varray: [93.994, 91.142, 91.911, 82.344]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.14553971092
varray: [93.994, 91.142, 91.911, 82.344]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.308
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.24025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1277.46375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.59025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.57125
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.57125
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -19.8402361726
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -20.81035
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [-6.244, -7.15, -7.366, -7.528]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-59.697, -60.991, -61.434, -62.541]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.072
varray: [-6.244, -7.15, -7.366, -7.528]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
dD_mean | -61.16575
varray: [-59.697, -60.991, -61.434, -62.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
d18O_stdev | 0.573306200908
varray: [-6.244, -7.15, -7.366, -7.528]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
dD_stdev | 1.17629853779
varray: [-59.697, -60.991, -61.434, -62.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.072
varray: [-6.244, -7.15, -7.366, -7.528]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.16575
varray: [-59.697, -60.991, -61.434, -62.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.573306200908
varray: [-6.244, -7.15, -7.366, -7.528]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.17629853779
varray: [-59.697, -60.991, -61.434, -62.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.89834600274
actual_array: [ 1.93 -15.84]
varray: [-6.244, -7.15, -7.366, -7.528]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.072
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.3626119288
actual_array: [ 18.1 -119. ]
varray: [-59.697, -60.991, -61.434, -62.541]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -61.16575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [40187.0, 40723.0, 41259.0, 41795.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.1201
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -22.96435
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.2015
varray: [0.0, 2.403, 0.0, 2.403]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: residuals
SAMP_residuals_STD | 1.38737269686
varray: [0.0, 2.403, 0.0, 2.403]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.5815
varray: [0.0, -1.163, 0.0, -1.163]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.671458363068
varray: [0.0, -1.163, 0.0, -1.163]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 901.1595
varray: [0.0, 1802.319, 0.0, 1802.319]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1040.56935982
varray: [0.0, 1802.319, 0.0, 1802.319]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -9.695
varray: [0.0, -19.39, 0.0, -19.39]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: slope_shift
SAMP_slope_shift_STD | 11.1948217196
varray: [0.0, -19.39, 0.0, -19.39]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: slope_shift
SAMP_d18O_STD | 0.471695169928
varray: [0.0, -0.817, 0.0, -0.817]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 24.1655728672
varray: [0.0, -41.856, 0.0, -41.856]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 45.9935
varray: [0.0, 91.987, 0.0, 91.987]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 53.1087192119
varray: [0.0, 91.987, 0.0, 91.987]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.2015
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5815
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 901.1595
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.695
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.67775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.67775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -58.2703829122
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -64.6646
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, -0.817, 0.0, -0.817]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, -41.856, 0.0, -41.856]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.4085
varray: [0.0, -0.817, 0.0, -0.817]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
dD_mean | -20.928
varray: [0.0, -41.856, 0.0, -41.856]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
d18O_stdev | 0.471695169928
varray: [0.0, -0.817, 0.0, -0.817]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
dD_stdev | 24.1655728672
varray: [0.0, -41.856, 0.0, -41.856]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.4085
varray: [0.0, -0.817, 0.0, -0.817]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.928
varray: [0.0, -41.856, 0.0, -41.856]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.471695169928
varray: [0.0, -0.817, 0.0, -0.817]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
measured_dD_stdev | 24.1655728672
varray: [0.0, -41.856, 0.0, -41.856]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.16600717522
actual_array: [ 1.93 -15.84]
varray: [0.0, -0.817, 0.0, -0.817]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -0.4085
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.1914767023
actual_array: [ 18.1 -119. ]
varray: [0.0, -41.856, 0.0, -41.856]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -20.928
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [42385.0, 42922.0, 43457.0, 44056.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.2266
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -399.2686
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.00825
varray: [0.0, 0.0, 2.057, 1.976]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: residuals
SAMP_residuals_STD | 1.16469634812
varray: [0.0, 0.0, 2.057, 1.976]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.02525
varray: [0.0, 0.0, -0.059, 0.16]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0940403282285
varray: [0.0, 0.0, -0.059, 0.16]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 634.408
varray: [0.0, 0.0, 1271.803, 1265.829]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 732.555319002
varray: [0.0, 0.0, 1271.803, 1265.829]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.87825
varray: [0.0, 0.0, 1.66, 1.853]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.01717201921
varray: [0.0, 0.0, 1.66, 1.853]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: slope_shift
SAMP_d18O_STD | 0.415657310774
varray: [0.0, 0.0, -0.693, -0.745]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 26.4427839684
varray: [0.0, 0.0, -45.694, -45.906]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 39.91975
varray: [0.0, 0.0, 81.964, 77.715]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 46.1279841338
varray: [0.0, 0.0, 81.964, 77.715]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.00825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.02525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 634.408
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.87825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.871
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.871
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -64.9822002258
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -70.73835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, -0.693, -0.745]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, -45.694, -45.906]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.3595
varray: [0.0, 0.0, -0.693, -0.745]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
dD_mean | -22.9
varray: [0.0, 0.0, -45.694, -45.906]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
d18O_stdev | 0.415657310774
varray: [0.0, 0.0, -0.693, -0.745]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
dD_stdev | 26.4427839684
varray: [0.0, 0.0, -45.694, -45.906]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.3595
varray: [0.0, 0.0, -0.693, -0.745]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.9
varray: [0.0, 0.0, -45.694, -45.906]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.415657310774
varray: [0.0, 0.0, -0.693, -0.745]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
measured_dD_stdev | 26.4427839684
varray: [0.0, 0.0, -45.694, -45.906]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.11650096939
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -0.693, -0.745]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -0.3595
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.0621945895
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -45.694, -45.906]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -22.9
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [44704.0, 45240.0, 45777.0, 46313.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.41985
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -666.0201
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [46900.0, 47498.0, 48097.0, 48695.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8435
varray: [3.344, 3.009, 2.573, 2.448]
inst: mean_all
operation: mean_all(residuals)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: residuals
SAMP_residuals_STD | 0.411291056715
varray: [3.344, 3.009, 2.573, 2.448]
inst: std_all
operation: std_all(residuals)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34075
varray: [0.416, 0.183, 0.49, 0.274]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.138177120634
varray: [0.416, 0.183, 0.49, 0.274]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.1835
varray: [1323.831, 1306.707, 1287.208, 1282.988]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 18.8485996739
varray: [1323.831, 1306.707, 1287.208, 1282.988]
inst: std_all
operation: std_all(baseline_shift)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.45725
varray: [2.392, 2.949, 2.331, 2.157]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.342619687117
varray: [2.392, 2.949, 2.331, 2.157]
inst: std_all
operation: std_all(slope_shift)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: slope_shift
SAMP_d18O_STD | 0.138312146972
varray: [-1.285, -1.248, -1.429, -1.549]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.65655345325
varray: [-59.05, -62.231, -64.139, -65.063]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.4785
varray: [126.598, 114.813, 100.822, 95.681]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 13.9861856725
varray: [126.598, 114.813, 100.822, 95.681]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8435
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.1835
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.45725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.24164278892
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1796
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [-1.285, -1.248, -1.429, -1.549]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-59.05, -62.231, -64.139, -65.063]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.37775
varray: [-1.285, -1.248, -1.429, -1.549]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
dD_mean | -62.62075
varray: [-59.05, -62.231, -64.139, -65.063]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
d18O_stdev | 0.138312146972
varray: [-1.285, -1.248, -1.429, -1.549]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
dD_stdev | 2.65655345325
varray: [-59.05, -62.231, -64.139, -65.063]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.37775
varray: [-1.285, -1.248, -1.429, -1.549]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.62075
varray: [-59.05, -62.231, -64.139, -65.063]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.138312146972
varray: [-1.285, -1.248, -1.429, -1.549]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.65655345325
varray: [-59.05, -62.231, -64.139, -65.063]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.14527023646
actual_array: [ 1.93 -15.84]
varray: [-1.285, -1.248, -1.429, -1.549]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.37775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.7428829865
actual_array: [ 18.1 -119. ]
varray: [-59.05, -62.231, -64.139, -65.063]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -62.62075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [49344.0, 49880.0, 50416.0, 50953.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.5846
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.2446
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.81125
varray: [0.0, 0.0, 1.563, 1.682]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: residuals
SAMP_residuals_STD | 0.938009728095
varray: [0.0, 0.0, 1.563, 1.682]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.068
varray: [0.0, 0.0, 0.037, 0.235]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.112691318802
varray: [0.0, 0.0, 0.037, 0.235]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 624.45225
varray: [0.0, 0.0, 1246.74, 1251.069]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 721.057515099
varray: [0.0, 0.0, 1246.74, 1251.069]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.17625
varray: [0.0, 0.0, 2.309, 2.396]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.35868082467
varray: [0.0, 0.0, 2.309, 2.396]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: slope_shift
SAMP_d18O_STD | 5.46384006446
varray: [0.0, 0.0, -9.426, -9.501]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 63.9336679588
varray: [0.0, 0.0, -110.536, -110.936]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 32.61
varray: [0.0, 0.0, 62.734, 67.706]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 37.7094541285
varray: [0.0, 0.0, 62.734, 67.706]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.81125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.068
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 624.45225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.17625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.068
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.068
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -71.8242597899
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -78.0481
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, -9.426, -9.501]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, -110.536, -110.936]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.73175
varray: [0.0, 0.0, -9.426, -9.501]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
dD_mean | -55.368
varray: [0.0, 0.0, -110.536, -110.936]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
d18O_stdev | 5.46384006446
varray: [0.0, 0.0, -9.426, -9.501]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
dD_stdev | 63.9336679588
varray: [0.0, 0.0, -110.536, -110.936]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.73175
varray: [0.0, 0.0, -9.426, -9.501]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.368
varray: [0.0, 0.0, -110.536, -110.936]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.46384006446
varray: [0.0, 0.0, -9.426, -9.501]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
measured_dD_stdev | 63.9336679588
varray: [0.0, 0.0, -110.536, -110.936]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.53391950899
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -9.426, -9.501]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -4.73175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.8626349355
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -110.536, -110.936]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -55.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [51540.0, 52076.0, 52612.0, 53147.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.61685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -675.97585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.8765
varray: [0.0, 0.0, 1.753, 1.753]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: residuals
SAMP_residuals_STD | 1.01209502189
varray: [0.0, 0.0, 1.753, 1.753]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.078
varray: [0.0, 0.0, -0.156, -0.156]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0900666419936
varray: [0.0, 0.0, -0.156, -0.156]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 628.15
varray: [0.0, 0.0, 1256.3, 1256.3]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 725.325143183
varray: [0.0, 0.0, 1256.3, 1256.3]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.155
varray: [0.0, 0.0, 2.31, 2.31]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.33367912183
varray: [0.0, 0.0, 2.31, 2.31]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: slope_shift
SAMP_d18O_STD | 3.38558197853
varray: [0.0, 0.0, -5.864, -5.864]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 49.4615975615
varray: [0.0, 0.0, -85.67, -85.67]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 35.2375
varray: [0.0, 0.0, 70.475, 70.475]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 40.6887602211
varray: [0.0, 0.0, 70.475, 70.475]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.8765
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.078
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 628.15
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.155
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.00275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.00275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -69.5580446297
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -75.4206
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, -5.864, -5.864]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, -85.67, -85.67]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.932
varray: [0.0, 0.0, -5.864, -5.864]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
dD_mean | -42.835
varray: [0.0, 0.0, -85.67, -85.67]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
d18O_stdev | 3.38558197853
varray: [0.0, 0.0, -5.864, -5.864]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
dD_stdev | 49.4615975615
varray: [0.0, 0.0, -85.67, -85.67]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.932
varray: [0.0, 0.0, -5.864, -5.864]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.835
varray: [0.0, 0.0, -85.67, -85.67]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.38558197853
varray: [0.0, 0.0, -5.864, -5.864]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
measured_dD_stdev | 49.4615975615
varray: [0.0, 0.0, -85.67, -85.67]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.71557677547
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -5.864, -5.864]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -2.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.9733310407
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -85.67, -85.67]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -42.835
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [53735.0, 54272.0, 54870.0, 55406.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.5516
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -672.2781
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [56056.0, 56654.0, 57252.0, 57851.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [58502.0, 59100.0, 59699.0, 60298.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.7125
varray: [2.326, 2.0, 1.455, 1.069]
inst: mean_all
operation: mean_all(residuals)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: residuals
SAMP_residuals_STD | 0.559594198207
varray: [2.326, 2.0, 1.455, 1.069]
inst: std_all
operation: std_all(residuals)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2315
varray: [0.404, 0.244, 0.065, 0.213]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.139016785557
varray: [0.404, 0.244, 0.065, 0.213]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1252.3835
varray: [1278.289, 1264.667, 1241.353, 1225.225]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 23.6730180797
varray: [1278.289, 1264.667, 1241.353, 1225.225]
inst: std_all
operation: std_all(baseline_shift)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.26375
varray: [2.887, 2.528, 2.059, 1.581]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.567549924383
varray: [2.887, 2.528, 2.059, 1.581]
inst: std_all
operation: std_all(slope_shift)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: slope_shift
SAMP_d18O_STD | 0.183874051097
varray: [-7.749, -7.732, -7.943, -8.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.21271005053
varray: [-67.837, -65.516, -67.203, -68.282]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 68.6415
varray: [91.75, 79.842, 58.427, 44.547]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 21.1686475162
varray: [91.75, 79.842, 58.427, 44.547]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.7125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2315
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1252.3835
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.26375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.16675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.16675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -40.5227055657
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -42.0166
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [-7.749, -7.732, -7.943, -8.122]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [-67.837, -65.516, -67.203, -68.282]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.8865
varray: [-7.749, -7.732, -7.943, -8.122]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
dD_mean | -67.2095
varray: [-67.837, -65.516, -67.203, -68.282]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
d18O_stdev | 0.183874051097
varray: [-7.749, -7.732, -7.943, -8.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
dD_stdev | 1.21271005053
varray: [-67.837, -65.516, -67.203, -68.282]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.8865
varray: [-7.749, -7.732, -7.943, -8.122]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.2095
varray: [-67.837, -65.516, -67.203, -68.282]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.183874051097
varray: [-7.749, -7.732, -7.943, -8.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.21271005053
varray: [-67.837, -65.516, -67.203, -68.282]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.7212603833
actual_array: [ 1.93 -15.84]
varray: [-7.749, -7.732, -7.943, -8.122]
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.8865
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.095954338
actual_array: [ 18.1 -119. ]
varray: [-67.837, -65.516, -67.203, -68.282]
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -67.2095
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [60944.0, 61480.0, 62016.0, 62552.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.7156
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -48.0446
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [63137.0, 63736.0, 64334.0, 64932.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [65581.0, 66179.0, 66777.0, 67375.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [68023.0, 68622.0, 69220.0, 69818.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [70467.0, 71065.0, 71663.0, 72261.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [72910.0, 73509.0, 74107.0, 74705.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [75356.0, 75954.0, 76551.0, 77150.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.87925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.87925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.6581
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.9529
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1657
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.72513444007
except: True
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 18.2770520058
except: True
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: [77796.0, 78394.0, 78992.0, 79590.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.4281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.29
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.425
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1185981021
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.29
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.711267019
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.256
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.427
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0842468573
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.256
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7131642988
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.427
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.939
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.182
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.76397201549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -6.939
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5321076205
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -58.182
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.955
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78013730719
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -6.955
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6383552895
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -58.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.251
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.278
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0791952036
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.251
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5718169533
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.278
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.349
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1782076153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.349
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8213092476
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.4
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.786
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2297344826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.4
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0537260237
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.786
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.399
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2287241519
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.399
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0233695468
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.394
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.747
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2236724982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.394
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0167290675
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.747
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.317
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.572
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1458770319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.317
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8507170846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.572
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.556
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.839449008716
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 3.556
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.08655782102
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 23.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.437
operation: col(d(18_16)Mean)
inst: col
dD_actual | 34.658
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.72955038293
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.437
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.5396209038
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 34.658
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.636
operation: col(d(18_16)Mean)
inst: col
dD_actual | 37.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.93060619844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.0013414502
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 37.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.689
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.378
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.98415372719
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.689
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.0685613402
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 38.378
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.763
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.257
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0589182013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.763
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.9024158143
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 39.257
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.771
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.53
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06700084715
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.771
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1613945076
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 39.53
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.847
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.897
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.14378598273
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.847
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.5095453517
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 39.897
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.068
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.19026119636
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.671762775
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 40.068
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.335
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.15085829785
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.9250496289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 40.335
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.908
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.555
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6354
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.357
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.908
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.555
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.20541615733
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.1337504074
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 40.555
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.444
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.202
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.3155119597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -11.444
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.061637584
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -87.202
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.028
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5753943815
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -12.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.074812676
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -103.028
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.962
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.254
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8491940097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -12.962
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.083764903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -107.254
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.924
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.990640312
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.66799354
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -108.924
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.175
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.825
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0643944554
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.522718092
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -109.825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.229
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.647
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1189523149
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.229
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.302500092
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -110.647
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.982
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1290556222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.620294459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -110.982
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.163
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.758
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0522704866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.163
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.407799121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -110.758
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.281
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.451
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1714895129
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.281
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.065206573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -111.451
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.243
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.586
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1330969451
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -13.243
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.19327296
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -111.586
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.101
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.55
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9379763252
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -8.101
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.3676061269
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -68.55
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.608
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4398832747
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.608
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8635974141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -62.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.511
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3923977303
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.690129855
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -61.511
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3368295401
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.214861264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -61.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.452
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.405
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2822716806
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.452
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6409341232
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.405
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.441
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.315
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2711580426
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.441
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5555565319
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.315
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.463
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.22
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2933853186
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.463
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4654357412
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.22
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.423
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.373
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2529720894
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.423
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6105776463
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.373
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.437
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.242
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2671167196
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.437
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4863058191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.508
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3388502016
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.508
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4255928653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.669
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.99493930624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 7.669
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.3550987492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -47.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.578
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.152
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.9029992097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 7.578
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.9658893032
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.152
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.308
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.05050749646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 7.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.113877128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.714
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.794
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.04040418915
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 7.714
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.5749161204
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.794
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.97
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.553
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.74363861203
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.97
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.5486904825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -84.553
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.424
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.81335143249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -2.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.3749558374
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -85.424
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.033
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.36269969241
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 8.033
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.5509330368
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -51.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.995
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.411
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.32430712462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 7.995
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.0575067446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -52.411
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 11.67
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.734
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.03727256187
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 11.67
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.2370390949
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -35.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.19
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.47999869231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.19
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.201285542
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -48.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.544
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.45878174696
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3377561449
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.544
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.171
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.707
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.46080240842
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.171
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.4923844489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.707
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.217
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.603
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.50727762206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 4.217
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3937258991
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -49.603
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.244
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.697
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.06179215729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -6.244
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9692970724
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.697
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.15
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.991
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.97715179978
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.15
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1968371059
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -60.991
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.366
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.434
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1953832377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.366
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6170845826
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -61.434
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.528
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3590568162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.528
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.6672289543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -62.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.817
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.856
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.57872727893
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -0.817
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.0446125764
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -41.856
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.817
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.856
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.57872727893
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -0.817
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.0446125764
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -41.856
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.693
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.694
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.45344626826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -0.693
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.6854925212
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -45.694
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.745
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.906
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.50598346628
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -0.745
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.8866041805
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -45.906
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.285
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.05
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.05156206114
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3555270556
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -59.05
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.248
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.01417982409
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.248
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.3731505846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -62.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.429
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.139
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.19704968644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.429
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.1831555181
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -64.139
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.549
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.063
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.31828937419
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -1.549
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.0596987878
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -65.063
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.536
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.276664544
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -9.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.197201063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -110.536
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.501
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.936
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3524393489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -9.501
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.576657023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -110.936
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.67786647942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -5.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.6083212533
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -85.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.67786647942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -5.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.6083212533
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -85.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.749
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.837
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5823399078
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.749
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.6912258766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -67.837
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.732
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.516
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5651642853
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.732
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.4894326635
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -65.516
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.203
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7783440696
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -7.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.0897881786
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -67.203
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.282
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9591932705
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: -8.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1133706331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: -68.282
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 09:32:36
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6354
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.357
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.243
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.586
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.9529
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1657
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.75328707152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.72513444007
measure_array: [ 4.6354 -12.9529]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.989774901519
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3383408281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.2770520058
measure_array: [ 37.357 -107.1657]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05414077316
tarray: []
source_col: d(D_H)Mean