## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
210331_20-202_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4923 | -76.1213 |  | 9.5783 |  | 0.25686 | 2.9648
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27274 | 3.1339
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.36777 | 4.4376
4 |  | EV | -10.2963 | -74.2451 |  |  |  | 0.10313 | 1.371
5 |  | CPER2-2030-3 | -3.996 | -85.0224 |  |  |  | 0.15595 | 0.52138
6 |  | CPER2-010-2 | -2.5665 | -83.6727 |  |  |  | 0.052959 | 0.14384
7 |  | CPER2-A-5 | -1.3261 | -76.4596 |  | 18.977 |  | 0.032715 | 0.4638 |  | -2.0871
8 |  | CPER2-E-5 | -0.47915 | -71.1262 |  | 26.3302 |  | 0.01372 | 0.15121 |  | -0.82555
9 |  | CPER2-D-1 | -0.74197 | -80.2918 | -4.3812 | 215.8012 | 21.4725 | 0.097373 | 0.67368 |  | -1.2723
10 | * | CPER2-E-2 | -1.5411 | -35.6403 |  |  |  | 0.77673 | 37.1805
11 |  | CPER2-D-3 | 1.684 | -66.8872 | -4.777 | 14.701 | 25.3053 | 0.006702 | 0.12149 |  | 2.2084
12 |  | CPER2-C-4 | -4.6761 | -84.8942 |  | 124.3055 |  | 0.075363 | 0.50858 |  | -3.5143
13 | * | EV | -0.84674 | -0.9506 |  |  |  | 0 | 0
14 |  | CPER2-C-1 | -5.2485 | -82.6503 |  | 5.4485 |  | 0.11441 | 0.44081
15 |  | CPER2-010-3 | -0.68765 | -76.6842 |  |  |  | 0.10441 | 0.25927
16 | * | CPER2-D-2 | 11.5648 | -20.6621 | -7.787 | 255.7702 | 53.4122 | 13.8504 | 21.126 |  | -28.4086
17 | * | CPER2-C-3 | -3.6439 | -43.1643 |  | 69.1017 |  | 3.1215 | 45.2426 |  | -32.0966
18 |  | CPER2-2030-1 | -3.0937 | -82.9183 |  |  |  | 0.089571 | 0.14489
19 |  | CPER2-E-4 | -1.3597 | -72.1012 |  | 17.9625 |  | 0.055632 | 0.60636 |  | -1.3501
20 |  | CPER2-A-1 | -0.96858 | -70.8595 |  | 19.6017 |  | 0.042625 | 0.13625 |  | -1.1666
21 | * | CPER2-B-3 | -0.84674 | -0.9506 |  |  |  | 0 | 0
22 |  | EV | -9.8796 | -73.0383 |  |  |  | 0.11147 | 0.19592
23 | * | CPER2-B-1 | -0.84674 | -0.9506 |  |  |  | 0 | 0
24 | * | CPER2-B-4 | -0.84674 | -0.9506 |  |  |  | 0 | 0
25 | * | CPER2-A-4 | -1.7195 | -37.7541 |  |  |  | 0.98077 | 39.4443
26 | * | CPER2-E-1 | -0.84674 | -0.9506 |  |  |  | 0 | 0
27 |  | CPER2-D-5 | 0.63295 | -58.8568 |  |  |  | 0.13566 | 0.71907
28 |  | CPER2-1020-3 | 0.55068 | -71.1967 |  |  |  | 0.019975 | 0.63146
29 |  | CPER2-A-3 | -2.7331 | -75.1055 |  | 7.8785 |  | 0.06343 | 0.13578
30 |  | CPER2-B-5 | -1.004 | -74.5294 |  | 10.052 |  | 0.043213 | 0.082168 |  | -0.4338
31 | * | EV | -5.4855 | -37.1618 |  |  |  | 5.1766 | 38.8095
32 |  | CPER2-3040-2 | -2.1461 | -79.6171 |  |  |  | 0.2017 | 0.29482
33 |  | CPER2-C-5 | -7.1649 | -90.1821 |  | 40.9845 |  | 0.158 | 0.60883 |  | -1.9716
34 |  | CPER2-1020-1 | -8.1919 | -96.9085 |  |  |  | 0.045989 | 0.3564
35 |  | CPER2-3040-1 | -6.3244 | -91.3758 |  |  |  | 0.038767 | 0.16023
36 |  | CPER2-deep-2 | -9.6685 | -97.107 |  |  |  | 0.11905 | 0.28681
37 |  | CPER2-B-2 | -4.1967 | -87.4064 |  | 10.6112 |  | 0.15258 | 0.52656 |  | -1.2928
38 |  | CPER2-1020-2 | -2.4979 | -83.1286 |  |  |  | 0.050763 | 0.15926
39 |  | CPER2-010-1 | -1.7969 | -80.8785 |  |  |  | 0.044784 | 0.076352
40 | * | EV | -0.84674 | -0.9506 |  |  |  | 0 | 0
41 | * | CPER2-2030-2 | -0.84674 | -0.9506 |  |  |  | 0 | 0
42 | * | CPER2-A-2 | -0.84674 | -0.9506 |  |  |  | 0 | 0
43 | * | CPER2-C-2 | -2.3694 | -41.7731 |  |  |  | 1.6991 | 43.7516
44 |  | CPER2-D-4 | 1.9155 | -76.281 | -5.8803 | 16.606 | 40.919 | 0.013675 | 0.090044 |  | 2.5792
45 |  | CPER2-E-3 | -1.4805 | -74.2135 |  | 41.9425 |  | 0.078462 | 0.50205 |  | -1.6446
46 |  | EV | -10.1243 | -73.3731 |  |  |  | 0.032332 | 0.017321

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -10.009 | -77.838 | -11.2035 | -84.8129 | -1.193 | 137.42 | 2021/03/31 10:50:07
1 |  | EV | -9.441 | -70.953 | -10.6158 | -77.395 | -1.623 | 105.312 | 2021/03/31 10:59:03
1 |  | EV | -9.335 | -69.587 | -10.5061 | -75.9233 | -2.101 | 103.618 | 2021/03/31 11:08:00
1 |  | EV | -9.271 | -69.119 | -10.4399 | -75.4191 | -1.936 | 101.852 | 2021/03/31 11:16:56
1 |  | EV | -9.229 | -68.795 | -10.3964 | -75.07 | -2.245 | 100.493 | 2021/03/31 11:25:53
1 |  | EV | -9.239 | -68.578 | -10.4068 | -74.8362 | -2.089 | 99.153 | 2021/03/31 11:34:49
1 |  | EV | -9.189 | -68.253 | -10.355 | -74.486 | -2.055 | 97.851 | 2021/03/31 11:43:46
1 |  | EV | -9.18 | -68.131 | -10.3457 | -74.3546 | -2.263 | 97.394 | 2021/03/31 11:52:43
1 |  | EV | -9.174 | -68.373 | -10.3395 | -74.6153 | -1.973 | 96.589 | 2021/03/31 12:01:39
1 |  | EV | -9.15 | -68.081 | -10.3147 | -74.3007 | -2.025 | 96.479 | 2021/03/31 12:10:36
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4923 | -76.1213 |  | 9.5783 |  | 0.25686 | 2.9648
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 1.963 | 9.404 | 1.1845 | 9.1812 | -2.141 | 95.549 | 2021/03/31 12:20:26
2 |  | PZ | 2.518 | 15.869 | 1.7587 | 16.1466 | -2.26 | 94.806 | 2021/03/31 12:29:23
2 |  | PZ | 2.698 | 17.621 | 1.945 | 18.0342 | -2.474 | 94.87 | 2021/03/31 12:38:19
2 |  | PZ | 2.744 | 18.154 | 1.9926 | 18.6084 | -2.392 | 94.531 | 2021/03/31 12:47:15
2 |  | PZ | 2.738 | 18.71 | 1.9864 | 19.2075 | -2.19 | 94.097 | 2021/03/31 12:56:13
2 |  | PZ | 2.802 | 19.038 | 2.0526 | 19.5608 | -2.147 | 94.04 | 2021/03/31 13:05:09
2 |  | PZ | 2.833 | 19.205 | 2.0847 | 19.7408 | -2.178 | 93.877 | 2021/03/31 13:14:06
2 |  | PZ | 2.835 | 19.564 | 2.0868 | 20.1275 | -2.124 | 93.923 | 2021/03/31 13:23:03
2 |  | PZ | 2.849 | 19.522 | 2.1013 | 20.0823 | -2.107 | 93.639 | 2021/03/31 13:31:59
2 |  | PZ | 2.855 | 19.734 | 2.1075 | 20.3107 | -2.032 | 93.704 | 2021/03/31 13:40:55
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27274 | 3.1339
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.569 | -97.88 | -14.8872 | -106.406 | -2.101 | 94.25 | 2021/03/31 13:50:43
3 |  | UT2 | -14.218 | -107.132 | -15.5588 | -116.3741 | -2.085 | 94.085 | 2021/03/31 13:59:40
3 |  | UT2 | -14.344 | -108.714 | -15.6891 | -118.0785 | -2.159 | 94.217 | 2021/03/31 14:08:37
3 |  | UT2 | -14.583 | -110.752 | -15.9364 | -120.2742 | -2.014 | 93.221 | 2021/03/31 14:17:34
3 |  | UT2 | -14.631 | -111.18 | -15.9861 | -120.7354 | -2.151 | 93.646 | 2021/03/31 14:26:31
3 |  | UT2 | -14.693 | -111.765 | -16.0503 | -121.3656 | -2.015 | 93.285 | 2021/03/31 14:35:27
3 |  | UT2 | -14.723 | -111.748 | -16.0813 | -121.3473 | -2.102 | 93.808 | 2021/03/31 14:44:23
3 |  | UT2 | -14.743 | -112.146 | -16.102 | -121.7761 | -2.018 | 93.771 | 2021/03/31 14:53:20
3 |  | UT2 | -14.697 | -112.188 | -16.0544 | -121.8214 | -2.058 | 93.718 | 2021/03/31 15:02:14
3 |  | UT2 | -14.697 | -112.188 | -16.0544 | -121.8214 | -2.058 | 93.718 | 2021/03/31 15:12:13
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.36777 | 4.4376
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.409 | -71.772 | -10.5827 | -78.2774 | -1.898 | 93.542 | 2021/03/31 15:31:01
4 |  | EV | -9.191 | -68.513 | -10.3571 | -74.7662 | -2.296 | 93.65 | 2021/03/31 15:39:57
4 |  | EV | -9.1 | -67.993 | -10.2629 | -74.2059 | -2.204 | 93.424 | 2021/03/31 15:48:53
4 |  | EV | -9.093 | -67.716 | -10.2557 | -73.9075 | -2.223 | 93.728 | 2021/03/31 15:57:48
4 |  | EV | -9.081 | -67.449 | -10.2433 | -73.6198 | -2.221 | 93.667 | 2021/03/31 16:06:46
4 |  | EV | -9.097 | -67.364 | -10.2598 | -73.5282 | -2.107 | 93.428 | 2021/03/31 16:15:42
4 |  | EV | -9.116 | -67.225 | -10.2795 | -73.3785 | -2.03 | 93.549 | 2021/03/31 16:24:38
4 |  | EV | -9.094 | -67.576 | -10.2567 | -73.7566 | -2.265 | 93.47 | 2021/03/31 16:33:35
4 |  | EV | -9.071 | -67.366 | -10.2329 | -73.5304 | -2.055 | 93.659 | 2021/03/31 16:42:32
4 |  | EV | -9.071 | -67.32 | -10.2329 | -73.4808 | -1.999 | 93.485 | 2021/03/31 16:51:29
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.2963 | -74.2451 |  |  |  | 0.10313 | 1.371
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | CPER2-2030-3 | -3.275 | -77.274 | -4.2355 | -84.2052 | -2.197 | 94.939 | 2021/03/31 17:01:20
5 |  | CPER2-2030-3 | -2.986 | -78.11 | -3.9365 | -85.1059 | -2.19 | 94.874 | 2021/03/31 17:10:16
5 |  | CPER2-2030-3 | -2.935 | -78.339 | -3.8837 | -85.3526 | -2.149 | 94.617 | 2021/03/31 17:19:12
5 |  | CPER2-2030-3 | -2.978 | -78.407 | -3.9282 | -85.4259 | -2.334 | 94.601 | 2021/03/31 17:28:09
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | CPER2-2030-3 | -3.996 | -85.0224 |  |  |  | 0.15595 | 0.52138
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | CPER2-010-2 | -1.739 | -76.889 | -2.6462 | -83.7904 | -1.992 | 94.127 | 2021/03/31 17:37:59
6 |  | CPER2-010-2 | -1.647 | -76.67 | -2.551 | -83.5545 | -2.035 | 93.889 | 2021/03/31 17:46:57
6 |  | CPER2-010-2 | -1.644 | -76.642 | -2.5479 | -83.5243 | -2.163 | 93.902 | 2021/03/31 17:55:53
6 |  | CPER2-010-2 | -1.618 | -76.918 | -2.521 | -83.8217 | -2.289 | 94.104 | 2021/03/31 18:04:51
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | CPER2-010-2 | -2.5665 | -83.6727 |  |  |  | 0.052959 | 0.14384
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | CPER2-A-5 | -0.512 | -70.769 | -1.3765 | -77.1968 | -2.272 | 113.084 | 2021/03/31 18:14:41
7 |  | CPER2-A-5 | -0.446 | -69.933 | -1.3082 | -76.2961 | -2.144 | 112.86 | 2021/03/31 18:23:39
7 |  | CPER2-A-5 | -0.443 | -69.74 | -1.3051 | -76.0881 | -2.145 | 113.182 | 2021/03/31 18:32:35
7 |  | CPER2-A-5 | -0.452 | -69.897 | -1.3144 | -76.2573 | -2.341 | 112.933 | 2021/03/31 18:41:32
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | CPER2-A-5 | -1.3261 | -76.4596 |  | 18.977 |  | 0.032715 | 0.4638 |  | -2.0871
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | CPER2-E-5 | 0.338 | -65.338 | -0.497 | -71.3454 | -1.912 | 119.677 | 2021/03/31 18:51:20
8 |  | CPER2-E-5 | 0.359 | -65.124 | -0.47527 | -71.1149 | -2.002 | 120.374 | 2021/03/31 19:00:15
8 |  | CPER2-E-5 | 0.371 | -65.103 | -0.46285 | -71.0922 | -2.153 | 120.702 | 2021/03/31 19:09:12
8 |  | CPER2-E-5 | 0.353 | -64.973 | -0.48148 | -70.9522 | -2.255 | 120.719 | 2021/03/31 19:18:08
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | CPER2-E-5 | -0.47915 | -71.1262 |  | 26.3302 |  | 0.01372 | 0.15121 |  | -0.82555
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | CPER2-D-1 | 0.247 | -72.67 | -0.59116 | -79.2449 | -4.142 | 308.872 | 2021/03/31 19:27:56
9 |  | CPER2-D-1 | 0.059 | -73.755 | -0.78569 | -80.4139 | -4.61 | 308.674 | 2021/03/31 19:36:53
9 |  | CPER2-D-1 | 0.044 | -73.937 | -0.80121 | -80.6099 | -4.42 | 310.03 | 2021/03/31 19:45:49
9 |  | CPER2-D-1 | 0.055 | -74.205 | -0.78983 | -80.8987 | -4.353 | 311.78 | 2021/03/31 19:54:45
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | CPER2-D-1 | -0.74197 | -80.2918 | -4.3812 | 215.8012 | 21.4725 | 0.097373 | 0.67368 |  | -1.2723
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | CPER2-E-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 20:04:33
10 |  | CPER2-E-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 20:14:32
10 |  | CPER2-E-2 | -1.409 | -64.831 | -2.3047 | -70.7992 | -2.102 | 121.343 | 2021/03/31 20:24:31
10 |  | CPER2-E-2 | -1.275 | -63.96 | -2.166 | -69.8608 | -2.269 | 119.891 | 2021/03/31 20:33:26
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 | * | CPER2-E-2 | -1.5411 | -35.6403 |  |  |  | 0.77673 | 37.1805
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | CPER2-D-3 | 2.439 | -61.022 | 1.677 | -66.6954 | -4.918 | 109.167 | 2021/03/31 20:43:15
11 |  | CPER2-D-3 | 2.451 | -61.238 | 1.6894 | -66.9281 | -4.639 | 108.642 | 2021/03/31 20:52:11
11 |  | CPER2-D-3 | 2.441 | -61.296 | 1.6791 | -66.9906 | -4.692 | 108.648 | 2021/03/31 21:01:07
11 |  | CPER2-D-3 | 2.452 | -61.244 | 1.6905 | -66.9346 | -4.859 | 108.498 | 2021/03/31 21:10:04
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | CPER2-D-3 | 1.684 | -66.8872 | -4.777 | 14.701 | 25.3053 | 0.006702 | 0.12149 |  | 2.2084
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | CPER2-C-4 | -3.6 | -77.166 | -4.5718 | -84.0889 | -2.328 | 217.578 | 2021/03/31 21:19:53
12 |  | CPER2-C-4 | -3.696 | -78.083 | -4.6712 | -85.0768 | -2.343 | 218.055 | 2021/03/31 21:28:50
12 |  | CPER2-C-4 | -3.728 | -78.099 | -4.7043 | -85.0941 | -2.506 | 218.5 | 2021/03/31 21:37:46
12 |  | CPER2-C-4 | -3.779 | -78.306 | -4.757 | -85.3171 | -2.238 | 219.24 | 2021/03/31 21:46:43
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | CPER2-C-4 | -4.6761 | -84.8942 |  | 124.3055 |  | 0.075363 | 0.50858 |  | -3.5143
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 21:56:33
13 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 22:05:30
13 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 22:14:26
13 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 22:23:22
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 | * | EV | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | CPER2-C-1 | -4.408 | -75.215 | -5.4079 | -81.9869 | -2.045 | 99.684 | 2021/03/31 22:33:12
14 |  | CPER2-C-1 | -4.274 | -75.962 | -5.2692 | -82.7917 | -2.167 | 99.221 | 2021/03/31 22:42:10
14 |  | CPER2-C-1 | -4.17 | -75.887 | -5.1616 | -82.7109 | -2.221 | 99.568 | 2021/03/31 22:51:07
14 |  | CPER2-C-1 | -4.164 | -76.259 | -5.1554 | -83.1117 | -1.992 | 99.472 | 2021/03/31 23:00:03
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | CPER2-C-1 | -5.2485 | -82.6503 |  | 5.4485 |  | 0.11441 | 0.44081
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | CPER2-010-3 | 0.011 | -70.672 | -0.83536 | -77.0923 | -2.193 | 94.105 | 2021/03/31 23:09:49
15 |  | CPER2-010-3 | 0.147 | -70.237 | -0.69463 | -76.6236 | -2.135 | 94.094 | 2021/03/31 23:18:46
15 |  | CPER2-010-3 | 0.205 | -70.171 | -0.63462 | -76.5525 | -2.216 | 93.999 | 2021/03/31 23:27:43
15 |  | CPER2-010-3 | 0.252 | -70.093 | -0.58598 | -76.4684 | -2.152 | 93.77 | 2021/03/31 23:36:39
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | CPER2-010-3 | -0.68765 | -76.6842 |  |  |  | 0.10441 | 0.25927
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | CPER2-D-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 23:46:27
16 |  | CPER2-D-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/03/31 23:55:23
16 |  | CPER2-D-2 | 23.958 | -36.482 | 23.9437 | -40.2561 | -15.762 | 697.54 | 2021/04/01 00:04:20
16 |  | CPER2-D-2 | 24.021 | -36.7 | 24.0089 | -40.491 | -15.386 | 701.692 | 2021/04/01 00:13:15
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 | * | CPER2-D-2 | 11.5648 | -20.6621 | -7.787 | 255.7702 | 53.4122 | 13.8504 | 21.126 |  | -28.4086
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | CPER2-C-3 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 00:23:03
17 |  | CPER2-C-3 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 00:31:59
17 |  | CPER2-C-3 | -5.377 | -78.324 | -6.4106 | -85.3365 | -2.477 | 327.385 | 2021/04/01 00:40:56
17 |  | CPER2-C-3 | -5.436 | -78.401 | -6.4716 | -85.4194 | -2.666 | 325.173 | 2021/04/01 00:49:54
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 | * | CPER2-C-3 | -3.6439 | -43.1643 |  | 69.1017 |  | 3.1215 | 45.2426 |  | -32.0966
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | CPER2-2030-1 | -2.299 | -76.197 | -3.2256 | -83.0449 | -2.29 | 96.766 | 2021/04/01 00:59:42
18 |  | CPER2-2030-1 | -2.168 | -76.051 | -3.0901 | -82.8876 | -2.052 | 95.495 | 2021/04/01 01:08:38
18 |  | CPER2-2030-1 | -2.102 | -76.184 | -3.0218 | -83.0309 | -2.304 | 94.612 | 2021/04/01 01:17:34
18 |  | CPER2-2030-1 | -2.117 | -75.886 | -3.0373 | -82.7098 | -2.012 | 93.863 | 2021/04/01 01:26:31
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | CPER2-2030-1 | -3.0937 | -82.9183 |  |  |  | 0.089571 | 0.14489
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | CPER2-E-4 | -0.576 | -66.921 | -1.4428 | -73.0509 | -2.388 | 112.249 | 2021/04/01 01:36:19
19 |  | CPER2-E-4 | -0.486 | -65.887 | -1.3496 | -71.9369 | -2.158 | 111.966 | 2021/04/01 01:45:15
19 |  | CPER2-E-4 | -0.472 | -65.811 | -1.3351 | -71.855 | -2.326 | 112.062 | 2021/04/01 01:54:11
19 |  | CPER2-E-4 | -0.449 | -65.539 | -1.3113 | -71.562 | -2.177 | 111.724 | 2021/04/01 02:03:07
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | CPER2-E-4 | -1.3597 | -72.1012 |  | 17.9625 |  | 0.055632 | 0.60636 |  | -1.3501
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | CPER2-A-1 | -0.135 | -65.065 | -0.98643 | -71.0513 | -2.22 | 114.045 | 2021/04/01 02:12:57
20 |  | CPER2-A-1 | -0.17 | -64.733 | -1.0226 | -70.6936 | -2.237 | 113.653 | 2021/04/01 02:21:53
20 |  | CPER2-A-1 | -0.081 | -64.873 | -0.93055 | -70.8444 | -2.166 | 113.599 | 2021/04/01 02:30:50
20 |  | CPER2-A-1 | -0.085 | -64.877 | -0.93469 | -70.8487 | -2.238 | 113.261 | 2021/04/01 02:39:46
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | CPER2-A-1 | -0.96858 | -70.8595 |  | 19.6017 |  | 0.042625 | 0.13625 |  | -1.1666
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | CPER2-B-3 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 02:49:34
21 |  | CPER2-B-3 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 02:59:33
21 |  | CPER2-B-3 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 03:09:33
21 |  | CPER2-B-3 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 03:19:32
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 | * | CPER2-B-3 | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.591 | -67.034 | -9.7362 | -73.1727 | -1.886 | 93.35 | 2021/04/01 03:30:21
22 |  | EV | -8.692 | -66.719 | -9.8408 | -72.8333 | -2.055 | 93.29 | 2021/04/01 03:39:16
22 |  | EV | -8.793 | -67.117 | -9.9453 | -73.2621 | -1.904 | 93.159 | 2021/04/01 03:48:12
22 |  | EV | -8.842 | -66.767 | -9.996 | -72.885 | -2.056 | 93.32 | 2021/04/01 03:57:08
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -9.8796 | -73.0383 |  |  |  | 0.11147 | 0.19592
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | CPER2-B-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 04:06:55
23 |  | CPER2-B-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 04:16:54
23 |  | CPER2-B-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 04:26:53
23 |  | CPER2-B-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 04:36:52
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 | * | CPER2-B-1 | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | CPER2-B-4 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 04:47:41
24 |  | CPER2-B-4 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 04:57:39
24 |  | CPER2-B-4 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 05:07:37
24 |  | CPER2-B-4 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 05:17:35
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 | * | CPER2-B-4 | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | CPER2-A-4 | -1.828 | -68.403 | -2.7383 | -74.6476 | -2.169 | 99.741 | 2021/04/01 05:28:25
25 |  | CPER2-A-4 | -1.546 | -68.236 | -2.4465 | -74.4677 | -2.094 | 99.955 | 2021/04/01 05:37:20
25 |  | CPER2-A-4 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 05:46:16
25 |  | CPER2-A-4 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 05:56:14
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 | * | CPER2-A-4 | -1.7195 | -37.7541 |  |  |  | 0.98077 | 39.4443
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | CPER2-E-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 06:07:04
26 |  | CPER2-E-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 06:17:02
26 |  | CPER2-E-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 06:27:01
26 |  | CPER2-E-1 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 06:36:59
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 | * | CPER2-E-1 | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | CPER2-D-5 | 1.228 | -54.804 | 0.42393 | -59.9961 | -2.287 | 95.692 | 2021/04/01 06:47:50
27 |  | CPER2-D-5 | 1.478 | -53.594 | 0.68261 | -58.6925 | -2.305 | 95.562 | 2021/04/01 06:56:46
27 |  | CPER2-D-5 | 1.496 | -53.302 | 0.70124 | -58.3779 | -2.389 | 95.643 | 2021/04/01 07:05:42
27 |  | CPER2-D-5 | 1.518 | -53.286 | 0.724 | -58.3607 | -2.367 | 95.569 | 2021/04/01 07:14:38
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | CPER2-D-5 | 0.63295 | -58.8568 |  |  |  | 0.13566 | 0.71907
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | CPER2-1020-3 | 1.349 | -64.269 | 0.54913 | -70.1937 | -2.044 | 93.967 | 2021/04/01 07:24:27
28 |  | CPER2-1020-3 | 1.334 | -65.358 | 0.53361 | -71.367 | -2.227 | 93.692 | 2021/04/01 07:33:24
28 |  | CPER2-1020-3 | 1.379 | -65.533 | 0.58017 | -71.5555 | -2.301 | 94.226 | 2021/04/01 07:42:19
28 |  | CPER2-1020-3 | 1.34 | -65.64 | 0.53982 | -71.6708 | -2.155 | 93.625 | 2021/04/01 07:51:15
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | CPER2-1020-3 | 0.55068 | -71.1967 |  |  |  | 0.019975 | 0.63146
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | CPER2-A-3 | -1.732 | -68.646 | -2.6389 | -74.9094 | -2.059 | 101.793 | 2021/04/01 08:01:00
29 |  | CPER2-A-3 | -1.836 | -68.9 | -2.7465 | -75.1831 | -2.165 | 101.853 | 2021/04/01 08:09:56
29 |  | CPER2-A-3 | -1.879 | -68.957 | -2.791 | -75.2445 | -2.11 | 101.944 | 2021/04/01 08:18:52
29 |  | CPER2-A-3 | -1.845 | -68.809 | -2.7558 | -75.0851 | -2.338 | 102.075 | 2021/04/01 08:27:48
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | CPER2-A-3 | -2.7331 | -75.1055 |  | 7.8785 |  | 0.06343 | 0.13578
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | CPER2-B-5 | -0.213 | -68.29 | -1.0671 | -74.5259 | -1.966 | 104.078 | 2021/04/01 08:37:34
30 |  | CPER2-B-5 | -0.142 | -68.188 | -0.99367 | -74.416 | -2.248 | 103.833 | 2021/04/01 08:46:30
30 |  | CPER2-B-5 | -0.142 | -68.388 | -0.99367 | -74.6315 | -2.308 | 104.315 | 2021/04/01 08:55:27
30 |  | CPER2-B-5 | -0.111 | -68.307 | -0.9616 | -74.5442 | -2.167 | 104.133 | 2021/04/01 09:04:23
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | CPER2-B-5 | -1.004 | -74.5294 |  | 10.052 |  | 0.043213 | 0.082168 |  | -0.4338
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 09:14:08
31 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 09:23:04
31 |  | EV | -8.938 | -67.205 | -10.0953 | -73.3569 | -2.117 | 93.536 | 2021/04/01 09:31:59
31 |  | EV | -8.994 | -67.235 | -10.1532 | -73.3892 | -2.179 | 93.639 | 2021/04/01 09:40:56
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 | * | EV | -5.4855 | -37.1618 |  |  |  | 5.1766 | 38.8095
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | CPER2-3040-2 | -1.54 | -72.576 | -2.4402 | -79.1436 | -2.189 | 93.664 | 2021/04/01 09:50:42
32 |  | CPER2-3040-2 | -1.255 | -73.149 | -2.1453 | -79.761 | -2.272 | 93.547 | 2021/04/01 09:59:39
32 |  | CPER2-3040-2 | -1.137 | -73.13 | -2.0232 | -79.7405 | -2.075 | 93.758 | 2021/04/01 10:08:35
32 |  | CPER2-3040-2 | -1.091 | -73.207 | -1.9756 | -79.8234 | -2.015 | 93.605 | 2021/04/01 10:17:31
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | CPER2-3040-2 | -2.1461 | -79.6171 |  |  |  | 0.2017 | 0.29482
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | CPER2-C-5 | -5.892 | -82.033 | -6.9435 | -89.3325 | -2.097 | 133.932 | 2021/04/01 10:27:19
33 |  | CPER2-C-5 | -6.082 | -82.652 | -7.1401 | -89.9994 | -2.221 | 135.171 | 2021/04/01 10:36:15
33 |  | CPER2-C-5 | -6.215 | -83.337 | -7.2777 | -90.7375 | -2.102 | 135.276 | 2021/04/01 10:45:11
33 |  | CPER2-C-5 | -6.235 | -83.264 | -7.2984 | -90.6588 | -2.248 | 135.71 | 2021/04/01 10:54:06
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | CPER2-C-5 | -7.1649 | -90.1821 |  | 40.9845 |  | 0.158 | 0.60883 |  | -1.9716
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | CPER2-1020-1 | -7.044 | -88.559 | -8.1355 | -96.3636 | -2.257 | 94.378 | 2021/04/01 11:03:54
34 |  | CPER2-1020-1 | -7.107 | -89.139 | -8.2007 | -96.9885 | -2.089 | 94.012 | 2021/04/01 11:12:50
34 |  | CPER2-1020-1 | -7.088 | -89.165 | -8.181 | -97.0165 | -2.294 | 94 | 2021/04/01 11:21:46
34 |  | CPER2-1020-1 | -7.155 | -89.396 | -8.2503 | -97.2654 | -2.033 | 93.685 | 2021/04/01 11:30:42
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | CPER2-1020-1 | -8.1919 | -96.9085 |  |  |  | 0.045989 | 0.3564
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | CPER2-3040-1 | -5.346 | -84.147 | -6.3785 | -91.6101 | -2.071 | 93.408 | 2021/04/01 11:40:29
35 |  | CPER2-3040-1 | -5.272 | -83.836 | -6.3019 | -91.2751 | -2.031 | 93.286 | 2021/04/01 11:49:26
35 |  | CPER2-3040-1 | -5.299 | -83.949 | -6.3299 | -91.3968 | -2.074 | 93.434 | 2021/04/01 11:58:23
35 |  | CPER2-3040-1 | -5.258 | -83.786 | -6.2874 | -91.2212 | -2.013 | 93.586 | 2021/04/01 12:07:19
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | CPER2-3040-1 | -6.3244 | -91.3758 |  |  |  | 0.038767 | 0.16023
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | CPER2-deep-2 | -8.353 | -88.824 | -9.49 | -96.6491 | -2.01 | 93.095 | 2021/04/01 12:17:02
36 |  | CPER2-deep-2 | -8.552 | -89.337 | -9.6959 | -97.2018 | -2.02 | 92.845 | 2021/04/01 12:25:58
36 |  | CPER2-deep-2 | -8.625 | -89.389 | -9.7714 | -97.2578 | -1.942 | 93.062 | 2021/04/01 12:34:54
36 |  | CPER2-deep-2 | -8.572 | -89.446 | -9.7166 | -97.3193 | -2.151 | 93.191 | 2021/04/01 12:43:51
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | CPER2-deep-2 | -9.6685 | -97.107 |  |  |  | 0.11905 | 0.28681
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | CPER2-B-2 | -3.459 | -81.027 | -4.4259 | -88.2487 | -2.259 | 104.82 | 2021/04/01 12:53:35
37 |  | CPER2-B-2 | -3.218 | -80.082 | -4.1766 | -87.2305 | -2.244 | 104.731 | 2021/04/01 13:02:31
37 |  | CPER2-B-2 | -3.136 | -79.899 | -4.0917 | -87.0334 | -2.178 | 104.602 | 2021/04/01 13:11:27
37 |  | CPER2-B-2 | -3.137 | -79.973 | -4.0927 | -87.1131 | -2.258 | 104.443 | 2021/04/01 13:20:23
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | CPER2-B-2 | -4.1967 | -87.4064 |  | 10.6112 |  | 0.15258 | 0.52656 |  | -1.2928
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | CPER2-1020-2 | -1.667 | -76.478 | -2.5717 | -83.3476 | -2.092 | 93.732 | 2021/04/01 13:30:07
38 |  | CPER2-1020-2 | -1.577 | -76.274 | -2.4785 | -83.1278 | -2.062 | 93.976 | 2021/04/01 13:39:03
38 |  | CPER2-1020-2 | -1.591 | -76.258 | -2.493 | -83.1106 | -2.068 | 93.778 | 2021/04/01 13:47:59
38 |  | CPER2-1020-2 | -1.548 | -76.089 | -2.4485 | -82.9285 | -2.125 | 93.875 | 2021/04/01 13:56:55
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | CPER2-1020-2 | -2.4979 | -83.1286 |  |  |  | 0.050763 | 0.15926
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | CPER2-010-1 | -0.954 | -74.291 | -1.8339 | -80.9913 | -2.198 | 94.013 | 2021/04/01 14:06:40
39 |  | CPER2-010-1 | -0.956 | -74.194 | -1.836 | -80.8868 | -2.231 | 93.716 | 2021/04/01 14:15:35
39 |  | CPER2-010-1 | -0.864 | -74.138 | -1.7408 | -80.8265 | -2.049 | 93.626 | 2021/04/01 14:24:31
39 |  | CPER2-010-1 | -0.899 | -74.122 | -1.777 | -80.8093 | -2.254 | 93.82 | 2021/04/01 14:33:27
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | CPER2-010-1 | -1.7969 | -80.8785 |  |  |  | 0.044784 | 0.076352
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 14:43:12
40 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 14:52:08
40 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 15:01:04
40 |  | EV |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 15:09:59
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 | * | EV | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | CPER2-2030-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 15:19:45
41 |  | CPER2-2030-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 15:29:44
41 |  | CPER2-2030-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 15:39:42
41 |  | CPER2-2030-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 15:49:41
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 | * | CPER2-2030-2 | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | CPER2-A-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 16:00:29
42 |  | CPER2-A-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 16:10:27
42 |  | CPER2-A-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 16:20:25
42 |  | CPER2-A-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 16:30:24
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 | * | CPER2-A-2 | -0.84674 | -0.9506 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | CPER2-C-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 16:41:11
43 |  | CPER2-C-2 |  |  | -0.84674 | -0.9506 |  |  | 2021/04/01 16:51:08
43 |  | CPER2-C-2 | -2.943 | -75.78 | -3.892 | -82.5956 | -2.119 | 92.998 | 2021/04/01 17:01:06
43 |  | CPER2-C-2 | -2.943 | -75.78 | -3.892 | -82.5956 | -2.119 | 92.998 | 2021/04/01 17:10:02
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 | * | CPER2-C-2 | -2.3694 | -41.7731 |  |  |  | 1.6991 | 43.7516
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | CPER2-D-4 | 2.684 | -69.798 | 1.9305 | -76.1506 | -5.546 | 110.124 | 2021/04/01 17:20:49
44 |  | CPER2-D-4 | 2.671 | -70.008 | 1.9171 | -76.3769 | -5.983 | 111.042 | 2021/04/01 17:29:45
44 |  | CPER2-D-4 | 2.672 | -69.91 | 1.9181 | -76.2713 | -5.883 | 110.786 | 2021/04/01 17:38:41
44 |  | CPER2-D-4 | 2.651 | -69.96 | 1.8964 | -76.3251 | -6.109 | 110.623 | 2021/04/01 17:47:38
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | CPER2-D-4 | 1.9155 | -76.281 | -5.8803 | 16.606 | 40.919 | 0.013675 | 0.090044 |  | 2.5792
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | CPER2-E-3 | -0.729 | -68.742 | -1.6011 | -75.0129 | -2.356 | 134.269 | 2021/04/01 17:57:23
45 |  | CPER2-E-3 | -0.558 | -67.874 | -1.4241 | -74.0777 | -2.144 | 135.646 | 2021/04/01 18:06:19
45 |  | CPER2-E-3 | -0.583 | -67.693 | -1.45 | -73.8827 | -2.138 | 137.093 | 2021/04/01 18:15:16
45 |  | CPER2-E-3 | -0.58 | -67.691 | -1.4469 | -73.8805 | -2.262 | 136.913 | 2021/04/01 18:24:12
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | CPER2-E-3 | -1.4805 | -74.2135 |  | 41.9425 |  | 0.078462 | 0.50205 |  | -1.6446
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | EV | -8.938 | -67.205 | -10.0953 | -73.3569 | -2.117 | 93.536 | 2021/04/01 18:33:58
46 |  | EV | -8.994 | -67.235 | -10.1532 | -73.3892 | -2.179 | 93.639 | 2021/04/01 18:42:54
46 |  | EV | -8.938 | -67.205 | -10.0953 | -73.3569 | -2.117 | 93.536 | 2021/04/01 18:51:49
46 |  | EV | -8.994 | -67.235 | -10.1532 | -73.3892 | -2.179 | 93.639 | 2021/04/01 19:00:44
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | EV | -10.1243 | -73.3731 |  |  |  | 0.032332 | 0.017321

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-15824 | 2021/03/31 10:50:07 | 1-01 | 1 | -10.009 | -77.838 | 13847 | -1 | 1 | EV | standard | H2O | 1617209713 | 0.212 | 0.899 | 55.451 | 0 | -0.001 | 1.344 | 137.42 | 30.875 | 2.337 | 67.041 | 0.865 | -1.193 | 0 | 0 | 1 | 45.312 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-15824 | 2021/03/31 10:59:03 | 1-01 | 2 | -9.441 | -70.953 | 12651 | -1 | 1 | EV | standard | H2O | 1617210248 | 0.19 | 0.96 | 43.811 | -0.001 | -0.007 | 1.068 | 105.312 | 29.187 | 2.15 | 59.647 | 0.858 | -1.623 | 0 | 0 | 1 | 44.5 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-15824 | 2021/03/31 11:08:00 | 1-01 | 3 | -9.335 | -69.587 | 12670 | -1 | 1 | EV | standard | H2O | 1617210785 | 0.193 | 0.839 | 51.928 | 0.001 | -0.004 | 1.269 | 103.618 | 29.273 | 2.152 | 57.083 | 0.858 | -2.101 | 0 | 0 | 1 | 44.125 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-15824 | 2021/03/31 11:16:56 | 1-01 | 4 | -9.271 | -69.119 | 12668 | 0 | 1 | EV | standard | H2O | 1617211321 | 0.195 | 0.961 | 52.953 | 0 | 0 | 1.296 | 101.852 | 28.979 | 2.156 | 58.729 | 0.858 | -1.936 | 0 | 0 | 1 | 44.062 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-15824 | 2021/03/31 11:25:53 | 1-01 | 5 | -9.229 | -68.795 | 12649 | 0 | 1 | EV | standard | H2O | 1617211858 | 0.194 | 1.041 | 49.233 | 0 | -0.003 | 1.197 | 100.493 | 29.358 | 2.165 | 56.563 | 0.858 | -2.245 | 0 | 0 | 1 | 43.812 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-15824 | 2021/03/31 11:34:49 | 1-01 | 6 | -9.239 | -68.578 | 12623 | 0 | 1 | EV | standard | H2O | 1617212394 | 0.21 | 0.908 | 45.023 | 0 | -0.005 | 1.097 | 99.153 | 28.996 | 2.168 | 58.184 | 0.858 | -2.089 | 0 | 0 | 1 | 43.938 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-15824 | 2021/03/31 11:43:46 | 1-01 | 7 | -9.189 | -68.253 | 12627 | 0 | 1 | EV | standard | H2O | 1617212931 | 0.178 | 0.875 | 50.047 | 0 | -0.004 | 1.231 | 97.851 | 28.987 | 2.158 | 58.463 | 0.858 | -2.055 | 0 | 0 | 1 | 43.938 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-15824 | 2021/03/31 11:52:43 | 1-01 | 8 | -9.18 | -68.131 | 12607 | 0 | 1 | EV | standard | H2O | 1617213467 | 0.19 | 0.971 | 46.675 | 0 | 0.001 | 1.137 | 97.394 | 29.09 | 2.165 | 56.995 | 0.858 | -2.263 | 0 | 0 | 1 | 43.875 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-15824 | 2021/03/31 12:01:39 | 1-01 | 9 | -9.174 | -68.373 | 12561 | 0 | 1 | EV | standard | H2O | 1617214004 | 0.204 | 0.873 | 49.013 | 0 | 0.001 | 1.196 | 96.589 | 28.879 | 2.139 | 58.318 | 0.858 | -1.973 | 0 | 0 | 1 | 43.688 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-15824 | 2021/03/31 12:10:36 | 1-01 | 10 | -9.15 | -68.081 | 12674 | 0 | 1 | EV | standard | H2O | 1617214541 | 0.19 | 0.908 | 50.346 | -0.001 | -0.001 | 1.224 | 96.479 | 29.053 | 2.165 | 58.459 | 0.858 | -2.025 | 0 | 0 | 1 | 43.812 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-15825 | 2021/03/31 12:20:26 | 1-02 | 1 | 1.963 | 9.404 | 12955 | -1 | 1 | PZ | standard | H2O | 1617215131 | 0.217 | 0.915 | 45.221 | 0 | -0.007 | 1.089 | 95.549 | 29.985 | 2.352 | 62.308 | 0.859 | -2.141 | 0 | 0 | 1 | 43.875 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-15825 | 2021/03/31 12:29:23 | 1-02 | 2 | 2.518 | 15.869 | 12580 | -1 | 1 | PZ | standard | H2O | 1617215668 | 0.22 | 0.944 | 46.903 | -0.001 | -0.005 | 1.145 | 94.806 | 29.312 | 2.307 | 60.275 | 0.858 | -2.26 | 0 | 0 | 1 | 43.812 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-15825 | 2021/03/31 12:38:19 | 1-02 | 3 | 2.698 | 17.621 | 12632 | -1 | 1 | PZ | standard | H2O | 1617216204 | 0.205 | 1.048 | 47.98 | 0 | -0.009 | 1.17 | 94.87 | 29.42 | 2.309 | 59.774 | 0.858 | -2.474 | 0 | 0 | 1 | 43.875 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-15825 | 2021/03/31 12:47:15 | 1-02 | 4 | 2.744 | 18.154 | 12618 | 0 | 1 | PZ | standard | H2O | 1617216740 | 0.229 | 0.984 | 49.583 | -0.001 | -0.003 | 1.213 | 94.531 | 29.279 | 2.304 | 59.726 | 0.859 | -2.392 | 0 | 0 | 1 | 43.688 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-15825 | 2021/03/31 12:56:13 | 1-02 | 5 | 2.738 | 18.71 | 12577 | 0 | 1 | PZ | standard | H2O | 1617217278 | 0.19 | 0.931 | 47.497 | 0 | -0.002 | 1.158 | 94.097 | 29.367 | 2.302 | 60.031 | 0.858 | -2.19 | 0 | 0 | 1 | 43.75 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-15825 | 2021/03/31 13:05:09 | 1-02 | 6 | 2.802 | 19.038 | 12574 | 0 | 1 | PZ | standard | H2O | 1617217814 | 0.214 | 1.146 | 47.12 | 0 | -0.01 | 1.151 | 94.04 | 28.788 | 2.31 | 60.833 | 0.858 | -2.147 | 0 | 0 | 1 | 43.688 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-15825 | 2021/03/31 13:14:06 | 1-02 | 7 | 2.833 | 19.205 | 12589 | 0 | 1 | PZ | standard | H2O | 1617218351 | 0.211 | 0.914 | 45.939 | -0.001 | -0.003 | 1.124 | 93.877 | 29.117 | 2.313 | 60.751 | 0.857 | -2.178 | 0 | 0 | 1 | 43.75 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-15825 | 2021/03/31 13:23:03 | 1-02 | 8 | 2.835 | 19.564 | 12611 | 0 | 1 | PZ | standard | H2O | 1617218887 | 0.205 | 1.004 | 51.16 | 0 | -0.001 | 1.259 | 93.923 | 28.836 | 2.305 | 61.877 | 0.858 | -2.124 | 0 | 0 | 1 | 43.875 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-15825 | 2021/03/31 13:31:59 | 1-02 | 9 | 2.849 | 19.522 | 12584 | 0 | 1 | PZ | standard | H2O | 1617219424 | 0.219 | 0.825 | 50.589 | 0 | -0.002 | 1.236 | 93.639 | 28.855 | 2.306 | 61.839 | 0.858 | -2.107 | 0 | 0 | 1 | 43.875 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-15825 | 2021/03/31 13:40:55 | 1-02 | 10 | 2.855 | 19.734 | 12637 | 0 | 1 | PZ | standard | H2O | 1617219959 | 0.197 | 0.869 | 51.851 | 0 | -0.002 | 1.27 | 93.704 | 28.981 | 2.327 | 62.403 | 0.86 | -2.032 | 0 | 0 | 1 | 44.062 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-15826 | 2021/03/31 13:50:43 | 1-03 | 1 | -13.569 | -97.88 | 12945 | -1 | 1 | UT2 | standard | H2O | 1617220548 | 0.189 | 0.991 | 44.747 | 0 | 0.009 | 1.091 | 94.25 | 30.106 | 2.164 | 58.625 | 0.858 | -2.101 | 0 | 0 | 1 | 44.125 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-15826 | 2021/03/31 13:59:40 | 1-03 | 2 | -14.218 | -107.132 | 12804 | -1 | 1 | UT2 | standard | H2O | 1617221084 | 0.188 | 0.951 | 48.095 | 0 | 0.01 | 1.178 | 94.085 | 29.542 | 2.138 | 57.862 | 0.858 | -2.085 | 0 | 0 | 1 | 44.312 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-15826 | 2021/03/31 14:08:37 | 1-03 | 3 | -14.344 | -108.714 | 12899 | -1 | 1 | UT2 | standard | H2O | 1617221622 | 0.201 | 0.887 | 46.97 | 0.001 | 0.004 | 1.139 | 94.217 | 30.234 | 2.153 | 57.213 | 0.859 | -2.159 | 0 | 0 | 1 | 44.562 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-15826 | 2021/03/31 14:17:34 | 1-03 | 4 | -14.583 | -110.752 | 12561 | 0 | 1 | UT2 | standard | H2O | 1617222159 | 0.209 | 0.9 | 47.861 | 0 | 0.008 | 1.167 | 93.221 | 29.331 | 2.096 | 57.021 | 0.858 | -2.014 | 0 | 0 | 1 | 44.5 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-15826 | 2021/03/31 14:26:31 | 1-03 | 5 | -14.631 | -111.18 | 12672 | 0 | 1 | UT2 | standard | H2O | 1617222696 | 0.207 | 0.836 | 48.605 | 0.001 | 0.006 | 1.179 | 93.646 | 29.596 | 2.107 | 56.498 | 0.858 | -2.151 | 0 | 0 | 1 | 44.75 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-15826 | 2021/03/31 14:35:27 | 1-03 | 6 | -14.693 | -111.765 | 12565 | 0 | 1 | UT2 | standard | H2O | 1617223232 | 0.201 | 0.87 | 49.964 | 0 | 0.003 | 1.229 | 93.285 | 29.281 | 2.091 | 56.14 | 0.859 | -2.015 | 0 | 0 | 1 | 44.625 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-15826 | 2021/03/31 14:44:23 | 1-03 | 7 | -14.723 | -111.748 | 12743 | 0 | 1 | UT2 | standard | H2O | 1617223769 | 0.179 | 0.762 | 49.579 | 0.001 | 0.004 | 1.212 | 93.808 | 29.66 | 2.107 | 56.593 | 0.858 | -2.102 | 0 | 0 | 1 | 44.812 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-15826 | 2021/03/31 14:53:20 | 1-03 | 8 | -14.743 | -112.146 | 12671 | 0 | 1 | UT2 | standard | H2O | 1617224306 | 0.198 | 0.813 | 49.268 | 0 | 0.001 | 1.195 | 93.771 | 28.784 | 2.095 | 57.839 | 0.857 | -2.018 | 0 | 0 | 1 | 44.625 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-15826 | 2021/03/31 15:02:14 | 1-03 | 9 | -14.697 | -112.188 | 12573 | 0 | 1 | UT2 | standard | H2O | 1617225438 | 0.2 | 0.906 | 43.55 | 0 | 0.005 | 1.057 | 93.718 | 28.23 | 2.069 | 57.679 | 0.858 | -2.058 | 0 | 0 | 1 | 44.5 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-15826 | 2021/03/31 15:12:13 | 1-03 | 10 | -14.697 | -112.188 | 12573 | 0 | 1 | UT2 | standard | H2O | 1617225438 | 0.2 | 0.906 | 43.55 | 0 | 0.005 | 1.057 | 93.718 | 28.23 | 2.069 | 57.679 | 0.858 | -2.058 | 0 | 0 | 1 | 44.5 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-15827 | 2021/03/31 15:31:01 | 1-04 | 1 | -9.409 | -71.772 | 12650 | -1 | 1 | EV | standard | H2O | 1617226565 | 0.223 | 0.935 | 47.539 | 0 | -0.002 | 1.161 | 93.542 | 28.969 | 2.156 | 59.076 | 0.857 | -1.898 | 0 | 0 | 1 | 44.25 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-15827 | 2021/03/31 15:39:57 | 1-04 | 2 | -9.191 | -68.513 | 12650 | -1 | 1 | EV | standard | H2O | 1617227101 | 0.182 | 0.948 | 52.315 | 0.001 | 0.002 | 1.283 | 93.65 | 29.667 | 2.158 | 56.856 | 0.858 | -2.296 | 0 | 0 | 1 | 44 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-15827 | 2021/03/31 15:48:53 | 1-04 | 3 | -9.1 | -67.993 | 12548 | -1 | 1 | EV | standard | H2O | 1617227638 | 0.222 | 0.919 | 53.023 | 0 | 0.002 | 1.296 | 93.424 | 29.088 | 2.154 | 57.381 | 0.858 | -2.204 | 0 | 0 | 1 | 44 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-15827 | 2021/03/31 15:57:48 | 1-04 | 4 | -9.093 | -67.716 | 12665 | 0 | 1 | EV | standard | H2O | 1617228174 | 0.192 | 0.879 | 49.383 | 0 | -0.001 | 1.207 | 93.728 | 28.985 | 2.172 | 58.559 | 0.858 | -2.223 | 0 | 0 | 1 | 43.812 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-15827 | 2021/03/31 16:06:46 | 1-04 | 5 | -9.081 | -67.449 | 12642 | 0 | 1 | EV | standard | H2O | 1617228710 | 0.208 | 0.908 | 48.793 | 0 | -0.003 | 1.194 | 93.667 | 29.204 | 2.167 | 57.581 | 0.858 | -2.221 | 0 | 0 | 1 | 43.938 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-15827 | 2021/03/31 16:15:42 | 1-04 | 6 | -9.097 | -67.364 | 12554 | 0 | 1 | EV | standard | H2O | 1617229247 | 0.192 | 0.821 | 48.858 | 0 | 0 | 1.186 | 93.428 | 28.592 | 2.156 | 58.648 | 0.858 | -2.107 | 0 | 0 | 1 | 43.938 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-15827 | 2021/03/31 16:24:38 | 1-04 | 7 | -9.116 | -67.225 | 12656 | 0 | 1 | EV | standard | H2O | 1617229783 | 0.195 | 0.858 | 48.626 | 0 | 0.002 | 1.183 | 93.549 | 29.077 | 2.173 | 59.191 | 0.858 | -2.03 | 0 | 0 | 1 | 43.875 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-15827 | 2021/03/31 16:33:35 | 1-04 | 8 | -9.094 | -67.576 | 12556 | 0 | 1 | EV | standard | H2O | 1617230320 | 0.18 | 0.746 | 49.763 | 0 | 0.001 | 1.216 | 93.47 | 29.083 | 2.15 | 58.013 | 0.858 | -2.265 | 0 | 0 | 1 | 43.938 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-15827 | 2021/03/31 16:42:32 | 1-04 | 9 | -9.071 | -67.366 | 12694 | 0 | 1 | EV | standard | H2O | 1617230858 | 0.208 | 1.075 | 47.794 | 0 | -0.002 | 1.159 | 93.659 | 29.192 | 2.174 | 58.63 | 0.858 | -2.055 | 0 | 0 | 1 | 43.938 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-15827 | 2021/03/31 16:51:29 | 1-04 | 10 | -9.071 | -67.32 | 12651 | 0 | 1 | EV | standard | H2O | 1617231394 | 0.211 | 0.777 | 52.251 | 0 | 0.001 | 1.278 | 93.485 | 29.158 | 2.178 | 59.106 | 0.858 | -1.999 | 0 | 0 | 1 | 43.75 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-15828 | 2021/03/31 17:01:20 | 1-05 | 1 | -3.275 | -77.274 | 12650 | -1 | 1 | CPER2-2030-3 | 20-202 | H2O | 1617231985 | 0.196 | 0.786 | 49.86 | -0.001 | 0.001 | 1.21 | 94.939 | 29.029 | 2.159 | 58.073 | 0.857 | -2.197 | 0 | 0 | 1 | 43.875 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-15828 | 2021/03/31 17:10:16 | 1-05 | 2 | -2.986 | -78.11 | 12609 | -1 | 1 | CPER2-2030-3 | 20-202 | H2O | 1617232521 | 0.187 | 0.802 | 44.704 | 0 | 0.001 | 1.088 | 94.874 | 28.876 | 2.146 | 58.124 | 0.858 | -2.19 | 0 | 0 | 1 | 43.688 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-15828 | 2021/03/31 17:19:12 | 1-05 | 3 | -2.935 | -78.339 | 12541 | -1 | 1 | CPER2-2030-3 | 20-202 | H2O | 1617233057 | 0.207 | 0.844 | 47.887 | -0.001 | 0.001 | 1.167 | 94.617 | 28.826 | 2.138 | 58.085 | 0.858 | -2.149 | 0 | 0 | 1 | 43.688 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-15828 | 2021/03/31 17:28:09 | 1-05 | 4 | -2.978 | -78.407 | 12483 | 0 | 1 | CPER2-2030-3 | 20-202 | H2O | 1617233594 | 0.181 | 0.829 | 46.928 | 0 | 0.004 | 1.152 | 94.601 | 28.667 | 2.133 | 56.942 | 0.858 | -2.334 | 0 | 0 | 1 | 43.562 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-15829 | 2021/03/31 17:37:59 | 1-06 | 1 | -1.739 | -76.889 | 12704 | -1 | 1 | CPER2-010-2 | 20-202 | H2O | 1617234184 | 0.193 | 0.908 | 47.67 | 0 | 0.001 | 1.155 | 94.127 | 29.034 | 2.168 | 59.61 | 0.858 | -1.992 | 0 | 0 | 1 | 43.375 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-15829 | 2021/03/31 17:46:57 | 1-06 | 2 | -1.647 | -76.67 | 12640 | -1 | 1 | CPER2-010-2 | 20-202 | H2O | 1617234722 | 0.21 | 1.059 | 46.158 | -0.001 | 0 | 1.12 | 93.889 | 29.017 | 2.156 | 58.952 | 0.858 | -2.035 | 0 | 0 | 1 | 43.5 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-15829 | 2021/03/31 17:55:53 | 1-06 | 3 | -1.644 | -76.642 | 12601 | -1 | 1 | CPER2-010-2 | 20-202 | H2O | 1617235258 | 0.205 | 0.786 | 45.823 | 0 | -0.001 | 1.125 | 93.902 | 28.973 | 2.16 | 58.065 | 0.858 | -2.163 | 0 | 0 | 1 | 43.375 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-15829 | 2021/03/31 18:04:51 | 1-06 | 4 | -1.618 | -76.918 | 12670 | 0 | 1 | CPER2-010-2 | 20-202 | H2O | 1617235796 | 0.217 | 0.895 | 50.165 | -0.001 | 0.001 | 1.223 | 94.104 | 29.226 | 2.161 | 57.458 | 0.857 | -2.289 | 0 | 0 | 1 | 43.5 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-15830 | 2021/03/31 18:14:41 | 1-07 | 1 | -0.512 | -70.769 | 12535 | -1 | 1 | CPER2-A-5 | 20-202 | H2O | 1617236387 | 0.219 | 0.964 | 51.977 | -0.001 | 0.006 | 1.264 | 113.084 | 28.228 | 2.151 | 57.037 | 0.858 | -2.272 | 0 | 0 | 1 | 43.438 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-15830 | 2021/03/31 18:23:39 | 1-07 | 2 | -0.446 | -69.933 | 12437 | -1 | 1 | CPER2-A-5 | 20-202 | H2O | 1617236924 | 0.192 | 0.868 | 48.39 | 0 | -0.004 | 1.184 | 112.86 | 27.567 | 2.158 | 57.505 | 0.858 | -2.144 | 0 | 0 | 1 | 43.562 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-15830 | 2021/03/31 18:32:35 | 1-07 | 3 | -0.443 | -69.74 | 12538 | -1 | 1 | CPER2-A-5 | 20-202 | H2O | 1617237460 | 0.213 | 0.825 | 51.107 | 0 | 0.002 | 1.245 | 113.182 | 28.303 | 2.15 | 57.039 | 0.858 | -2.145 | 0 | 0 | 1 | 43.562 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-15830 | 2021/03/31 18:41:32 | 1-07 | 4 | -0.452 | -69.897 | 12443 | 0 | 1 | CPER2-A-5 | 20-202 | H2O | 1617237997 | 0.212 | 0.848 | 49.709 | 0 | 0.001 | 1.219 | 112.933 | 28.145 | 2.148 | 56.664 | 0.857 | -2.341 | 0 | 0 | 1 | 43.75 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-15831 | 2021/03/31 18:51:20 | 1-08 | 1 | 0.338 | -65.338 | 12378 | -1 | 1 | CPER2-E-5 | 20-202 | H2O | 1617238584 | 0.202 | 0.846 | 48.907 | 0 | -0.002 | 1.197 | 119.677 | 27.013 | 2.151 | 59.001 | 0.857 | -1.912 | 0 | 0 | 1 | 43.75 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-15831 | 2021/03/31 19:00:15 | 1-08 | 2 | 0.359 | -65.124 | 12471 | -1 | 1 | CPER2-E-5 | 20-202 | H2O | 1617239120 | 0.217 | 0.904 | 46.111 | -0.001 | -0.001 | 1.123 | 120.374 | 27.115 | 2.154 | 58.963 | 0.859 | -2.002 | 0 | 0 | 1 | 43.875 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-15831 | 2021/03/31 19:09:12 | 1-08 | 3 | 0.371 | -65.103 | 12562 | -1 | 1 | CPER2-E-5 | 20-202 | H2O | 1617239657 | 0.215 | 0.884 | 48.413 | -0.001 | 0.001 | 1.183 | 120.702 | 27.941 | 2.163 | 57.457 | 0.858 | -2.153 | 0 | 0 | 1 | 43.812 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-15831 | 2021/03/31 19:18:08 | 1-08 | 4 | 0.353 | -64.973 | 12558 | 0 | 1 | CPER2-E-5 | 20-202 | H2O | 1617240193 | 0.215 | 0.959 | 46.46 | 0 | 0.006 | 1.132 | 120.719 | 27.95 | 2.17 | 57.87 | 0.859 | -2.255 | 0 | 0 | 1 | 43.875 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-15832 | 2021/03/31 19:27:56 | 1-09 | 1 | 0.247 | -72.67 | 12663 | -1 | 1 | CPER2-D-1 | 20-202 | H2O | 1617240781 | 0.206 | 1.153 | 47.19 | 0 | 0.004 | 1.146 | 308.872 | 17.281 | 2.699 | 59.405 | 0.859 | -4.142 | 0 | 0 | 1 | 44 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-15832 | 2021/03/31 19:36:53 | 1-09 | 2 | 0.059 | -73.755 | 12439 | -1 | 1 | CPER2-D-1 | 20-202 | H2O | 1617241318 | 0.204 | 0.987 | 52.314 | -0.001 | 0.002 | 1.274 | 308.674 | 17.21 | 2.665 | 56.474 | 0.86 | -4.61 | 0 | 0 | 1 | 43.875 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-15832 | 2021/03/31 19:45:49 | 1-09 | 3 | 0.044 | -73.937 | 12462 | -1 | 1 | CPER2-D-1 | 20-202 | H2O | 1617241855 | 0.217 | 1.043 | 50.08 | -0.001 | 0 | 1.22 | 310.03 | 16.884 | 2.687 | 57.803 | 0.86 | -4.42 | 0 | 0 | 1 | 44.188 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-15832 | 2021/03/31 19:54:45 | 1-09 | 4 | 0.055 | -74.205 | 12569 | 0 | 1 | CPER2-D-1 | 20-202 | H2O | 1617242390 | 0.227 | 1.011 | 52.253 | 0 | 0.001 | 1.284 | 311.78 | 17.364 | 2.693 | 57.822 | 0.859 | -4.353 | 0 | 0 | 1 | 44.375 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-15833 | 2021/03/31 20:04:33 | 1-10 | 1 | [5 empty] | CPER2-E-2 | 20-202 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-15833 | 2021/03/31 20:14:32 | 1-10 | 2 | [5 empty] | CPER2-E-2 | 20-202 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-15833 | 2021/03/31 20:24:31 | 1-10 | 3 | -1.409 | -64.831 | 12484 | -1 | 1 | CPER2-E-2 | 20-202 | H2O | 1617244176 | 0.206 | 0.783 | 41.765 | 0 | 0.001 | 1.019 | 121.343 | 27.472 | 2.175 | 58.455 | 0.859 | -2.102 | 0 | 0 | 1 | 44.75 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-15833 | 2021/03/31 20:33:26 | 1-10 | 4 | -1.275 | -63.96 | 12671 | 0 | 1 | CPER2-E-2 | 20-202 | H2O | 1617244711 | 0.189 | 0.94 | 44.408 | -0.001 | 0.002 | 1.094 | 119.891 | 28.276 | 2.196 | 57.795 | 0.858 | -2.269 | 0 | 0 | 1 | 44.375 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-15834 | 2021/03/31 20:43:15 | 1-11 | 1 | 2.439 | -61.022 | 12641 | -1 | 1 | CPER2-D-3 | 20-202 | H2O | 1617245301 | 0.208 | 0.896 | 44.86 | -0.001 | -0.002 | 1.086 | 109.167 | 23.292 | 2.779 | 61.382 | 0.858 | -4.918 | 0 | 0 | 1 | 44.375 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-15834 | 2021/03/31 20:52:11 | 1-11 | 2 | 2.451 | -61.238 | 12494 | -1 | 1 | CPER2-D-3 | 20-202 | H2O | 1617245836 | 0.2 | 0.913 | 43.969 | 0 | -0.002 | 1.075 | 108.642 | 22.633 | 2.76 | 61.57 | 0.858 | -4.639 | 0 | 0 | 1 | 44.312 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-15834 | 2021/03/31 21:01:07 | 1-11 | 3 | 2.441 | -61.296 | 12524 | -1 | 1 | CPER2-D-3 | 20-202 | H2O | 1617246372 | 0.205 | 1.088 | 45.523 | -0.001 | 0.001 | 1.112 | 108.648 | 22.7 | 2.772 | 61.066 | 0.858 | -4.692 | 0 | 0 | 1 | 44.188 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-15834 | 2021/03/31 21:10:04 | 1-11 | 4 | 2.452 | -61.244 | 12549 | 0 | 1 | CPER2-D-3 | 20-202 | H2O | 1617246908 | 0.186 | 0.761 | 47.644 | -0.001 | -0.002 | 1.166 | 108.498 | 23.125 | 2.772 | 61.409 | 0.858 | -4.859 | 0 | 0 | 1 | 44 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-15835 | 2021/03/31 21:19:53 | 1-12 | 1 | -3.6 | -77.166 | 12503 | -1 | 1 | CPER2-C-4 | 20-202 | H2O | 1617247498 | 0.21 | 0.952 | 47.072 | 0 | 0.001 | 1.157 | 217.578 | 24.137 | 2.173 | 55.856 | 0.859 | -2.328 | 0 | 0 | 1 | 44.062 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-15835 | 2021/03/31 21:28:50 | 1-12 | 2 | -3.696 | -78.083 | 12471 | -1 | 1 | CPER2-C-4 | 20-202 | H2O | 1617248035 | 0.233 | 1.047 | 47.496 | 0 | 0.003 | 1.159 | 218.055 | 24.297 | 2.164 | 54.906 | 0.858 | -2.343 | 0 | 0 | 1 | 44.125 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-15835 | 2021/03/31 21:37:46 | 1-12 | 3 | -3.728 | -78.099 | 12484 | -1 | 1 | CPER2-C-4 | 20-202 | H2O | 1617248571 | 0.216 | 0.934 | 52.376 | 0 | 0.002 | 1.28 | 218.5 | 24.445 | 2.16 | 55.208 | 0.859 | -2.506 | 0 | 0 | 1 | 43.938 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-15835 | 2021/03/31 21:46:43 | 1-12 | 4 | -3.779 | -78.306 | 12537 | 0 | 1 | CPER2-C-4 | 20-202 | H2O | 1617249108 | 0.183 | 0.925 | 50.247 | 0 | 0.001 | 1.227 | 219.24 | 24.236 | 2.19 | 56.566 | 0.859 | -2.238 | 0 | 0 | 1 | 44 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-15836 | 2021/03/31 21:56:33 | 1-13 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-15836 | 2021/03/31 22:05:30 | 1-13 | 2 | [5 empty] | EV | standard | [18 empty] | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-15836 | 2021/03/31 22:14:26 | 1-13 | 3 | [5 empty] | EV | standard | [18 empty] | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-15836 | 2021/03/31 22:23:22 | 1-13 | 4 | [5 empty] | EV | standard | [18 empty] | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-15837 | 2021/03/31 22:33:12 | 1-14 | 1 | -4.408 | -75.215 | 12683 | -1 | 1 | CPER2-C-1 | 20-202 | H2O | 1617251897 | 0.21 | 0.89 | 54.181 | 0 | 0.001 | 1.322 | 99.684 | 28.655 | 2.167 | 59.205 | 0.858 | -2.045 | 0 | 0 | 1 | 43.812 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-15837 | 2021/03/31 22:42:10 | 1-14 | 2 | -4.274 | -75.962 | 12472 | -1 | 1 | CPER2-C-1 | 20-202 | H2O | 1617252434 | 0.194 | 0.818 | 49.927 | 0 | 0.002 | 1.223 | 99.221 | 28.352 | 2.143 | 56.835 | 0.858 | -2.167 | 0 | 0 | 1 | 43.812 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-15837 | 2021/03/31 22:51:07 | 1-14 | 3 | -4.17 | -75.887 | 12646 | -1 | 1 | CPER2-C-1 | 20-202 | H2O | 1617252972 | 0.201 | 0.848 | 48.494 | 0 | -0.001 | 1.175 | 99.568 | 29.095 | 2.155 | 57.828 | 0.859 | -2.221 | 0 | 0 | 1 | 43.812 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-15837 | 2021/03/31 23:00:03 | 1-14 | 4 | -4.164 | -76.259 | 12584 | 0 | 1 | CPER2-C-1 | 20-202 | H2O | 1617253508 | 0.22 | 0.903 | 48.273 | 0 | 0.002 | 1.182 | 99.472 | 28.173 | 2.152 | 59.472 | 0.858 | -1.992 | 0 | 0 | 1 | 43.812 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-15838 | 2021/03/31 23:09:49 | 1-15 | 1 | 0.011 | -70.672 | 12668 | -1 | 1 | CPER2-010-3 | 20-202 | H2O | 1617254094 | 0.202 | 0.9 | 46.683 | 0 | 0 | 1.139 | 94.105 | 28.995 | 2.189 | 59.082 | 0.858 | -2.193 | 0 | 0 | 1 | 43.938 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-15838 | 2021/03/31 23:18:46 | 1-15 | 2 | 0.147 | -70.237 | 12603 | -1 | 1 | CPER2-010-3 | 20-202 | H2O | 1617254631 | 0.215 | 0.857 | 50.167 | 0 | -0.002 | 1.227 | 94.094 | 28.615 | 2.172 | 58.414 | 0.857 | -2.135 | 0 | 0 | 1 | 43.688 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-15838 | 2021/03/31 23:27:43 | 1-15 | 3 | 0.205 | -70.171 | 12605 | -1 | 1 | CPER2-010-3 | 20-202 | H2O | 1617255168 | 0.198 | 0.83 | 49.535 | 0 | 0.002 | 1.205 | 93.999 | 28.741 | 2.162 | 59.027 | 0.857 | -2.216 | 0 | 0 | 1 | 43.75 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-15838 | 2021/03/31 23:36:39 | 1-15 | 4 | 0.252 | -70.093 | 12495 | 0 | 1 | CPER2-010-3 | 20-202 | H2O | 1617255705 | 0.224 | 0.913 | 45.72 | 0 | -0.001 | 1.111 | 93.77 | 28.512 | 2.141 | 57.89 | 0.858 | -2.152 | 0 | 0 | 1 | 43.875 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-15839 | 2021/03/31 23:46:27 | 1-16 | 1 | [5 empty] | CPER2-D-2 | 20-202 | [18 empty] | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-15839 | 2021/03/31 23:55:23 | 1-16 | 2 | [5 empty] | CPER2-D-2 | 20-202 | [18 empty] | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-15839 | 2021/04/01 00:04:20 | 1-16 | 3 | 23.958 | -36.482 | 12041 | -1 | 1 | CPER2-D-2 | 20-202 | H2O | 1617257365 | 0.322 | 2.211 | 48.664 | -0.003 | -0.002 | 1.2 | 697.54 | -18.641 | 6.757 | 59.772 | 0.861 | -15.762 | 0 | 0 | 1 | 43.812 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-15839 | 2021/04/01 00:13:15 | 1-16 | 4 | 24.021 | -36.7 | 12105 | 0 | 1 | CPER2-D-2 | 20-202 | H2O | 1617257900 | 0.318 | 2.008 | 45.434 | -0.003 | -0.001 | 1.12 | 701.692 | -19.287 | 6.812 | 63.187 | 0.862 | -15.386 | 0 | 0 | 1 | 43.875 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-15840 | 2021/04/01 00:23:03 | 1-17 | 1 | [5 empty] | CPER2-C-3 | 20-202 | [18 empty] | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-15840 | 2021/04/01 00:31:59 | 1-17 | 2 | [5 empty] | CPER2-C-3 | 20-202 | [18 empty] | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-15840 | 2021/04/01 00:40:56 | 1-17 | 3 | -5.377 | -78.324 | 12459 | -1 | 1 | CPER2-C-3 | 20-202 | H2O | 1617259562 | 0.225 | 0.93 | 50.206 | 0 | 0.001 | 1.224 | 327.385 | 20.053 | 2.233 | 54.835 | 0.859 | -2.477 | 0 | 0 | 1 | 43.938 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-15840 | 2021/04/01 00:49:54 | 1-17 | 4 | -5.436 | -78.401 | 12364 | 0 | 1 | CPER2-C-3 | 20-202 | H2O | 1617260099 | 0.224 | 0.848 | 51.616 | 0 | 0.001 | 1.252 | 325.173 | 20.055 | 2.202 | 53.372 | 0.859 | -2.666 | 0 | 0 | 1 | 43.875 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-15841 | 2021/04/01 00:59:42 | 1-18 | 1 | -2.299 | -76.197 | 12604 | -1 | 1 | CPER2-2030-1 | 20-202 | H2O | 1617260687 | 0.2 | 0.933 | 51.367 | 0 | 0.001 | 1.257 | 96.766 | 29.013 | 2.167 | 57.464 | 0.858 | -2.29 | 0 | 0 | 1 | 44 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-15841 | 2021/04/01 01:08:38 | 1-18 | 2 | -2.168 | -76.051 | 12669 | -1 | 1 | CPER2-2030-1 | 20-202 | H2O | 1617261223 | 0.172 | 0.847 | 45.701 | 0 | 0 | 1.11 | 95.495 | 28.818 | 2.153 | 59.252 | 0.858 | -2.052 | 0 | 0 | 1 | 43.875 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-15841 | 2021/04/01 01:17:34 | 1-18 | 3 | -2.102 | -76.184 | 12564 | -1 | 1 | CPER2-2030-1 | 20-202 | H2O | 1617261759 | 0.18 | 0.852 | 53.857 | 0 | 0 | 1.318 | 94.612 | 29.001 | 2.147 | 56.632 | 0.858 | -2.304 | 0 | 0 | 1 | 43.875 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-15841 | 2021/04/01 01:26:31 | 1-18 | 4 | -2.117 | -75.886 | 12480 | 0 | 1 | CPER2-2030-1 | 20-202 | H2O | 1617262295 | 0.194 | 0.905 | 48.042 | 0 | 0 | 1.174 | 93.863 | 28.452 | 2.145 | 58.826 | 0.858 | -2.012 | 0 | 0 | 1 | 43.75 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-15842 | 2021/04/01 01:36:19 | 1-19 | 1 | -0.576 | -66.921 | 12565 | -1 | 1 | CPER2-E-4 | 20-202 | H2O | 1617262884 | 0.204 | 0.895 | 47.47 | 0 | 0 | 1.161 | 112.249 | 28.209 | 2.173 | 57.065 | 0.858 | -2.388 | 0 | 0 | 1 | 43.562 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-15842 | 2021/04/01 01:45:15 | 1-19 | 2 | -0.486 | -65.887 | 12518 | -1 | 1 | CPER2-E-4 | 20-202 | H2O | 1617263420 | 0.203 | 0.795 | 44.589 | 0 | -0.001 | 1.085 | 111.966 | 28.107 | 2.175 | 58.142 | 0.859 | -2.158 | 0 | 0 | 1 | 43.375 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-15842 | 2021/04/01 01:54:11 | 1-19 | 3 | -0.472 | -65.811 | 12509 | -1 | 1 | CPER2-E-4 | 20-202 | H2O | 1617263956 | 0.213 | 0.997 | 49.286 | 0 | -0.002 | 1.208 | 112.062 | 27.849 | 2.175 | 58.057 | 0.858 | -2.326 | 0 | 0 | 1 | 43.438 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-15842 | 2021/04/01 02:03:07 | 1-19 | 4 | -0.449 | -65.539 | 12468 | 0 | 1 | CPER2-E-4 | 20-202 | H2O | 1617264492 | 0.216 | 0.94 | 50.64 | 0 | -0.003 | 1.24 | 111.724 | 27.896 | 2.164 | 57.929 | 0.857 | -2.177 | 0 | 0 | 1 | 43.438 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-15843 | 2021/04/01 02:12:57 | 1-20 | 1 | -0.135 | -65.065 | 12608 | -1 | 1 | CPER2-A-1 | 20-202 | H2O | 1617265082 | 0.216 | 0.942 | 48.429 | 0 | 0.001 | 1.181 | 114.045 | 27.9 | 2.198 | 58.46 | 0.858 | -2.22 | 0 | 0 | 1 | 43.438 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-15843 | 2021/04/01 02:21:53 | 1-20 | 2 | -0.17 | -64.733 | 12475 | -1 | 1 | CPER2-A-1 | 20-202 | H2O | 1617265617 | 0.214 | 0.889 | 46.112 | 0 | -0.001 | 1.123 | 113.653 | 27.406 | 2.169 | 57.669 | 0.857 | -2.237 | 0 | 0 | 1 | 43.312 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-15843 | 2021/04/01 02:30:50 | 1-20 | 3 | -0.081 | -64.873 | 12490 | -1 | 1 | CPER2-A-1 | 20-202 | H2O | 1617266155 | 0.215 | 0.937 | 49.128 | 0.001 | -0.002 | 1.201 | 113.599 | 27.77 | 2.176 | 58.106 | 0.858 | -2.166 | 0 | 0 | 1 | 43.438 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-15843 | 2021/04/01 02:39:46 | 1-20 | 4 | -0.085 | -64.877 | 12410 | 0 | 1 | CPER2-A-1 | 20-202 | H2O | 1617266691 | 0.22 | 0.837 | 46.741 | 0 | -0.002 | 1.142 | 113.261 | 27.609 | 2.157 | 57.692 | 0.858 | -2.238 | 0 | 0 | 1 | 43.438 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-15844 | 2021/04/01 02:49:34 | 1-21 | 1 | [5 empty] | CPER2-B-3 | 20-202 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-15844 | 2021/04/01 02:59:33 | 1-21 | 2 | [5 empty] | CPER2-B-3 | 20-202 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-15844 | 2021/04/01 03:09:33 | 1-21 | 3 | [5 empty] | CPER2-B-3 | 20-202 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-15844 | 2021/04/01 03:19:32 | 1-21 | 4 | [5 empty] | CPER2-B-3 | 20-202 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-15845 | 2021/04/01 03:30:21 | 1-22 | 1 | -8.591 | -67.034 | 12350 | -1 | 1 | EV | standard | H2O | 1617269726 | 0.198 | 0.789 | 38.686 | 0.001 | 0 | 0.943 | 93.35 | 27.983 | 2.124 | 58.174 | 0.857 | -1.886 | 0 | 0 | 1 | 43.938 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-15845 | 2021/04/01 03:39:16 | 1-22 | 2 | -8.692 | -66.719 | 12431 | -1 | 1 | EV | standard | H2O | 1617270261 | 0.189 | 0.907 | 40.155 | 0 | 0.001 | 0.968 | 93.29 | 28.523 | 2.149 | 58.202 | 0.857 | -2.055 | 0 | 0 | 1 | 43.562 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-15845 | 2021/04/01 03:48:12 | 1-22 | 3 | -8.793 | -67.117 | 12446 | -1 | 1 | EV | standard | H2O | 1617270797 | 0.205 | 0.941 | 48.039 | 0 | -0.001 | 1.174 | 93.159 | 28.443 | 2.149 | 59.068 | 0.858 | -1.904 | 0 | 0 | 1 | 43.5 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-15845 | 2021/04/01 03:57:08 | 1-22 | 4 | -8.842 | -66.767 | 12470 | 0 | 1 | EV | standard | H2O | 1617271333 | 0.19 | 0.91 | 48.648 | 0.001 | 0.002 | 1.188 | 93.32 | 28.398 | 2.152 | 58.1 | 0.858 | -2.056 | 0 | 0 | 1 | 43.375 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-15846 | 2021/04/01 04:06:55 | 1-23 | 1 | [5 empty] | CPER2-B-1 | 20-202 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-15846 | 2021/04/01 04:16:54 | 1-23 | 2 | [5 empty] | CPER2-B-1 | 20-202 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-15846 | 2021/04/01 04:26:53 | 1-23 | 3 | [5 empty] | CPER2-B-1 | 20-202 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-15846 | 2021/04/01 04:36:52 | 1-23 | 4 | [5 empty] | CPER2-B-1 | 20-202 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-15847 | 2021/04/01 04:47:41 | 1-24 | 1 | [5 empty] | CPER2-B-4 | 20-202 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-15847 | 2021/04/01 04:57:39 | 1-24 | 2 | [5 empty] | CPER2-B-4 | 20-202 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-15847 | 2021/04/01 05:07:37 | 1-24 | 3 | [5 empty] | CPER2-B-4 | 20-202 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-15847 | 2021/04/01 05:17:35 | 1-24 | 4 | [5 empty] | CPER2-B-4 | 20-202 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-15848 | 2021/04/01 05:28:25 | 1-25 | 1 | -1.828 | -68.403 | 12309 | -1 | 1 | CPER2-A-4 | 20-202 | H2O | 1617276810 | 0.216 | 0.8 | 37.991 | 0 | 0.001 | 0.927 | 99.741 | 27.856 | 2.117 | 57.308 | 0.857 | -2.169 | 0 | 0 | 1 | 44.188 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-15848 | 2021/04/01 05:37:20 | 1-25 | 2 | -1.546 | -68.236 | 12471 | -1 | 1 | CPER2-A-4 | 20-202 | H2O | 1617277346 | 0.209 | 0.817 | 41.09 | 0 | -0.004 | 1.002 | 99.955 | 28.073 | 2.145 | 58.641 | 0.858 | -2.094 | 0 | 0 | 1 | 43.875 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-15848 | 2021/04/01 05:46:16 | 1-25 | 3 | [5 empty] | CPER2-A-4 | 20-202 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-15848 | 2021/04/01 05:56:14 | 1-25 | 4 | [5 empty] | CPER2-A-4 | 20-202 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-15849 | 2021/04/01 06:07:04 | 1-26 | 1 | [5 empty] | CPER2-E-1 | 20-202 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-15849 | 2021/04/01 06:17:02 | 1-26 | 2 | [5 empty] | CPER2-E-1 | 20-202 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-15849 | 2021/04/01 06:27:01 | 1-26 | 3 | [5 empty] | CPER2-E-1 | 20-202 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-15849 | 2021/04/01 06:36:59 | 1-26 | 4 | [5 empty] | CPER2-E-1 | 20-202 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-15850 | 2021/04/01 06:47:50 | 1-27 | 1 | 1.228 | -54.804 | 12574 | -1 | 1 | CPER2-D-5 | 20-202 | H2O | 1617281575 | 0.23 | 0.924 | 39.616 | -0.001 | 0.001 | 0.964 | 95.692 | 28.39 | 2.198 | 58.776 | 0.858 | -2.287 | 0 | 0 | 1 | 43.812 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-15850 | 2021/04/01 06:56:46 | 1-27 | 2 | 1.478 | -53.594 | 12582 | -1 | 1 | CPER2-D-5 | 20-202 | H2O | 1617282111 | 0.203 | 0.937 | 43.582 | 0 | -0.003 | 1.058 | 95.562 | 28.536 | 2.214 | 59.092 | 0.858 | -2.305 | 0 | 0 | 1 | 43.75 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-15850 | 2021/04/01 07:05:42 | 1-27 | 3 | 1.496 | -53.302 | 12667 | -1 | 1 | CPER2-D-5 | 20-202 | H2O | 1617282648 | 0.197 | 0.797 | 48.121 | 0 | -0.002 | 1.178 | 95.643 | 28.988 | 2.218 | 59.257 | 0.858 | -2.389 | 0 | 0 | 1 | 43.688 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-15850 | 2021/04/01 07:14:38 | 1-27 | 4 | 1.518 | -53.286 | 12617 | 0 | 1 | CPER2-D-5 | 20-202 | H2O | 1617283183 | 0.202 | 0.911 | 45.145 | 0 | -0.004 | 1.097 | 95.569 | 28.732 | 2.208 | 58.539 | 0.858 | -2.367 | 0 | 0 | 1 | 43.562 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-15851 | 2021/04/01 07:24:27 | 1-28 | 1 | 1.349 | -64.269 | 12731 | -1 | 1 | CPER2-1020-3 | 20-202 | H2O | 1617283772 | 0.226 | 0.99 | 46.957 | 0 | 0.002 | 1.137 | 93.967 | 28.903 | 2.196 | 60.013 | 0.858 | -2.044 | 0 | 0 | 1 | 43.625 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-15851 | 2021/04/01 07:33:24 | 1-28 | 2 | 1.334 | -65.358 | 12572 | -1 | 1 | CPER2-1020-3 | 20-202 | H2O | 1617284309 | 0.233 | 0.942 | 51.049 | 0 | 0.001 | 1.252 | 93.692 | 28.766 | 2.16 | 57.702 | 0.857 | -2.227 | 0 | 0 | 1 | 43.562 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-15851 | 2021/04/01 07:42:19 | 1-28 | 3 | 1.379 | -65.533 | 12839 | -1 | 1 | CPER2-1020-3 | 20-202 | H2O | 1617284844 | 0.183 | 0.812 | 49.411 | 0.001 | 0 | 1.209 | 94.226 | 29.713 | 2.208 | 58.478 | 0.858 | -2.301 | 0 | 0 | 1 | 43.625 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-15851 | 2021/04/01 07:51:15 | 1-28 | 4 | 1.34 | -65.64 | 12611 | 0 | 1 | CPER2-1020-3 | 20-202 | H2O | 1617285380 | 0.203 | 0.944 | 47.37 | 0 | -0.001 | 1.158 | 93.625 | 28.848 | 2.179 | 59.167 | 0.857 | -2.155 | 0 | 0 | 1 | 43.688 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-15852 | 2021/04/01 08:01:00 | 1-29 | 1 | -1.732 | -68.646 | 12637 | -1 | 1 | CPER2-A-3 | 20-202 | H2O | 1617285965 | 0.191 | 0.882 | 51.211 | 0.001 | 0.005 | 1.246 | 101.793 | 28.53 | 2.176 | 59.477 | 0.858 | -2.059 | 0 | 0 | 1 | 43.75 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-15852 | 2021/04/01 08:09:56 | 1-29 | 2 | -1.836 | -68.9 | 12636 | -1 | 1 | CPER2-A-3 | 20-202 | H2O | 1617286501 | 0.19 | 0.761 | 48.332 | 0 | 0 | 1.178 | 101.853 | 28.964 | 2.193 | 58.497 | 0.858 | -2.165 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-15852 | 2021/04/01 08:18:52 | 1-29 | 3 | -1.879 | -68.957 | 12616 | -1 | 1 | CPER2-A-3 | 20-202 | H2O | 1617287037 | 0.194 | 0.858 | 45.405 | 0 | 0 | 1.109 | 101.944 | 28.39 | 2.171 | 59.205 | 0.858 | -2.11 | 0 | 0 | 1 | 43.812 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-15852 | 2021/04/01 08:27:48 | 1-29 | 4 | -1.845 | -68.809 | 12599 | 0 | 1 | CPER2-A-3 | 20-202 | H2O | 1617287574 | 0.219 | 0.817 | 43.831 | 0.001 | -0.003 | 1.071 | 102.075 | 28.425 | 2.167 | 57.651 | 0.858 | -2.338 | 0 | 0 | 1 | 43.875 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-15853 | 2021/04/01 08:37:34 | 1-30 | 1 | -0.213 | -68.29 | 12659 | -1 | 1 | CPER2-B-5 | 20-202 | H2O | 1617288159 | 0.217 | 0.836 | 53.517 | 0 | 0.003 | 1.31 | 104.078 | 28.763 | 2.178 | 59.313 | 0.858 | -1.966 | 0 | 0 | 1 | 43.688 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-15853 | 2021/04/01 08:46:30 | 1-30 | 2 | -0.142 | -68.188 | 12489 | -1 | 1 | CPER2-B-5 | 20-202 | H2O | 1617288695 | 0.194 | 0.868 | 46.873 | 0 | 0 | 1.145 | 103.833 | 27.871 | 2.167 | 58.551 | 0.858 | -2.248 | 0 | 0 | 1 | 43.875 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-15853 | 2021/04/01 08:55:27 | 1-30 | 3 | -0.142 | -68.388 | 12620 | -1 | 1 | CPER2-B-5 | 20-202 | H2O | 1617289232 | 0.2 | 0.942 | 50.73 | 0 | 0.004 | 1.236 | 104.315 | 28.285 | 2.176 | 58.327 | 0.858 | -2.308 | 0 | 0 | 1 | 43.75 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-15853 | 2021/04/01 09:04:23 | 1-30 | 4 | -0.111 | -68.307 | 12612 | 0 | 1 | CPER2-B-5 | 20-202 | H2O | 1617289768 | 0.224 | 0.894 | 50.365 | 0 | 0 | 1.224 | 104.133 | 28.326 | 2.17 | 58.667 | 0.858 | -2.167 | 0 | 0 | 1 | 43.875 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-15854 | 2021/04/01 09:14:08 | 1-31 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-15854 | 2021/04/01 09:23:04 | 1-31 | 2 | [5 empty] | EV | standard | [18 empty] | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-15854 | 2021/04/01 09:31:59 | 1-31 | 3 | -8.938 | -67.205 | 12663 | -1 | 1 | EV | standard | H2O | 1617291425 | 0.184 | 0.859 | 45.403 | 0 | -0.002 | 1.107 | 93.536 | 29.333 | 2.168 | 58.652 | 0.858 | -2.117 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-15854 | 2021/04/01 09:40:56 | 1-31 | 4 | -8.994 | -67.235 | 12664 | 0 | 1 | EV | standard | H2O | 1617291960 | 0.202 | 0.813 | 51.771 | 0 | 0.002 | 1.268 | 93.639 | 29.031 | 2.167 | 59.268 | 0.858 | -2.179 | 0 | 0 | 1 | 43.875 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-15855 | 2021/04/01 09:50:42 | 1-32 | 1 | -1.54 | -72.576 | 12608 | -1 | 1 | CPER2-3040-2 | 20-202 | H2O | 1617292547 | 0.199 | 0.901 | 45.409 | -0.001 | 0.001 | 1.105 | 93.664 | 28.97 | 2.152 | 58.089 | 0.858 | -2.189 | 0 | 0 | 1 | 43.938 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-15855 | 2021/04/01 09:59:39 | 1-32 | 2 | -1.255 | -73.149 | 12518 | -1 | 1 | CPER2-3040-2 | 20-202 | H2O | 1617293084 | 0.215 | 0.927 | 52.238 | 0 | 0.004 | 1.274 | 93.547 | 28.587 | 2.146 | 57.338 | 0.858 | -2.272 | 0 | 0 | 1 | 43.75 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-15855 | 2021/04/01 10:08:35 | 1-32 | 3 | -1.137 | -73.13 | 12691 | -1 | 1 | CPER2-3040-2 | 20-202 | H2O | 1617293620 | 0.186 | 0.754 | 46.336 | 0 | -0.003 | 1.129 | 93.758 | 29.166 | 2.172 | 59.006 | 0.858 | -2.075 | 0 | 0 | 1 | 43.75 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-15855 | 2021/04/01 10:17:31 | 1-32 | 4 | -1.091 | -73.207 | 12630 | 0 | 1 | CPER2-3040-2 | 20-202 | H2O | 1617294156 | 0.203 | 0.835 | 48.449 | 0 | 0 | 1.183 | 93.605 | 29.067 | 2.169 | 59.161 | 0.857 | -2.015 | 0 | 0 | 1 | 43.688 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-15856 | 2021/04/01 10:27:19 | 1-33 | 1 | -5.892 | -82.033 | 12423 | -1 | 1 | CPER2-C-5 | 20-202 | H2O | 1617294743 | 0.247 | 0.9 | 52.327 | 0 | 0.003 | 1.287 | 133.932 | 26.801 | 2.129 | 56.998 | 0.858 | -2.097 | 0 | 0 | 1 | 43.938 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-15856 | 2021/04/01 10:36:15 | 1-33 | 2 | -6.082 | -82.652 | 12582 | -1 | 1 | CPER2-C-5 | 20-202 | H2O | 1617295280 | 0.214 | 0.812 | 47.777 | 0 | 0.001 | 1.169 | 135.171 | 27.689 | 2.155 | 56.35 | 0.857 | -2.221 | 0 | 0 | 1 | 43.75 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-15856 | 2021/04/01 10:45:11 | 1-33 | 3 | -6.215 | -83.337 | 12594 | -1 | 1 | CPER2-C-5 | 20-202 | H2O | 1617295816 | 0.201 | 0.9 | 50.268 | 0 | 0.002 | 1.219 | 135.276 | 27.26 | 2.149 | 57.994 | 0.858 | -2.102 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-15856 | 2021/04/01 10:54:06 | 1-33 | 4 | -6.235 | -83.264 | 12669 | 0 | 1 | CPER2-C-5 | 20-202 | H2O | 1617296351 | 0.198 | 0.737 | 49.955 | -0.001 | 0.002 | 1.22 | 135.71 | 27.977 | 2.164 | 57.365 | 0.858 | -2.248 | 0 | 0 | 1 | 43.688 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-15857 | 2021/04/01 11:03:54 | 1-34 | 1 | -7.044 | -88.559 | 12687 | -1 | 1 | CPER2-1020-1 | 20-202 | H2O | 1617296939 | 0.201 | 0.845 | 48.462 | 0.001 | 0.001 | 1.188 | 94.378 | 29.15 | 2.146 | 56.518 | 0.858 | -2.257 | 0 | 0 | 1 | 43.562 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-15857 | 2021/04/01 11:12:50 | 1-34 | 2 | -7.107 | -89.139 | 12687 | -1 | 1 | CPER2-1020-1 | 20-202 | H2O | 1617297474 | 0.207 | 0.844 | 50.231 | 0 | 0.004 | 1.227 | 94.012 | 29.188 | 2.135 | 57.804 | 0.858 | -2.089 | 0 | 0 | 1 | 43.688 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-15857 | 2021/04/01 11:21:46 | 1-34 | 3 | -7.088 | -89.165 | 12707 | -1 | 1 | CPER2-1020-1 | 20-202 | H2O | 1617298011 | 0.213 | 0.831 | 50.402 | 0 | 0.002 | 1.23 | 94 | 29.211 | 2.158 | 57.819 | 0.858 | -2.294 | 0 | 0 | 1 | 43.5 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-15857 | 2021/04/01 11:30:42 | 1-34 | 4 | -7.155 | -89.396 | 12662 | 0 | 1 | CPER2-1020-1 | 20-202 | H2O | 1617298547 | 0.211 | 0.957 | 50.121 | 0.001 | 0.003 | 1.22 | 93.685 | 28.928 | 2.152 | 58.692 | 0.858 | -2.033 | 0 | 0 | 1 | 43.625 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-15858 | 2021/04/01 11:40:29 | 1-35 | 1 | -5.346 | -84.147 | 12520 | -1 | 1 | CPER2-3040-1 | 20-202 | H2O | 1617299134 | 0.222 | 0.867 | 50.013 | 0.001 | 0.001 | 1.22 | 93.408 | 28.502 | 2.118 | 58.167 | 0.857 | -2.071 | 0 | 0 | 1 | 43.438 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-15858 | 2021/04/01 11:49:26 | 1-35 | 2 | -5.272 | -83.836 | 12492 | -1 | 1 | CPER2-3040-1 | 20-202 | H2O | 1617299671 | 0.193 | 0.846 | 47.777 | 0.001 | 0.001 | 1.168 | 93.286 | 28.671 | 2.119 | 57.579 | 0.858 | -2.031 | 0 | 0 | 1 | 43.562 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-15858 | 2021/04/01 11:58:23 | 1-35 | 3 | -5.299 | -83.949 | 12599 | -1 | 1 | CPER2-3040-1 | 20-202 | H2O | 1617300208 | 0.192 | 0.944 | 51.567 | -0.001 | 0.004 | 1.268 | 93.434 | 28.951 | 2.14 | 58.059 | 0.859 | -2.074 | 0 | 0 | 1 | 43.75 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-15858 | 2021/04/01 12:07:19 | 1-35 | 4 | -5.258 | -83.786 | 12637 | 0 | 1 | CPER2-3040-1 | 20-202 | H2O | 1617300745 | 0.21 | 0.819 | 45.5 | 0 | 0 | 1.103 | 93.586 | 28.681 | 2.154 | 59.063 | 0.858 | -2.013 | 0 | 0 | 1 | 44.125 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-15859 | 2021/04/01 12:17:02 | 1-36 | 1 | -8.353 | -88.824 | 12436 | -1 | 1 | CPER2-deep-2 | 20-202 | H2O | 1617301327 | 0.196 | 0.88 | 48.837 | 0.001 | 0.003 | 1.201 | 93.095 | 28.484 | 2.103 | 57.706 | 0.858 | -2.01 | 0 | 0 | 1 | 44.062 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-15859 | 2021/04/01 12:25:58 | 1-36 | 2 | -8.552 | -89.337 | 12336 | -1 | 1 | CPER2-deep-2 | 20-202 | H2O | 1617301863 | 0.188 | 0.766 | 49.38 | 0 | 0.001 | 1.206 | 92.845 | 28.34 | 2.098 | 57.004 | 0.858 | -2.02 | 0 | 0 | 1 | 44 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-15859 | 2021/04/01 12:34:54 | 1-36 | 3 | -8.625 | -89.389 | 12443 | -1 | 1 | CPER2-deep-2 | 20-202 | H2O | 1617302400 | 0.219 | 0.838 | 50.177 | 0 | 0.001 | 1.232 | 93.062 | 28.398 | 2.115 | 57.713 | 0.859 | -1.942 | 0 | 0 | 1 | 44.062 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-15859 | 2021/04/01 12:43:51 | 1-36 | 4 | -8.572 | -89.446 | 12434 | 0 | 1 | CPER2-deep-2 | 20-202 | H2O | 1617302936 | 0.226 | 0.924 | 50.993 | 0.001 | 0.002 | 1.245 | 93.191 | 28.328 | 2.104 | 56.989 | 0.858 | -2.151 | 0 | 0 | 1 | 43.938 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-15860 | 2021/04/01 12:53:35 | 1-37 | 1 | -3.459 | -81.027 | 12814 | -1 | 1 | CPER2-B-2 | 20-202 | H2O | 1617303520 | 0.201 | 0.899 | 50.732 | 0 | 0.001 | 1.23 | 104.82 | 29.225 | 2.172 | 58.305 | 0.859 | -2.259 | 0 | 0 | 1 | 43.875 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-15860 | 2021/04/01 13:02:31 | 1-37 | 2 | -3.218 | -80.082 | 12727 | -1 | 1 | CPER2-B-2 | 20-202 | H2O | 1617304055 | 0.207 | 0.838 | 50.711 | 0 | 0.001 | 1.237 | 104.731 | 29.042 | 2.176 | 57.918 | 0.859 | -2.244 | 0 | 0 | 1 | 43.812 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-15860 | 2021/04/01 13:11:27 | 1-37 | 3 | -3.136 | -79.899 | 12660 | -1 | 1 | CPER2-B-2 | 20-202 | H2O | 1617304592 | 0.208 | 0.942 | 47.942 | 0 | 0 | 1.171 | 104.602 | 28.604 | 2.163 | 57.876 | 0.859 | -2.178 | 0 | 0 | 1 | 43.688 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-15860 | 2021/04/01 13:20:23 | 1-37 | 4 | -3.137 | -79.973 | 12607 | 0 | 1 | CPER2-B-2 | 20-202 | H2O | 1617305128 | 0.193 | 0.888 | 48.016 | 0 | 0.002 | 1.165 | 104.443 | 28.685 | 2.156 | 57.323 | 0.858 | -2.258 | 0 | 0 | 1 | 43.625 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-15861 | 2021/04/01 13:30:07 | 1-38 | 1 | -1.667 | -76.478 | 12509 | -1 | 1 | CPER2-1020-2 | 20-202 | H2O | 1617305712 | 0.224 | 0.962 | 48.818 | -0.001 | 0.004 | 1.183 | 93.732 | 28.548 | 2.148 | 57.795 | 0.857 | -2.092 | 0 | 0 | 1 | 43.562 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-15861 | 2021/04/01 13:39:03 | 1-38 | 2 | -1.577 | -76.274 | 12626 | -1 | 1 | CPER2-1020-2 | 20-202 | H2O | 1617306248 | 0.204 | 0.903 | 46.867 | -0.001 | 0.004 | 1.145 | 93.976 | 28.57 | 2.156 | 59.136 | 0.859 | -2.062 | 0 | 0 | 1 | 43.562 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-15861 | 2021/04/01 13:47:59 | 1-38 | 3 | -1.591 | -76.258 | 12617 | -1 | 1 | CPER2-1020-2 | 20-202 | H2O | 1617306784 | 0.205 | 0.819 | 45.323 | 0.001 | 0 | 1.107 | 93.778 | 29.058 | 2.161 | 58.681 | 0.858 | -2.068 | 0 | 0 | 1 | 43.562 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-15861 | 2021/04/01 13:56:55 | 1-38 | 4 | -1.548 | -76.089 | 12711 | 0 | 1 | CPER2-1020-2 | 20-202 | H2O | 1617307320 | 0.191 | 0.905 | 48.584 | 0 | 0.001 | 1.19 | 93.875 | 29.598 | 2.181 | 58.253 | 0.858 | -2.125 | 0 | 0 | 1 | 43.562 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-15862 | 2021/04/01 14:06:40 | 1-39 | 1 | -0.954 | -74.291 | 12767 | -1 | 1 | CPER2-010-1 | 20-202 | H2O | 1617307905 | 0.203 | 0.874 | 51.571 | 0 | 0.002 | 1.266 | 94.013 | 29.239 | 2.185 | 58.903 | 0.857 | -2.198 | 0 | 0 | 1 | 43.562 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-15862 | 2021/04/01 14:15:35 | 1-39 | 2 | -0.956 | -74.194 | 12699 | -1 | 1 | CPER2-010-1 | 20-202 | H2O | 1617308440 | 0.209 | 0.903 | 50.872 | 0 | 0 | 1.244 | 93.716 | 29.504 | 2.175 | 57.527 | 0.858 | -2.231 | 0 | 0 | 1 | 43.625 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-15862 | 2021/04/01 14:24:31 | 1-39 | 3 | -0.864 | -74.138 | 12682 | -1 | 1 | CPER2-010-1 | 20-202 | H2O | 1617308975 | 0.204 | 0.92 | 47.697 | 0 | 0.002 | 1.171 | 93.626 | 29.167 | 2.171 | 58.454 | 0.858 | -2.049 | 0 | 0 | 1 | 43.562 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-15862 | 2021/04/01 14:33:27 | 1-39 | 4 | -0.899 | -74.122 | 12723 | 0 | 1 | CPER2-010-1 | 20-202 | H2O | 1617309512 | 0.2 | 0.75 | 53.942 | 0 | 0.001 | 1.32 | 93.82 | 29.151 | 2.179 | 59.383 | 0.858 | -2.254 | 0 | 0 | 1 | 43.625 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-15863 | 2021/04/01 14:43:12 | 1-40 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-15863 | 2021/04/01 14:52:08 | 1-40 | 2 | [5 empty] | EV | standard | [18 empty] | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-15863 | 2021/04/01 15:01:04 | 1-40 | 3 | [5 empty] | EV | standard | [18 empty] | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-15863 | 2021/04/01 15:09:59 | 1-40 | 4 | [5 empty] | EV | standard | [18 empty] | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-15864 | 2021/04/01 15:19:45 | 1-41 | 1 | [5 empty] | CPER2-2030-2 | 20-202 | H2O | [17 empty] | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-15864 | 2021/04/01 15:29:44 | 1-41 | 2 | [5 empty] | CPER2-2030-2 | 20-202 | H2O | [17 empty] | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-15864 | 2021/04/01 15:39:42 | 1-41 | 3 | [5 empty] | CPER2-2030-2 | 20-202 | H2O | [17 empty] | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-15864 | 2021/04/01 15:49:41 | 1-41 | 4 | [5 empty] | CPER2-2030-2 | 20-202 | H2O | [17 empty] | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-15865 | 2021/04/01 16:00:29 | 1-42 | 1 | [5 empty] | CPER2-A-2 | 20-202 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-15865 | 2021/04/01 16:10:27 | 1-42 | 2 | [5 empty] | CPER2-A-2 | 20-202 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-15865 | 2021/04/01 16:20:25 | 1-42 | 3 | [5 empty] | CPER2-A-2 | 20-202 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-15865 | 2021/04/01 16:30:24 | 1-42 | 4 | [5 empty] | CPER2-A-2 | 20-202 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-15866 | 2021/04/01 16:41:11 | 1-43 | 1 | [5 empty] | CPER2-C-2 | 20-202 | H2O | [17 empty] | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-15866 | 2021/04/01 16:51:08 | 1-43 | 2 | [5 empty] | CPER2-C-2 | 20-202 | H2O | [17 empty] | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-15866 | 2021/04/01 17:01:06 | 1-43 | 3 | -2.943 | -75.78 | 12212 | -1 | 1 | CPER2-C-2 | 20-202 | H2O | 1617318372 | 0.213 | 0.774 | 36.859 | 0 | 0.001 | 0.893 | 92.998 | 27.85 | 2.086 | 56.966 | 0.858 | -2.119 | 0 | 0 | 1 | 44 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-15866 | 2021/04/01 17:10:02 | 1-43 | 4 | -2.943 | -75.78 | 12212 | -1 | 1 | CPER2-C-2 | 20-202 | H2O | 1617318372 | 0.213 | 0.774 | 36.859 | 0 | 0.001 | 0.893 | 92.998 | 27.85 | 2.086 | 56.966 | 0.858 | -2.119 | 0 | 0 | 1 | 44 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-15867 | 2021/04/01 17:20:49 | 1-44 | 1 | 2.684 | -69.798 | 12524 | -1 | 1 | CPER2-D-4 | 20-202 | H2O | 1617319554 | 0.205 | 0.966 | 41.935 | 0 | -0.001 | 1.026 | 110.124 | 20.672 | 3.112 | 62.785 | 0.858 | -5.546 | 0 | 0 | 1 | 43.938 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-15867 | 2021/04/01 17:29:45 | 1-44 | 2 | 2.671 | -70.008 | 12716 | -1 | 1 | CPER2-D-4 | 20-202 | H2O | 1617320090 | 0.207 | 1.05 | 46.845 | 0 | -0.002 | 1.143 | 111.042 | 21.308 | 3.137 | 61.95 | 0.859 | -5.983 | 0 | 0 | 1 | 43.875 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-15867 | 2021/04/01 17:38:41 | 1-44 | 3 | 2.672 | -69.91 | 12623 | -1 | 1 | CPER2-D-4 | 20-202 | H2O | 1617320626 | 0.212 | 0.966 | 48.637 | 0 | -0.001 | 1.188 | 110.786 | 20.926 | 3.119 | 61.95 | 0.858 | -5.883 | 0 | 0 | 1 | 43.812 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-15867 | 2021/04/01 17:47:38 | 1-44 | 4 | 2.651 | -69.96 | 12568 | 0 | 1 | CPER2-D-4 | 20-202 | H2O | 1617321163 | 0.209 | 0.927 | 43.498 | -0.001 | 0.001 | 1.064 | 110.623 | 21.005 | 3.096 | 60.225 | 0.858 | -6.109 | 0 | 0 | 1 | 43.812 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-15868 | 2021/04/01 17:57:23 | 1-45 | 1 | -0.729 | -68.742 | 12224 | -1 | 1 | CPER2-E-3 | 20-202 | H2O | 1617321748 | 0.187 | 0.866 | 45.251 | 0 | -0.002 | 1.11 | 134.269 | 26.46 | 2.13 | 55.32 | 0.858 | -2.356 | 0 | 0 | 1 | 43.812 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-15868 | 2021/04/01 18:06:19 | 1-45 | 2 | -0.558 | -67.874 | 12439 | -1 | 1 | CPER2-E-3 | 20-202 | H2O | 1617322285 | 0.218 | 0.848 | 50.069 | 0 | -0.002 | 1.215 | 135.646 | 26.828 | 2.147 | 57.796 | 0.859 | -2.144 | 0 | 0 | 1 | 43.812 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-15868 | 2021/04/01 18:15:16 | 1-45 | 3 | -0.583 | -67.693 | 12732 | -1 | 1 | CPER2-E-3 | 20-202 | H2O | 1617322821 | 0.207 | 0.751 | 46.375 | 0 | 0.002 | 1.134 | 137.093 | 27.704 | 2.201 | 58.818 | 0.858 | -2.138 | 0 | 0 | 1 | 43.812 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-15868 | 2021/04/01 18:24:12 | 1-45 | 4 | -0.58 | -67.691 | 12659 | 0 | 1 | CPER2-E-3 | 20-202 | H2O | 1617323357 | 0.21 | 0.832 | 44.182 | 0 | 0 | 1.072 | 136.913 | 27.769 | 2.187 | 58.081 | 0.858 | -2.262 | 0 | 0 | 1 | 43.812 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-15869 | 2021/04/01 18:33:58 | 1-46 | 1 | -8.938 | -67.205 | 12663 | -1 | 1 | EV | standard | H2O | 1617291425 | 0.184 | 0.859 | 45.403 | 0 | -0.002 | 1.107 | 93.536 | 29.333 | 2.168 | 58.652 | 0.858 | -2.117 | 0 | 0 | 1 | 43.688 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-15869 | 2021/04/01 18:42:54 | 1-46 | 2 | -8.994 | -67.235 | 12664 | 0 | 1 | EV | standard | H2O | 1617291960 | 0.202 | 0.813 | 51.771 | 0 | 0.002 | 1.268 | 93.639 | 29.031 | 2.167 | 59.268 | 0.858 | -2.179 | 0 | 0 | 1 | 43.875 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-15869 | 2021/04/01 18:51:49 | 1-46 | 3 | -8.938 | -67.205 | 12663 | -1 | 1 | EV | standard | H2O | 1617291425 | 0.184 | 0.859 | 45.403 | 0 | -0.002 | 1.107 | 93.536 | 29.333 | 2.168 | 58.652 | 0.858 | -2.117 | 0 | 0 | 1 | 43.688 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-15869 | 2021/04/01 19:00:44 | 1-46 | 4 | -8.994 | -67.235 | 12664 | 0 | 1 | EV | standard | H2O | 1617291960 | 0.202 | 0.813 | 51.771 | 0 | 0.002 | 1.268 | 93.639 | 29.031 | 2.167 | 59.268 | 0.858 | -2.179 | 0 | 0 | 1 | 43.875 | 1 | 46 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 94.03775
varray: [95.549, 94.806, 94.87, 94.531, 94.097, 94.04, 93.877, 93.923, 93.639, 93.704, 94.25, 94.085, 94.217, 93.221, 93.646, 93.285, 93.808, 93.771, 93.718, 93.718]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_shift
baseline_shift_STD | 0.558095905265
varray: [95.549, 94.806, 94.87, 94.531, 94.097, 94.04, 93.877, 93.923, 93.639, 93.704, 94.25, 94.085, 94.217, 93.221, 93.646, 93.285, 93.808, 93.771, 93.718, 93.718]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_shift
ch4_ppm_MEAN | -2.1403
varray: [-2.141, -2.26, -2.474, -2.392, -2.19, -2.147, -2.178, -2.124, -2.107, -2.032, -2.101, -2.085, -2.159, -2.014, -2.151, -2.015, -2.102, -2.018, -2.058, -2.058]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.1198170096
varray: [-2.141, -2.26, -2.474, -2.392, -2.19, -2.147, -2.178, -2.124, -2.107, -2.032, -2.101, -2.085, -2.159, -2.014, -2.151, -2.015, -2.102, -2.018, -2.058, -2.058]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: ch4_ppm
residuals_MEAN | 2.2112
varray: [2.352, 2.307, 2.309, 2.304, 2.302, 2.31, 2.313, 2.305, 2.306, 2.327, 2.164, 2.138, 2.153, 2.096, 2.107, 2.091, 2.107, 2.095, 2.069, 2.069]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: residuals
residuals_STD | 0.107873024773
varray: [2.352, 2.307, 2.309, 2.304, 2.302, 2.31, 2.313, 2.305, 2.306, 2.327, 2.164, 2.138, 2.153, 2.096, 2.107, 2.091, 2.107, 2.095, 2.069, 2.069]
inst: std_standards
operation: std_standards(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: residuals
slope_shift_MEAN | 29.2467
varray: [29.985, 29.312, 29.42, 29.279, 29.367, 28.788, 29.117, 28.836, 28.855, 28.981, 30.106, 29.542, 30.234, 29.331, 29.596, 29.281, 29.66, 28.784, 28.23, 28.23]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: slope_shift
slope_shift_STD | 0.542787454767
varray: [29.985, 29.312, 29.42, 29.279, 29.367, 28.788, 29.117, 28.836, 28.855, 28.981, 30.106, 29.542, 30.234, 29.331, 29.596, 29.281, 29.66, 28.784, 28.23, 28.23]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: slope_shift
baseline_curvature_MEAN | 59.1483
varray: [62.308, 60.275, 59.774, 59.726, 60.031, 60.833, 60.751, 61.877, 61.839, 62.403, 58.625, 57.862, 57.213, 57.021, 56.498, 56.14, 56.593, 57.839, 57.679, 57.679]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_curvature
baseline_curvature_STD | 2.08096420213
varray: [62.308, 60.275, 59.774, 59.726, 60.031, 60.833, 60.751, 61.877, 61.839, 62.403, 58.625, 57.862, 57.213, 57.021, 56.498, 56.14, 56.593, 57.839, 57.679, 57.679]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.2112
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 94.03775
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 2.1403
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 29.2467
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.16180953716
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.37300953716
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.215746049546
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.42694604955
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.323619074319
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.53481907432
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.22112
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.3594510288
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 2.4997510288
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.279047952633
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 94.3167979526
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 94.5958459053
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 1.6742877158
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 95.7120377158
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 10.0457262948
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 104.083476295
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.71393727384
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 31.9606372738
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 6.24289260638
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 65.3911926064
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -3.54393958544
actual_array: [ 1.93 -15.84]
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15, 1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855, -13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697, -9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071, -3.275, -2.986, -2.935, -2.978, -1.739, -1.647, -1.644, -1.618, -0.512, -0.446, -0.443, -0.452, 0.338, 0.359, 0.371, 0.353, 0.247, 0.059, 0.044, 0.055, 0.0, 0.0, -1.409, -1.275, 2.439, 2.451, 2.441, 2.452, -3.6, -3.696, -3.728, -3.779, 0.0, 0.0, 0.0, 0.0, -4.408, -4.274, -4.17, -4.164, 0.011, 0.147, 0.205, 0.252, 0.0, 0.0, 23.958, 24.021, 0.0, 0.0, -5.377, -5.436, -2.299, -2.168, -2.102, -2.117, -0.576, -0.486, -0.472, -0.449, -0.135, -0.17, -0.081, -0.085, 0.0, 0.0, 0.0, 0.0, -8.591, -8.692, -8.793, -8.842, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -1.828, -1.546, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 1.228, 1.478, 1.496, 1.518, 1.349, 1.334, 1.379, 1.34, -1.732, -1.836, -1.879, -1.845, -0.213, -0.142, -0.142, -0.111, 0.0, 0.0, -8.938, -8.994, -1.54, -1.255, -1.137, -1.091, -5.892, -6.082, -6.215, -6.235, -7.044, -7.107, -7.088, -7.155, -5.346, -5.272, -5.299, -5.258, -8.353, -8.552, -8.625, -8.572, -3.459, -3.218, -3.136, -3.137, -1.667, -1.577, -1.591, -1.548, -0.954, -0.956, -0.864, -0.899, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -2.943, -2.943, 2.684, 2.671, 2.672, 2.651, -0.729, -0.558, -0.583, -0.58, -8.938, -8.994, -8.938, -8.994]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.60662980769
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0, 5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0, 16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0, 22274.0, 22810.0, 23346.0, 23883.0, 24473.0, 25011.0, 25547.0, 26085.0, 26675.0, 27213.0, 27749.0, 28286.0, 28874.0, 29409.0, 29946.0, 30482.0, 31070.0, 31607.0, 32143.0, 32679.0, 33267.0, 33866.0, 34465.0, 35000.0, 35589.0, 36125.0, 36661.0, 37198.0, 37787.0, 38324.0, 38860.0, 39397.0, 39987.0, 40524.0, 41060.0, 41596.0, 42186.0, 42724.0, 43261.0, 43797.0, 44383.0, 44920.0, 45457.0, 45993.0, 46581.0, 47117.0, 47654.0, 48189.0, 48777.0, 49313.0, 49850.0, 50388.0, 50976.0, 51512.0, 52048.0, 52585.0, 53173.0, 53709.0, 54245.0, 54781.0, 55371.0, 55907.0, 56444.0, 56980.0, 57568.0, 58167.0, 58767.0, 59366.0, 60015.0, 60550.0, 61086.0, 61622.0, 62209.0, 62808.0, 63407.0, 64006.0, 64655.0, 65253.0, 65851.0, 66449.0, 67099.0, 67634.0, 68170.0, 68768.0, 69418.0, 70016.0, 70615.0, 71213.0, 71864.0, 72400.0, 72936.0, 73472.0, 74061.0, 74598.0, 75133.0, 75669.0, 76254.0, 76790.0, 77326.0, 77862.0, 78448.0, 78984.0, 79521.0, 80057.0, 80642.0, 81178.0, 81713.0, 82250.0, 82836.0, 83373.0, 83909.0, 84445.0, 85033.0, 85569.0, 86105.0, 86640.0, 87228.0, 87764.0, 88300.0, 88836.0, 89423.0, 89960.0, 90497.0, 91033.0, 91616.0, 92152.0, 92688.0, 93225.0, 93809.0, 94345.0, 94881.0, 95417.0, 96001.0, 96537.0, 97073.0, 97609.0, 98194.0, 98729.0, 99265.0, 99801.0, 100386.0, 100922.0, 101458.0, 101993.0, 102579.0, 103178.0, 103776.0, 104375.0, 105023.0, 105621.0, 106219.0, 106818.0, 107465.0, 108062.0, 108660.0, 109196.0, 109843.0, 110379.0, 110915.0, 111452.0, 112037.0, 112573.0, 113110.0, 113646.0, 114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -59.0133164078
actual_array: [ 18.1 -119. ]
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081, 9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734, -97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188, -71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32, -77.274, -78.11, -78.339, -78.407, -76.889, -76.67, -76.642, -76.918, -70.769, -69.933, -69.74, -69.897, -65.338, -65.124, -65.103, -64.973, -72.67, -73.755, -73.937, -74.205, 0.0, 0.0, -64.831, -63.96, -61.022, -61.238, -61.296, -61.244, -77.166, -78.083, -78.099, -78.306, 0.0, 0.0, 0.0, 0.0, -75.215, -75.962, -75.887, -76.259, -70.672, -70.237, -70.171, -70.093, 0.0, 0.0, -36.482, -36.7, 0.0, 0.0, -78.324, -78.401, -76.197, -76.051, -76.184, -75.886, -66.921, -65.887, -65.811, -65.539, -65.065, -64.733, -64.873, -64.877, 0.0, 0.0, 0.0, 0.0, -67.034, -66.719, -67.117, -66.767, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -68.403, -68.236, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -54.804, -53.594, -53.302, -53.286, -64.269, -65.358, -65.533, -65.64, -68.646, -68.9, -68.957, -68.809, -68.29, -68.188, -68.388, -68.307, 0.0, 0.0, -67.205, -67.235, -72.576, -73.149, -73.13, -73.207, -82.033, -82.652, -83.337, -83.264, -88.559, -89.139, -89.165, -89.396, -84.147, -83.836, -83.949, -83.786, -88.824, -89.337, -89.389, -89.446, -81.027, -80.082, -79.899, -79.973, -76.478, -76.274, -76.258, -76.089, -74.291, -74.194, -74.138, -74.122, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -75.78, -75.78, -69.798, -70.008, -69.91, -69.96, -68.742, -67.874, -67.693, -67.691, -67.205, -67.235, -67.205, -67.235]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -53.8917692308
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0, 5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0, 10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0, 16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0, 22274.0, 22810.0, 23346.0, 23883.0, 24473.0, 25011.0, 25547.0, 26085.0, 26675.0, 27213.0, 27749.0, 28286.0, 28874.0, 29409.0, 29946.0, 30482.0, 31070.0, 31607.0, 32143.0, 32679.0, 33267.0, 33866.0, 34465.0, 35000.0, 35589.0, 36125.0, 36661.0, 37198.0, 37787.0, 38324.0, 38860.0, 39397.0, 39987.0, 40524.0, 41060.0, 41596.0, 42186.0, 42724.0, 43261.0, 43797.0, 44383.0, 44920.0, 45457.0, 45993.0, 46581.0, 47117.0, 47654.0, 48189.0, 48777.0, 49313.0, 49850.0, 50388.0, 50976.0, 51512.0, 52048.0, 52585.0, 53173.0, 53709.0, 54245.0, 54781.0, 55371.0, 55907.0, 56444.0, 56980.0, 57568.0, 58167.0, 58767.0, 59366.0, 60015.0, 60550.0, 61086.0, 61622.0, 62209.0, 62808.0, 63407.0, 64006.0, 64655.0, 65253.0, 65851.0, 66449.0, 67099.0, 67634.0, 68170.0, 68768.0, 69418.0, 70016.0, 70615.0, 71213.0, 71864.0, 72400.0, 72936.0, 73472.0, 74061.0, 74598.0, 75133.0, 75669.0, 76254.0, 76790.0, 77326.0, 77862.0, 78448.0, 78984.0, 79521.0, 80057.0, 80642.0, 81178.0, 81713.0, 82250.0, 82836.0, 83373.0, 83909.0, 84445.0, 85033.0, 85569.0, 86105.0, 86640.0, 87228.0, 87764.0, 88300.0, 88836.0, 89423.0, 89960.0, 90497.0, 91033.0, 91616.0, 92152.0, 92688.0, 93225.0, 93809.0, 94345.0, 94881.0, 95417.0, 96001.0, 96537.0, 97073.0, 97609.0, 98194.0, 98729.0, 99265.0, 99801.0, 100386.0, 100922.0, 101458.0, 101993.0, 102579.0, 103178.0, 103776.0, 104375.0, 105023.0, 105621.0, 106219.0, 106818.0, 107465.0, 108062.0, 108660.0, 109196.0, 109843.0, 110379.0, 110915.0, 111452.0, 112037.0, 112573.0, 113110.0, 113646.0, 114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1755
varray: [2.337, 2.15, 2.152, 2.156, 2.165, 2.168, 2.158, 2.165, 2.139, 2.165]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: residuals
SAMP_residuals_STD | 0.0574364382987
varray: [2.337, 2.15, 2.152, 2.156, 2.165, 2.168, 2.158, 2.165, 2.139, 2.165]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.9503
varray: [-1.193, -1.623, -2.101, -1.936, -2.245, -2.089, -2.055, -2.263, -1.973, -2.025]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.320497547226
varray: [-1.193, -1.623, -2.101, -1.936, -2.245, -2.089, -2.055, -2.263, -1.973, -2.025]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 103.6161
varray: [137.42, 105.312, 103.618, 101.852, 100.493, 99.153, 97.851, 97.394, 96.589, 96.479]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 12.2503244261
varray: [137.42, 105.312, 103.618, 101.852, 100.493, 99.153, 97.851, 97.394, 96.589, 96.479]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.2677
varray: [30.875, 29.187, 29.273, 28.979, 29.358, 28.996, 28.987, 29.09, 28.879, 29.053]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.583364770581
varray: [30.875, 29.187, 29.273, 28.979, 29.358, 28.996, 28.987, 29.09, 28.879, 29.053]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_d18O_STD | 0.256856661281
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.96484636289
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.9482
varray: [67.041, 59.647, 57.083, 58.729, 56.563, 58.184, 58.463, 56.995, 58.318, 58.459]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.9901469975
varray: [67.041, 59.647, 57.083, 58.729, 56.563, 58.184, 58.463, 56.995, 58.318, 58.459]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1755
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.9503
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 103.6161
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.2677
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0357
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0357
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.61450795948
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.2001
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.3217
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_mean | -69.7708
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
d18O_stdev | 0.256856661281
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_stdev | 2.96484636289
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.3217
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.7708
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.256856661281
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.96484636289
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4923302452
actual_array: [ 1.93 -15.84]
varray: [-10.009, -9.441, -9.335, -9.271, -9.229, -9.239, -9.189, -9.18, -9.174, -9.15]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.3217
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.1213020053
actual_array: [ 18.1 -119. ]
varray: [-77.838, -70.953, -69.587, -69.119, -68.795, -68.578, -68.253, -68.131, -68.373, -68.081]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.7708
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [1.0, 537.0, 1074.0, 1610.0, 2147.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.86225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 9.57835
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3135
varray: [2.352, 2.307, 2.309, 2.304, 2.302, 2.31, 2.313, 2.305, 2.306, 2.327]
inst: mean_all
operation: mean_all(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: residuals
SAMP_residuals_STD | 0.0152406619862
varray: [2.352, 2.307, 2.309, 2.304, 2.302, 2.31, 2.313, 2.305, 2.306, 2.327]
inst: std_all
operation: std_all(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.2045
varray: [-2.141, -2.26, -2.474, -2.392, -2.19, -2.147, -2.178, -2.124, -2.107, -2.032]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.135343390431
varray: [-2.141, -2.26, -2.474, -2.392, -2.19, -2.147, -2.178, -2.124, -2.107, -2.032]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 94.3036
varray: [95.549, 94.806, 94.87, 94.531, 94.097, 94.04, 93.877, 93.923, 93.639, 93.704]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.616096348886
varray: [95.549, 94.806, 94.87, 94.531, 94.097, 94.04, 93.877, 93.923, 93.639, 93.704]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.194
varray: [29.985, 29.312, 29.42, 29.279, 29.367, 28.788, 29.117, 28.836, 28.855, 28.981]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.363173359277
varray: [29.985, 29.312, 29.42, 29.279, 29.367, 28.788, 29.117, 28.836, 28.855, 28.981]
inst: std_all
operation: std_all(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: slope_shift
SAMP_d18O_STD | 0.272742959987
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.13390127654
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.9817
varray: [62.308, 60.275, 59.774, 59.726, 60.031, 60.833, 60.751, 61.877, 61.839, 62.403]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.04507320849
varray: [62.308, 60.275, 59.774, 59.726, 60.031, 60.833, 60.751, 61.877, 61.839, 62.403]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3135
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.2045
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 94.3036
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.194
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1023
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1023
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.626447178
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8334
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.6835
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.6821
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.6835
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
dD_mean | 17.6821
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
d18O_stdev | 0.272742959987
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
dD_stdev | 3.13390127654
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.6835
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.6821
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.272742959987
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.13390127654
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [1.963, 2.518, 2.698, 2.744, 2.738, 2.802, 2.833, 2.835, 2.849, 2.855]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.6835
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [9.404, 15.869, 17.621, 18.154, 18.71, 19.038, 19.205, 19.564, 19.522, 19.734]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 17.6821
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [5420.0, 5957.0, 6493.0, 7029.0, 7567.0, 8103.0, 8640.0, 9177.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.72425
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.26585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1089
varray: [2.164, 2.138, 2.153, 2.096, 2.107, 2.091, 2.107, 2.095, 2.069, 2.069]
inst: mean_all
operation: mean_all(residuals)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: residuals
SAMP_residuals_STD | 0.0328276644852
varray: [2.164, 2.138, 2.153, 2.096, 2.107, 2.091, 2.107, 2.095, 2.069, 2.069]
inst: std_all
operation: std_all(residuals)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.0761
varray: [-2.101, -2.085, -2.159, -2.014, -2.151, -2.015, -2.102, -2.018, -2.058, -2.058]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0532007727596
varray: [-2.101, -2.085, -2.159, -2.014, -2.151, -2.015, -2.102, -2.018, -2.058, -2.058]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.7719
varray: [94.25, 94.085, 94.217, 93.221, 93.646, 93.285, 93.808, 93.771, 93.718, 93.718]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.347731202064
varray: [94.25, 94.085, 94.217, 93.221, 93.646, 93.285, 93.808, 93.771, 93.718, 93.718]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.2994
varray: [30.106, 29.542, 30.234, 29.331, 29.596, 29.281, 29.66, 28.784, 28.23, 28.23]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.695632917246
varray: [30.106, 29.542, 30.234, 29.331, 29.596, 29.281, 29.66, 28.784, 28.23, 28.23]
inst: std_all
operation: std_all(slope_shift)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: slope_shift
SAMP_d18O_STD | 0.367767378168
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.43755728727
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.3149
varray: [58.625, 57.862, 57.213, 57.021, 56.498, 56.14, 56.593, 57.839, 57.679, 57.679]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.76165863599
varray: [58.625, 57.862, 57.213, 57.021, 56.498, 56.14, 56.593, 57.839, 57.679, 57.679]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1089
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.0761
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.7719
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.2994
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1023
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1023
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.626447178
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8334
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.4898
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
dD_mean | -109.5693
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
d18O_stdev | 0.367767378168
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
dD_stdev | 4.43755728727
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.4898
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
measured_dD_mean | -109.5693
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.367767378168
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.43755728727
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.569, -14.218, -14.344, -14.583, -14.631, -14.693, -14.723, -14.743, -14.697, -14.697]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.4898
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-97.88, -107.132, -108.714, -110.752, -111.18, -111.765, -111.748, -112.146, -112.188, -112.188]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -109.5693
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [10837.0, 11374.0, 11911.0, 12448.0, 12985.0, 13521.0, 14057.0, 14594.0, 15128.0, 15727.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.92885
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.26585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1638
varray: [2.156, 2.158, 2.154, 2.172, 2.167, 2.156, 2.173, 2.15, 2.174, 2.178]
inst: mean_all
operation: mean_all(residuals)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: residuals
SAMP_residuals_STD | 0.0100531918646
varray: [2.156, 2.158, 2.154, 2.172, 2.167, 2.156, 2.173, 2.15, 2.174, 2.178]
inst: std_all
operation: std_all(residuals)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.1298
varray: [-1.898, -2.296, -2.204, -2.223, -2.221, -2.107, -2.03, -2.265, -2.055, -1.999]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.131371229727
varray: [-1.898, -2.296, -2.204, -2.223, -2.221, -2.107, -2.03, -2.265, -2.055, -1.999]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.5602
varray: [93.542, 93.65, 93.424, 93.728, 93.667, 93.428, 93.549, 93.47, 93.659, 93.485]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.109373773009
varray: [93.542, 93.65, 93.424, 93.728, 93.667, 93.428, 93.549, 93.47, 93.659, 93.485]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.1015
varray: [28.969, 29.667, 29.088, 28.985, 29.204, 28.592, 29.077, 29.083, 29.192, 29.158]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.265313755559
varray: [28.969, 29.667, 29.088, 28.985, 29.204, 28.592, 29.077, 29.083, 29.192, 29.158]
inst: std_all
operation: std_all(slope_shift)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: slope_shift
SAMP_d18O_STD | 0.10312996116
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.37097161005
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.3041
varray: [59.076, 56.856, 57.381, 58.559, 57.581, 58.648, 59.191, 58.013, 58.63, 59.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.807715358279
varray: [59.076, 56.856, 57.381, 58.559, 57.581, 58.648, 59.191, 58.013, 58.63, 59.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1638
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.1298
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.5602
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.1015
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0474
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0474
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.14363241679
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.8442
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1323
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
dD_mean | -68.0294
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
d18O_stdev | 0.10312996116
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
dD_stdev | 1.37097161005
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1323
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.0294
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.10312996116
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.37097161005
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2963493912
actual_array: [ 1.93 -15.84]
varray: [-9.409, -9.191, -9.1, -9.093, -9.081, -9.097, -9.116, -9.094, -9.071, -9.071]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.1323
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.2451266548
actual_array: [ 18.1 -119. ]
varray: [-71.772, -68.513, -67.993, -67.716, -67.449, -67.364, -67.225, -67.576, -67.366, -67.32]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.0294
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [16855.0, 17391.0, 17927.0, 18462.0, 19000.0, 19536.0, 20072.0, 20609.0, 21146.0, 21683.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.87395
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.47755
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.144
varray: [2.159, 2.146, 2.138, 2.133]
inst: mean_all
operation: mean_all(residuals)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: residuals
SAMP_residuals_STD | 0.0113431330181
varray: [2.159, 2.146, 2.138, 2.133]
inst: std_all
operation: std_all(residuals)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.2175
varray: [-2.197, -2.19, -2.149, -2.334]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0805005175967
varray: [-2.197, -2.19, -2.149, -2.334]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 94.75775
varray: [94.939, 94.874, 94.617, 94.601]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.173922156917
varray: [94.939, 94.874, 94.617, 94.601]
inst: std_all
operation: std_all(baseline_shift)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.8495
varray: [29.029, 28.876, 28.826, 28.667]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.149198972293
varray: [29.029, 28.876, 28.826, 28.667]
inst: std_all
operation: std_all(slope_shift)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: slope_shift
SAMP_d18O_STD | 0.155949778241
varray: [-3.275, -2.986, -2.935, -2.978]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.521383096517
varray: [-77.274, -78.11, -78.339, -78.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.806
varray: [58.073, 58.124, 58.085, 56.942]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.576411311478
varray: [58.073, 58.124, 58.085, 56.942]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.144
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.2175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 94.75775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.8495
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0672
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0672
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.03907380608
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3423
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-3.275, -2.986, -2.935, -2.978]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-77.274, -78.11, -78.339, -78.407]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.0435
varray: [-3.275, -2.986, -2.935, -2.978]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
dD_mean | -78.0325
varray: [-77.274, -78.11, -78.339, -78.407]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
d18O_stdev | 0.155949778241
varray: [-3.275, -2.986, -2.935, -2.978]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
dD_stdev | 0.521383096517
varray: [-77.274, -78.11, -78.339, -78.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.0435
varray: [-3.275, -2.986, -2.935, -2.978]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.0325
varray: [-77.274, -78.11, -78.339, -78.407]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.155949778241
varray: [-3.275, -2.986, -2.935, -2.978]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.521383096517
varray: [-77.274, -78.11, -78.339, -78.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.99598918088
actual_array: [ 1.93 -15.84]
varray: [-3.275, -2.986, -2.935, -2.978]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.0435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -85.0224148418
actual_array: [ 18.1 -119. ]
varray: [-77.274, -78.11, -78.339, -78.407]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.0325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [22274.0, 22810.0, 23346.0, 23883.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.89375
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.72
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16125
varray: [2.168, 2.156, 2.16, 2.161]
inst: mean_all
operation: mean_all(residuals)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: residuals
SAMP_residuals_STD | 0.00499165971062
varray: [2.168, 2.156, 2.16, 2.161]
inst: std_all
operation: std_all(residuals)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.11975
varray: [-1.992, -2.035, -2.163, -2.289]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.134187369003
varray: [-1.992, -2.035, -2.163, -2.289]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 94.0055
varray: [94.127, 93.889, 93.902, 94.104]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.127474180392
varray: [94.127, 93.889, 93.902, 94.104]
inst: std_all
operation: std_all(baseline_shift)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.0625
varray: [29.034, 29.017, 28.973, 29.226]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.111989582849
varray: [29.034, 29.017, 28.973, 29.226]
inst: std_all
operation: std_all(slope_shift)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0529591037185
varray: [-1.739, -1.647, -1.644, -1.618]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.143838740725
varray: [-76.889, -76.67, -76.642, -76.918]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.52125
varray: [59.61, 58.952, 58.065, 57.458]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.950366025978
varray: [59.61, 58.952, 58.065, 57.458]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.11975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 94.0055
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.0625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04995
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04995
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.25895441389
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.62705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-1.739, -1.647, -1.644, -1.618]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-76.889, -76.67, -76.642, -76.918]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.662
varray: [-1.739, -1.647, -1.644, -1.618]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
dD_mean | -76.77975
varray: [-76.889, -76.67, -76.642, -76.918]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0529591037185
varray: [-1.739, -1.647, -1.644, -1.618]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
dD_stdev | 0.143838740725
varray: [-76.889, -76.67, -76.642, -76.918]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.662
varray: [-1.739, -1.647, -1.644, -1.618]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
measured_dD_mean | -76.77975
varray: [-76.889, -76.67, -76.642, -76.918]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0529591037185
varray: [-1.739, -1.647, -1.644, -1.618]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.143838740725
varray: [-76.889, -76.67, -76.642, -76.918]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.56648786197
actual_array: [ 1.93 -15.84]
varray: [-1.739, -1.647, -1.644, -1.618]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.662
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.6727084731
actual_array: [ 18.1 -119. ]
varray: [-76.889, -76.67, -76.642, -76.918]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.77975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [24473.0, 25011.0, 25547.0, 26085.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.8765
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.03225
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15175
varray: [2.151, 2.158, 2.15, 2.148]
inst: mean_all
operation: mean_all(residuals)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: residuals
SAMP_residuals_STD | 0.00434932945023
varray: [2.151, 2.158, 2.15, 2.148]
inst: std_all
operation: std_all(residuals)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.2255
varray: [-2.272, -2.144, -2.145, -2.341]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0976814550806
varray: [-2.272, -2.144, -2.145, -2.341]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 113.01475
varray: [113.084, 112.86, 113.182, 112.933]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.145371650148
varray: [113.084, 112.86, 113.182, 112.933]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.06075
varray: [28.228, 27.567, 28.303, 28.145]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.335432432342
varray: [28.228, 27.567, 28.303, 28.145]
inst: std_all
operation: std_all(slope_shift)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0327146756059
varray: [-0.512, -0.446, -0.443, -0.452]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.463799076469
varray: [-70.769, -69.933, -69.74, -69.897]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.06125
varray: [57.037, 57.505, 57.039, 56.664]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.344385805941
varray: [57.037, 57.505, 57.039, 56.664]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.2255
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 113.01475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.06075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05945
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05945
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.6885853835
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.08705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-0.512, -0.446, -0.443, -0.452]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-70.769, -69.933, -69.74, -69.897]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.46325
varray: [-0.512, -0.446, -0.443, -0.452]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
dD_mean | -70.08475
varray: [-70.769, -69.933, -69.74, -69.897]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0327146756059
varray: [-0.512, -0.446, -0.443, -0.452]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
dD_stdev | 0.463799076469
varray: [-70.769, -69.933, -69.74, -69.897]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.46325
varray: [-0.512, -0.446, -0.443, -0.452]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.08475
varray: [-70.769, -69.933, -69.74, -69.897]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0327146756059
varray: [-0.512, -0.446, -0.443, -0.452]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.463799076469
varray: [-70.769, -69.933, -69.74, -69.897]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.32608633751
actual_array: [ 1.93 -15.84]
varray: [-0.512, -0.446, -0.443, -0.452]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.46325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.459550111
actual_array: [ 18.1 -119. ]
varray: [-70.769, -69.933, -69.74, -69.897]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.08475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [26675.0, 27213.0, 27749.0, 28286.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.886
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 18.977
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1595
varray: [2.151, 2.154, 2.163, 2.17]
inst: mean_all
operation: mean_all(residuals)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: residuals
SAMP_residuals_STD | 0.00866025403784
varray: [2.151, 2.154, 2.163, 2.17]
inst: std_all
operation: std_all(residuals)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.0805
varray: [-1.912, -2.002, -2.153, -2.255]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.153037032555
varray: [-1.912, -2.002, -2.153, -2.255]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 120.368
varray: [119.677, 120.374, 120.702, 120.719]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.487262420741
varray: [119.677, 120.374, 120.702, 120.719]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.50475
varray: [27.013, 27.115, 27.941, 27.95]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.510648199187
varray: [27.013, 27.115, 27.941, 27.95]
inst: std_all
operation: std_all(slope_shift)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: slope_shift
SAMP_d18O_STD | 0.013720422734
varray: [0.338, 0.359, 0.371, 0.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.151213976426
varray: [-65.338, -65.124, -65.103, -64.973]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.32275
varray: [59.001, 58.963, 57.457, 57.87]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.779839459975
varray: [59.001, 58.963, 57.457, 57.87]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1595
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.0805
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 120.368
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.50475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0517
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0517
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.33809696093
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.82555
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.338, 0.359, 0.371, 0.353]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-65.338, -65.124, -65.103, -64.973]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.35525
varray: [0.338, 0.359, 0.371, 0.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
dD_mean | -65.1345
varray: [-65.338, -65.124, -65.103, -64.973]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
d18O_stdev | 0.013720422734
varray: [0.338, 0.359, 0.371, 0.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
dD_stdev | 0.151213976426
varray: [-65.338, -65.124, -65.103, -64.973]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.35525
varray: [0.338, 0.359, 0.371, 0.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.1345
varray: [-65.338, -65.124, -65.103, -64.973]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.013720422734
varray: [0.338, 0.359, 0.371, 0.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.151213976426
varray: [-65.338, -65.124, -65.103, -64.973]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.479146902459
actual_array: [ 1.93 -15.84]
varray: [0.338, 0.359, 0.371, 0.353]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.35525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.1261763721
actual_array: [ 18.1 -119. ]
varray: [-65.338, -65.124, -65.103, -64.973]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.1345
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [28874.0, 29409.0, 29946.0, 30482.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.87825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 26.33025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.686
varray: [2.699, 2.665, 2.687, 2.693]
inst: mean_all
operation: mean_all(residuals)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: residuals
SAMP_residuals_STD | 0.0148323969742
varray: [2.699, 2.665, 2.687, 2.693]
inst: std_all
operation: std_all(residuals)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.38125
varray: [-4.142, -4.61, -4.42, -4.353]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.193103383364
varray: [-4.142, -4.61, -4.42, -4.353]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 309.839
varray: [308.872, 308.674, 310.03, 311.78]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.42551464391
varray: [308.872, 308.674, 310.03, 311.78]
inst: std_all
operation: std_all(baseline_shift)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 17.18475
varray: [17.281, 17.21, 16.884, 17.364]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.21014498963
varray: [17.281, 17.21, 16.884, 17.364]
inst: std_all
operation: std_all(slope_shift)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0973734221096
varray: [0.247, 0.059, 0.044, 0.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.673683098101
varray: [-72.67, -73.755, -73.937, -74.205]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.876
varray: [59.405, 56.474, 57.803, 57.822]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.19884527776
varray: [59.405, 56.474, 57.803, 57.822]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.686
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.38125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 309.839
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 17.18475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.4748
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4748
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 21.4725036179
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2723
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.247, 0.059, 0.044, 0.055]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-72.67, -73.755, -73.937, -74.205]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.10125
varray: [0.247, 0.059, 0.044, 0.055]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
dD_mean | -73.64175
varray: [-72.67, -73.755, -73.937, -74.205]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0973734221096
varray: [0.247, 0.059, 0.044, 0.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
dD_stdev | 0.673683098101
varray: [-72.67, -73.755, -73.937, -74.205]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.10125
varray: [0.247, 0.059, 0.044, 0.055]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.64175
varray: [-72.67, -73.755, -73.937, -74.205]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0973734221096
varray: [0.247, 0.059, 0.044, 0.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.673683098101
varray: [-72.67, -73.755, -73.937, -74.205]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.741972334962
actual_array: [ 1.93 -15.84]
varray: [0.247, 0.059, 0.044, 0.055]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.10125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.2918435082
actual_array: [ 18.1 -119. ]
varray: [-72.67, -73.755, -73.937, -74.205]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.64175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [31070.0, 31607.0, 32143.0, 32679.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.35175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 215.80125
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.09275
varray: [0.0, 0.0, 2.175, 2.196]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: residuals
SAMP_residuals_STD | 1.26182813806
varray: [0.0, 0.0, 2.175, 2.196]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.09275
varray: [0.0, 0.0, -2.102, -2.269]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.26363955172
varray: [0.0, 0.0, -2.102, -2.269]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.3085
varray: [0.0, 0.0, 121.343, 119.891]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 69.6407802967
varray: [0.0, 0.0, 121.343, 119.891]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.937
varray: [0.0, 0.0, 27.472, 28.276]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: slope_shift
SAMP_slope_shift_STD | 16.0964083364
varray: [0.0, 0.0, 27.472, 28.276]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: slope_shift
SAMP_d18O_STD | 0.776732901325
varray: [0.0, 0.0, -1.409, -1.275]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 37.1804596562
varray: [0.0, 0.0, -64.831, -63.96]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 29.0625
varray: [0.0, 0.0, 58.455, 57.795]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.5595660729
varray: [0.0, 0.0, 58.455, 57.795]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.09275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.09275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.3085
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.937
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.11845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.11845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.5811324168
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -30.0858
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, -1.409, -1.275]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, -64.831, -63.96]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.671
varray: [0.0, 0.0, -1.409, -1.275]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
dD_mean | -32.19775
varray: [0.0, 0.0, -64.831, -63.96]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
d18O_stdev | 0.776732901325
varray: [0.0, 0.0, -1.409, -1.275]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
dD_stdev | 37.1804596562
varray: [0.0, 0.0, -64.831, -63.96]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.671
varray: [0.0, 0.0, -1.409, -1.275]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.19775
varray: [0.0, 0.0, -64.831, -63.96]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.776732901325
varray: [0.0, 0.0, -1.409, -1.275]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
measured_dD_stdev | 37.1804596562
varray: [0.0, 0.0, -64.831, -63.96]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.54105477689
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -1.409, -1.275]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.671
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.6402923268
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -64.831, -63.96]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -32.19775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [33267.0, 33866.0, 34465.0, 35000.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 92.945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -33.72925
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.77075
varray: [2.779, 2.76, 2.772, 2.772]
inst: mean_all
operation: mean_all(residuals)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: residuals
SAMP_residuals_STD | 0.00788986691903
varray: [2.779, 2.76, 2.772, 2.772]
inst: std_all
operation: std_all(residuals)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.777
varray: [-4.918, -4.639, -4.692, -4.859]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132757925062
varray: [-4.918, -4.639, -4.692, -4.859]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 108.73875
varray: [109.167, 108.642, 108.648, 108.498]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.293799676651
varray: [109.167, 108.642, 108.648, 108.498]
inst: std_all
operation: std_all(baseline_shift)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.9375
varray: [23.292, 22.633, 22.7, 23.125]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.321430656078
varray: [23.292, 22.633, 22.7, 23.125]
inst: std_all
operation: std_all(slope_shift)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00670198975429
varray: [2.439, 2.451, 2.441, 2.452]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.121490740388
varray: [-61.022, -61.238, -61.296, -61.244]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.35675
varray: [61.382, 61.57, 61.066, 61.409]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.210854412648
varray: [61.382, 61.57, 61.066, 61.409]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.77075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.777
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 108.73875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.9375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.55955
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.55955
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 25.30526411
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.20845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [2.439, 2.451, 2.441, 2.452]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-61.022, -61.238, -61.296, -61.244]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.44575
varray: [2.439, 2.451, 2.441, 2.452]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
dD_mean | -61.2
varray: [-61.022, -61.238, -61.296, -61.244]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00670198975429
varray: [2.439, 2.451, 2.441, 2.452]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
dD_stdev | 0.121490740388
varray: [-61.022, -61.238, -61.296, -61.244]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.44575
varray: [2.439, 2.451, 2.441, 2.452]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.2
varray: [-61.022, -61.238, -61.296, -61.244]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00670198975429
varray: [2.439, 2.451, 2.441, 2.452]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.121490740388
varray: [-61.022, -61.238, -61.296, -61.244]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.6839891867
actual_array: [ 1.93 -15.84]
varray: [2.439, 2.451, 2.441, 2.452]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.44575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.887166428
actual_array: [ 18.1 -119. ]
varray: [-61.022, -61.238, -61.296, -61.244]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -61.2
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [35589.0, 36125.0, 36661.0, 37198.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.267
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 14.701
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17175
varray: [2.173, 2.164, 2.16, 2.19]
inst: mean_all
operation: mean_all(residuals)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: residuals
SAMP_residuals_STD | 0.0133260396718
varray: [2.173, 2.164, 2.16, 2.19]
inst: std_all
operation: std_all(residuals)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.35375
varray: [-2.328, -2.343, -2.506, -2.238]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.111589650058
varray: [-2.328, -2.343, -2.506, -2.238]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 218.34325
varray: [217.578, 218.055, 218.5, 219.24]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.7065
varray: [217.578, 218.055, 218.5, 219.24]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 24.27875
varray: [24.137, 24.297, 24.445, 24.236]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.12896091139
varray: [24.137, 24.297, 24.445, 24.236]
inst: std_all
operation: std_all(slope_shift)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0753630103787
varray: [-3.6, -3.696, -3.728, -3.779]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.508577427733
varray: [-77.166, -78.083, -78.099, -78.306]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.634
varray: [55.856, 54.906, 55.208, 56.566]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.736968565589
varray: [55.856, 54.906, 55.208, 56.566]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.35375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 218.34325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 24.27875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03945
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03945
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.78409913169
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.5143
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-3.6, -3.696, -3.728, -3.779]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-77.166, -78.083, -78.099, -78.306]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.70075
varray: [-3.6, -3.696, -3.728, -3.779]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
dD_mean | -77.9135
varray: [-77.166, -78.083, -78.099, -78.306]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0753630103787
varray: [-3.6, -3.696, -3.728, -3.779]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
dD_stdev | 0.508577427733
varray: [-77.166, -78.083, -78.099, -78.306]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.70075
varray: [-3.6, -3.696, -3.728, -3.779]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.9135
varray: [-77.166, -78.083, -78.099, -78.306]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0753630103787
varray: [-3.6, -3.696, -3.728, -3.779]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.508577427733
varray: [-77.166, -78.083, -78.099, -78.306]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.67607585613
actual_array: [ 1.93 -15.84]
varray: [-3.6, -3.696, -3.728, -3.779]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.70075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -84.8942048575
actual_array: [ 18.1 -119. ]
varray: [-77.166, -78.083, -78.099, -78.306]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -77.9135
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [37787.0, 38324.0, 38860.0, 39397.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.866
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 124.3055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [39987.0, 40524.0, 41060.0, 41596.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15425
varray: [2.167, 2.143, 2.155, 2.152]
inst: mean_all
operation: mean_all(residuals)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: residuals
SAMP_residuals_STD | 0.0099121138008
varray: [2.167, 2.143, 2.155, 2.152]
inst: std_all
operation: std_all(residuals)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.10625
varray: [-2.045, -2.167, -2.221, -1.992]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105928828308
varray: [-2.045, -2.167, -2.221, -1.992]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 99.48625
varray: [99.684, 99.221, 99.568, 99.472]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.196933787519
varray: [99.684, 99.221, 99.568, 99.472]
inst: std_all
operation: std_all(baseline_shift)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.56875
varray: [28.655, 28.352, 29.095, 28.173]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.403309951113
varray: [28.655, 28.352, 29.095, 28.173]
inst: std_all
operation: std_all(slope_shift)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: slope_shift
SAMP_d18O_STD | 0.114414451302
varray: [-4.408, -4.274, -4.17, -4.164]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.440808631948
varray: [-75.215, -75.962, -75.887, -76.259]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.335
varray: [59.205, 56.835, 57.828, 59.472]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23243904515
varray: [59.205, 56.835, 57.828, 59.472]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.10625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 99.48625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.56875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05695
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05695
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.57552460203
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.8133
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-4.408, -4.274, -4.17, -4.164]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-75.215, -75.962, -75.887, -76.259]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.254
varray: [-4.408, -4.274, -4.17, -4.164]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
dD_mean | -75.83075
varray: [-75.215, -75.962, -75.887, -76.259]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
d18O_stdev | 0.114414451302
varray: [-4.408, -4.274, -4.17, -4.164]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
dD_stdev | 0.440808631948
varray: [-75.215, -75.962, -75.887, -76.259]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.254
varray: [-4.408, -4.274, -4.17, -4.164]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
measured_dD_mean | -75.83075
varray: [-75.215, -75.962, -75.887, -76.259]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.114414451302
varray: [-4.408, -4.274, -4.17, -4.164]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.440808631948
varray: [-75.215, -75.962, -75.887, -76.259]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.24854896846
actual_array: [ 1.93 -15.84]
varray: [-4.408, -4.274, -4.17, -4.164]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -4.254
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.650260783
actual_array: [ 18.1 -119. ]
varray: [-75.215, -75.962, -75.887, -76.259]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.83075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [42186.0, 42724.0, 43261.0, 43797.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.8835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.4485
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.166
varray: [2.189, 2.172, 2.162, 2.141]
inst: mean_all
operation: mean_all(residuals)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: residuals
SAMP_residuals_STD | 0.0200499376558
varray: [2.189, 2.172, 2.162, 2.141]
inst: std_all
operation: std_all(residuals)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.174
varray: [-2.193, -2.135, -2.216, -2.152]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0371034589583
varray: [-2.193, -2.135, -2.216, -2.152]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.992
varray: [94.105, 94.094, 93.999, 93.77]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.155462750094
varray: [94.105, 94.094, 93.999, 93.77]
inst: std_all
operation: std_all(baseline_shift)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.71575
varray: [28.995, 28.615, 28.741, 28.512]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.208392858163
varray: [28.995, 28.615, 28.741, 28.512]
inst: std_all
operation: std_all(slope_shift)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: slope_shift
SAMP_d18O_STD | 0.104407455034
varray: [0.011, 0.147, 0.205, 0.252]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.25926868303
varray: [-70.672, -70.237, -70.171, -70.093]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.60325
varray: [59.082, 58.414, 59.027, 57.89]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.563709957928
varray: [59.082, 58.414, 59.027, 57.89]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.166
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.174
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.992
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.71575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0452
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0452
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.04413892909
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.54505
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.011, 0.147, 0.205, 0.252]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-70.672, -70.237, -70.171, -70.093]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.15375
varray: [0.011, 0.147, 0.205, 0.252]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
dD_mean | -70.29325
varray: [-70.672, -70.237, -70.171, -70.093]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
d18O_stdev | 0.104407455034
varray: [0.011, 0.147, 0.205, 0.252]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
dD_stdev | 0.25926868303
varray: [-70.672, -70.237, -70.171, -70.093]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.15375
varray: [0.011, 0.147, 0.205, 0.252]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.29325
varray: [-70.672, -70.237, -70.171, -70.093]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.104407455034
varray: [0.011, 0.147, 0.205, 0.252]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.25926868303
varray: [-70.672, -70.237, -70.171, -70.093]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.687648180606
actual_array: [ 1.93 -15.84]
varray: [0.011, 0.147, 0.205, 0.252]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.15375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.6841869323
actual_array: [ 18.1 -119. ]
varray: [-70.672, -70.237, -70.171, -70.093]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.29325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [44383.0, 44920.0, 45457.0, 45993.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.87175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.04575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.39225
varray: [0.0, 0.0, 6.757, 6.812]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: residuals
SAMP_residuals_STD | 3.91709725647
varray: [0.0, 0.0, 6.757, 6.812]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -7.787
varray: [0.0, 0.0, -15.762, -15.386]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 8.99296324912
varray: [0.0, 0.0, -15.762, -15.386]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 349.808
varray: [0.0, 0.0, 697.54, 701.692]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 403.927042508
varray: [0.0, 0.0, 697.54, 701.692]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -9.482
varray: [0.0, 0.0, -18.641, -19.287]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: slope_shift
SAMP_slope_shift_STD | 10.9520462928
varray: [0.0, 0.0, -18.641, -19.287]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: slope_shift
SAMP_d18O_STD | 13.850368163
varray: [0.0, 0.0, 23.958, 24.021]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 21.1260111632
varray: [0.0, 0.0, -36.482, -36.7]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 30.73975
varray: [0.0, 0.0, 59.772, 63.187]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 35.5225751354
varray: [0.0, 0.0, 59.772, 63.187]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.39225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 7.787
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 349.808
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.482
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.18105
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.18105
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 53.4121743849
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -28.40855
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 23.958, 24.021]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, -36.482, -36.7]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 11.99475
varray: [0.0, 0.0, 23.958, 24.021]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
dD_mean | -18.2955
varray: [0.0, 0.0, -36.482, -36.7]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
d18O_stdev | 13.850368163
varray: [0.0, 0.0, 23.958, 24.021]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
dD_stdev | 21.1260111632
varray: [0.0, 0.0, -36.482, -36.7]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
measured_d18O_mean | 11.99475
varray: [0.0, 0.0, 23.958, 24.021]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
measured_dD_mean | -18.2955
varray: [0.0, 0.0, -36.482, -36.7]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 13.850368163
varray: [0.0, 0.0, 23.958, 24.021]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
measured_dD_stdev | 21.1260111632
varray: [0.0, 0.0, -36.482, -36.7]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 11.5647768047
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 23.958, 24.021]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 11.99475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -20.6620801029
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -36.482, -36.7]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -18.2955
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [46581.0, 47117.0, 47654.0, 48189.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 90.6455
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 255.77025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.10875
varray: [0.0, 0.0, 2.233, 2.202]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: residuals
SAMP_residuals_STD | 1.2803367721
varray: [0.0, 0.0, 2.233, 2.202]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.28575
varray: [0.0, 0.0, -2.477, -2.666]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.4866598748
varray: [0.0, 0.0, -2.477, -2.666]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 163.1395
varray: [0.0, 0.0, 327.385, 325.173]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 188.379432983
varray: [0.0, 0.0, 327.385, 325.173]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 10.027
varray: [0.0, 0.0, 20.053, 20.055]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: slope_shift
SAMP_slope_shift_STD | 11.5781823271
varray: [0.0, 0.0, 20.053, 20.055]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: slope_shift
SAMP_d18O_STD | 3.1215371614
varray: [0.0, 0.0, -5.377, -5.436]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 45.2426213901
varray: [0.0, 0.0, -78.324, -78.401]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 27.05175
varray: [0.0, 0.0, 54.835, 53.372]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 31.2423798536
varray: [0.0, 0.0, 54.835, 53.372]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.10875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.28575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 163.1395
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 10.027
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.10245
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.10245
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -49.8575434153
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -32.09655
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, -5.377, -5.436]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, -78.324, -78.401]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.70325
varray: [0.0, 0.0, -5.377, -5.436]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
dD_mean | -39.18125
varray: [0.0, 0.0, -78.324, -78.401]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
d18O_stdev | 3.1215371614
varray: [0.0, 0.0, -5.377, -5.436]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
dD_stdev | 45.2426213901
varray: [0.0, 0.0, -78.324, -78.401]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.70325
varray: [0.0, 0.0, -5.377, -5.436]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.18125
varray: [0.0, 0.0, -78.324, -78.401]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.1215371614
varray: [0.0, 0.0, -5.377, -5.436]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
measured_dD_stdev | 45.2426213901
varray: [0.0, 0.0, -78.324, -78.401]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.64391692336
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -5.377, -5.436]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.70325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.1642790963
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -78.324, -78.401]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -39.18125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [48777.0, 49313.0, 49850.0, 50388.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 92.929
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 69.10175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.153
varray: [2.167, 2.153, 2.147, 2.145]
inst: mean_all
operation: mean_all(residuals)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: residuals
SAMP_residuals_STD | 0.00993310961717
varray: [2.167, 2.153, 2.147, 2.145]
inst: std_all
operation: std_all(residuals)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.1645
varray: [-2.29, -2.052, -2.304, -2.012]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.153972941346
varray: [-2.29, -2.052, -2.304, -2.012]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 95.184
varray: [96.766, 95.495, 94.612, 93.863]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.24788754835
varray: [96.766, 95.495, 94.612, 93.863]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.821
varray: [29.013, 28.818, 29.001, 28.452]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.261683014351
varray: [29.013, 28.818, 29.001, 28.452]
inst: std_all
operation: std_all(slope_shift)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0895712007288
varray: [-2.299, -2.168, -2.102, -2.117]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.144891913738
varray: [-76.197, -76.051, -76.184, -75.886]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.0435
varray: [57.464, 59.252, 56.632, 58.826]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.21118825952
varray: [57.464, 59.252, 56.632, 58.826]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.153
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.1645
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 95.184
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.821
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0582
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0582
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.63205499276
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1048
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-2.299, -2.168, -2.102, -2.117]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-76.197, -76.051, -76.184, -75.886]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.1715
varray: [-2.299, -2.168, -2.102, -2.117]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
dD_mean | -76.0795
varray: [-76.197, -76.051, -76.184, -75.886]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0895712007288
varray: [-2.299, -2.168, -2.102, -2.117]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
dD_stdev | 0.144891913738
varray: [-76.197, -76.051, -76.184, -75.886]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.1715
varray: [-2.299, -2.168, -2.102, -2.117]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
measured_dD_mean | -76.0795
varray: [-76.197, -76.051, -76.184, -75.886]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0895712007288
varray: [-2.299, -2.168, -2.102, -2.117]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.144891913738
varray: [-76.197, -76.051, -76.184, -75.886]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.09369084567
actual_array: [ 1.93 -15.84]
varray: [-2.299, -2.168, -2.102, -2.117]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.1715
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.918262746
actual_array: [ 18.1 -119. ]
varray: [-76.197, -76.051, -76.184, -75.886]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.0795
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [50976.0, 51512.0, 52048.0, 52585.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.88475
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.14625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17175
varray: [2.173, 2.175, 2.175, 2.164]
inst: mean_all
operation: mean_all(residuals)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: residuals
SAMP_residuals_STD | 0.0052519837522
varray: [2.173, 2.175, 2.175, 2.164]
inst: std_all
operation: std_all(residuals)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.26225
varray: [-2.388, -2.158, -2.326, -2.177]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.112565166311
varray: [-2.388, -2.158, -2.326, -2.177]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 112.00025
varray: [112.249, 111.966, 112.062, 111.724]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.218461857846
varray: [112.249, 111.966, 112.062, 111.724]
inst: std_all
operation: std_all(baseline_shift)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.01525
varray: [28.209, 28.107, 27.849, 27.896]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.171091349869
varray: [28.209, 28.107, 27.849, 27.896]
inst: std_all
operation: std_all(slope_shift)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0556319752181
varray: [-0.576, -0.486, -0.472, -0.449]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.60635825494
varray: [-66.921, -65.887, -65.811, -65.539]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.79825
varray: [57.065, 58.142, 58.057, 57.929]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.496610763073
varray: [57.065, 58.142, 58.057, 57.929]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.26225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 112.00025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.01525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03945
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03945
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.78409913169
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.35005
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-0.576, -0.486, -0.472, -0.449]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-66.921, -65.887, -65.811, -65.539]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.49575
varray: [-0.576, -0.486, -0.472, -0.449]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
dD_mean | -66.0395
varray: [-66.921, -65.887, -65.811, -65.539]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0556319752181
varray: [-0.576, -0.486, -0.472, -0.449]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
dD_stdev | 0.60635825494
varray: [-66.921, -65.887, -65.811, -65.539]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.49575
varray: [-0.576, -0.486, -0.472, -0.449]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.0395
varray: [-66.921, -65.887, -65.811, -65.539]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0556319752181
varray: [-0.576, -0.486, -0.472, -0.449]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.60635825494
varray: [-66.921, -65.887, -65.811, -65.539]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.35971557592
actual_array: [ 1.93 -15.84]
varray: [-0.576, -0.486, -0.472, -0.449]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.49575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.1012186899
actual_array: [ 18.1 -119. ]
varray: [-66.921, -65.887, -65.811, -65.539]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -66.0395
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [53173.0, 53709.0, 54245.0, 54781.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.866
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 17.9625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.175
varray: [2.198, 2.169, 2.176, 2.157]
inst: mean_all
operation: mean_all(residuals)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: residuals
SAMP_residuals_STD | 0.0172240142437
varray: [2.198, 2.169, 2.176, 2.157]
inst: std_all
operation: std_all(residuals)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.21525
varray: [-2.22, -2.237, -2.166, -2.238]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0338563140345
varray: [-2.22, -2.237, -2.166, -2.238]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 113.6395
varray: [114.045, 113.653, 113.599, 113.261]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.32120346615
varray: [114.045, 113.653, 113.599, 113.261]
inst: std_all
operation: std_all(baseline_shift)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.67125
varray: [27.9, 27.406, 27.77, 27.609]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.213159369174
varray: [27.9, 27.406, 27.77, 27.609]
inst: std_all
operation: std_all(slope_shift)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: slope_shift
SAMP_d18O_STD | 0.042625305473
varray: [-0.135, -0.17, -0.081, -0.085]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.136254663529
varray: [-65.065, -64.733, -64.873, -64.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.98175
varray: [58.46, 57.669, 58.106, 57.692]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.376797359333
varray: [58.46, 57.669, 58.106, 57.692]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.21525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 113.6395
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.67125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0362
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0362
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.63712011577
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.16655
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-0.135, -0.17, -0.081, -0.085]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-65.065, -64.733, -64.873, -64.877]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.11775
varray: [-0.135, -0.17, -0.081, -0.085]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
dD_mean | -64.887
varray: [-65.065, -64.733, -64.873, -64.877]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
d18O_stdev | 0.042625305473
varray: [-0.135, -0.17, -0.081, -0.085]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
dD_stdev | 0.136254663529
varray: [-65.065, -64.733, -64.873, -64.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.11775
varray: [-0.135, -0.17, -0.081, -0.085]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.887
varray: [-65.065, -64.733, -64.873, -64.877]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.042625305473
varray: [-0.135, -0.17, -0.081, -0.085]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.136254663529
varray: [-65.065, -64.733, -64.873, -64.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.968581664561
actual_array: [ 1.93 -15.84]
varray: [-0.135, -0.17, -0.081, -0.085]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.11775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.8595211526
actual_array: [ 18.1 -119. ]
varray: [-65.065, -64.733, -64.873, -64.877]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [55371.0, 55907.0, 56444.0, 56980.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.86275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 19.60175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [57568.0, 58167.0, 58767.0, 59366.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1435
varray: [2.124, 2.149, 2.149, 2.152]
inst: mean_all
operation: mean_all(residuals)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: residuals
SAMP_residuals_STD | 0.0130766968306
varray: [2.124, 2.149, 2.149, 2.152]
inst: std_all
operation: std_all(residuals)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.97525
varray: [-1.886, -2.055, -1.904, -2.056]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0929565310598
varray: [-1.886, -2.055, -1.904, -2.056]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.27975
varray: [93.35, 93.29, 93.159, 93.32]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0841442214296
varray: [93.35, 93.29, 93.159, 93.32]
inst: std_all
operation: std_all(baseline_shift)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.33675
varray: [27.983, 28.523, 28.443, 28.398]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.241432357677
varray: [27.983, 28.523, 28.443, 28.398]
inst: std_all
operation: std_all(slope_shift)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111470474417
varray: [-8.591, -8.692, -8.793, -8.842]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.19591898836
varray: [-67.034, -66.719, -67.117, -66.767]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.386
varray: [58.174, 58.202, 59.068, 58.1]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.456698295742
varray: [58.174, 58.202, 59.068, 58.1]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1435
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.97525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.27975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.33675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0677
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0677
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.06168596237
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7623
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-8.591, -8.692, -8.793, -8.842]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-67.034, -66.719, -67.117, -66.767]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.7295
varray: [-8.591, -8.692, -8.793, -8.842]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
dD_mean | -66.90925
varray: [-67.034, -66.719, -67.117, -66.767]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111470474417
varray: [-8.591, -8.692, -8.793, -8.842]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
dD_stdev | 0.19591898836
varray: [-67.034, -66.719, -67.117, -66.767]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.7295
varray: [-8.591, -8.692, -8.793, -8.842]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.90925
varray: [-67.034, -66.719, -67.117, -66.767]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111470474417
varray: [-8.591, -8.692, -8.793, -8.842]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.19591898836
varray: [-67.034, -66.719, -67.117, -66.767]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.87955378407
actual_array: [ 1.93 -15.84]
varray: [-8.591, -8.692, -8.793, -8.842]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.7295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.0382828401
actual_array: [ 18.1 -119. ]
varray: [-67.034, -66.719, -67.117, -66.767]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -66.90925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [60015.0, 60550.0, 61086.0, 61622.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.89425
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.758
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [62209.0, 62808.0, 63407.0, 64006.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [64655.0, 65253.0, 65851.0, 66449.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.0655
varray: [2.117, 2.145, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: residuals
SAMP_residuals_STD | 1.23038652463
varray: [2.117, 2.145, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.06575
varray: [-2.169, -2.094, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.23100294476
varray: [-2.169, -2.094, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 49.924
varray: [99.741, 99.955, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 57.6473358795
varray: [99.741, 99.955, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.98225
varray: [27.856, 28.073, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: slope_shift
SAMP_slope_shift_STD | 16.1455546488
varray: [27.856, 28.073, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: slope_shift
SAMP_d18O_STD | 0.980770275515
varray: [-1.828, -1.546, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 39.4443406365
varray: [-68.403, -68.236, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 28.98725
varray: [57.308, 58.641, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.476016762
varray: [57.308, 58.641, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.0655
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.06575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 49.924
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.98225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1457
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1457
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.8134949349
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -30.16105
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-1.828, -1.546, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-68.403, -68.236, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.8435
varray: [-1.828, -1.546, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
dD_mean | -34.15975
varray: [-68.403, -68.236, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
d18O_stdev | 0.980770275515
varray: [-1.828, -1.546, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
dD_stdev | 39.4443406365
varray: [-68.403, -68.236, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.8435
varray: [-1.828, -1.546, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
measured_dD_mean | -34.15975
varray: [-68.403, -68.236, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.980770275515
varray: [-1.828, -1.546, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
measured_dD_stdev | 39.4443406365
varray: [-68.403, -68.236, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.71954842692
actual_array: [ 1.93 -15.84]
varray: [-1.828, -1.546, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.8435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.7541409761
actual_array: [ 18.1 -119. ]
varray: [-68.403, -68.236, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -34.15975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [67099.0, 67634.0, 68170.0, 68768.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 92.97225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -44.11375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [69418.0, 70016.0, 70615.0, 71213.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2095
varray: [2.198, 2.214, 2.218, 2.208]
inst: mean_all
operation: mean_all(residuals)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: residuals
SAMP_residuals_STD | 0.00869865890047
varray: [2.198, 2.214, 2.218, 2.208]
inst: std_all
operation: std_all(residuals)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.337
varray: [-2.287, -2.305, -2.389, -2.367]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0487442304278
varray: [-2.287, -2.305, -2.389, -2.367]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 95.6165
varray: [95.692, 95.562, 95.643, 95.569]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0622602066599
varray: [95.692, 95.562, 95.643, 95.569]
inst: std_all
operation: std_all(baseline_shift)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.6615
varray: [28.39, 28.536, 28.988, 28.732]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.258866117263
varray: [28.39, 28.536, 28.988, 28.732]
inst: std_all
operation: std_all(slope_shift)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: slope_shift
SAMP_d18O_STD | 0.13565642877
varray: [1.228, 1.478, 1.496, 1.518]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.719074173272
varray: [-54.804, -53.594, -53.302, -53.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.916
varray: [58.776, 59.092, 59.257, 58.539]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.320928860237
varray: [58.776, 59.092, 59.257, 58.539]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2095
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.337
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 95.6165
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.6615
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0017
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0017
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0768813314038
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.2323
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [1.228, 1.478, 1.496, 1.518]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-54.804, -53.594, -53.302, -53.286]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.43
varray: [1.228, 1.478, 1.496, 1.518]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
dD_mean | -53.7465
varray: [-54.804, -53.594, -53.302, -53.286]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
d18O_stdev | 0.13565642877
varray: [1.228, 1.478, 1.496, 1.518]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
dD_stdev | 0.719074173272
varray: [-54.804, -53.594, -53.302, -53.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.43
varray: [1.228, 1.478, 1.496, 1.518]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.7465
varray: [-54.804, -53.594, -53.302, -53.286]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.13565642877
varray: [1.228, 1.478, 1.496, 1.518]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.719074173272
varray: [-54.804, -53.594, -53.302, -53.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.632946143141
actual_array: [ 1.93 -15.84]
varray: [1.228, 1.478, 1.496, 1.518]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.43
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.8568040902
actual_array: [ 18.1 -119. ]
varray: [-54.804, -53.594, -53.302, -53.286]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -53.7465
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [71864.0, 72400.0, 72936.0, 73472.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.82825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.57875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.18575
varray: [2.196, 2.16, 2.208, 2.179]
inst: mean_all
operation: mean_all(residuals)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: residuals
SAMP_residuals_STD | 0.0208865985742
varray: [2.196, 2.16, 2.208, 2.179]
inst: std_all
operation: std_all(residuals)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.18175
varray: [-2.044, -2.227, -2.301, -2.155]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109481733636
varray: [-2.044, -2.227, -2.301, -2.155]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.8775
varray: [93.967, 93.692, 94.226, 93.625]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.275456590651
varray: [93.967, 93.692, 94.226, 93.625]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.0575
varray: [28.903, 28.766, 29.713, 28.848]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.440610561229
varray: [28.903, 28.766, 29.713, 28.848]
inst: std_all
operation: std_all(slope_shift)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0199749843554
varray: [1.349, 1.334, 1.379, 1.34]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.631457045253
varray: [-64.269, -65.358, -65.533, -65.64]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.84
varray: [60.013, 57.702, 58.478, 59.167]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.98470740155
varray: [60.013, 57.702, 58.478, 59.167]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.18575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.18175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.8775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.0575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.15095875543
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3083
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [1.349, 1.334, 1.379, 1.34]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-64.269, -65.358, -65.533, -65.64]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.3505
varray: [1.349, 1.334, 1.379, 1.34]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
dD_mean | -65.2
varray: [-64.269, -65.358, -65.533, -65.64]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0199749843554
varray: [1.349, 1.334, 1.379, 1.34]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
dD_stdev | 0.631457045253
varray: [-64.269, -65.358, -65.533, -65.64]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.3505
varray: [1.349, 1.334, 1.379, 1.34]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.2
varray: [-64.269, -65.358, -65.533, -65.64]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0199749843554
varray: [1.349, 1.334, 1.379, 1.34]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.631457045253
varray: [-64.269, -65.358, -65.533, -65.64]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.550683852259
actual_array: [ 1.93 -15.84]
varray: [1.349, 1.334, 1.379, 1.34]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.3505
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.1967457332
actual_array: [ 18.1 -119. ]
varray: [-64.269, -65.358, -65.533, -65.64]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.2
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [74061.0, 74598.0, 75133.0, 75669.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.852
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.16025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17675
varray: [2.176, 2.193, 2.171, 2.167]
inst: mean_all
operation: mean_all(residuals)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: residuals
SAMP_residuals_STD | 0.0114418821296
varray: [2.176, 2.193, 2.171, 2.167]
inst: std_all
operation: std_all(residuals)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.168
varray: [-2.059, -2.165, -2.11, -2.338]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.121317764569
varray: [-2.059, -2.165, -2.11, -2.338]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 101.91625
varray: [101.793, 101.853, 101.944, 102.075]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.122695761948
varray: [101.793, 101.853, 101.944, 102.075]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.57725
varray: [28.53, 28.964, 28.39, 28.425]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.26460709867
varray: [28.53, 28.964, 28.39, 28.425]
inst: std_all
operation: std_all(slope_shift)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0634297511688
varray: [-1.732, -1.836, -1.879, -1.845]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.135781687523
varray: [-68.646, -68.9, -68.957, -68.809]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.7075
varray: [59.477, 58.497, 59.205, 57.651]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.816524953691
varray: [59.477, 58.497, 59.205, 57.651]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.168
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 101.91625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.57725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03445
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03445
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.55797756874
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.4408
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-1.732, -1.836, -1.879, -1.845]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-68.646, -68.9, -68.957, -68.809]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.823
varray: [-1.732, -1.836, -1.879, -1.845]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
dD_mean | -68.828
varray: [-68.646, -68.9, -68.957, -68.809]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0634297511688
varray: [-1.732, -1.836, -1.879, -1.845]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
dD_stdev | 0.135781687523
varray: [-68.646, -68.9, -68.957, -68.809]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.823
varray: [-1.732, -1.836, -1.879, -1.845]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.828
varray: [-68.646, -68.9, -68.957, -68.809]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0634297511688
varray: [-1.732, -1.836, -1.879, -1.845]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.135781687523
varray: [-68.646, -68.9, -68.957, -68.809]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.73308193533
actual_array: [ 1.93 -15.84]
varray: [-1.732, -1.836, -1.879, -1.845]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.823
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.1055341631
actual_array: [ 18.1 -119. ]
varray: [-68.646, -68.9, -68.957, -68.809]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.828
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [76254.0, 76790.0, 77326.0, 77862.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.861
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.8785
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17275
varray: [2.178, 2.167, 2.176, 2.17]
inst: mean_all
operation: mean_all(residuals)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: residuals
SAMP_residuals_STD | 0.00512347538298
varray: [2.178, 2.167, 2.176, 2.17]
inst: std_all
operation: std_all(residuals)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.17225
varray: [-1.966, -2.248, -2.308, -2.167]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149145063613
varray: [-1.966, -2.248, -2.308, -2.167]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 104.08975
varray: [104.078, 103.833, 104.315, 104.133]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.198885855036
varray: [104.078, 103.833, 104.315, 104.133]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.31125
varray: [28.763, 27.871, 28.285, 28.326]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.364602409025
varray: [28.763, 27.871, 28.285, 28.326]
inst: std_all
operation: std_all(slope_shift)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0432126524682
varray: [-0.213, -0.142, -0.142, -0.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0821680189206
varray: [-68.29, -68.188, -68.388, -68.307]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.7145
varray: [59.313, 58.551, 58.327, 58.667]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.423220588661
varray: [59.313, 58.551, 58.327, 58.667]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.17225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 104.08975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.31125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.7388748191
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.4338
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-0.213, -0.142, -0.142, -0.111]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-68.29, -68.188, -68.388, -68.307]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.152
varray: [-0.213, -0.142, -0.142, -0.111]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
dD_mean | -68.29325
varray: [-68.29, -68.188, -68.388, -68.307]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0432126524682
varray: [-0.213, -0.142, -0.142, -0.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
dD_stdev | 0.0821680189206
varray: [-68.29, -68.188, -68.388, -68.307]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.152
varray: [-0.213, -0.142, -0.142, -0.111]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.29325
varray: [-68.29, -68.188, -68.388, -68.307]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0432126524682
varray: [-0.213, -0.142, -0.142, -0.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0821680189206
varray: [-68.29, -68.188, -68.388, -68.307]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.00402170812
actual_array: [ 1.93 -15.84]
varray: [-0.213, -0.142, -0.142, -0.111]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.5293972797
actual_array: [ 18.1 -119. ]
varray: [-68.29, -68.188, -68.388, -68.307]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.29325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [78448.0, 78984.0, 79521.0, 80057.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 10.052
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.08375
varray: [0.0, 0.0, 2.168, 2.167]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: residuals
SAMP_residuals_STD | 1.25140677506
varray: [0.0, 0.0, 2.168, 2.167]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.074
varray: [0.0, 0.0, -2.117, -2.179]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.24040665375
varray: [0.0, 0.0, -2.117, -2.179]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.79375
varray: [0.0, 0.0, 93.536, 93.639]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 54.0327846798
varray: [0.0, 0.0, 93.536, 93.639]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 14.591
varray: [0.0, 0.0, 29.333, 29.031]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: slope_shift
SAMP_slope_shift_STD | 16.848686655
varray: [0.0, 0.0, 29.333, 29.031]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: slope_shift
SAMP_d18O_STD | 5.17657299765
varray: [0.0, 0.0, -8.938, -8.994]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 38.8094870274
varray: [0.0, 0.0, -67.205, -67.235]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 29.48
varray: [0.0, 0.0, 58.652, 59.268]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 34.0415007895
varray: [0.0, 0.0, 58.652, 59.268]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.08375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.074
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.79375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.591
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.12745
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.12745
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.9881512301
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -29.6683
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, -8.938, -8.994]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, -67.205, -67.235]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.483
varray: [0.0, 0.0, -8.938, -8.994]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
dD_mean | -33.61
varray: [0.0, 0.0, -67.205, -67.235]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
d18O_stdev | 5.17657299765
varray: [0.0, 0.0, -8.938, -8.994]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
dD_stdev | 38.8094870274
varray: [0.0, 0.0, -67.205, -67.235]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.483
varray: [0.0, 0.0, -8.938, -8.994]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
measured_dD_mean | -33.61
varray: [0.0, 0.0, -67.205, -67.235]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.17657299765
varray: [0.0, 0.0, -8.938, -8.994]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
measured_dD_stdev | 38.8094870274
varray: [0.0, 0.0, -67.205, -67.235]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.48550575603
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -8.938, -8.994]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -4.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.1618431703
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -67.205, -67.235]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -33.61
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [80642.0, 81178.0, 81713.0, 82250.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 92.954
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -47.244
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15975
varray: [2.152, 2.146, 2.172, 2.169]
inst: mean_all
operation: mean_all(residuals)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: residuals
SAMP_residuals_STD | 0.0127115433104
varray: [2.152, 2.146, 2.172, 2.169]
inst: std_all
operation: std_all(residuals)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.13775
varray: [-2.189, -2.272, -2.075, -2.015]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114970648428
varray: [-2.189, -2.272, -2.075, -2.015]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.6435
varray: [93.664, 93.547, 93.758, 93.605]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0900462843949
varray: [93.664, 93.547, 93.758, 93.605]
inst: std_all
operation: std_all(baseline_shift)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.9475
varray: [28.97, 28.587, 29.166, 29.067]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.253304164987
varray: [28.97, 28.587, 29.166, 29.067]
inst: std_all
operation: std_all(slope_shift)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: slope_shift
SAMP_d18O_STD | 0.201695108187
varray: [-1.54, -1.255, -1.137, -1.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.294824806736
varray: [-72.576, -73.149, -73.13, -73.207]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.3985
varray: [58.089, 57.338, 59.006, 59.161]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.850668952453
varray: [58.089, 57.338, 59.006, 59.161]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.13775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.6435
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.9475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05145
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05145
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.32679088278
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7498
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-1.54, -1.255, -1.137, -1.091]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-72.576, -73.149, -73.13, -73.207]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.25575
varray: [-1.54, -1.255, -1.137, -1.091]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
dD_mean | -73.0155
varray: [-72.576, -73.149, -73.13, -73.207]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
d18O_stdev | 0.201695108187
varray: [-1.54, -1.255, -1.137, -1.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
dD_stdev | 0.294824806736
varray: [-72.576, -73.149, -73.13, -73.207]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.25575
varray: [-1.54, -1.255, -1.137, -1.091]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.0155
varray: [-72.576, -73.149, -73.13, -73.207]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.201695108187
varray: [-1.54, -1.255, -1.137, -1.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.294824806736
varray: [-72.576, -73.149, -73.13, -73.207]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.14612238184
actual_array: [ 1.93 -15.84]
varray: [-1.54, -1.255, -1.137, -1.091]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.25575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.6171249982
actual_array: [ 18.1 -119. ]
varray: [-72.576, -73.149, -73.13, -73.207]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.0155
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [82836.0, 83373.0, 83909.0, 84445.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.878
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.39425
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14925
varray: [2.129, 2.155, 2.149, 2.164]
inst: mean_all
operation: mean_all(residuals)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: residuals
SAMP_residuals_STD | 0.0148408220797
varray: [2.129, 2.155, 2.149, 2.164]
inst: std_all
operation: std_all(residuals)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.167
varray: [-2.097, -2.221, -2.102, -2.248]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0787443119639
varray: [-2.097, -2.221, -2.102, -2.248]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 135.02225
varray: [133.932, 135.171, 135.276, 135.71]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.763361153059
varray: [133.932, 135.171, 135.276, 135.71]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.43175
varray: [26.801, 27.689, 27.26, 27.977]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.513425992719
varray: [26.801, 27.689, 27.26, 27.977]
inst: std_all
operation: std_all(slope_shift)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: slope_shift
SAMP_d18O_STD | 0.158002109691
varray: [-5.892, -6.082, -6.215, -6.235]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.608826466792
varray: [-82.033, -82.652, -83.337, -83.264]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.17675
varray: [56.998, 56.35, 57.994, 57.365]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.687702152098
varray: [56.998, 56.35, 57.994, 57.365]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.167
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 135.02225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.43175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06195
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06195
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.80164616498
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.97155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-5.892, -6.082, -6.215, -6.235]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-82.033, -82.652, -83.337, -83.264]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.106
varray: [-5.892, -6.082, -6.215, -6.235]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
dD_mean | -82.8215
varray: [-82.033, -82.652, -83.337, -83.264]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
d18O_stdev | 0.158002109691
varray: [-5.892, -6.082, -6.215, -6.235]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
dD_stdev | 0.608826466792
varray: [-82.033, -82.652, -83.337, -83.264]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.106
varray: [-5.892, -6.082, -6.215, -6.235]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
measured_dD_mean | -82.8215
varray: [-82.033, -82.652, -83.337, -83.264]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.158002109691
varray: [-5.892, -6.082, -6.215, -6.235]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.608826466792
varray: [-82.033, -82.652, -83.337, -83.264]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.16489818497
actual_array: [ 1.93 -15.84]
varray: [-5.892, -6.082, -6.215, -6.235]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -6.106
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.182058665
actual_array: [ 18.1 -119. ]
varray: [-82.033, -82.652, -83.337, -83.264]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -82.8215
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [85033.0, 85569.0, 86105.0, 86640.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.8885
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 40.9845
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14775
varray: [2.146, 2.135, 2.158, 2.152]
inst: mean_all
operation: mean_all(residuals)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: residuals
SAMP_residuals_STD | 0.00981070843517
varray: [2.146, 2.135, 2.158, 2.152]
inst: std_all
operation: std_all(residuals)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.16825
varray: [-2.257, -2.089, -2.294, -2.033]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.126836837446
varray: [-2.257, -2.089, -2.294, -2.033]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 94.01875
varray: [94.378, 94.012, 94.0, 93.685]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.28334122538
varray: [94.378, 94.012, 94.0, 93.685]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.11925
varray: [29.15, 29.188, 29.211, 28.928]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.129957364803
varray: [29.15, 29.188, 29.211, 28.928]
inst: std_all
operation: std_all(slope_shift)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0459891291503
varray: [-7.044, -7.107, -7.088, -7.155]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.356404615571
varray: [-88.559, -89.139, -89.165, -89.396]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.70825
varray: [56.518, 57.804, 57.819, 58.692]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.895524566944
varray: [56.518, 57.804, 57.819, 58.692]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.16825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 94.01875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.11925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06345
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06345
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.86948263386
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.44005
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-7.044, -7.107, -7.088, -7.155]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-88.559, -89.139, -89.165, -89.396]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.0985
varray: [-7.044, -7.107, -7.088, -7.155]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
dD_mean | -89.06475
varray: [-88.559, -89.139, -89.165, -89.396]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0459891291503
varray: [-7.044, -7.107, -7.088, -7.155]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
dD_stdev | 0.356404615571
varray: [-88.559, -89.139, -89.165, -89.396]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.0985
varray: [-7.044, -7.107, -7.088, -7.155]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
measured_dD_mean | -89.06475
varray: [-88.559, -89.139, -89.165, -89.396]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0459891291503
varray: [-7.044, -7.107, -7.088, -7.155]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.356404615571
varray: [-88.559, -89.139, -89.165, -89.396]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.19188338875
actual_array: [ 1.93 -15.84]
varray: [-7.044, -7.107, -7.088, -7.155]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -7.0985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -96.9085039143
actual_array: [ 18.1 -119. ]
varray: [-88.559, -89.139, -89.165, -89.396]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.06475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [87228.0, 87764.0, 88300.0, 88836.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.89
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.019
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13275
varray: [2.118, 2.119, 2.14, 2.154]
inst: mean_all
operation: mean_all(residuals)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: residuals
SAMP_residuals_STD | 0.0174236429409
varray: [2.118, 2.119, 2.14, 2.154]
inst: std_all
operation: std_all(residuals)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.04725
varray: [-2.071, -2.031, -2.074, -2.013]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0300929116792
varray: [-2.071, -2.031, -2.074, -2.013]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.4285
varray: [93.408, 93.286, 93.434, 93.586]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.123238251637
varray: [93.408, 93.286, 93.434, 93.586]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.70125
varray: [28.502, 28.671, 28.951, 28.681]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.18565267751
varray: [28.502, 28.671, 28.951, 28.681]
inst: std_all
operation: std_all(slope_shift)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0387674691806
varray: [-5.346, -5.272, -5.299, -5.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.160230042959
varray: [-84.147, -83.836, -83.949, -83.786]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.217
varray: [58.167, 57.579, 58.059, 59.063]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.619199483204
varray: [58.167, 57.579, 58.059, 59.063]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.04725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.4285
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.70125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.54784732272
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.9313
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-5.346, -5.272, -5.299, -5.258]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-84.147, -83.836, -83.949, -83.786]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.29375
varray: [-5.346, -5.272, -5.299, -5.258]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
dD_mean | -83.9295
varray: [-84.147, -83.836, -83.949, -83.786]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0387674691806
varray: [-5.346, -5.272, -5.299, -5.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
dD_stdev | 0.160230042959
varray: [-84.147, -83.836, -83.949, -83.786]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.29375
varray: [-5.346, -5.272, -5.299, -5.258]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
measured_dD_mean | -83.9295
varray: [-84.147, -83.836, -83.949, -83.786]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0387674691806
varray: [-5.346, -5.272, -5.299, -5.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.160230042959
varray: [-84.147, -83.836, -83.949, -83.786]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.32442591115
actual_array: [ 1.93 -15.84]
varray: [-5.346, -5.272, -5.299, -5.258]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.29375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.3758121325
actual_array: [ 18.1 -119. ]
varray: [-84.147, -83.836, -83.949, -83.786]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -83.9295
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [89423.0, 89960.0, 90497.0, 91033.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.60925
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.105
varray: [2.103, 2.098, 2.115, 2.104]
inst: mean_all
operation: mean_all(residuals)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: residuals
SAMP_residuals_STD | 0.00716472842007
varray: [2.103, 2.098, 2.115, 2.104]
inst: std_all
operation: std_all(residuals)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.03075
varray: [-2.01, -2.02, -1.942, -2.151]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0873360368538
varray: [-2.01, -2.02, -1.942, -2.151]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.04825
varray: [93.095, 92.845, 93.062, 93.191]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.146130934439
varray: [93.095, 92.845, 93.062, 93.191]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3875
varray: [28.484, 28.34, 28.398, 28.328]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0712249956125
varray: [28.484, 28.34, 28.398, 28.328]
inst: std_all
operation: std_all(slope_shift)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: slope_shift
SAMP_d18O_STD | 0.119053209393
varray: [-8.353, -8.552, -8.625, -8.572]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.286808879453
varray: [-88.824, -89.337, -89.389, -89.446]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.353
varray: [57.706, 57.004, 57.713, 56.989]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.411706205928
varray: [57.706, 57.004, 57.713, 56.989]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.105
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.03075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.04825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1062
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1062
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.80282199711
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7953
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-8.353, -8.552, -8.625, -8.572]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-88.824, -89.337, -89.389, -89.446]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.5255
varray: [-8.353, -8.552, -8.625, -8.572]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
dD_mean | -89.249
varray: [-88.824, -89.337, -89.389, -89.446]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
d18O_stdev | 0.119053209393
varray: [-8.353, -8.552, -8.625, -8.572]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
dD_stdev | 0.286808879453
varray: [-88.824, -89.337, -89.389, -89.446]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.5255
varray: [-8.353, -8.552, -8.625, -8.572]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
measured_dD_mean | -89.249
varray: [-88.824, -89.337, -89.389, -89.446]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.119053209393
varray: [-8.353, -8.552, -8.625, -8.572]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.286808879453
varray: [-88.824, -89.337, -89.389, -89.446]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.66846564143
actual_array: [ 1.93 -15.84]
varray: [-8.353, -8.552, -8.625, -8.572]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.5255
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -97.107013911
actual_array: [ 18.1 -119. ]
varray: [-88.824, -89.337, -89.389, -89.446]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [91616.0, 92152.0, 92688.0, 93225.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.93275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9895
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16675
varray: [2.172, 2.176, 2.163, 2.156]
inst: mean_all
operation: mean_all(residuals)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: residuals
SAMP_residuals_STD | 0.00899536917901
varray: [2.172, 2.176, 2.163, 2.156]
inst: std_all
operation: std_all(residuals)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.23475
varray: [-2.259, -2.244, -2.178, -2.258]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0384480168539
varray: [-2.259, -2.244, -2.178, -2.258]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 104.649
varray: [104.82, 104.731, 104.602, 104.443]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.163920712541
varray: [104.82, 104.731, 104.602, 104.443]
inst: std_all
operation: std_all(baseline_shift)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.889
varray: [29.225, 29.042, 28.604, 28.685]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.293908148917
varray: [29.225, 29.042, 28.604, 28.685]
inst: std_all
operation: std_all(slope_shift)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: slope_shift
SAMP_d18O_STD | 0.152583310577
varray: [-3.459, -3.218, -3.136, -3.137]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.5265588761
varray: [-81.027, -80.082, -79.899, -79.973]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.8555
varray: [58.305, 57.918, 57.876, 57.323]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.404117557154
varray: [58.305, 57.918, 57.876, 57.323]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.23475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 104.649
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.889
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04445
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04445
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.01022069465
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2928
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-3.459, -3.218, -3.136, -3.137]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-81.027, -80.082, -79.899, -79.973]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.2375
varray: [-3.459, -3.218, -3.136, -3.137]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
dD_mean | -80.24525
varray: [-81.027, -80.082, -79.899, -79.973]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
d18O_stdev | 0.152583310577
varray: [-3.459, -3.218, -3.136, -3.137]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
dD_stdev | 0.5265588761
varray: [-81.027, -80.082, -79.899, -79.973]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.2375
varray: [-3.459, -3.218, -3.136, -3.137]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
measured_dD_mean | -80.24525
varray: [-81.027, -80.082, -79.899, -79.973]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.152583310577
varray: [-3.459, -3.218, -3.136, -3.137]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.5265588761
varray: [-81.027, -80.082, -79.899, -79.973]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.19672986555
actual_array: [ 1.93 -15.84]
varray: [-3.459, -3.218, -3.136, -3.137]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.2375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.4064202437
actual_array: [ 18.1 -119. ]
varray: [-81.027, -80.082, -79.899, -79.973]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -80.24525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [93809.0, 94345.0, 94881.0, 95417.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.871
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 10.61125
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1615
varray: [2.148, 2.156, 2.161, 2.181]
inst: mean_all
operation: mean_all(residuals)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: residuals
SAMP_residuals_STD | 0.014059397806
varray: [2.148, 2.156, 2.161, 2.181]
inst: std_all
operation: std_all(residuals)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.08675
varray: [-2.092, -2.062, -2.068, -2.125]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0286050694808
varray: [-2.092, -2.062, -2.068, -2.125]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.84025
varray: [93.732, 93.976, 93.778, 93.875]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.108364739038
varray: [93.732, 93.976, 93.778, 93.875]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.9435
varray: [28.548, 28.57, 29.058, 29.598]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.495783218756
varray: [28.548, 28.57, 29.058, 29.598]
inst: std_all
operation: std_all(slope_shift)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0507633397903
varray: [-1.667, -1.577, -1.591, -1.548]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.159263670266
varray: [-76.478, -76.274, -76.258, -76.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.46625
varray: [57.795, 59.136, 58.681, 58.253]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.574669397712
varray: [57.795, 59.136, 58.681, 58.253]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1615
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.08675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.84025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.9435
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0497
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0497
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.24764833575
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.68205
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-1.667, -1.577, -1.591, -1.548]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-76.478, -76.274, -76.258, -76.089]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.59575
varray: [-1.667, -1.577, -1.591, -1.548]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
dD_mean | -76.27475
varray: [-76.478, -76.274, -76.258, -76.089]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0507633397903
varray: [-1.667, -1.577, -1.591, -1.548]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
dD_stdev | 0.159263670266
varray: [-76.478, -76.274, -76.258, -76.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.59575
varray: [-1.667, -1.577, -1.591, -1.548]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
measured_dD_mean | -76.27475
varray: [-76.478, -76.274, -76.258, -76.089]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0507633397903
varray: [-1.667, -1.577, -1.591, -1.548]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.159263670266
varray: [-76.478, -76.274, -76.258, -76.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.4979359529
actual_array: [ 1.93 -15.84]
varray: [-1.667, -1.577, -1.591, -1.548]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.59575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.1286240859
actual_array: [ 18.1 -119. ]
varray: [-76.478, -76.274, -76.258, -76.089]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.27475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [96001.0, 96537.0, 97073.0, 97609.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.87625
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.1975
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1775
varray: [2.185, 2.175, 2.171, 2.179]
inst: mean_all
operation: mean_all(residuals)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: residuals
SAMP_residuals_STD | 0.00597215762239
varray: [2.185, 2.175, 2.171, 2.179]
inst: std_all
operation: std_all(residuals)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.183
varray: [-2.198, -2.231, -2.049, -2.254]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0922424341974
varray: [-2.198, -2.231, -2.049, -2.254]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.79375
varray: [94.013, 93.716, 93.626, 93.82]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.166277629283
varray: [94.013, 93.716, 93.626, 93.82]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.26525
varray: [29.239, 29.504, 29.167, 29.151]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.163703746648
varray: [29.239, 29.504, 29.167, 29.151]
inst: std_all
operation: std_all(slope_shift)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0447837396086
varray: [-0.954, -0.956, -0.864, -0.899]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0763517081232
varray: [-74.291, -74.194, -74.138, -74.122]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.56675
varray: [58.903, 57.527, 58.454, 59.383]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.790173135036
varray: [58.903, 57.527, 58.454, 59.383]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.183
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.79375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.26525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0337
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0337
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.5240593343
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.58155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-0.954, -0.956, -0.864, -0.899]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-74.291, -74.194, -74.138, -74.122]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.91825
varray: [-0.954, -0.956, -0.864, -0.899]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
dD_mean | -74.18625
varray: [-74.291, -74.194, -74.138, -74.122]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0447837396086
varray: [-0.954, -0.956, -0.864, -0.899]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
dD_stdev | 0.0763517081232
varray: [-74.291, -74.194, -74.138, -74.122]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.91825
varray: [-0.954, -0.956, -0.864, -0.899]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
measured_dD_mean | -74.18625
varray: [-74.291, -74.194, -74.138, -74.122]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0447837396086
varray: [-0.954, -0.956, -0.864, -0.899]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0763517081232
varray: [-74.291, -74.194, -74.138, -74.122]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.79689567526
actual_array: [ 1.93 -15.84]
varray: [-0.954, -0.956, -0.864, -0.899]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.91825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.8784849911
actual_array: [ 18.1 -119. ]
varray: [-74.291, -74.194, -74.138, -74.122]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -74.18625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [98194.0, 98729.0, 99265.0, 99801.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.86025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.244
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [100386.0, 100922.0, 101458.0, 101993.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [102579.0, 103178.0, 103776.0, 104375.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.2112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.2112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.1483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [105023.0, 105621.0, 106219.0, 106818.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 94.03775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -94.03775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.043
varray: [0.0, 0.0, 2.086, 2.086]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: residuals
SAMP_residuals_STD | 1.20435266153
varray: [0.0, 0.0, 2.086, 2.086]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0595
varray: [0.0, 0.0, -2.119, -2.119]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.22340522041
varray: [0.0, 0.0, -2.119, -2.119]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.499
varray: [0.0, 0.0, 92.998, 92.998]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 53.6924203341
varray: [0.0, 0.0, 92.998, 92.998]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.925
varray: [0.0, 0.0, 27.85, 27.85]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: slope_shift
SAMP_slope_shift_STD | 16.0792049969
varray: [0.0, 0.0, 27.85, 27.85]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: slope_shift
SAMP_d18O_STD | 1.69914184223
varray: [0.0, 0.0, -2.943, -2.943]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 43.7516033992
varray: [0.0, 0.0, -75.78, -75.78]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 28.483
varray: [0.0, 0.0, 56.966, 56.966]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 32.8893354347
varray: [0.0, 0.0, 56.966, 56.966]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.043
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0595
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.499
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1682
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1682
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -52.8310419682
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -30.6653
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [0.0, 0.0, -2.943, -2.943]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [0.0, 0.0, -75.78, -75.78]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.4715
varray: [0.0, 0.0, -2.943, -2.943]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
dD_mean | -37.89
varray: [0.0, 0.0, -75.78, -75.78]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
d18O_stdev | 1.69914184223
varray: [0.0, 0.0, -2.943, -2.943]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
dD_stdev | 43.7516033992
varray: [0.0, 0.0, -75.78, -75.78]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.4715
varray: [0.0, 0.0, -2.943, -2.943]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.89
varray: [0.0, 0.0, -75.78, -75.78]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.69914184223
varray: [0.0, 0.0, -2.943, -2.943]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
measured_dD_stdev | 43.7516033992
varray: [0.0, 0.0, -75.78, -75.78]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.36936878759
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -2.943, -2.943]
fit_rsq: 1.0
fit_intercept: 0.818307597074
except: True
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.4715
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.7730930269
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -75.78, -75.78]
fit_rsq: 1.0
fit_intercept: 0.882316338439
except: True
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -37.89
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [107465.0, 108062.0, 108660.0, 109196.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 92.99475
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -47.53875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.116
varray: [3.112, 3.137, 3.119, 3.096]
inst: mean_all
operation: mean_all(residuals)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: residuals
SAMP_residuals_STD | 0.0169901932498
varray: [3.112, 3.137, 3.119, 3.096]
inst: std_all
operation: std_all(residuals)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -5.88025
varray: [-5.546, -5.983, -5.883, -6.109]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.241256951541
varray: [-5.546, -5.983, -5.883, -6.109]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 110.64375
varray: [110.124, 111.042, 110.786, 110.623]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.387043817502
varray: [110.124, 111.042, 110.786, 110.623]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.97775
varray: [20.672, 21.308, 20.926, 21.005]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.26202337682
varray: [20.672, 21.308, 20.926, 21.005]
inst: std_all
operation: std_all(slope_shift)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0136747943312
varray: [2.684, 2.671, 2.672, 2.651]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0900444334759
varray: [-69.798, -70.008, -69.91, -69.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.7275
varray: [62.785, 61.95, 61.95, 60.225]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.07623185234
varray: [62.785, 61.95, 61.95, 60.225]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.116
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 5.88025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 110.64375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.97775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.9048
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.9048
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 40.9189580318
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.5792
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [2.684, 2.671, 2.672, 2.651]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-69.798, -70.008, -69.91, -69.96]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.6695
varray: [2.684, 2.671, 2.672, 2.651]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
dD_mean | -69.919
varray: [-69.798, -70.008, -69.91, -69.96]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0136747943312
varray: [2.684, 2.671, 2.672, 2.651]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
dD_stdev | 0.0900444334759
varray: [-69.798, -70.008, -69.91, -69.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.6695
varray: [2.684, 2.671, 2.672, 2.651]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.919
varray: [-69.798, -70.008, -69.91, -69.96]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0136747943312
varray: [2.684, 2.671, 2.672, 2.651]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0900444334759
varray: [-69.798, -70.008, -69.91, -69.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.91551355884
actual_array: [ 1.93 -15.84]
varray: [2.684, 2.671, 2.672, 2.651]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.6695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.2809719186
actual_array: [ 18.1 -119. ]
varray: [-69.798, -70.008, -69.91, -69.96]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.919
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [109843.0, 110379.0, 110915.0, 111452.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 90.92175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 16.606
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16625
varray: [2.13, 2.147, 2.201, 2.187]
inst: mean_all
operation: mean_all(residuals)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: residuals
SAMP_residuals_STD | 0.0332803745973
varray: [2.13, 2.147, 2.201, 2.187]
inst: std_all
operation: std_all(residuals)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.225
varray: [-2.356, -2.144, -2.138, -2.262]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.104339190464
varray: [-2.356, -2.144, -2.138, -2.262]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 135.98025
varray: [134.269, 135.646, 137.093, 136.913]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31000442111
varray: [134.269, 135.646, 137.093, 136.913]
inst: std_all
operation: std_all(baseline_shift)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.19025
varray: [26.46, 26.828, 27.704, 27.769]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.648942922503
varray: [26.46, 26.828, 27.704, 27.769]
inst: std_all
operation: std_all(slope_shift)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0784623051747
varray: [-0.729, -0.558, -0.583, -0.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.502052454099
varray: [-68.742, -67.874, -67.693, -67.691]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.50375
varray: [55.32, 57.796, 58.818, 58.081]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.51818364612
varray: [55.32, 57.796, 58.818, 58.081]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 135.98025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.19025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04495
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04495
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.03283285094
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.64455
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-0.729, -0.558, -0.583, -0.58]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-68.742, -67.874, -67.693, -67.691]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.6125
varray: [-0.729, -0.558, -0.583, -0.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
dD_mean | -68.0
varray: [-68.742, -67.874, -67.693, -67.691]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0784623051747
varray: [-0.729, -0.558, -0.583, -0.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
dD_stdev | 0.502052454099
varray: [-68.742, -67.874, -67.693, -67.691]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.6125
varray: [-0.729, -0.558, -0.583, -0.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.0
varray: [-68.742, -67.874, -67.693, -67.691]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0784623051747
varray: [-0.729, -0.558, -0.583, -0.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.502052454099
varray: [-68.742, -67.874, -67.693, -67.691]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.48052214775
actual_array: [ 1.93 -15.84]
varray: [-0.729, -0.558, -0.583, -0.58]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.6125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.2134512469
actual_array: [ 18.1 -119. ]
varray: [-68.742, -67.874, -67.693, -67.691]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [112037.0, 112573.0, 113110.0, 113646.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.8715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 41.9425
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1675
varray: [2.168, 2.167, 2.168, 2.167]
inst: mean_all
operation: mean_all(residuals)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: residuals
SAMP_residuals_STD | 0.00057735026919
varray: [2.168, 2.167, 2.168, 2.167]
inst: std_all
operation: std_all(residuals)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.148
varray: [-2.117, -2.179, -2.117, -2.179]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0357957166898
varray: [-2.117, -2.179, -2.117, -2.179]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.5875
varray: [93.536, 93.639, 93.536, 93.639]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0594670777265
varray: [93.536, 93.639, 93.536, 93.639]
inst: std_all
operation: std_all(baseline_shift)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.182
varray: [29.333, 29.031, 29.333, 29.031]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.174359781295
varray: [29.333, 29.031, 29.333, 29.031]
inst: std_all
operation: std_all(slope_shift)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0323316150746
varray: [-8.938, -8.994, -8.938, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0173205080757
varray: [-67.205, -67.235, -67.205, -67.235]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.96
varray: [58.652, 59.268, 58.652, 59.268]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.355647765821
varray: [58.652, 59.268, 58.652, 59.268]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.148
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.5875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.182
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0437
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0437
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.9763024602
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.1883
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4898
varray: [-8.938, -8.994, -8.938, -8.994]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.5693
varray: [-67.205, -67.235, -67.205, -67.235]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.966
varray: [-8.938, -8.994, -8.938, -8.994]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
dD_mean | -67.22
varray: [-67.205, -67.235, -67.205, -67.235]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0323316150746
varray: [-8.938, -8.994, -8.938, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
dD_stdev | 0.0173205080757
varray: [-67.205, -67.235, -67.205, -67.235]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.966
varray: [-8.938, -8.994, -8.938, -8.994]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.22
varray: [-67.205, -67.235, -67.205, -67.235]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0323316150746
varray: [-8.938, -8.994, -8.938, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0173205080757
varray: [-67.205, -67.235, -67.205, -67.235]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1242711651
actual_array: [ 1.93 -15.84]
varray: [-8.938, -8.994, -8.938, -8.994]
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.3730832824
actual_array: [ 18.1 -119. ]
varray: [-67.205, -67.235, -67.205, -67.235]
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.22
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: [114232.0, 114768.0, 115303.0, 115838.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 91.87025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.45025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.009
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.838
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2035110317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -10.009
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.8128615481
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -77.838
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.441
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.953
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6157754188
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.441
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.394998169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.953
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.335
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.587
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5060923643
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.335
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9232768362
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.587
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.271
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.119
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4398686333
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.271
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4190560575
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.119
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.229
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.795
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3964093098
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.229
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0699801338
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.795
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.578
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4067567678
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8361854565
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.578
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.189
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3550194779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.189
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.486032138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.18
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.131
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3457067657
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.18
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3545899691
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.131
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.174
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.373
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.339498291
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.174
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6153195171
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.373
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.15
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3146643918
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.15
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3007202278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.963
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.404
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.18446565308
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.963
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 9.18121788837
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 9.404
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.518
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.869
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.75874957055
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.518
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 16.1465754404
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 15.869
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.698
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94500381406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.698
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.0341711761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 17.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.744
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.154
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.99260212073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.744
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.6084226185
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 18.154
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.738
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.71
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.98639364595
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.738
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.207454142
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 18.71
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.038
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05261737697
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.560839645
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 19.038
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.833
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08469449669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.833
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.740764581
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 19.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.835
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.564
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08676398828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.835
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1275493236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 19.564
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.849
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.522
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10125042945
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.849
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.0822987409
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 19.522
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.855
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.734
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.6835
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.6821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.855
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.734
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10745890423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.855
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3107064441
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 19.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.569
operation: col(d(18_16)Mean)
inst: col
dD_actual | -97.88
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.8872060699
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -13.569
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -106.406008657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -97.88
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.218
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.132
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5587560923
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.218
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.37406559
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -107.132
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.344
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6891340628
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.344
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.078504205
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -108.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.583
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9364383083
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.583
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.274234861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -110.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.631
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.18
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9861061066
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.631
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.735359847
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -111.18
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.693
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.765
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.050260346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.693
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.36563582
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -111.765
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.723
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0813027199
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.723
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.347320108
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -111.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.743
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.146
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1019976359
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.743
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.776123249
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -112.146
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.697
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.188
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0543993292
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.697
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.821373832
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -112.188
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.697
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.188
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0543993292
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -14.697
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.821373832
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -112.188
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.409
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.772
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5826635533
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.409
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2773845317
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -71.772
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.191
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.513
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3570889695
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.191
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7661547928
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.513
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.1
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.262927102
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.1
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2059094831
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.093
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.716
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2556838814
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.093
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9074711162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.716
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.449
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2432669318
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6198066976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.449
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.097
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2598228646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.097
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5282281374
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.225
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2794830347
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3784702565
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.094
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.576
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2567186272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.094
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7566358405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.576
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2329194738
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.530382927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.32
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2329194738
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -9.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.480822765
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.32
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.275
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.274
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.23553283294
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.205210866
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -77.274
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.986
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.11
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.93649129754
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.986
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.1059129408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.11
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.935
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.88371926188
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.3526363561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.978
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.407
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.92821333116
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.978
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.4258992043
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.407
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.739
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.889
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.64616328836
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.739
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.7904138579
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.889
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.647
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.55096667501
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.647
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.554464391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.644
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.642
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.54786243762
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.644
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5242973358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.642
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.52095904689
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.8216583079
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.512
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.769
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.37653019513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.512
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.1967575209
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.769
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.933
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.30823697251
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2960554461
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.933
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.443
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.74
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.30513273512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.443
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0881182447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.74
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.452
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.897
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.31444544729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.452
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2572692324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.897
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.338
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.338
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.496996267462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.338
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.3454262193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.338
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.359
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.124
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.475266605719
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.359
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1148637265
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.124
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.371
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.103
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.462849656152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.371
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0922384351
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.103
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.973
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.481475080503
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9521771077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.973
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.247
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.591158135012
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.247
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.2448850857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -72.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.059
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.755
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.785690344896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.059
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.4138584723
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.755
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.937
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.801211531855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6099443307
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.055
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.789829328085
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.055
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.8986861441
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -74.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.409
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.30469717527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.409
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.7991870423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.275
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.96
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.16604123843
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8607761486
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -63.96
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.439
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.022
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.67700465257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.439
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.6953901489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -61.022
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.451
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.238
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.68942160214
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.451
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.9281074314
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -61.238
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.441
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.296
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.67907414417
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.441
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.9905963314
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -61.296
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.452
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.244
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.69045634794
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.452
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.9345718004
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -61.244
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.6
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.166
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.57182521705
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.6
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.0888522248
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -77.166
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.696
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.083
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.67116081359
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.696
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.0768232805
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.083
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.728
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.099
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7042726791
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.728
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.0940615978
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.099
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.779
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.306
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.75704471476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.779
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.3170823268
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.306
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.408
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.215
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.40789982123
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -4.408
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.9868549187
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.215
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.274
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.962
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.2692438844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -4.274
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.7916688539
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.962
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.17
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.16163032149
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -4.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.710864242
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.164
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.259
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.1554218467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -4.164
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.1116551174
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.259
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.011
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.835358143164
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.011
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0922502228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.694632714737
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.6235834733
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.205
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.634617458497
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.205
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.5524754148
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.252
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.093
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.585984406026
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.252
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.4684386184
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.093
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 23.958
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.482
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 23.9436994637
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 23.958
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.2561211114
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -36.482
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 24.021
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 24.0088884489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 24.021
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.4909931836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -36.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.377
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.41056849877
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.377
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.3364754337
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.436
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.401
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.47161850081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.436
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.4194348353
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -78.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.299
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.22562093482
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.299
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.0448566381
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.168
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.051
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.09006923538
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.168
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.8875569935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.051
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.184
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.02177601276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.0308505054
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.184
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.886
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.03729719972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.7097868471
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.886
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.576
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.921
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.44275392615
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.576
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0509422293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -66.921
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.486
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.3496268044
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.486
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9369159789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.472
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.811
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.33514036324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.472
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8550339721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.811
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.449
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.539
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.3113412099
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.449
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5619825794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.539
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.135
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.065
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.986431029563
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.135
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0512974317
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.17
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.733
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.02264713247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.6936023494
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.733
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.873
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.930554756512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.844437625
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.873
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.085
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.934693739701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.085
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8487472044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.034
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73624149115
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.1726878447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.034
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.692
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.719
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.84075081667
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.692
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8333084744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -66.719
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.793
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.117
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.9452601422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.793
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2621116153
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.117
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.842
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.767
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99596268626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.842
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8850234261
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -66.767
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.828
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.403
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.73825566432
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.828
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6476413619
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.546
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.236
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.44645734949
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.546
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4677164259
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.236
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.804
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.423927492095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.996149119
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -54.804
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.478
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.594
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.682613941409
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.478
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.6925013792
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -53.594
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.496
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.302
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.70123936576
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.496
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.3779020899
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -53.302
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.518
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.724003773299
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.518
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.3606637727
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -53.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.349
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.269
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.549131733563
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.349
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.19369115
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -64.269
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.334
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.358
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.533610546604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.334
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.3669741158
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.379
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.533
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.580174107481
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.379
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5555182104
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.533
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.34
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.64
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.539819021388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 1.34
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.6707994568
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -65.64
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.732
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.646
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.63892006778
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.732
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9094483047
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.646
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.836
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.9
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.74653363069
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.836
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1831065906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.9
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.879
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.957
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.79102769998
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.879
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2445180957
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.957
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.845
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.809
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.75584634287
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.845
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0850636614
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.809
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.213
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.29
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.06714120175
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.213
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5258957465
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.29
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.188
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.993674250144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4160014742
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.188
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.388
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.993674250144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6314804395
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.307
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.961597130429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5442114586
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.307
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.938
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0952982828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.938
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.35692236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.235
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1532440474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3892442048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.235
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.576
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.44024887471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1436099721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -72.576
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.255
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.149
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.14534632249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.255
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7609572075
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.149
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.137
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.13
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.02324631841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.137
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7404867058
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.13
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.091
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.207
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.97564801174
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.091
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.8234461075
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -73.207
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.892
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.033
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.94346258436
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.892
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.3325328444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -82.033
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.082
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.652
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.14006428584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -6.082
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.9994402419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -82.652
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.215
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.337
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.27768547687
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -6.215
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7374556979
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -83.337
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.235
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.264
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.29838039282
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -6.235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.6588058756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -83.264
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.559
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.1354897428
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -7.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.3636114809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -88.559
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.107
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.139
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.20067872803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -7.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.9885004802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.139
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.088
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.165
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.18101855788
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -7.088
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.0165127456
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.165
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.155
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.25034652629
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -7.155
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.2653909505
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.346
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.37849137906
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.346
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.6101455072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -84.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.272
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.836
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.30192019006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.272
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.2750757163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -83.836
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.299
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.949
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.32985832659
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.299
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.3968213316
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -83.949
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.786
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.2874337489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -5.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.2212059749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -83.786
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.824
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.48997199141
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.6491211099
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -88.824
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.552
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.337
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69588640506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.552
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.2018246558
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.337
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.389
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.77142284826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.2578491867
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.389
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.446
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.716581321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.3192606918
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -89.446
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.459
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.027
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.42592605964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.2486736492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -81.027
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.218
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.1765523225
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.218
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.2305355383
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -80.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.136
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.899
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.09170316713
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.136
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0333722851
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -79.899
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.137
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.973
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.09273791292
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -3.137
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.1130995022
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -79.973
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.667
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.478
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.57166159096
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.667
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.3476045843
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.478
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.577
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.274
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.47853446921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.577
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.1278160397
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.274
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.49302091037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.1105777225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.548
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.44852684108
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -1.548
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9284979969
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -76.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.954
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.291
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.83388783752
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.954
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.9913420992
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -74.291
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.956
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.194
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.83595732911
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.956
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.886834801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -74.194
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.138
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74076071576
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.8265006908
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -74.138
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.899
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.122
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.77697681867
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.899
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.8092623735
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -74.122
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.846740346934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.950603058198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.78
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.89199722826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5955829956
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.78
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.78
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.89199722826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -2.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5955829956
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -75.78
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.684
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.798
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.9305173729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.684
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1506071446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.798
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.671
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.91706567753
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.671
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.3768600581
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -70.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.672
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.91
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.91810042333
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.672
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2712753651
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.91
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.651
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.96
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.89637076159
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: 2.651
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.3251451065
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -69.96
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.729
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.742
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.60107003313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.729
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.012878208
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -68.742
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.558
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.4241285018
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.558
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0776994988
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.583
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.693
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.44999714673
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.583
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8826910352
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.693
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.44689290934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -0.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8805362456
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:00:51
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.938
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0952982828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.938
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.35692236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.235
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1532440474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3892442048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.235
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.938
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0952982828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.938
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.35692236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.6835
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.6821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.697
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.188
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4898
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.5693
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.235
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1532440474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.818307597074
measure_array: [ 2.6835 -14.4898]
inst: fit_type1
actual_mean: -8.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.966420934159
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3892442048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.882316338439
measure_array: [ 17.6821 -109.5693]
inst: fit_type1
actual_mean: -67.235
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.92816484318
tarray: []
source_col: d(D_H)Mean