## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
210528_20-212_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.5608 | -76.7483 |  | 27.432 |  | 0.21452 | 2.4471
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.38939 | 5.0395
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.59271 | 7.7234
4 |  | EV | -10.483 | -76.1061 |  |  |  | 0.16881 | 2.3927
5 |  | WOOD2-deep-3 | -11.6629 | -90.1586 |  |  |  | 0.078737 | 1.2703
6 |  | WOOD2-deep-2 | -12.6288 | -95.7753 |  |  |  | 0.084235 | 0.46233
7 |  | WOOD2-1020-1 | -9.4817 | -83.193 |  |  |  | 0.051227 | 0.67242
8 |  | WOOD2-C-3 | -6.6838 | -68.2193 |  |  |  | 0.15898 | 1.2608
9 | * | WOOD2-B-3 | 4.0351 | -55.411 | -8.7845 | 197.7644 | 33.1676 | 7.7506 | 45.9281 | -53.7905 | -46.3754
10 | * | WOOD2-5060-1 | -7.9129 | -59.0075 |  |  |  | 5.7689 | 50.2526
11 |  | WOOD2-E-2 | -7.07 | -81.4607 |  |  |  | 0.27648 | 1.5884
12 |  | WOOD2-C-2 | -6.5909 | -70.7598 |  |  |  | 0.019279 | 0.58998
13 |  | EV | -10.2562 | -74.7757 |  |  |  | 0.17573 | 0.31703
14 |  | WOOD2-C-5 | -3.6797 | -60.8818 |  |  |  | 0.31558 | 1.0944
15 |  | WOOD2-E-4 | -4.5374 | -71.7314 |  |  |  | 0.046629 | 0.93237
16 |  | WOOD2-010-2 | -8.4163 | -78.0544 |  |  |  | 0.21204 | 0.69591
17 |  | WOOD2-E-3 | -7.8872 | -82.3732 |  |  |  | 0.038108 | 0.30181
18 | * | WOOD2-D-1 | -0.65145 | -44.2731 | -3.7625 |  |  | 2.4477 | 32.5313 | -27.9872
19 | * | WOOD2-3040-1 | -6.399 | -52.3043 |  |  |  | 4.0559 | 42.1902
20 |  | WOOD2-1020-3 | -9.7541 | -82.2836 |  |  |  | 0.04771 | 0.51258
21 |  | WOOD2-deep-1 | -9.9615 | -82.8839 |  |  |  | 0.072113 | 0.63324
22 |  | EV | -10.3965 | -75.4686 |  |  |  | 0.027801 | 0.55132
23 | * | WOOD2-D-5 | -3.9892 | -58.1709 | -2.7348 |  |  | 1.3291 | 49.2466 | -16.029
24 | * | WOOD2-A-5 | -3.3749 | -42.0626 |  |  |  | 0.63477 | 29.8728
25 |  | WOOD2-A-1 | 3.2932 | -61.3075 |  | 19.3161 |  | 0.331 | 0.48076
26 |  | WOOD2-A-3 | -3.0493 | -68.8513 |  | 55.2384 |  | 0.29491 | 0.67122
27 |  | WOOD2-2030-3 | -11.1736 | -89.6095 |  |  |  | 0.38833 | 1.7299
28 |  | WOOD2-010-3 | -8.5063 | -78.7166 |  |  |  | 0.15111 | 1.0585
29 |  | WOOD2-E-5 | -3.4358 | -62.3953 |  |  |  | 0.22208 | 1.2866
30 |  | WOOD2-1020-2 | -9.8352 | -81.7278 |  |  |  | 0.33228 | 1.7034
31 | * | EV | -8.8432 | -63.7263 |  |  |  | 3.3024 | 27.0826
32 |  | WOOD2-E-1 | -6.4465 | -66.9005 |  | 779.2396 |  | 0.15843 | 0.61326 | -27.0673 | -10.4444
33 |  | WOOD2-A-2 | -2.5268 | -65.3263 |  | 14.9849 |  | 0.19775 | 0.12708
34 |  | WOOD2-5060-3 | -12.7145 | -98.9518 |  |  |  | 0.49087 | 2.6853
35 |  | WOOD2-2030-2 | -10.8315 | -87.2394 |  |  |  | 0.11214 | 1.1913
36 |  | WOOD2-C-1 | -6.6468 | -71.7208 |  |  |  | 0.1717 | 1.0069
37 |  | WOOD2-3040-2 | -11.4014 | -91.3056 |  |  |  | 0.20002 | 1.5072
38 | * | WOOD2-2030-1 | -6.2881 | -50.5827 |  |  |  | 3.9303 | 40.1197
39 | * | WOOD2-010-1 | -7.0756 | -65.8492 |  |  |  | 2.7842 | 33.7662
40 | * | EV | -6.5853 | -48.4495 |  |  |  | 4.2666 | 37.5537
41 | * | WOOD2-6070-2 | -7.6491 | -56.7318 |  |  |  | 5.4704 | 47.5153
42 | * | WOOD2-A-4 | -2.8146 | -17.2267 |  |  |  | 0 | 0
43 | * | WOOD2-C-4 | -2.8146 | -17.2267 |  |  |  | 0 | 0
44 | * | WOOD2-3040-3 | -2.8146 | -17.2267 |  |  |  | 0 | 0
45 | * | WOOD2-B-1 | -2.8146 | -17.2267 |  |  |  | 0 | 0
46 | * | WOOD2-D-4 | -2.8146 | -17.2267 |  |  |  | 0 | 0
47 | * | WOOD2-B-2 | 1.2632 | -26.8844 | -2.8285 |  |  | 7.992 | 20.1195 | -18.4085
48 | * | WOOD2-D-3 | -3.6601 | -32.1776 |  |  |  | 1.657 | 31.1465
49 |  | EV | -10.6521 | -79.4268 |  |  |  | 0.36875 | 1.8571
50 | * | WOOD2-B-4 | 3.1021 | -46.6086 | -3.9005 |  |  | 6.695 | 35.3396 | -29.5815
51 | * | WOOD2-B-5 | -1.2405 | -61.0971 | -2.8265 |  |  | 1.7811 | 52.7658 | -19.825
52 | * | WOOD2-D-2 | -4.5311 | -37.5555 |  |  |  | 3.364 | 42.35
53 | * | EV | -6.2689 | -48.6975 |  |  |  | 3.9677 | 37.8943

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -8.16 | -68.464 | -11.1417 | -82.9547 | 0.304 | 1473.246 | 2021/05/28 11:02:23
1 |  | EV | -7.714 | -63.583 | -10.6866 | -78.2688 | 0.286 | 1326.796 | 2021/05/28 11:11:21
1 |  | EV | -7.58 | -61.977 | -10.5498 | -76.727 | 0.447 | 1321.334 | 2021/05/28 11:20:17
1 |  | EV | -7.447 | -61.101 | -10.4141 | -75.886 | 0.255 | 1319.81 | 2021/05/28 11:29:13
1 |  | EV | -7.498 | -61.094 | -10.4662 | -75.8792 | 0.287 | 1315.777 | 2021/05/28 11:38:09
1 |  | EV | -7.55 | -61.128 | -10.5192 | -75.9119 | 0.401 | 1314.831 | 2021/05/28 11:47:07
1 |  | EV | -7.532 | -60.837 | -10.5009 | -75.6325 | 0.598 | 1306.933 | 2021/05/28 11:56:04
1 |  | EV | -7.465 | -60.652 | -10.4325 | -75.4549 | 0.193 | 1308.593 | 2021/05/28 12:05:01
1 |  | EV | -7.468 | -60.578 | -10.4355 | -75.3839 | 0.493 | 1304.161 | 2021/05/28 12:13:57
1 |  | EV | -7.493 | -60.578 | -10.4611 | -75.3839 | 0.427 | 1304.593 | 2021/05/28 12:22:54
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.5608 | -76.7483 |  | 27.432 |  | 0.21452 | 2.4471
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.628 | 23.67 | 0.88768 | 5.4974 | 0.145 | 1306.447 | 2021/05/28 12:32:43
2 |  | PZ | 4.381 | 33.473 | 1.6561 | 14.9086 | 0.408 | 1305.021 | 2021/05/28 12:41:39
2 |  | PZ | 4.707 | 36.657 | 1.9888 | 17.9654 | 0.169 | 1304.098 | 2021/05/28 12:50:36
2 |  | PZ | 4.755 | 37.837 | 2.0378 | 19.0982 | 0.416 | 1304.758 | 2021/05/28 12:59:34
2 |  | PZ | 4.782 | 38.308 | 2.0653 | 19.5504 | 0.42 | 1304.04 | 2021/05/28 13:08:30
2 |  | PZ | 4.789 | 38.691 | 2.0725 | 19.9181 | 0.215 | 1305.636 | 2021/05/28 13:17:27
2 |  | PZ | 4.804 | 39.136 | 2.0878 | 20.3453 | 0.457 | 1303.065 | 2021/05/28 13:26:24
2 |  | PZ | 4.823 | 39.507 | 2.1072 | 20.7015 | 0.259 | 1303.233 | 2021/05/28 13:35:21
2 |  | PZ | 4.861 | 40.015 | 2.1459 | 21.1892 | 0.474 | 1304.46 | 2021/05/28 13:44:17
2 |  | PZ | 4.964 | 40.678 | 2.251 | 21.8257 | 0.602 | 1306.753 | 2021/05/28 13:53:14
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.38939 | 5.0395
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -11.155 | -85.367 | -14.1981 | -99.1822 | 0.261 | 1301.616 | 2021/05/28 14:03:03
3 |  | UT2 | -12.556 | -102.233 | -15.6277 | -115.3742 | 0.215 | 1298.393 | 2021/05/28 14:12:00
3 |  | UT2 | -12.751 | -105.566 | -15.8267 | -118.574 | 0.602 | 1302.623 | 2021/05/28 14:20:57
3 |  | UT2 | -12.877 | -107.638 | -15.9553 | -120.5632 | 0.334 | 1299.355 | 2021/05/28 14:29:55
3 |  | UT2 | -12.973 | -108.64 | -16.0533 | -121.5252 | 0.376 | 1298.891 | 2021/05/28 14:38:51
3 |  | UT2 | -12.996 | -109.484 | -16.0768 | -122.3355 | 0.528 | 1298.056 | 2021/05/28 14:47:48
3 |  | UT2 | -13.055 | -109.673 | -16.137 | -122.5169 | 0.334 | 1299.88 | 2021/05/28 14:56:44
3 |  | UT2 | -13.116 | -110.486 | -16.1992 | -123.2974 | 0.345 | 1296.795 | 2021/05/28 15:05:42
3 |  | UT2 | -13.152 | -111.122 | -16.2359 | -123.908 | 0.487 | 1296.449 | 2021/05/28 15:14:39
3 |  | UT2 | -13.009 | -109.888 | -16.09 | -122.7233 | 0.405 | 1303.938 | 2021/05/28 15:23:36
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.59271 | 7.7234
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -7.94 | -67.516 | -10.9172 | -82.0446 | 0.459 | 1305.039 | 2021/05/28 15:33:26
4 |  | EV | -7.654 | -63.208 | -10.6254 | -77.9088 | 0.472 | 1296.686 | 2021/05/28 15:42:23
4 |  | EV | -7.511 | -61.198 | -10.4794 | -75.9791 | 0.356 | 1301.129 | 2021/05/28 15:51:19
4 |  | EV | -7.398 | -60.314 | -10.3641 | -75.1304 | 0.45 | 1300.832 | 2021/05/28 16:00:16
4 |  | EV | -7.511 | -60.835 | -10.4794 | -75.6306 | 0.254 | 1298.954 | 2021/05/28 16:09:12
4 |  | EV | -7.455 | -60.319 | -10.4223 | -75.1352 | 0.406 | 1298.563 | 2021/05/28 16:18:10
4 |  | EV | -7.426 | -60.011 | -10.3927 | -74.8395 | 0.566 | 1300.911 | 2021/05/28 16:27:07
4 |  | EV | -7.408 | -59.755 | -10.3743 | -74.5937 | 0.305 | 1298.801 | 2021/05/28 16:36:03
4 |  | EV | -7.388 | -59.938 | -10.3539 | -74.7694 | 0.653 | 1301.011 | 2021/05/28 16:45:00
4 |  | EV | -7.454 | -60.209 | -10.4213 | -75.0296 | 0.577 | 1301.466 | 2021/05/28 16:53:57
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.483 | -76.1061 |  |  |  | 0.16881 | 2.3927
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | WOOD2-deep-3 | -8.56 | -74.081 | -11.5499 | -88.3472 | 0.588 | 1302.435 | 2021/05/28 17:03:46
5 |  | WOOD2-deep-3 | -8.744 | -76.38 | -11.7377 | -90.5544 | 0.202 | 1302.06 | 2021/05/28 17:12:43
5 |  | WOOD2-deep-3 | -8.678 | -76.595 | -11.6703 | -90.7608 | 0.321 | 1301.702 | 2021/05/28 17:21:40
5 |  | WOOD2-deep-3 | -8.701 | -76.815 | -11.6938 | -90.972 | 0.562 | 1301.112 | 2021/05/28 17:30:35
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | WOOD2-deep-3 | -11.6629 | -90.1586 |  |  |  | 0.078737 | 1.2703
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | WOOD2-deep-2 | -9.514 | -81.142 | -12.5234 | -95.1261 | 0.275 | 1299.002 | 2021/05/28 17:40:25
6 |  | WOOD2-deep-2 | -9.664 | -82.044 | -12.6765 | -95.992 | 0.168 | 1300.272 | 2021/05/28 17:49:23
6 |  | WOOD2-deep-2 | -9.587 | -81.918 | -12.5979 | -95.8711 | 0.605 | 1299.65 | 2021/05/28 17:58:20
6 |  | WOOD2-deep-2 | -9.704 | -82.169 | -12.7173 | -96.112 | 0.218 | 1299.801 | 2021/05/28 18:07:17
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | WOOD2-deep-2 | -12.6288 | -95.7753 |  |  |  | 0.084235 | 0.46233
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | WOOD2-1020-1 | -6.61 | -69.679 | -9.56 | -84.1212 | 0.541 | 1297.767 | 2021/05/28 18:17:07
7 |  | WOOD2-1020-1 | -6.511 | -68.503 | -9.4589 | -82.9922 | 0.421 | 1297.004 | 2021/05/28 18:26:02
7 |  | WOOD2-1020-1 | -6.507 | -68.12 | -9.4549 | -82.6245 | 0.404 | 1297.06 | 2021/05/28 18:34:58
7 |  | WOOD2-1020-1 | -6.505 | -68.547 | -9.4528 | -83.0344 | 0.362 | 1288.667 | 2021/05/28 18:43:54
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | WOOD2-1020-1 | -9.4817 | -83.193 |  |  |  | 0.051227 | 0.67242
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | WOOD2-C-3 | -4.029 | -54.984 | -6.9261 | -70.0134 | 0.496 | 1304.307 | 2021/05/28 18:53:41
8 |  | WOOD2-C-3 | -3.729 | -52.766 | -6.62 | -67.884 | 0.496 | 1309.008 | 2021/05/28 19:02:37
8 |  | WOOD2-C-3 | -3.714 | -52.368 | -6.6047 | -67.5019 | 0.482 | 1304.71 | 2021/05/28 19:11:33
8 |  | WOOD2-C-3 | -3.694 | -52.343 | -6.5843 | -67.4779 | 0.457 | 1303.659 | 2021/05/28 19:20:31
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | WOOD2-C-3 | -6.6838 | -68.2193 |  |  |  | 0.15898 | 1.2608
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | WOOD2-B-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/28 19:30:19
9 |  | WOOD2-B-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/28 19:39:16
9 |  | WOOD2-B-3 | 13.423 | -79.114 | 10.8833 | -93.1791 | -17.822 | 2997.737 | 2021/05/28 19:48:12
9 |  | WOOD2-B-3 | 13.426 | -79.981 | 10.8863 | -94.0115 | -17.316 | 3002.022 | 2021/05/28 19:57:08
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 | * | WOOD2-B-3 | 4.0351 | -55.411 | -8.7845 | 197.7644 | 33.1676 | 7.7506 | 45.9281 | -53.7905 | -46.3754
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | WOOD2-5060-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/28 20:06:57
10 |  | WOOD2-5060-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/28 20:16:57
10 |  | WOOD2-5060-1 | -9.992 | -87.04 | -13.0112 | -100.7884 | 0.182 | 1307.731 | 2021/05/28 20:25:54
10 |  | WOOD2-5060-1 | -9.992 | -87.04 | -13.0112 | -100.7884 | 0.182 | 1307.731 | 2021/05/28 20:34:50
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 | * | WOOD2-5060-1 | -7.9129 | -59.0075 |  |  |  | 5.7689 | 50.2526
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | WOOD2-E-2 | -4.577 | -69.212 | -7.4853 | -83.6728 | 0.401 | 1301.521 | 2021/05/28 20:44:40
11 |  | WOOD2-E-2 | -4.1 | -66.666 | -6.9986 | -81.2286 | 0.332 | 1303.477 | 2021/05/28 20:53:37
11 |  | WOOD2-E-2 | -4.033 | -66.069 | -6.9302 | -80.6554 | 0.094 | 1302.046 | 2021/05/28 21:02:34
11 |  | WOOD2-E-2 | -3.97 | -65.684 | -6.8659 | -80.2858 | 0.441 | 1301.015 | 2021/05/28 21:11:30
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | WOOD2-E-2 | -7.07 | -81.4607 |  |  |  | 0.27648 | 1.5884
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | WOOD2-C-2 | -3.682 | -56.606 | -6.572 | -71.5706 | 0.22 | 1308.918 | 2021/05/28 21:21:20
12 |  | WOOD2-C-2 | -3.714 | -55.698 | -6.6047 | -70.6989 | 0.293 | 1305.78 | 2021/05/28 21:30:17
12 |  | WOOD2-C-2 | -3.686 | -55.476 | -6.5761 | -70.4857 | 0.254 | 1304.774 | 2021/05/28 21:39:12
12 |  | WOOD2-C-2 | -3.72 | -55.266 | -6.6108 | -70.2841 | 0.479 | 1305.173 | 2021/05/28 21:48:09
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | WOOD2-C-2 | -6.5909 | -70.7598 |  |  |  | 0.019279 | 0.58998
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -7.043 | -59.512 | -10.0018 | -74.3605 | 0.498 | 1296.617 | 2021/05/28 21:57:59
13 |  | EV | -7.317 | -60.208 | -10.2815 | -75.0286 | 0.322 | 1295.722 | 2021/05/28 22:06:55
13 |  | EV | -7.356 | -59.905 | -10.3213 | -74.7378 | 0.677 | 1299.09 | 2021/05/28 22:15:51
13 |  | EV | -7.453 | -60.153 | -10.4202 | -74.9758 | 0.441 | 1296.671 | 2021/05/28 22:24:48
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.2562 | -74.7757 |  |  |  | 0.17573 | 0.31703
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD2-C-5 | -1.304 | -47.051 | -4.1453 | -62.3974 | 0.105 | 1303.098 | 2021/05/28 22:34:38
14 |  | WOOD2-C-5 | -0.806 | -45.361 | -3.6371 | -60.775 | 0.24 | 1302.912 | 2021/05/28 22:43:35
14 |  | WOOD2-C-5 | -0.679 | -44.815 | -3.5075 | -60.2508 | 0.199 | 1302.155 | 2021/05/28 22:52:32
14 |  | WOOD2-C-5 | -0.602 | -44.662 | -3.4289 | -60.1039 | 0.427 | 1301.533 | 2021/05/28 23:01:28
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | WOOD2-C-5 | -3.6797 | -60.8818 |  |  |  | 0.31558 | 1.0944
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD2-E-4 | -1.62 | -55.398 | -4.4678 | -70.4109 | 0.163 | 1299.824 | 2021/05/28 23:11:14
15 |  | WOOD2-E-4 | -1.697 | -56.995 | -4.5464 | -71.944 | 0.275 | 1299.142 | 2021/05/28 23:20:11
15 |  | WOOD2-E-4 | -1.715 | -57.373 | -4.5647 | -72.3069 | 0.546 | 1298.285 | 2021/05/28 23:29:07
15 |  | WOOD2-E-4 | -1.721 | -57.328 | -4.5709 | -72.2637 | 0.465 | 1298.879 | 2021/05/28 23:38:05
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | WOOD2-E-4 | -4.5374 | -71.7314 |  |  |  | 0.046629 | 0.93237
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | WOOD2-010-2 | -5.176 | -62.331 | -8.0966 | -77.0668 | 0.364 | 1303.044 | 2021/05/28 23:47:52
16 |  | WOOD2-010-2 | -5.549 | -63.558 | -8.4772 | -78.2448 | 0.447 | 1300.633 | 2021/05/28 23:56:48
16 |  | WOOD2-010-2 | -5.593 | -63.703 | -8.5221 | -78.384 | 0.385 | 1300.373 | 2021/05/29 00:05:44
16 |  | WOOD2-010-2 | -5.639 | -63.847 | -8.5691 | -78.5222 | 0.275 | 1300.743 | 2021/05/29 00:14:39
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | WOOD2-010-2 | -8.4163 | -78.0544 |  |  |  | 0.21204 | 0.69591
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | WOOD2-E-3 | -5.019 | -67.412 | -7.9364 | -81.9448 | 0.318 | 1298.801 | 2021/05/29 00:24:27
17 |  | WOOD2-E-3 | -4.929 | -67.951 | -7.8446 | -82.4622 | 0.581 | 1297.895 | 2021/05/29 00:33:23
17 |  | WOOD2-E-3 | -4.979 | -67.996 | -7.8956 | -82.5054 | 0.327 | 1298.116 | 2021/05/29 00:42:19
17 |  | WOOD2-E-3 | -4.956 | -68.074 | -7.8721 | -82.5803 | 0.519 | 1298.167 | 2021/05/29 00:51:15
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | WOOD2-E-3 | -7.8872 | -82.3732 |  |  |  | 0.038108 | 0.30181
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | WOOD2-D-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 01:01:03
18 |  | WOOD2-D-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 01:09:59
18 |  | WOOD2-D-1 | 4.24 | -56.62 | 1.5122 | -71.584 | -7.377 | 2316.093 | 2021/05/29 01:18:56
18 |  | WOOD2-D-1 | 4.239 | -56.069 | 1.5112 | -71.055 | -7.673 | 2396.582 | 2021/05/29 01:27:52
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 | * | WOOD2-D-1 | -0.65145 | -44.2731 | -3.7625 |  |  | 2.4477 | 32.5313 | -27.9872
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | WOOD2-3040-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 01:37:40
19 |  | WOOD2-3040-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 01:46:37
19 |  | WOOD2-3040-1 | -7.039 | -73.043 | -9.9978 | -87.3507 | 0.427 | 1302.207 | 2021/05/29 01:55:33
19 |  | WOOD2-3040-1 | -7.011 | -73.108 | -9.9692 | -87.4131 | 0.599 | 1299.465 | 2021/05/29 02:04:31
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 | * | WOOD2-3040-1 | -6.399 | -52.3043 |  |  |  | 4.0559 | 42.1902
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | WOOD2-1020-3 | -6.87 | -68.495 | -9.8253 | -82.9845 | 0.337 | 1295.905 | 2021/05/29 02:14:19
20 |  | WOOD2-1020-3 | -6.783 | -67.722 | -9.7365 | -82.2424 | 0.241 | 1297.287 | 2021/05/29 02:23:15
20 |  | WOOD2-1020-3 | -6.762 | -67.515 | -9.7151 | -82.0436 | 0.261 | 1297.444 | 2021/05/29 02:32:11
20 |  | WOOD2-1020-3 | -6.786 | -67.328 | -9.7396 | -81.8641 | 0.274 | 1298.267 | 2021/05/29 02:41:07
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | WOOD2-1020-3 | -9.7541 | -82.2836 |  |  |  | 0.04771 | 0.51258
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | WOOD2-deep-1 | -7.081 | -69.13 | -10.0406 | -83.5941 | 0.393 | 1281.943 | 2021/05/29 02:50:56
21 |  | WOOD2-deep-1 | -6.908 | -67.583 | -9.8641 | -82.1089 | 0.46 | 1301.573 | 2021/05/29 02:59:52
21 |  | WOOD2-deep-1 | -7 | -68.394 | -9.958 | -82.8875 | 0.226 | 1290.409 | 2021/05/29 03:08:49
21 |  | WOOD2-deep-1 | -7.025 | -68.454 | -9.9835 | -82.9451 | 0.403 | 1291.678 | 2021/05/29 03:17:45
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | WOOD2-deep-1 | -9.9615 | -82.8839 |  |  |  | 0.072113 | 0.63324
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -7.435 | -61.438 | -10.4019 | -76.2095 | 0.438 | 1296.706 | 2021/05/29 03:27:30
22 |  | EV | -7.464 | -60.664 | -10.4315 | -75.4664 | 0.218 | 1296.059 | 2021/05/29 03:36:27
22 |  | EV | -7.397 | -60.181 | -10.3631 | -75.0027 | 0.258 | 1299.623 | 2021/05/29 03:45:23
22 |  | EV | -7.423 | -60.382 | -10.3896 | -75.1957 | 0.467 | 1293.283 | 2021/05/29 03:54:20
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.3965 | -75.4686 |  |  |  | 0.027801 | 0.55132
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | WOOD2-D-5 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 04:04:06
23 |  | WOOD2-D-5 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 04:13:01
23 |  | WOOD2-D-5 | -2.315 | -85.079 | -5.177 | -98.9057 | -5.478 | 1825.968 | 2021/05/29 04:21:57
23 |  | WOOD2-D-5 | -2.289 | -85.515 | -5.1505 | -99.3243 | -5.461 | 1822.257 | 2021/05/29 04:30:54
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 | * | WOOD2-D-5 | -3.9892 | -58.1709 | -2.7348 |  |  | 1.3291 | 49.2466 | -16.029
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | WOOD2-A-5 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 04:40:41
24 |  | WOOD2-A-5 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 04:49:37
24 |  | WOOD2-A-5 | -1.138 | -52.033 | -3.9759 | -67.1803 | 0.367 | 1382.561 | 2021/05/29 04:58:32
24 |  | WOOD2-A-5 | -1.058 | -51.446 | -3.8943 | -66.6168 | 0.428 | 1381.323 | 2021/05/29 05:07:28
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 | * | WOOD2-A-5 | -3.3749 | -42.0626 |  |  |  | 0.63477 | 29.8728
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | WOOD2-A-1 | 5.508 | -46.58 | 2.8062 | -61.9453 | 0.094 | 1323.541 | 2021/05/29 05:17:16
25 |  | WOOD2-A-1 | 6.024 | -45.959 | 3.3327 | -61.3491 | 0.128 | 1320.623 | 2021/05/29 05:26:13
25 |  | WOOD2-A-1 | 6.164 | -45.562 | 3.4756 | -60.9679 | 0.367 | 1320.615 | 2021/05/29 05:35:10
25 |  | WOOD2-A-1 | 6.245 | -45.562 | 3.5583 | -60.9679 | 0.023 | 1321.187 | 2021/05/29 05:44:05
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | WOOD2-A-1 | 3.2932 | -61.3075 |  | 19.3161 |  | 0.331 | 0.48076
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | WOOD2-A-3 | 0.199 | -52.832 | -2.6115 | -67.9474 | 0.315 | 1359.552 | 2021/05/29 05:53:53
26 |  | WOOD2-A-3 | -0.272 | -53.754 | -3.0922 | -68.8326 | 0.318 | 1357.135 | 2021/05/29 06:02:48
26 |  | WOOD2-A-3 | -0.413 | -54.293 | -3.2361 | -69.35 | 0.373 | 1356.133 | 2021/05/29 06:11:43
26 |  | WOOD2-A-3 | -0.434 | -54.215 | -3.2575 | -69.2751 | 0.109 | 1356.835 | 2021/05/29 06:20:40
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | WOOD2-A-3 | -3.0493 | -68.8513 |  | 55.2384 |  | 0.29491 | 0.67122
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | WOOD2-2030-3 | -7.629 | -72.894 | -10.5998 | -87.2077 | 0.211 | 1298.063 | 2021/05/29 06:30:28
27 |  | WOOD2-2030-3 | -8.247 | -75.65 | -11.2305 | -89.8535 | 0.332 | 1295.921 | 2021/05/29 06:39:25
27 |  | WOOD2-2030-3 | -8.395 | -76.266 | -11.3815 | -90.4449 | 0.241 | 1295.544 | 2021/05/29 06:48:21
27 |  | WOOD2-2030-3 | -8.494 | -76.773 | -11.4826 | -90.9317 | 0.175 | 1293.581 | 2021/05/29 06:57:18
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | WOOD2-2030-3 | -11.1736 | -89.6095 |  |  |  | 0.38833 | 1.7299
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | WOOD2-010-3 | -5.804 | -65.628 | -8.7375 | -80.232 | 0.298 | 1293.822 | 2021/05/29 07:07:06
28 |  | WOOD2-010-3 | -5.51 | -63.663 | -8.4374 | -78.3456 | 0.3 | 1297.565 | 2021/05/29 07:16:02
28 |  | WOOD2-010-3 | -5.497 | -63.524 | -8.4242 | -78.2121 | 0.497 | 1295.932 | 2021/05/29 07:24:59
28 |  | WOOD2-010-3 | -5.499 | -63.383 | -8.4262 | -78.0768 | 0.661 | 1295.436 | 2021/05/29 07:33:55
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | WOOD2-010-3 | -8.5063 | -78.7166 |  |  |  | 0.15111 | 1.0585
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | WOOD2-E-5 | -0.921 | -48.92 | -3.7545 | -64.1917 | 0.352 | 1297.674 | 2021/05/29 07:43:39
29 |  | WOOD2-E-5 | -0.607 | -46.848 | -3.434 | -62.2025 | 0.416 | 1297.177 | 2021/05/29 07:52:35
29 |  | WOOD2-E-5 | -0.487 | -46.333 | -3.3116 | -61.7081 | 0.283 | 1297.177 | 2021/05/29 08:01:32
29 |  | WOOD2-E-5 | -0.42 | -46.094 | -3.2432 | -61.4787 | 0.216 | 1297.775 | 2021/05/29 08:10:28
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | WOOD2-E-5 | -3.4358 | -62.3953 |  |  |  | 0.22208 | 1.2866
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | WOOD2-1020-2 | -6.387 | -64.698 | -9.3324 | -79.3392 | 0.354 | 1296.627 | 2021/05/29 08:20:13
30 |  | WOOD2-1020-2 | -6.98 | -67.501 | -9.9376 | -82.0302 | 0.393 | 1295.779 | 2021/05/29 08:29:08
30 |  | WOOD2-1020-2 | -7.05 | -68.107 | -10.009 | -82.612 | 0.139 | 1295.829 | 2021/05/29 08:38:04
30 |  | WOOD2-1020-2 | -7.102 | -68.438 | -10.0621 | -82.9298 | 0.485 | 1294.524 | 2021/05/29 08:47:00
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | WOOD2-1020-2 | -9.8352 | -81.7278 |  |  |  | 0.33228 | 1.7034
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 08:56:01
31 |  | EV | -7.386 | -61.513 | -10.3519 | -76.2815 | 0.275 | 1293.754 | 2021/05/29 09:05:59
31 |  | EV | -7.382 | -60.52 | -10.3478 | -75.3282 | 0.388 | 1294.751 | 2021/05/29 09:14:56
31 |  | EV | -7.385 | -60.162 | -10.3508 | -74.9845 | 0.185 | 1295.07 | 2021/05/29 09:23:52
31 |  | EV | -7.385 | -59.981 | -10.3508 | -74.8107 | 0.522 | 1295.025 | 2021/05/29 09:32:48
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 | * | EV | -8.8432 | -63.7263 |  |  |  | 3.3024 | 27.0826
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | WOOD2-E-1 | -3.795 | -52.639 | -6.6873 | -67.7621 | -0.49 | 2080.669 | 2021/05/29 09:42:34
32 |  | WOOD2-E-1 | -3.493 | -51.615 | -6.3791 | -66.779 | -0.658 | 2084.666 | 2021/05/29 09:51:29
32 |  | WOOD2-E-1 | -3.494 | -51.425 | -6.3802 | -66.5966 | -0.482 | 2081.129 | 2021/05/29 10:00:26
32 |  | WOOD2-E-1 | -3.454 | -51.287 | -6.3393 | -66.4641 | -0.805 | 2079.196 | 2021/05/29 10:09:22
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | WOOD2-E-1 | -6.4465 | -66.9005 |  | 779.2396 |  | 0.15843 | 0.61326 | -27.0673 | -10.4444
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | WOOD2-A-2 | -0.004 | -50.25 | -2.8187 | -65.4686 | 0.306 | 1322.598 | 2021/05/29 10:19:08
33 |  | WOOD2-A-2 | 0.333 | -50.162 | -2.4748 | -65.3841 | 0.378 | 1317.238 | 2021/05/29 10:28:04
33 |  | WOOD2-A-2 | 0.348 | -50.021 | -2.4595 | -65.2487 | 0.293 | 1315.406 | 2021/05/29 10:37:00
33 |  | WOOD2-A-2 | 0.451 | -49.974 | -2.3544 | -65.2036 | 0.338 | 1313.399 | 2021/05/29 10:45:56
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | WOOD2-A-2 | -2.5268 | -65.3263 |  | 14.9849 |  | 0.19775 | 0.12708
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | WOOD2-5060-3 | -8.981 | -81.241 | -11.9795 | -95.2211 | 0.097 | 1293.078 | 2021/05/29 10:55:42
34 |  | WOOD2-5060-3 | -9.802 | -85.471 | -12.8173 | -99.2821 | 0.228 | 1294.244 | 2021/05/29 11:04:38
34 |  | WOOD2-5060-3 | -9.979 | -86.633 | -12.998 | -100.3976 | 0.184 | 1294.749 | 2021/05/29 11:13:33
34 |  | WOOD2-5060-3 | -10.043 | -87.163 | -13.0633 | -100.9065 | 0.654 | 1293.636 | 2021/05/29 11:22:29
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | WOOD2-5060-3 | -12.7145 | -98.9518 |  |  |  | 0.49087 | 2.6853
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | WOOD2-2030-2 | -8.022 | -74.639 | -11.0009 | -88.8829 | 0.395 | 1289.9 | 2021/05/29 11:32:17
35 |  | WOOD2-2030-2 | -7.819 | -72.834 | -10.7937 | -87.1501 | 0.496 | 1288.733 | 2021/05/29 11:41:13
35 |  | WOOD2-2030-2 | -7.807 | -72.178 | -10.7815 | -86.5203 | 0.356 | 1293.392 | 2021/05/29 11:50:09
35 |  | WOOD2-2030-2 | -7.776 | -72.057 | -10.7499 | -86.4041 | 0.358 | 1291.966 | 2021/05/29 11:59:05
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | WOOD2-2030-2 | -10.8315 | -87.2394 |  |  |  | 0.11214 | 1.1913
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | WOOD2-C-1 | -3.989 | -58.139 | -6.8853 | -73.0423 | 0.13 | 1300.751 | 2021/05/29 12:08:48
36 |  | WOOD2-C-1 | -3.773 | -56.858 | -6.6649 | -71.8125 | 0.184 | 1296.713 | 2021/05/29 12:17:45
36 |  | WOOD2-C-1 | -3.66 | -56.191 | -6.5496 | -71.1722 | -0.019 | 1296.841 | 2021/05/29 12:26:40
36 |  | WOOD2-C-1 | -3.599 | -55.862 | -6.4873 | -70.8563 | 0.159 | 1297.097 | 2021/05/29 12:35:36
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | WOOD2-C-1 | -6.6468 | -71.7208 |  |  |  | 0.1717 | 1.0069
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | WOOD2-3040-2 | -8.119 | -74.979 | -11.0999 | -89.2094 | 0.253 | 1289.986 | 2021/05/29 12:45:20
37 |  | WOOD2-3040-2 | -8.464 | -77.338 | -11.4519 | -91.4741 | 0.582 | 1292.112 | 2021/05/29 12:54:16
37 |  | WOOD2-3040-2 | -8.538 | -78.157 | -11.5275 | -92.2604 | 0.349 | 1291.37 | 2021/05/29 13:03:11
37 |  | WOOD2-3040-2 | -8.537 | -78.176 | -11.5264 | -92.2786 | 0.275 | 1292.734 | 2021/05/29 13:12:07
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | WOOD2-3040-2 | -11.4014 | -91.3056 |  |  |  | 0.20002 | 1.5072
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | WOOD2-2030-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 13:21:51
38 |  | WOOD2-2030-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 13:30:47
38 |  | WOOD2-2030-1 | -6.812 | -69.344 | -9.7661 | -83.7995 | 0.179 | 1295.471 | 2021/05/29 13:39:42
38 |  | WOOD2-2030-1 | -6.803 | -69.634 | -9.7569 | -84.078 | 0.328 | 1289.782 | 2021/05/29 13:48:38
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 | * | WOOD2-2030-1 | -6.2881 | -50.5827 |  |  |  | 3.9303 | 40.1197
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | WOOD2-010-1 | -5.489 | -67.125 | -8.416 | -81.6692 | 0.01 | 1240.884 | 2021/05/29 13:58:22
39 |  | WOOD2-010-1 | -5.605 | -67.997 | -8.5344 | -82.5064 | 0.191 | 1222.773 | 2021/05/29 14:07:17
39 |  | WOOD2-010-1 | -5.608 | -67.464 | -8.5375 | -81.9947 | 0.056 | 1226.81 | 2021/05/29 14:16:12
39 |  | WOOD2-010-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 14:25:08
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 | * | WOOD2-010-1 | -7.0756 | -65.8492 |  |  |  | 2.7842 | 33.7662
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.39 | -65.045 | -10.3559 | -79.6723 | 0.081 | 1235.645 | 2021/05/29 14:35:54
40 |  | EV |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 14:44:50
40 |  | EV |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 14:54:48
40 |  | EV | -7.39 | -65.045 | -10.3559 | -79.6723 | 0.081 | 1235.645 | 2021/05/29 15:04:46
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 | * | EV | -6.5853 | -48.4495 |  |  |  | 4.2666 | 37.5537
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | WOOD2-6070-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 15:15:33
41 |  | WOOD2-6070-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 15:24:29
41 |  | WOOD2-6070-2 | -9.475 | -82.299 | -12.4836 | -96.2368 | -0.057 | 1233.731 | 2021/05/29 15:34:27
41 |  | WOOD2-6070-2 | -9.475 | -82.299 | -12.4836 | -96.2368 | -0.057 | 1233.731 | 2021/05/29 15:44:25
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 | * | WOOD2-6070-2 | -7.6491 | -56.7318 |  |  |  | 5.4704 | 47.5153
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | WOOD2-A-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 15:55:13
42 |  | WOOD2-A-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 16:05:12
42 |  | WOOD2-A-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 16:15:11
42 |  | WOOD2-A-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 16:25:10
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 | * | WOOD2-A-4 | -2.8146 | -17.2267 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | WOOD2-C-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 16:35:56
43 |  | WOOD2-C-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 16:45:53
43 |  | WOOD2-C-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 16:55:52
43 |  | WOOD2-C-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 17:05:50
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 | * | WOOD2-C-4 | -2.8146 | -17.2267 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | WOOD2-3040-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 17:16:38
44 |  | WOOD2-3040-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 17:26:36
44 |  | WOOD2-3040-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 17:36:34
44 |  | WOOD2-3040-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 17:46:32
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 | * | WOOD2-3040-3 | -2.8146 | -17.2267 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | WOOD2-B-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 17:57:20
45 |  | WOOD2-B-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 18:07:18
45 |  | WOOD2-B-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 18:17:15
45 |  | WOOD2-B-1 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 18:27:14
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 | * | WOOD2-B-1 | -2.8146 | -17.2267 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | WOOD2-D-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 18:38:03
46 |  | WOOD2-D-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 18:48:01
46 |  | WOOD2-D-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 18:57:59
46 |  | WOOD2-D-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 19:07:57
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 | * | WOOD2-D-4 | -2.8146 | -17.2267 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | WOOD2-B-2 | 15.984 | -40.239 | 13.4967 | -55.8577 | -11.314 | 2515.371 | 2021/05/29 19:18:46
47 |  | WOOD2-B-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 19:27:41
47 |  | WOOD2-B-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 19:37:40
47 |  | WOOD2-B-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 19:47:37
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 | * | WOOD2-B-2 | 1.2632 | -26.8844 | -2.8285 |  |  | 7.992 | 20.1195 | -18.4085
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | WOOD2-D-3 | -3.314 | -62.293 | -6.1965 | -77.0303 | -0.628 | 1296.325 | 2021/05/29 19:58:26
48 |  | WOOD2-D-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 20:07:21
48 |  | WOOD2-D-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 20:17:20
48 |  | WOOD2-D-3 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 20:27:17
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 | * | WOOD2-D-3 | -3.6601 | -32.1776 |  |  |  | 1.657 | 31.1465
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -7.154 | -62.178 | -10.1151 | -76.9199 | 0.175 | 1268.564 | 2021/05/29 20:38:05
49 |  | EV | -7.783 | -65.113 | -10.757 | -79.7376 | 0.107 | 1248.116 | 2021/05/29 20:47:00
49 |  | EV | -7.768 | -65.298 | -10.7417 | -79.9152 | -0.033 | 1234.185 | 2021/05/29 20:55:56
49 |  | EV | -8.016 | -66.568 | -10.9948 | -81.1345 | 0.092 | 1229.853 | 2021/05/29 21:04:52
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.6521 | -79.4268 |  |  |  | 0.36875 | 1.8571
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | WOOD2-B-4 | 11.596 | -61.21 | 9.0189 | -75.9906 | -7.801 | 2363.818 | 2021/05/29 21:14:37
50 |  | WOOD2-B-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 21:23:32
50 |  | WOOD2-B-4 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 21:33:30
50 |  | WOOD2-B-4 | 11.596 | -61.21 | 9.0189 | -75.9906 | -7.801 | 2363.818 | 2021/05/29 21:43:28
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 | * | WOOD2-B-4 | 3.1021 | -46.6086 | -3.9005 |  |  | 6.695 | 35.3396 | -29.5815
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | WOOD2-B-5 | 3.085 | -91.393 | 0.33356 | -104.9674 | -5.653 | 1969.628 | 2021/05/29 21:54:16
51 |  | WOOD2-B-5 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 22:03:11
51 |  | WOOD2-B-5 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 22:13:08
51 |  | WOOD2-B-5 | 3.085 | -91.393 | 0.33356 | -104.9674 | -5.653 | 1969.628 | 2021/05/29 22:23:03
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 | * | WOOD2-B-5 | -1.2405 | -61.0971 | -2.8265 |  |  | 1.7811 | 52.7658 | -19.825
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
52 |  | WOOD2-D-2 | -6.728 | -84.7 | -9.6804 | -98.5419 | -0.324 | 1268.271 | 2021/05/29 22:33:51
52 |  | WOOD2-D-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 22:42:47
52 |  | WOOD2-D-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 22:52:45
52 |  | WOOD2-D-2 |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 23:02:43
Sample 52 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
52 | * | WOOD2-D-2 | -4.5311 | -37.5555 |  |  |  | 3.364 | 42.35
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
53 |  | EV | -7.605 | -67.755 | -10.5754 | -82.274 | 0.08 | 1250.344 | 2021/05/29 23:13:32
53 |  | EV | -5.935 | -63.368 | -8.8712 | -78.0624 | 0.021 | 1224.515 | 2021/05/29 23:22:28
53 |  | EV |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 23:31:23
53 |  | EV |  |  | -2.8146 | -17.2267 |  |  | 2021/05/29 23:41:21
Sample 53 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
53 | * | EV | -6.2689 | -48.6975 |  |  |  | 3.9677 | 37.8943

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-16535 | 2021/05/28 11:02:23 | 1-01 | 1 | -8.16 | -68.464 | 22093 | -1 | 1 | EV | standard | H2O | 1622221649 | 0.344 | 0.722 | 130.527 | 0 | 0.001 | 3.055 | 1473.246 | -2.421 | 3.592 | 133.479 | 0.874 | 0.304 | 0 | 0 | 1 | 44.688 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-16535 | 2021/05/28 11:11:21 | 1-01 | 2 | -7.714 | -63.583 | 18082 | -1 | 1 | EV | standard | H2O | 1622222188 | 0.301 | 0.772 | 118.734 | 0 | 0 | 2.827 | 1326.796 | 0.892 | 2.896 | 111.09 | 0.86 | 0.286 | 0 | 0 | 1 | 44.062 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-16535 | 2021/05/28 11:20:17 | 1-01 | 3 | -7.58 | -61.977 | 18246 | -1 | 1 | EV | standard | H2O | 1622222724 | 0.31 | 0.722 | 116.796 | 0 | 0 | 2.772 | 1321.334 | 1.259 | 2.922 | 112.919 | 0.862 | 0.447 | 0 | 0 | 1 | 43.812 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-16535 | 2021/05/28 11:29:13 | 1-01 | 4 | -7.447 | -61.101 | 18473 | 0 | 1 | EV | standard | H2O | 1622223260 | 0.328 | 0.82 | 117.66 | 0 | 0.004 | 2.801 | 1319.81 | 1.646 | 2.946 | 112.586 | 0.866 | 0.255 | 0 | 0 | 1 | 43.688 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-16535 | 2021/05/28 11:38:09 | 1-01 | 5 | -7.498 | -61.094 | 18210 | 0 | 1 | EV | standard | H2O | 1622223796 | 0.293 | 0.754 | 115.048 | 0 | 0.003 | 2.729 | 1315.777 | 1.592 | 2.917 | 111.334 | 0.868 | 0.287 | 0 | 0 | 1 | 43.688 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-16535 | 2021/05/28 11:47:07 | 1-01 | 6 | -7.55 | -61.128 | 18298 | 0 | 1 | EV | standard | H2O | 1622224333 | 0.338 | 0.822 | 110.585 | 0 | 0.002 | 2.64 | 1314.831 | 1.53 | 2.964 | 112.902 | 0.863 | 0.401 | 0 | 0 | 1 | 43.625 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-16535 | 2021/05/28 11:56:04 | 1-01 | 7 | -7.532 | -60.837 | 17993 | 0 | 1 | EV | standard | H2O | 1622224871 | 0.303 | 0.741 | 111.738 | 0 | 0.001 | 2.649 | 1306.933 | 2.207 | 2.912 | 110.901 | 0.863 | 0.598 | 0 | 0 | 1 | 43.625 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-16535 | 2021/05/28 12:05:01 | 1-01 | 8 | -7.465 | -60.652 | 18083 | 0 | 1 | EV | standard | H2O | 1622225407 | 0.333 | 0.793 | 111.158 | -0.001 | 0.004 | 2.644 | 1308.593 | 1.985 | 2.863 | 110.159 | 0.862 | 0.193 | 0 | 0 | 1 | 43.812 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-16535 | 2021/05/28 12:13:57 | 1-01 | 9 | -7.468 | -60.578 | 17927 | 0 | 1 | EV | standard | H2O | 1622225944 | 0.283 | 0.694 | 111.797 | 0 | 0.004 | 2.656 | 1304.161 | 2.143 | 2.86 | 110.263 | 0.862 | 0.493 | 0 | 0 | 1 | 43.75 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-16535 | 2021/05/28 12:22:54 | 1-01 | 10 | -7.493 | -60.578 | 18006 | 0 | 1 | EV | standard | H2O | 1622226480 | 0.317 | 0.675 | 109.322 | 0 | 0.002 | 2.587 | 1304.593 | 2.059 | 2.871 | 110.332 | 0.865 | 0.427 | 0 | 0 | 1 | 43.75 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-16536 | 2021/05/28 12:32:43 | 1-02 | 1 | 3.628 | 23.67 | 18830 | -1 | 1 | PZ | standard | H2O | 1622227070 | 0.318 | 0.788 | 114.216 | -0.001 | -0.009 | 2.703 | 1306.447 | 2.011 | 3.027 | 114.787 | 0.865 | 0.145 | 0 | 0 | 1 | 43.812 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-16536 | 2021/05/28 12:41:39 | 1-02 | 2 | 4.381 | 33.473 | 17970 | -1 | 1 | PZ | standard | H2O | 1622227606 | 0.301 | 0.811 | 108.275 | -0.001 | -0.008 | 2.584 | 1305.021 | 1.905 | 2.919 | 111.167 | 0.863 | 0.408 | 0 | 0 | 1 | 43.812 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-16536 | 2021/05/28 12:50:36 | 1-02 | 3 | 4.707 | 36.657 | 18114 | -1 | 1 | PZ | standard | H2O | 1622228143 | 0.303 | 0.779 | 112.301 | -0.001 | -0.005 | 2.662 | 1304.098 | 1.404 | 2.866 | 111.211 | 0.869 | 0.169 | 0 | 0 | 1 | 43.812 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-16536 | 2021/05/28 12:59:34 | 1-02 | 4 | 4.755 | 37.837 | 18194 | 0 | 1 | PZ | standard | H2O | 1622228680 | 0.326 | 0.798 | 108.634 | 0 | -0.006 | 2.574 | 1304.758 | 2.084 | 2.969 | 111.515 | 0.873 | 0.416 | 0 | 0 | 1 | 43.812 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-16536 | 2021/05/28 13:08:30 | 1-02 | 5 | 4.782 | 38.308 | 18113 | 0 | 1 | PZ | standard | H2O | 1622229217 | 0.324 | 0.759 | 105.875 | -0.001 | -0.006 | 2.523 | 1304.04 | 1.657 | 2.953 | 111.855 | 0.864 | 0.42 | 0 | 0 | 1 | 43.75 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-16536 | 2021/05/28 13:17:27 | 1-02 | 6 | 4.789 | 38.691 | 18217 | 0 | 1 | PZ | standard | H2O | 1622229754 | 0.305 | 0.737 | 109.8 | 0 | -0.003 | 2.612 | 1305.636 | 1.11 | 2.957 | 112.501 | 0.861 | 0.215 | 0 | 0 | 1 | 43.688 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-16536 | 2021/05/28 13:26:24 | 1-02 | 7 | 4.804 | 39.136 | 18145 | 0 | 1 | PZ | standard | H2O | 1622230290 | 0.332 | 0.721 | 107.585 | -0.001 | -0.001 | 2.551 | 1303.065 | 1.948 | 2.936 | 111.647 | 0.864 | 0.457 | 0 | 0 | 1 | 43.75 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-16536 | 2021/05/28 13:35:21 | 1-02 | 8 | 4.823 | 39.507 | 18237 | 0 | 1 | PZ | standard | H2O | 1622230827 | 0.332 | 0.811 | 109.308 | 0 | 0.002 | 2.6 | 1303.233 | 1.517 | 2.93 | 112.179 | 0.866 | 0.259 | 0 | 0 | 1 | 43.75 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-16536 | 2021/05/28 13:44:17 | 1-02 | 9 | 4.861 | 40.015 | 18481 | 0 | 1 | PZ | standard | H2O | 1622231363 | 0.308 | 0.764 | 110.545 | -0.001 | -0.001 | 2.616 | 1304.46 | 1.58 | 3.02 | 114.278 | 0.865 | 0.474 | 0 | 0 | 1 | 43.75 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-16536 | 2021/05/28 13:53:14 | 1-02 | 10 | 4.964 | 40.678 | 18792 | 0 | 1 | PZ | standard | H2O | 1622231900 | 0.331 | 0.701 | 110.65 | 0 | -0.001 | 2.613 | 1306.753 | 1.927 | 3.095 | 116.014 | 0.866 | 0.602 | 0 | 0 | 1 | 43.812 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-16537 | 2021/05/28 14:03:03 | 1-03 | 1 | -11.155 | -85.367 | 18413 | -1 | 1 | UT2 | standard | H2O | 1622232490 | 0.318 | 1.119 | 108.199 | 0.001 | 0.021 | 2.559 | 1301.616 | 2.649 | 2.912 | 111.718 | 0.871 | 0.261 | 0 | 0 | 1 | 43.75 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-16537 | 2021/05/28 14:12:00 | 1-03 | 2 | -12.556 | -102.233 | 17782 | -1 | 1 | UT2 | standard | H2O | 1622233027 | 0.319 | 1.038 | 110.256 | 0.001 | 0.018 | 2.634 | 1298.393 | 2.668 | 2.789 | 107.704 | 0.862 | 0.215 | 0 | 0 | 1 | 43.75 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-16537 | 2021/05/28 14:20:57 | 1-03 | 3 | -12.751 | -105.566 | 18336 | -1 | 1 | UT2 | standard | H2O | 1622233564 | 0.298 | 0.872 | 110.226 | 0 | 0.011 | 2.629 | 1302.623 | 2.388 | 2.932 | 113.22 | 0.863 | 0.602 | 0 | 0 | 1 | 43.75 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-16537 | 2021/05/28 14:29:55 | 1-03 | 4 | -12.877 | -107.638 | 17879 | 0 | 1 | UT2 | standard | H2O | 1622234101 | 0.295 | 0.827 | 105.466 | 0 | 0.01 | 2.528 | 1299.355 | 2.147 | 2.813 | 110.054 | 0.864 | 0.334 | 0 | 0 | 1 | 43.812 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-16537 | 2021/05/28 14:38:51 | 1-03 | 5 | -12.973 | -108.64 | 17887 | 0 | 1 | UT2 | standard | H2O | 1622234638 | 0.316 | 0.804 | 106.915 | 0 | 0.009 | 2.545 | 1298.891 | 2.587 | 2.862 | 109.384 | 0.86 | 0.376 | 0 | 0 | 1 | 43.75 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-16537 | 2021/05/28 14:47:48 | 1-03 | 6 | -12.996 | -109.484 | 17863 | 0 | 1 | UT2 | standard | H2O | 1622235174 | 0.31 | 0.793 | 106.092 | 0 | 0.007 | 2.527 | 1298.056 | 2.745 | 2.892 | 110.004 | 0.864 | 0.528 | 0 | 0 | 1 | 43.812 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-16537 | 2021/05/28 14:56:44 | 1-03 | 7 | -13.055 | -109.673 | 18063 | 0 | 1 | UT2 | standard | H2O | 1622235711 | 0.303 | 0.741 | 107.558 | -0.001 | 0.004 | 2.56 | 1299.88 | 2.408 | 2.847 | 110.755 | 0.86 | 0.334 | 0 | 0 | 1 | 43.812 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-16537 | 2021/05/28 15:05:42 | 1-03 | 8 | -13.116 | -110.486 | 17751 | 0 | 1 | UT2 | standard | H2O | 1622236248 | 0.327 | 0.81 | 107.866 | 0 | 0.006 | 2.571 | 1296.795 | 2.598 | 2.842 | 108.335 | 0.863 | 0.345 | 0 | 0 | 1 | 43.812 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-16537 | 2021/05/28 15:14:39 | 1-03 | 9 | -13.152 | -111.122 | 17770 | 0 | 1 | UT2 | standard | H2O | 1622236785 | 0.312 | 0.831 | 107.133 | 0 | 0.009 | 2.552 | 1296.449 | 2.661 | 2.867 | 109.106 | 0.864 | 0.487 | 0 | 0 | 1 | 43.75 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-16537 | 2021/05/28 15:23:36 | 1-03 | 10 | -13.009 | -109.888 | 18578 | 0 | 1 | UT2 | standard | H2O | 1622237323 | 0.337 | 0.784 | 109.527 | 0 | 0.006 | 2.592 | 1303.938 | 2.324 | 2.912 | 113.709 | 0.865 | 0.405 | 0 | 0 | 1 | 43.75 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-16538 | 2021/05/28 15:33:26 | 1-04 | 1 | -7.94 | -67.516 | 18940 | -1 | 1 | EV | standard | H2O | 1622237912 | 0.317 | 0.765 | 109.086 | 0 | -0.001 | 2.596 | 1305.039 | 2.506 | 3.057 | 116.386 | 0.865 | 0.459 | 0 | 0 | 1 | 43.75 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-16538 | 2021/05/28 15:42:23 | 1-04 | 2 | -7.654 | -63.208 | 17780 | -1 | 1 | EV | standard | H2O | 1622238449 | 0.317 | 0.693 | 109.008 | 0 | 0.004 | 2.586 | 1296.686 | 2.914 | 2.833 | 108.521 | 0.863 | 0.472 | 0 | 0 | 1 | 43.812 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-16538 | 2021/05/28 15:51:19 | 1-04 | 3 | -7.511 | -61.198 | 18375 | -1 | 1 | EV | standard | H2O | 1622238986 | 0.303 | 0.715 | 109.811 | 0 | 0.001 | 2.61 | 1301.129 | 2.549 | 2.949 | 112.599 | 0.861 | 0.356 | 0 | 0 | 1 | 43.75 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-16538 | 2021/05/28 16:00:16 | 1-04 | 4 | -7.398 | -60.314 | 18369 | 0 | 1 | EV | standard | H2O | 1622239522 | 0.319 | 0.718 | 108.713 | 0 | -0.001 | 2.574 | 1300.832 | 2.191 | 2.925 | 112.798 | 0.864 | 0.45 | 0 | 0 | 1 | 43.812 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-16538 | 2021/05/28 16:09:12 | 1-04 | 5 | -7.511 | -60.835 | 18090 | 0 | 1 | EV | standard | H2O | 1622240059 | 0.327 | 0.749 | 106.836 | 0 | 0.006 | 2.532 | 1298.954 | 2.247 | 2.89 | 110.448 | 0.863 | 0.254 | 0 | 0 | 1 | 43.875 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-16538 | 2021/05/28 16:18:10 | 1-04 | 6 | -7.455 | -60.319 | 18090 | 0 | 1 | EV | standard | H2O | 1622240597 | 0.292 | 0.725 | 107.43 | -0.001 | 0.002 | 2.545 | 1298.563 | 2.549 | 2.893 | 111.066 | 0.863 | 0.406 | 0 | 0 | 1 | 43.812 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-16538 | 2021/05/28 16:27:07 | 1-04 | 7 | -7.426 | -60.011 | 18437 | 0 | 1 | EV | standard | H2O | 1622241134 | 0.313 | 0.728 | 109.912 | 0 | 0.001 | 2.611 | 1300.911 | 2.547 | 2.965 | 113.138 | 0.862 | 0.566 | 0 | 0 | 1 | 43.812 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-16538 | 2021/05/28 16:36:03 | 1-04 | 8 | -7.408 | -59.755 | 18129 | 0 | 1 | EV | standard | H2O | 1622241670 | 0.315 | 0.692 | 106.968 | 0 | 0.003 | 2.536 | 1298.801 | 2.568 | 2.837 | 109.815 | 0.863 | 0.305 | 0 | 0 | 1 | 43.875 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-16538 | 2021/05/28 16:45:00 | 1-04 | 9 | -7.388 | -59.938 | 18483 | 0 | 1 | EV | standard | H2O | 1622242206 | 0.297 | 0.742 | 107.053 | 0.001 | 0.003 | 2.548 | 1301.011 | 2.55 | 3.018 | 114.374 | 0.865 | 0.653 | 0 | 0 | 1 | 43.875 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-16538 | 2021/05/28 16:53:57 | 1-04 | 10 | -7.454 | -60.209 | 18544 | 0 | 1 | EV | standard | H2O | 1622242743 | 0.289 | 0.675 | 107.484 | 0.001 | 0.004 | 2.554 | 1301.466 | 2.242 | 2.998 | 114.87 | 0.866 | 0.577 | 0 | 0 | 1 | 43.875 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-16539 | 2021/05/28 17:03:46 | 1-05 | 1 | -8.56 | -74.081 | 18688 | -1 | 1 | WOOD2-deep-3 | 20-212 | H2O | 1622243333 | 0.322 | 0.766 | 110.247 | 0 | 0.002 | 2.613 | 1302.435 | 2.532 | 3.019 | 115.268 | 0.868 | 0.588 | 0 | 0 | 1 | 43.812 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-16539 | 2021/05/28 17:12:43 | 1-05 | 2 | -8.744 | -76.38 | 18582 | -1 | 1 | WOOD2-deep-3 | 20-212 | H2O | 1622243869 | 0.338 | 0.7 | 108.445 | 0.002 | 0.005 | 2.571 | 1302.06 | 2.575 | 2.947 | 112.379 | 0.863 | 0.202 | 0 | 0 | 1 | 43.812 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-16539 | 2021/05/28 17:21:40 | 1-05 | 3 | -8.678 | -76.595 | 18568 | -1 | 1 | WOOD2-deep-3 | 20-212 | H2O | 1622244406 | 0.319 | 0.764 | 110.545 | 0 | 0.006 | 2.617 | 1301.702 | 2.574 | 2.949 | 113.462 | 0.865 | 0.321 | 0 | 0 | 1 | 43.812 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-16539 | 2021/05/28 17:30:35 | 1-05 | 4 | -8.701 | -76.815 | 18510 | 0 | 1 | WOOD2-deep-3 | 20-212 | H2O | 1622244942 | 0.303 | 0.809 | 109.097 | 0 | 0.005 | 2.57 | 1301.112 | 2.971 | 2.961 | 113.069 | 0.864 | 0.562 | 0 | 0 | 1 | 43.812 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-16540 | 2021/05/28 17:40:25 | 1-06 | 1 | -9.514 | -81.142 | 18187 | -1 | 1 | WOOD2-deep-2 | 20-212 | H2O | 1622245531 | 0.315 | 0.805 | 105.699 | -0.001 | 0.004 | 2.529 | 1299.002 | 2.381 | 2.874 | 110.948 | 0.866 | 0.275 | 0 | 0 | 1 | 43.875 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-16540 | 2021/05/28 17:49:23 | 1-06 | 2 | -9.664 | -82.044 | 18357 | -1 | 1 | WOOD2-deep-2 | 20-212 | H2O | 1622246069 | 0.296 | 0.8 | 110.583 | 0.001 | 0.008 | 2.62 | 1300.272 | 2.526 | 2.913 | 111.323 | 0.864 | 0.168 | 0 | 0 | 1 | 43.812 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-16540 | 2021/05/28 17:58:20 | 1-06 | 3 | -9.587 | -81.918 | 18308 | -1 | 1 | WOOD2-deep-2 | 20-212 | H2O | 1622246607 | 0.295 | 0.713 | 106.803 | 0 | 0.002 | 2.533 | 1299.65 | 2.754 | 2.945 | 111.99 | 0.867 | 0.605 | 0 | 0 | 1 | 43.875 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-16540 | 2021/05/28 18:07:17 | 1-06 | 4 | -9.704 | -82.169 | 18316 | 0 | 1 | WOOD2-deep-2 | 20-212 | H2O | 1622247143 | 0.297 | 0.783 | 108.425 | 0.001 | 0.006 | 2.578 | 1299.801 | 2.87 | 2.89 | 110.918 | 0.866 | 0.218 | 0 | 0 | 1 | 43.875 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-16541 | 2021/05/28 18:17:07 | 1-07 | 1 | -6.61 | -69.679 | 18031 | -1 | 1 | WOOD2-1020-1 | 20-212 | H2O | 1622247734 | 0.314 | 0.755 | 105.704 | 0 | 0.003 | 2.519 | 1297.767 | 2.75 | 2.881 | 110.243 | 0.863 | 0.541 | 0 | 0 | 1 | 43.938 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-16541 | 2021/05/28 18:26:02 | 1-07 | 2 | -6.511 | -68.503 | 17911 | -1 | 1 | WOOD2-1020-1 | 20-212 | H2O | 1622248269 | 0.309 | 0.73 | 107.475 | -0.002 | 0 | 2.551 | 1297.004 | 2.374 | 2.892 | 109.852 | 0.862 | 0.421 | 0 | 0 | 1 | 43.875 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-16541 | 2021/05/28 18:34:58 | 1-07 | 3 | -6.507 | -68.12 | 17919 | -1 | 1 | WOOD2-1020-1 | 20-212 | H2O | 1622248805 | 0.275 | 0.759 | 110.197 | -0.001 | 0.005 | 2.632 | 1297.06 | 2.597 | 2.829 | 109.203 | 0.862 | 0.404 | 0 | 0 | 1 | 43.875 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-16541 | 2021/05/28 18:43:54 | 1-07 | 4 | -6.505 | -68.547 | 16721 | 0 | 1 | WOOD2-1020-1 | 20-212 | H2O | 1622249340 | 0.305 | 0.78 | 103.922 | -0.001 | 0.003 | 2.485 | 1288.667 | 2.106 | 2.641 | 102.291 | 0.862 | 0.362 | 0 | 0 | 1 | 43.75 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-16542 | 2021/05/28 18:53:41 | 1-08 | 1 | -4.029 | -54.984 | 17991 | -1 | 1 | WOOD2-C-3 | 20-212 | H2O | 1622249928 | 0.317 | 0.722 | 103.44 | -0.001 | 0.001 | 2.466 | 1304.307 | 2.098 | 2.859 | 110.458 | 0.864 | 0.496 | 0 | 0 | 1 | 43.875 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-16542 | 2021/05/28 19:02:37 | 1-08 | 2 | -3.729 | -52.766 | 18597 | -1 | 1 | WOOD2-C-3 | 20-212 | H2O | 1622250464 | 0.33 | 0.705 | 111.539 | 0 | 0.004 | 2.648 | 1309.008 | 2.095 | 3.016 | 114.416 | 0.862 | 0.496 | 0 | 0 | 1 | 43.875 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-16542 | 2021/05/28 19:11:33 | 1-08 | 3 | -3.714 | -52.368 | 18039 | -1 | 1 | WOOD2-C-3 | 20-212 | H2O | 1622251000 | 0.325 | 0.76 | 105.126 | -0.002 | 0.002 | 2.497 | 1304.71 | 2.288 | 2.922 | 110.771 | 0.867 | 0.482 | 0 | 0 | 1 | 43.875 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-16542 | 2021/05/28 19:20:31 | 1-08 | 4 | -3.694 | -52.343 | 17874 | 0 | 1 | WOOD2-C-3 | 20-212 | H2O | 1622251537 | 0.291 | 0.767 | 105.945 | 0 | 0.004 | 2.511 | 1303.659 | 1.723 | 2.877 | 110.233 | 0.863 | 0.457 | 0 | 0 | 1 | 43.875 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-16543 | 2021/05/28 19:30:19 | 1-09 | 1 | [5 empty] | WOOD2-B-3 | 20-212 | [18 empty] | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-16543 | 2021/05/28 19:39:16 | 1-09 | 2 | [5 empty] | WOOD2-B-3 | 20-212 | [18 empty] | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-16543 | 2021/05/28 19:48:12 | 1-09 | 3 | 13.423 | -79.114 | 16863 | -1 | 1 | WOOD2-B-3 | 20-212 | H2O | 1622253199 | 0.571 | 2.058 | 103.415 | -0.003 | 0.001 | 2.474 | 2997.737 | -107.027 | 7.767 | 128.568 | 0.88 | -17.822 | 0 | 0 | 1 | 43.875 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-16543 | 2021/05/28 19:57:08 | 1-09 | 4 | 13.426 | -79.981 | 16854 | 0 | 1 | WOOD2-B-3 | 20-212 | H2O | 1622253735 | 0.618 | 2.029 | 102.42 | -0.005 | 0.003 | 2.446 | 3002.022 | -108.135 | 7.771 | 132.159 | 0.881 | -17.316 | 0 | 0 | 1 | 43.938 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-16544 | 2021/05/28 20:06:57 | 1-10 | 1 | [5 empty] | WOOD2-5060-1 | 20-212 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-16544 | 2021/05/28 20:16:57 | 1-10 | 2 | [5 empty] | WOOD2-5060-1 | 20-212 | [18 empty] | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-16544 | 2021/05/28 20:25:54 | 1-10 | 3 | -9.992 | -87.04 | 18023 | -1 | 1 | WOOD2-5060-1 | 20-212 | H2O | 1622255460 | 0.325 | 0.721 | 104.972 | 0.001 | 0.001 | 2.495 | 1307.731 | 1.376 | 2.854 | 110.89 | 0.861 | 0.182 | 0 | 0 | 1 | 43.938 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-16544 | 2021/05/28 20:34:50 | 1-10 | 4 | -9.992 | -87.04 | 18023 | -1 | 1 | WOOD2-5060-1 | 20-212 | H2O | 1622255460 | 0.325 | 0.721 | 104.972 | 0.001 | 0.001 | 2.495 | 1307.731 | 1.376 | 2.854 | 110.89 | 0.861 | 0.182 | 0 | 0 | 1 | 43.938 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-16545 | 2021/05/28 20:44:40 | 1-11 | 1 | -4.577 | -69.212 | 18249 | -1 | 1 | WOOD2-E-2 | 20-212 | H2O | 1622256586 | 0.293 | 0.786 | 104.353 | -0.001 | 0.002 | 2.49 | 1301.521 | 1.978 | 2.961 | 112.406 | 0.863 | 0.401 | 0 | 0 | 1 | 43.938 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-16545 | 2021/05/28 20:53:37 | 1-11 | 2 | -4.1 | -66.666 | 18392 | -1 | 1 | WOOD2-E-2 | 20-212 | H2O | 1622257123 | 0.293 | 0.691 | 107.486 | -0.001 | 0.002 | 2.549 | 1303.477 | 1.911 | 2.952 | 112.925 | 0.863 | 0.332 | 0 | 0 | 1 | 43.938 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-16545 | 2021/05/28 21:02:34 | 1-11 | 3 | -4.033 | -66.069 | 18275 | -1 | 1 | WOOD2-E-2 | 20-212 | H2O | 1622257660 | 0.319 | 0.768 | 111.115 | -0.001 | 0 | 2.647 | 1302.046 | 2.184 | 2.872 | 111.356 | 0.864 | 0.094 | 0 | 0 | 1 | 43.875 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-16545 | 2021/05/28 21:11:30 | 1-11 | 4 | -3.97 | -65.684 | 18230 | 0 | 1 | WOOD2-E-2 | 20-212 | H2O | 1622258197 | 0.303 | 0.722 | 108.825 | 0 | 0.003 | 2.583 | 1301.015 | 2.619 | 2.916 | 111.353 | 0.862 | 0.441 | 0 | 0 | 1 | 43.875 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-16546 | 2021/05/28 21:21:20 | 1-12 | 1 | -3.682 | -56.606 | 18614 | -1 | 1 | WOOD2-C-2 | 20-212 | H2O | 1622258786 | 0.323 | 0.806 | 108.427 | 0 | 0.002 | 2.561 | 1308.918 | 2.008 | 2.956 | 113.551 | 0.862 | 0.22 | 0 | 0 | 1 | 43.875 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-16546 | 2021/05/28 21:30:17 | 1-12 | 2 | -3.714 | -55.698 | 18169 | -1 | 1 | WOOD2-C-2 | 20-212 | H2O | 1622259324 | 0.301 | 0.747 | 108.783 | 0 | 0.002 | 2.582 | 1305.78 | 1.872 | 2.871 | 111.712 | 0.862 | 0.293 | 0 | 0 | 1 | 43.875 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-16546 | 2021/05/28 21:39:12 | 1-12 | 3 | -3.686 | -55.476 | 18041 | -1 | 1 | WOOD2-C-2 | 20-212 | H2O | 1622259859 | 0.306 | 0.766 | 106.929 | -0.001 | -0.001 | 2.541 | 1304.774 | 1.892 | 2.889 | 110.876 | 0.861 | 0.254 | 0 | 0 | 1 | 43.875 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-16546 | 2021/05/28 21:48:09 | 1-12 | 4 | -3.72 | -55.266 | 18086 | 0 | 1 | WOOD2-C-2 | 20-212 | H2O | 1622260395 | 0.308 | 0.759 | 107.095 | 0 | 0 | 2.541 | 1305.173 | 1.327 | 2.948 | 112.054 | 0.863 | 0.479 | 0 | 0 | 1 | 43.938 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-16547 | 2021/05/28 21:57:59 | 1-13 | 1 | -7.043 | -59.512 | 17846 | -1 | 1 | EV | standard | H2O | 1622260985 | 0.305 | 0.773 | 110.062 | 0 | 0.002 | 2.618 | 1296.617 | 2.267 | 2.863 | 109.788 | 0.86 | 0.498 | 0 | 0 | 1 | 43.938 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-16547 | 2021/05/28 22:06:55 | 1-13 | 2 | -7.317 | -60.208 | 17748 | -1 | 1 | EV | standard | H2O | 1622261521 | 0.281 | 0.798 | 106.041 | 0 | 0.002 | 2.523 | 1295.722 | 2.732 | 2.795 | 108.202 | 0.86 | 0.322 | 0 | 0 | 1 | 43.875 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-16547 | 2021/05/28 22:15:51 | 1-13 | 3 | -7.356 | -59.905 | 18250 | -1 | 1 | EV | standard | H2O | 1622262058 | 0.316 | 0.78 | 108.322 | -0.001 | 0.005 | 2.572 | 1299.09 | 2.809 | 2.935 | 112.723 | 0.862 | 0.677 | 0 | 0 | 1 | 43.938 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-16547 | 2021/05/28 22:24:48 | 1-13 | 4 | -7.453 | -60.153 | 17882 | 0 | 1 | EV | standard | H2O | 1622262594 | 0.34 | 0.866 | 107.466 | 0 | 0.004 | 2.566 | 1296.671 | 1.936 | 2.874 | 110.955 | 0.862 | 0.441 | 0 | 0 | 1 | 43.938 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-16548 | 2021/05/28 22:34:38 | 1-14 | 1 | -1.304 | -47.051 | 18226 | -1 | 1 | WOOD2-C-5 | 20-212 | H2O | 1622263185 | 0.303 | 0.771 | 107.967 | -0.002 | -0.002 | 2.568 | 1303.098 | 1.724 | 2.907 | 111.683 | 0.86 | 0.105 | 0 | 0 | 1 | 43.875 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-16548 | 2021/05/28 22:43:35 | 1-14 | 2 | -0.806 | -45.361 | 18202 | -1 | 1 | WOOD2-C-5 | 20-212 | H2O | 1622263722 | 0.265 | 0.788 | 107.723 | 0 | -0.001 | 2.574 | 1302.912 | 1.983 | 2.911 | 110.511 | 0.862 | 0.24 | 0 | 0 | 1 | 43.875 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-16548 | 2021/05/28 22:52:32 | 1-14 | 3 | -0.679 | -44.815 | 18084 | -1 | 1 | WOOD2-C-5 | 20-212 | H2O | 1622264258 | 0.31 | 0.741 | 109.368 | -0.001 | 0 | 2.602 | 1302.155 | 1.846 | 2.85 | 110.157 | 0.861 | 0.199 | 0 | 0 | 1 | 43.938 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-16548 | 2021/05/28 23:01:28 | 1-14 | 4 | -0.602 | -44.662 | 18054 | 0 | 1 | WOOD2-C-5 | 20-212 | H2O | 1622264794 | 0.327 | 0.711 | 106.914 | 0 | 0 | 2.543 | 1301.533 | 2.164 | 2.869 | 111.125 | 0.861 | 0.427 | 0 | 0 | 1 | 43.688 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-16549 | 2021/05/28 23:11:14 | 1-15 | 1 | -1.62 | -55.398 | 18203 | -1 | 1 | WOOD2-E-4 | 20-212 | H2O | 1622265381 | 0.307 | 0.764 | 106.953 | 0 | 0.002 | 2.547 | 1299.824 | 1.857 | 2.911 | 111.046 | 0.86 | 0.163 | 0 | 0 | 1 | 43.75 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-16549 | 2021/05/28 23:20:11 | 1-15 | 2 | -1.697 | -56.995 | 18132 | -1 | 1 | WOOD2-E-4 | 20-212 | H2O | 1622265917 | 0.297 | 0.724 | 109.415 | 0 | 0.003 | 2.609 | 1299.142 | 2.326 | 2.852 | 110.377 | 0.864 | 0.275 | 0 | 0 | 1 | 43.812 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-16549 | 2021/05/28 23:29:07 | 1-15 | 3 | -1.715 | -57.373 | 18027 | -1 | 1 | WOOD2-E-4 | 20-212 | H2O | 1622266454 | 0.327 | 0.736 | 108.415 | -0.001 | 0.005 | 2.574 | 1298.285 | 2.012 | 2.941 | 111.852 | 0.863 | 0.546 | 0 | 0 | 1 | 43.812 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-16549 | 2021/05/28 23:38:05 | 1-15 | 4 | -1.721 | -57.328 | 18101 | 0 | 1 | WOOD2-E-4 | 20-212 | H2O | 1622266991 | 0.328 | 0.678 | 106.53 | 0 | 0.003 | 2.544 | 1298.879 | 2.091 | 2.916 | 111.106 | 0.867 | 0.465 | 0 | 0 | 1 | 43.812 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-16550 | 2021/05/28 23:47:52 | 1-16 | 1 | -5.176 | -62.331 | 18695 | -1 | 1 | WOOD2-010-2 | 20-212 | H2O | 1622267578 | 0.32 | 0.714 | 109.097 | 0 | 0.004 | 2.586 | 1303.044 | 2.646 | 2.978 | 113.839 | 0.864 | 0.364 | 0 | 0 | 1 | 43.75 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-16550 | 2021/05/28 23:56:48 | 1-16 | 2 | -5.549 | -63.558 | 18381 | -1 | 1 | WOOD2-010-2 | 20-212 | H2O | 1622268115 | 0.304 | 0.754 | 110.769 | 0 | 0.003 | 2.623 | 1300.633 | 2.103 | 2.944 | 113.181 | 0.864 | 0.447 | 0 | 0 | 1 | 43.812 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-16550 | 2021/05/29 00:05:44 | 1-16 | 3 | -5.593 | -63.703 | 18325 | -1 | 1 | WOOD2-010-2 | 20-212 | H2O | 1622268650 | 0.287 | 0.727 | 108.714 | -0.001 | 0.002 | 2.567 | 1300.373 | 2.128 | 2.955 | 112.163 | 0.866 | 0.385 | 0 | 0 | 1 | 43.812 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-16550 | 2021/05/29 00:14:39 | 1-16 | 4 | -5.639 | -63.847 | 18381 | 0 | 1 | WOOD2-010-2 | 20-212 | H2O | 1622269186 | 0.343 | 0.686 | 108.289 | -0.001 | 0.002 | 2.584 | 1300.743 | 2.617 | 2.918 | 110.957 | 0.863 | 0.275 | 0 | 0 | 1 | 43.812 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-16551 | 2021/05/29 00:24:27 | 1-17 | 1 | -5.019 | -67.412 | 18117 | -1 | 1 | WOOD2-E-3 | 20-212 | H2O | 1622269773 | 0.32 | 0.808 | 106.777 | 0 | 0.004 | 2.553 | 1298.801 | 2.343 | 2.882 | 110.854 | 0.861 | 0.318 | 0 | 0 | 1 | 43.875 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-16551 | 2021/05/29 00:33:23 | 1-17 | 2 | -4.929 | -67.951 | 18010 | -1 | 1 | WOOD2-E-3 | 20-212 | H2O | 1622270309 | 0.307 | 0.774 | 110.345 | 0.001 | 0.003 | 2.611 | 1297.895 | 2.493 | 2.882 | 110.946 | 0.865 | 0.581 | 0 | 0 | 1 | 43.812 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-16551 | 2021/05/29 00:42:19 | 1-17 | 3 | -4.979 | -67.996 | 18028 | -1 | 1 | WOOD2-E-3 | 20-212 | H2O | 1622270845 | 0.287 | 0.698 | 109.234 | -0.001 | 0.002 | 2.594 | 1298.116 | 2.247 | 2.875 | 110.641 | 0.862 | 0.327 | 0 | 0 | 1 | 43.875 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-16551 | 2021/05/29 00:51:15 | 1-17 | 4 | -4.956 | -68.074 | 18035 | 0 | 1 | WOOD2-E-3 | 20-212 | H2O | 1622271381 | 0.29 | 0.714 | 111.035 | 0 | 0.003 | 2.632 | 1298.167 | 1.822 | 2.943 | 111.747 | 0.863 | 0.519 | 0 | 0 | 1 | 43.812 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-16552 | 2021/05/29 01:01:03 | 1-18 | 1 | [5 empty] | WOOD2-D-1 | 20-212 | [18 empty] | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-16552 | 2021/05/29 01:09:59 | 1-18 | 2 | [5 empty] | WOOD2-D-1 | 20-212 | [18 empty] | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-16552 | 2021/05/29 01:18:56 | 1-18 | 3 | 4.24 | -56.62 | 14474 | -1 | 1 | WOOD2-D-1 | 20-212 | H2O | 1622273042 | 0.488 | 1.506 | 94.611 | -0.003 | -0.007 | 2.291 | 2316.093 | -53.957 | 3.595 | 95.098 | 0.876 | -7.377 | 0 | 0 | 1 | 43.875 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-16552 | 2021/05/29 01:27:52 | 1-18 | 4 | 4.239 | -56.069 | 15462 | 0 | 1 | WOOD2-D-1 | 20-212 | H2O | 1622273579 | 0.491 | 1.649 | 97.992 | -0.002 | 0.004 | 2.351 | 2396.582 | -57.992 | 3.836 | 102.88 | 0.867 | -7.673 | 0 | 0 | 1 | 43.812 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-16553 | 2021/05/29 01:37:40 | 1-19 | 1 | [5 empty] | WOOD2-3040-1 | 20-212 | [18 empty] | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-16553 | 2021/05/29 01:46:37 | 1-19 | 2 | [5 empty] | WOOD2-3040-1 | 20-212 | [18 empty] | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-16553 | 2021/05/29 01:55:33 | 1-19 | 3 | -7.039 | -73.043 | 17899 | -1 | 1 | WOOD2-3040-1 | 20-212 | H2O | 1622275240 | 0.285 | 0.742 | 107.022 | 0 | 0.005 | 2.553 | 1302.207 | 2.521 | 2.836 | 109.74 | 0.863 | 0.427 | 0 | 0 | 1 | 43.812 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-16553 | 2021/05/29 02:04:31 | 1-19 | 4 | -7.011 | -73.108 | 17804 | 0 | 1 | WOOD2-3040-1 | 20-212 | H2O | 1622275777 | 0.292 | 0.775 | 106.074 | 0 | 0.006 | 2.522 | 1299.465 | 2.374 | 2.878 | 110.378 | 0.861 | 0.599 | 0 | 0 | 1 | 43.875 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-16554 | 2021/05/29 02:14:19 | 1-20 | 1 | -6.87 | -68.495 | 17631 | -1 | 1 | WOOD2-1020-3 | 20-212 | H2O | 1622276365 | 0.32 | 0.756 | 105.219 | 0 | 0.001 | 2.508 | 1295.905 | 2.311 | 2.821 | 108.092 | 0.864 | 0.337 | 0 | 0 | 1 | 43.875 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-16554 | 2021/05/29 02:23:15 | 1-20 | 2 | -6.783 | -67.722 | 17731 | -1 | 1 | WOOD2-1020-3 | 20-212 | H2O | 1622276901 | 0.32 | 0.737 | 107.034 | -0.001 | 0.002 | 2.553 | 1297.287 | 3.056 | 2.756 | 106.064 | 0.863 | 0.241 | 0 | 0 | 1 | 43.938 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-16554 | 2021/05/29 02:32:11 | 1-20 | 3 | -6.762 | -67.515 | 17801 | -1 | 1 | WOOD2-1020-3 | 20-212 | H2O | 1622277437 | 0.312 | 0.688 | 103.944 | -0.001 | -0.001 | 2.467 | 1297.444 | 2.748 | 2.811 | 107.716 | 0.862 | 0.261 | 0 | 0 | 1 | 43.938 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-16554 | 2021/05/29 02:41:07 | 1-20 | 4 | -6.786 | -67.328 | 17950 | 0 | 1 | WOOD2-1020-3 | 20-212 | H2O | 1622277974 | 0.313 | 0.711 | 108.814 | -0.002 | 0.002 | 2.579 | 1298.267 | 2.57 | 2.83 | 109.323 | 0.864 | 0.274 | 0 | 0 | 1 | 43.812 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-16555 | 2021/05/29 02:50:56 | 1-21 | 1 | -7.081 | -69.13 | 15769 | -1 | 1 | WOOD2-deep-1 | 20-212 | H2O | 1622278563 | 0.317 | 0.748 | 100.178 | -0.001 | 0.003 | 2.401 | 1281.943 | 2.739 | 2.453 | 96.033 | 0.865 | 0.393 | 0 | 0 | 1 | 43.875 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-16555 | 2021/05/29 02:59:52 | 1-21 | 2 | -6.908 | -67.583 | 18464 | -1 | 1 | WOOD2-deep-1 | 20-212 | H2O | 1622279099 | 0.333 | 0.752 | 107.466 | 0.001 | 0.006 | 2.55 | 1301.573 | 2.271 | 2.94 | 113.195 | 0.863 | 0.46 | 0 | 0 | 1 | 43.875 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-16555 | 2021/05/29 03:08:49 | 1-21 | 3 | -7 | -68.394 | 16935 | -1 | 1 | WOOD2-deep-1 | 20-212 | H2O | 1622279635 | 0.33 | 0.783 | 104.498 | -0.001 | 0.004 | 2.502 | 1290.409 | 2.74 | 2.643 | 102.591 | 0.865 | 0.226 | 0 | 0 | 1 | 43.938 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-16555 | 2021/05/29 03:17:45 | 1-21 | 4 | -7.025 | -68.454 | 17127 | 0 | 1 | WOOD2-deep-1 | 20-212 | H2O | 1622280172 | 0.351 | 0.68 | 104.524 | -0.001 | 0.001 | 2.493 | 1291.678 | 2.219 | 2.736 | 105.408 | 0.868 | 0.403 | 0 | 0 | 1 | 43.875 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-16556 | 2021/05/29 03:27:30 | 1-22 | 1 | -7.435 | -61.438 | 17878 | -1 | 1 | EV | standard | H2O | 1622280757 | 0.336 | 0.731 | 106.376 | 0 | 0 | 2.531 | 1296.706 | 2.267 | 2.835 | 109.749 | 0.86 | 0.438 | 0 | 0 | 1 | 43.875 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-16556 | 2021/05/29 03:36:27 | 1-22 | 2 | -7.464 | -60.664 | 17803 | -1 | 1 | EV | standard | H2O | 1622281293 | 0.332 | 0.751 | 106.137 | 0 | 0.001 | 2.52 | 1296.059 | 2.69 | 2.841 | 108.244 | 0.868 | 0.218 | 0 | 0 | 1 | 43.875 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-16556 | 2021/05/29 03:45:23 | 1-22 | 3 | -7.397 | -60.181 | 18289 | -1 | 1 | EV | standard | H2O | 1622281829 | 0.291 | 0.813 | 110.405 | 0.001 | 0.007 | 2.622 | 1299.623 | 3.103 | 2.874 | 110.047 | 0.873 | 0.258 | 0 | 0 | 1 | 43.812 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-16556 | 2021/05/29 03:54:20 | 1-22 | 4 | -7.423 | -60.382 | 17402 | 0 | 1 | EV | standard | H2O | 1622282366 | 0.342 | 0.758 | 104.722 | 0 | 0.002 | 2.482 | 1293.283 | 2.38 | 2.774 | 106.613 | 0.863 | 0.467 | 0 | 0 | 1 | 43.75 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-16557 | 2021/05/29 04:04:06 | 1-23 | 1 | [5 empty] | WOOD2-D-5 | 20-212 | [18 empty] | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-16557 | 2021/05/29 04:13:01 | 1-23 | 2 | [5 empty] | WOOD2-D-5 | 20-212 | [18 empty] | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-16557 | 2021/05/29 04:21:57 | 1-23 | 3 | -2.315 | -85.079 | 17748 | -1 | 1 | WOOD2-D-5 | 20-212 | H2O | 1622284024 | 0.359 | 1.165 | 110.067 | -0.002 | 0.006 | 2.604 | 1825.968 | -32.306 | 3.503 | 116.814 | 0.875 | -5.478 | 0 | 0 | 1 | 43.875 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-16557 | 2021/05/29 04:30:54 | 1-23 | 4 | -2.289 | -85.515 | 17646 | 0 | 1 | WOOD2-D-5 | 20-212 | H2O | 1622284560 | 0.373 | 1.05 | 106.748 | -0.001 | 0.003 | 2.548 | 1822.257 | -31.81 | 3.498 | 116.049 | 0.872 | -5.461 | 0 | 0 | 1 | 43.875 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-16558 | 2021/05/29 04:40:41 | 1-24 | 1 | [5 empty] | WOOD2-A-5 | 20-212 | [18 empty] | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-16558 | 2021/05/29 04:49:37 | 1-24 | 2 | [5 empty] | WOOD2-A-5 | 20-212 | [18 empty] | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-16558 | 2021/05/29 04:58:32 | 1-24 | 3 | -1.138 | -52.033 | 17834 | -1 | 1 | WOOD2-A-5 | 20-212 | H2O | 1622286219 | 0.335 | 0.825 | 107.636 | -0.001 | -0.002 | 2.569 | 1382.561 | -1.233 | 2.841 | 109.592 | 0.863 | 0.367 | 0 | 0 | 1 | 43.875 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-16558 | 2021/05/29 05:07:28 | 1-24 | 4 | -1.058 | -51.446 | 17800 | 0 | 1 | WOOD2-A-5 | 20-212 | H2O | 1622286755 | 0.354 | 0.818 | 104.843 | -0.001 | 0 | 2.492 | 1381.323 | -1.44 | 2.879 | 110.552 | 0.866 | 0.428 | 0 | 0 | 1 | 43.75 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-16559 | 2021/05/29 05:17:16 | 1-25 | 1 | 5.508 | -46.58 | 18036 | -1 | 1 | WOOD2-A-1 | 20-212 | H2O | 1622287342 | 0.325 | 0.746 | 108.828 | -0.003 | 0.001 | 2.586 | 1323.541 | 0.881 | 2.869 | 111.101 | 0.861 | 0.094 | 0 | 0 | 1 | 43.812 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-16559 | 2021/05/29 05:26:13 | 1-25 | 2 | 6.024 | -45.959 | 17762 | -1 | 1 | WOOD2-A-1 | 20-212 | H2O | 1622287879 | 0.342 | 0.701 | 106.277 | -0.001 | -0.001 | 2.525 | 1320.623 | 1.417 | 2.847 | 108.098 | 0.862 | 0.128 | 0 | 0 | 1 | 43.875 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-16559 | 2021/05/29 05:35:10 | 1-25 | 3 | 6.164 | -45.562 | 17793 | -1 | 1 | WOOD2-A-1 | 20-212 | H2O | 1622288417 | 0.312 | 0.744 | 106.072 | 0 | -0.001 | 2.527 | 1320.615 | 1.096 | 2.844 | 109.56 | 0.864 | 0.367 | 0 | 0 | 1 | 43.812 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-16559 | 2021/05/29 05:44:05 | 1-25 | 4 | 6.245 | -45.562 | 17858 | 0 | 1 | WOOD2-A-1 | 20-212 | H2O | 1622288952 | 0.32 | 0.776 | 108.793 | -0.001 | 0.001 | 2.581 | 1321.187 | 1.184 | 2.811 | 107.912 | 0.863 | 0.023 | 0 | 0 | 1 | 43.812 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-16560 | 2021/05/29 05:53:53 | 1-26 | 1 | 0.199 | -52.832 | 18043 | -1 | 1 | WOOD2-A-3 | 20-212 | H2O | 1622289539 | 0.302 | 0.742 | 109.4 | 0 | 0.004 | 2.585 | 1359.552 | 0.168 | 2.868 | 109.958 | 0.866 | 0.315 | 0 | 0 | 1 | 43.75 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-16560 | 2021/05/29 06:02:48 | 1-26 | 2 | -0.272 | -53.754 | 17767 | -1 | 1 | WOOD2-A-3 | 20-212 | H2O | 1622290075 | 0.303 | 0.703 | 105.657 | -0.001 | 0.001 | 2.524 | 1357.135 | -0.297 | 2.806 | 109.088 | 0.862 | 0.318 | 0 | 0 | 1 | 43.75 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-16560 | 2021/05/29 06:11:43 | 1-26 | 3 | -0.413 | -54.293 | 17702 | -1 | 1 | WOOD2-A-3 | 20-212 | H2O | 1622290610 | 0.278 | 0.806 | 102.976 | -0.001 | 0.002 | 2.436 | 1356.133 | 0.444 | 2.84 | 108.213 | 0.87 | 0.373 | 0 | 0 | 1 | 43.75 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-16560 | 2021/05/29 06:20:40 | 1-26 | 4 | -0.434 | -54.215 | 17729 | 0 | 1 | WOOD2-A-3 | 20-212 | H2O | 1622291147 | 0.335 | 0.805 | 105.413 | -0.001 | -0.001 | 2.49 | 1356.835 | -0.314 | 2.785 | 108.186 | 0.872 | 0.109 | 0 | 0 | 1 | 43.875 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-16561 | 2021/05/29 06:30:28 | 1-27 | 1 | -7.629 | -72.894 | 17864 | -1 | 1 | WOOD2-2030-3 | 20-212 | H2O | 1622291735 | 0.328 | 0.745 | 105.773 | 0.001 | 0.007 | 2.507 | 1298.063 | 1.77 | 2.828 | 109.416 | 0.869 | 0.211 | 0 | 0 | 1 | 43.812 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-16561 | 2021/05/29 06:39:25 | 1-27 | 2 | -8.247 | -75.65 | 17685 | -1 | 1 | WOOD2-2030-3 | 20-212 | H2O | 1622292272 | 0.302 | 0.739 | 104.715 | 0.001 | 0.004 | 2.496 | 1295.921 | 2.526 | 2.798 | 107.686 | 0.864 | 0.332 | 0 | 0 | 1 | 43.625 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-16561 | 2021/05/29 06:48:21 | 1-27 | 3 | -8.395 | -76.266 | 17675 | -1 | 1 | WOOD2-2030-3 | 20-212 | H2O | 1622292807 | 0.306 | 0.781 | 107.603 | 0 | 0.003 | 2.559 | 1295.544 | 2.829 | 2.794 | 107.436 | 0.86 | 0.241 | 0 | 0 | 1 | 43.688 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-16561 | 2021/05/29 06:57:18 | 1-27 | 4 | -8.494 | -76.773 | 17370 | 0 | 1 | WOOD2-2030-3 | 20-212 | H2O | 1622293344 | 0.298 | 0.749 | 106.64 | 0 | 0.001 | 2.56 | 1293.581 | 2.229 | 2.757 | 105.656 | 0.863 | 0.175 | 0 | 0 | 1 | 43.688 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-16562 | 2021/05/29 07:07:06 | 1-28 | 1 | -5.804 | -65.628 | 17483 | -1 | 1 | WOOD2-010-3 | 20-212 | H2O | 1622293933 | 0.3 | 0.709 | 106.296 | 0 | -0.001 | 2.534 | 1293.822 | 2.188 | 2.768 | 107.613 | 0.864 | 0.298 | 0 | 0 | 1 | 43.688 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-16562 | 2021/05/29 07:16:02 | 1-28 | 2 | -5.51 | -63.663 | 17952 | -1 | 1 | WOOD2-010-3 | 20-212 | H2O | 1622294469 | 0.312 | 0.719 | 109.058 | 0.001 | 0.002 | 2.583 | 1297.565 | 2.246 | 2.843 | 109.55 | 0.865 | 0.3 | 0 | 0 | 1 | 43.688 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-16562 | 2021/05/29 07:24:59 | 1-28 | 3 | -5.497 | -63.524 | 17751 | -1 | 1 | WOOD2-010-3 | 20-212 | H2O | 1622295005 | 0.322 | 0.728 | 108.449 | -0.001 | 0.002 | 2.582 | 1295.932 | 2.205 | 2.831 | 109.184 | 0.865 | 0.497 | 0 | 0 | 1 | 43.75 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-16562 | 2021/05/29 07:33:55 | 1-28 | 4 | -5.499 | -63.383 | 17714 | 0 | 1 | WOOD2-010-3 | 20-212 | H2O | 1622295541 | 0.328 | 0.742 | 105.515 | 0 | 0.002 | 2.515 | 1295.436 | 2.165 | 2.85 | 109.597 | 0.869 | 0.661 | 0 | 0 | 1 | 43.625 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-16563 | 2021/05/29 07:43:39 | 1-29 | 1 | -0.921 | -48.92 | 17783 | -1 | 1 | WOOD2-E-5 | 20-212 | H2O | 1622296125 | 0.312 | 0.729 | 106.377 | -0.001 | 0.002 | 2.526 | 1297.674 | 1.937 | 2.849 | 110.045 | 0.867 | 0.352 | 0 | 0 | 1 | 43.625 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-16563 | 2021/05/29 07:52:35 | 1-29 | 2 | -0.607 | -46.848 | 17752 | -1 | 1 | WOOD2-E-5 | 20-212 | H2O | 1622296661 | 0.354 | 0.716 | 105.325 | -0.002 | -0.002 | 2.512 | 1297.177 | 2.358 | 2.821 | 108.876 | 0.872 | 0.416 | 0 | 0 | 1 | 43.625 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-16563 | 2021/05/29 08:01:32 | 1-29 | 3 | -0.487 | -46.333 | 17717 | -1 | 1 | WOOD2-E-5 | 20-212 | H2O | 1622297199 | 0.333 | 0.712 | 108.236 | -0.001 | 0.003 | 2.568 | 1297.177 | 1.803 | 2.806 | 109.243 | 0.868 | 0.283 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-16563 | 2021/05/29 08:10:28 | 1-29 | 4 | -0.42 | -46.094 | 17776 | 0 | 1 | WOOD2-E-5 | 20-212 | H2O | 1622297735 | 0.315 | 0.804 | 109.025 | 0 | 0.002 | 2.593 | 1297.775 | 1.608 | 2.819 | 109.694 | 0.866 | 0.216 | 0 | 0 | 1 | 43.75 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-16564 | 2021/05/29 08:20:13 | 1-30 | 1 | -6.387 | -64.698 | 17860 | -1 | 1 | WOOD2-1020-2 | 20-212 | H2O | 1622298320 | 0.346 | 0.733 | 105.609 | 0 | 0.004 | 2.497 | 1296.627 | 2.596 | 2.824 | 108.99 | 0.866 | 0.354 | 0 | 0 | 1 | 43.75 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-16564 | 2021/05/29 08:29:08 | 1-30 | 2 | -6.98 | -67.501 | 17771 | -1 | 1 | WOOD2-1020-2 | 20-212 | H2O | 1622298855 | 0.302 | 0.749 | 108.577 | 0 | 0.002 | 2.565 | 1295.779 | 2.82 | 2.815 | 108.446 | 0.867 | 0.393 | 0 | 0 | 1 | 43.812 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-16564 | 2021/05/29 08:38:04 | 1-30 | 3 | -7.05 | -68.107 | 17725 | -1 | 1 | WOOD2-1020-2 | 20-212 | H2O | 1622299391 | 0.325 | 0.74 | 106.871 | -0.001 | -0.001 | 2.535 | 1295.829 | 2.483 | 2.776 | 107.23 | 0.868 | 0.139 | 0 | 0 | 1 | 43.812 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-16564 | 2021/05/29 08:47:00 | 1-30 | 4 | -7.102 | -68.438 | 17564 | 0 | 1 | WOOD2-1020-2 | 20-212 | H2O | 1622299926 | 0.322 | 0.74 | 106.422 | -0.001 | 0.003 | 2.534 | 1294.524 | 2.34 | 2.81 | 107.693 | 0.87 | 0.485 | 0 | 0 | 1 | 43.75 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-16565 | 2021/05/29 08:56:01 | 1-31 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-16565 | 2021/05/29 09:05:59 | 1-31 | 2 | -7.386 | -61.513 | 17457 | -1 | 1 | EV | standard | H2O | 1622301065 | 0.319 | 0.724 | 97.959 | -0.001 | 0 | 2.338 | 1293.754 | 2.688 | 2.744 | 105.151 | 0.862 | 0.275 | 0 | 0 | 1 | 43.812 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-16565 | 2021/05/29 09:14:56 | 1-31 | 3 | -7.382 | -60.52 | 17626 | -1 | 1 | EV | standard | H2O | 1622301603 | 0.312 | 0.781 | 104.267 | 0 | 0.004 | 2.471 | 1294.751 | 2.486 | 2.81 | 107.788 | 0.864 | 0.388 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-16565 | 2021/05/29 09:23:52 | 1-31 | 4 | -7.385 | -60.162 | 17648 | 0 | 1 | EV | standard | H2O | 1622302138 | 0.313 | 0.781 | 102.741 | 0 | 0.001 | 2.441 | 1295.07 | 2.332 | 2.759 | 107.612 | 0.862 | 0.185 | 0 | 0 | 1 | 43.75 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-16565 | 2021/05/29 09:32:48 | 1-31 | 5 | -7.385 | -59.981 | 17679 | 0 | 1 | EV | standard | H2O | 1622302675 | 0.325 | 0.679 | 105.292 | 0 | 0.002 | 2.508 | 1295.025 | 2.572 | 2.803 | 107.854 | 0.864 | 0.522 | 0 | 0 | 1 | 43.75 | 1 | 31 | 1 | GJB_Rinse | 0
150 | P-16566 | 2021/05/29 09:42:34 | 1-32 | 1 | -3.795 | -52.639 | 17417 | -1 | 1 | WOOD2-E-1 | 20-212 | H2O | 1622303260 | 0.443 | 1.157 | 105.181 | -0.003 | 0.002 | 2.5 | 2080.669 | -26.872 | 2.78 | 101.836 | 0.871 | -0.49 | 0 | 0 | 1 | 43.75 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-16566 | 2021/05/29 09:51:29 | 1-32 | 2 | -3.493 | -51.615 | 17357 | -1 | 1 | WOOD2-E-1 | 20-212 | H2O | 1622303796 | 0.42 | 1.123 | 104.699 | 0 | 0.004 | 2.481 | 2084.666 | -27.352 | 2.753 | 101.962 | 0.873 | -0.658 | 0 | 0 | 1 | 43.75 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-16566 | 2021/05/29 10:00:26 | 1-32 | 3 | -3.494 | -51.425 | 17273 | -1 | 1 | WOOD2-E-1 | 20-212 | H2O | 1622304332 | 0.434 | 1.163 | 100.486 | 0 | 0.005 | 2.408 | 2081.129 | -27.194 | 2.717 | 101.57 | 0.869 | -0.482 | 0 | 0 | 1 | 43.688 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-16566 | 2021/05/29 10:09:22 | 1-32 | 4 | -3.454 | -51.287 | 17201 | 0 | 1 | WOOD2-E-1 | 20-212 | H2O | 1622304868 | 0.418 | 1.028 | 104.984 | -0.001 | -0.001 | 2.5 | 2079.196 | -26.851 | 2.708 | 99.083 | 0.87 | -0.805 | 0 | 0 | 1 | 43.688 | 1 | 32 | 1 | GJB_Rinse | 0
154 | P-16567 | 2021/05/29 10:19:08 | 1-33 | 1 | -0.004 | -50.25 | 17724 | -1 | 1 | WOOD2-A-2 | 20-212 | H2O | 1622305455 | 0.297 | 0.76 | 106.73 | 0 | 0.002 | 2.535 | 1322.598 | 1.033 | 2.845 | 109.089 | 0.866 | 0.306 | 0 | 0 | 1 | 43.75 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-16567 | 2021/05/29 10:28:04 | 1-33 | 2 | 0.333 | -50.162 | 17824 | -1 | 1 | WOOD2-A-2 | 20-212 | H2O | 1622305991 | 0.295 | 0.72 | 105.632 | 0 | 0.002 | 2.511 | 1317.238 | 1.034 | 2.854 | 110.366 | 0.861 | 0.378 | 0 | 0 | 1 | 43.75 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-16567 | 2021/05/29 10:37:00 | 1-33 | 3 | 0.348 | -50.021 | 17813 | -1 | 1 | WOOD2-A-2 | 20-212 | H2O | 1622306526 | 0.313 | 0.764 | 104.248 | 0 | 0.002 | 2.489 | 1315.406 | 0.94 | 2.855 | 110.152 | 0.863 | 0.293 | 0 | 0 | 1 | 43.75 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-16567 | 2021/05/29 10:45:56 | 1-33 | 4 | 0.451 | -49.974 | 17696 | 0 | 1 | WOOD2-A-2 | 20-212 | H2O | 1622307062 | 0.319 | 0.726 | 108.034 | 0 | 0 | 2.577 | 1313.399 | 1.814 | 2.828 | 108.632 | 0.863 | 0.338 | 0 | 0 | 1 | 43.75 | 1 | 33 | 1 | GJB_Rinse | 0
158 | P-16568 | 2021/05/29 10:55:42 | 1-34 | 1 | -8.981 | -81.241 | 17213 | -1 | 1 | WOOD2-5060-3 | 20-212 | H2O | 1622307648 | 0.3 | 0.783 | 102.049 | 0 | 0.008 | 2.443 | 1293.078 | 2.526 | 2.64 | 104.073 | 0.865 | 0.097 | 0 | 0 | 1 | 43.75 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-16568 | 2021/05/29 11:04:38 | 1-34 | 2 | -9.802 | -85.471 | 17396 | -1 | 1 | WOOD2-5060-3 | 20-212 | H2O | 1622308185 | 0.308 | 0.793 | 102.154 | 0 | 0.007 | 2.441 | 1294.244 | 3.282 | 2.733 | 104.389 | 0.862 | 0.228 | 0 | 0 | 1 | 43.688 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-16568 | 2021/05/29 11:13:33 | 1-34 | 3 | -9.979 | -86.633 | 17470 | -1 | 1 | WOOD2-5060-3 | 20-212 | H2O | 1622308720 | 0.295 | 0.759 | 103.946 | -0.001 | 0.005 | 2.47 | 1294.749 | 2.552 | 2.783 | 105.983 | 0.862 | 0.184 | 0 | 0 | 1 | 43.75 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-16568 | 2021/05/29 11:22:29 | 1-34 | 4 | -10.043 | -87.163 | 17386 | 0 | 1 | WOOD2-5060-3 | 20-212 | H2O | 1622309255 | 0.332 | 0.714 | 105.365 | 0.001 | 0.004 | 2.496 | 1293.636 | 2.282 | 2.771 | 108.04 | 0.864 | 0.654 | 0 | 0 | 1 | 43.812 | 1 | 34 | 1 | GJB_Rinse | 0
162 | P-16569 | 2021/05/29 11:32:17 | 1-35 | 1 | -8.022 | -74.639 | 16840 | -1 | 1 | WOOD2-2030-2 | 20-212 | H2O | 1622309843 | 0.314 | 0.779 | 101.242 | 0 | 0.001 | 2.434 | 1289.9 | 1.912 | 2.637 | 104.119 | 0.861 | 0.395 | 0 | 0 | 1 | 43.75 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-16569 | 2021/05/29 11:41:13 | 1-35 | 2 | -7.819 | -72.834 | 16684 | -1 | 1 | WOOD2-2030-2 | 20-212 | H2O | 1622310380 | 0.316 | 0.748 | 101.763 | -0.001 | 0.001 | 2.436 | 1288.733 | 2.582 | 2.641 | 102.118 | 0.86 | 0.496 | 0 | 0 | 1 | 43.812 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-16569 | 2021/05/29 11:50:09 | 1-35 | 3 | -7.807 | -72.178 | 17311 | -1 | 1 | WOOD2-2030-2 | 20-212 | H2O | 1622310915 | 0.329 | 0.755 | 101.408 | 0 | 0.008 | 2.425 | 1293.392 | 2.123 | 2.757 | 105.934 | 0.86 | 0.356 | 0 | 0 | 1 | 43.75 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-16569 | 2021/05/29 11:59:05 | 1-35 | 4 | -7.776 | -72.057 | 17120 | 0 | 1 | WOOD2-2030-2 | 20-212 | H2O | 1622311451 | 0.325 | 0.739 | 103.969 | 0 | 0.003 | 2.485 | 1291.966 | 1.898 | 2.72 | 105.537 | 0.861 | 0.358 | 0 | 0 | 1 | 43.812 | 1 | 35 | 1 | GJB_Rinse | 0
166 | P-16570 | 2021/05/29 12:08:48 | 1-36 | 1 | -3.989 | -58.139 | 17792 | -1 | 1 | WOOD2-C-1 | 20-212 | H2O | 1622312035 | 0.303 | 0.753 | 105.114 | -0.001 | 0 | 2.51 | 1300.751 | 1.884 | 2.827 | 108.888 | 0.863 | 0.13 | 0 | 0 | 1 | 43.812 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-16570 | 2021/05/29 12:17:45 | 1-36 | 2 | -3.773 | -56.858 | 17268 | -1 | 1 | WOOD2-C-1 | 20-212 | H2O | 1622312571 | 0.314 | 0.761 | 103.023 | -0.001 | 0.002 | 2.463 | 1296.713 | 2.112 | 2.745 | 105.709 | 0.861 | 0.184 | 0 | 0 | 1 | 43.812 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-16570 | 2021/05/29 12:26:40 | 1-36 | 3 | -3.66 | -56.191 | 17264 | -1 | 1 | WOOD2-C-1 | 20-212 | H2O | 1622313107 | 0.32 | 0.722 | 102.761 | -0.001 | -0.001 | 2.469 | 1296.841 | 2.226 | 2.692 | 104.654 | 0.862 | -0.019 | 0 | 0 | 1 | 43.812 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-16570 | 2021/05/29 12:35:36 | 1-36 | 4 | -3.599 | -55.862 | 17294 | 0 | 1 | WOOD2-C-1 | 20-212 | H2O | 1622313643 | 0.27 | 0.714 | 105.303 | 0 | 0.002 | 2.509 | 1297.097 | 1.666 | 2.753 | 106.514 | 0.864 | 0.159 | 0 | 0 | 1 | 43.812 | 1 | 36 | 1 | GJB_Rinse | 0
170 | P-16571 | 2021/05/29 12:45:20 | 1-37 | 1 | -8.119 | -74.979 | 16881 | -1 | 1 | WOOD2-3040-2 | 20-212 | H2O | 1622314227 | 0.302 | 0.71 | 101.242 | 0 | 0.003 | 2.436 | 1289.986 | 1.987 | 2.669 | 103.686 | 0.861 | 0.253 | 0 | 0 | 1 | 43.812 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-16571 | 2021/05/29 12:54:16 | 1-37 | 2 | -8.464 | -77.338 | 17231 | -1 | 1 | WOOD2-3040-2 | 20-212 | H2O | 1622314762 | 0.311 | 0.749 | 105.164 | 0 | 0.003 | 2.518 | 1292.112 | 2.528 | 2.746 | 106.433 | 0.86 | 0.582 | 0 | 0 | 1 | 43.75 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-16571 | 2021/05/29 13:03:11 | 1-37 | 3 | -8.538 | -78.157 | 17107 | -1 | 1 | WOOD2-3040-2 | 20-212 | H2O | 1622315298 | 0.288 | 0.707 | 102.203 | -0.001 | 0 | 2.446 | 1291.37 | 2.327 | 2.719 | 105.087 | 0.861 | 0.349 | 0 | 0 | 1 | 43.875 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-16571 | 2021/05/29 13:12:07 | 1-37 | 4 | -8.537 | -78.176 | 17309 | 0 | 1 | WOOD2-3040-2 | 20-212 | H2O | 1622315834 | 0.328 | 0.788 | 104.326 | 0 | 0.005 | 2.49 | 1292.734 | 2.488 | 2.737 | 105.899 | 0.863 | 0.275 | 0 | 0 | 1 | 43.812 | 1 | 37 | 1 | GJB_Rinse | 0
174 | P-16572 | 2021/05/29 13:21:51 | 1-38 | 1 | [5 empty] | WOOD2-2030-1 | 20-212 | [18 empty] | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-16572 | 2021/05/29 13:30:47 | 1-38 | 2 | [5 empty] | WOOD2-2030-1 | 20-212 | [18 empty] | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-16572 | 2021/05/29 13:39:42 | 1-38 | 3 | -6.812 | -69.344 | 17696 | -1 | 1 | WOOD2-2030-1 | 20-212 | H2O | 1622316954 | 0.317 | 0.791 | 104.61 | 0 | 0.003 | 2.492 | 1295.471 | 2.276 | 2.791 | 108.227 | 0.859 | 0.179 | 0 | 0 | 1 | 43.875 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-16572 | 2021/05/29 13:48:38 | 1-38 | 4 | -6.803 | -69.634 | 16922 | -1 | 1 | WOOD2-2030-1 | 20-212 | H2O | 1622317488 | 0.325 | 0.763 | 101.848 | -0.001 | 0 | 2.458 | 1289.782 | 2.517 | 2.665 | 103.515 | 0.864 | 0.328 | 0 | 0 | 1 | 43.875 | 1 | 38 | 1 | GJB_Rinse | 0
178 | P-16573 | 2021/05/29 13:58:22 | 1-39 | 1 | -5.489 | -67.125 | 9731 | -1 | 1 | WOOD2-010-1 | 20-212 | H2O | 1622318610 | 0.298 | 0.944 | 76.921 | -0.001 | 0.001 | 1.969 | 1240.884 | 1.844 | 1.491 | 59.759 | 0.855 | 0.01 | 0 | 0 | 1 | 43.875 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-16573 | 2021/05/29 14:07:17 | 1-39 | 2 | -5.605 | -67.997 | 6910 | -1 | 1 | WOOD2-010-1 | 20-212 | H2O | 1622319156 | 0.362 | 1.274 | 57.686 | -0.002 | -0.004 | 1.796 | 1222.773 | 1.234 | 1.054 | 43.853 | 0.855 | 0.191 | 0 | 0 | 1 | 43.875 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-16573 | 2021/05/29 14:16:12 | 1-39 | 3 | -5.608 | -67.464 | 7545 | -1 | 1 | WOOD2-010-1 | 20-212 | H2O | 1622319685 | 0.359 | 1.252 | 64.959 | -0.002 | -0.004 | 1.801 | 1226.81 | 1.395 | 1.146 | 47.529 | 0.855 | 0.056 | 0 | 0 | 1 | 43.812 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-16573 | 2021/05/29 14:25:08 | 1-39 | 4 | [5 empty] | WOOD2-010-1 | 20-212 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
182 | P-16574 | 2021/05/29 14:35:54 | 1-40 | 1 | -7.39 | -65.045 | 8925 | -1 | 1 | EV | standard | H2O | 1622320863 | 0.318 | 1.055 | 68.421 | -0.001 | 0 | 1.763 | 1235.645 | 1.689 | 1.344 | 54.795 | 0.855 | 0.081 | 0 | 0 | 1 | 43.938 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-16574 | 2021/05/29 14:44:50 | 1-40 | 2 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-16574 | 2021/05/29 14:54:48 | 1-40 | 3 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-16574 | 2021/05/29 15:04:46 | 1-40 | 4 | -7.39 | -65.045 | 8925 | -1 | 1 | EV | standard | H2O | 1622320863 | 0.318 | 1.055 | 68.421 | -0.001 | 0 | 1.763 | 1235.645 | 1.689 | 1.344 | 54.795 | 0.855 | 0.081 | 0 | 0 | 1 | 43.938 | 1 | 40 | 1 | GJB_Rinse | 0
186 | P-16575 | 2021/05/29 15:15:33 | 1-41 | 1 | [5 empty] | WOOD2-6070-2 | 20-212 | [18 empty] | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-16575 | 2021/05/29 15:24:29 | 1-41 | 2 | [5 empty] | WOOD2-6070-2 | 20-212 | H2O | [17 empty] | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-16575 | 2021/05/29 15:34:27 | 1-41 | 3 | -9.475 | -82.299 | 8634 | -1 | 1 | WOOD2-6070-2 | 20-212 | H2O | 1622323242 | 0.322 | 1.096 | 64.486 | 0 | 0.004 | 1.681 | 1233.731 | 1.969 | 1.284 | 52.415 | 0.856 | -0.057 | 0 | 0 | 1 | 43.938 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-16575 | 2021/05/29 15:44:25 | 1-41 | 4 | -9.475 | -82.299 | 8634 | -1 | 1 | WOOD2-6070-2 | 20-212 | H2O | 1622323242 | 0.322 | 1.096 | 64.486 | 0 | 0.004 | 1.681 | 1233.731 | 1.969 | 1.284 | 52.415 | 0.856 | -0.057 | 0 | 0 | 1 | 43.938 | 1 | 41 | 1 | GJB_Rinse | 0
190 | P-16576 | 2021/05/29 15:55:13 | 1-42 | 1 | [5 empty] | WOOD2-A-4 | 20-212 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-16576 | 2021/05/29 16:05:12 | 1-42 | 2 | [5 empty] | WOOD2-A-4 | 20-212 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-16576 | 2021/05/29 16:15:11 | 1-42 | 3 | [5 empty] | WOOD2-A-4 | 20-212 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-16576 | 2021/05/29 16:25:10 | 1-42 | 4 | [5 empty] | WOOD2-A-4 | 20-212 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
194 | P-16577 | 2021/05/29 16:35:56 | 1-43 | 1 | [5 empty] | WOOD2-C-4 | 20-212 | H2O | [17 empty] | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-16577 | 2021/05/29 16:45:53 | 1-43 | 2 | [5 empty] | WOOD2-C-4 | 20-212 | H2O | [17 empty] | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-16577 | 2021/05/29 16:55:52 | 1-43 | 3 | [5 empty] | WOOD2-C-4 | 20-212 | H2O | [17 empty] | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-16577 | 2021/05/29 17:05:50 | 1-43 | 4 | [5 empty] | WOOD2-C-4 | 20-212 | H2O | [17 empty] | 1 | 43 | 1 | GJB_Rinse | 0
198 | P-16578 | 2021/05/29 17:16:38 | 1-44 | 1 | [5 empty] | WOOD2-3040-3 | 20-212 | H2O | [17 empty] | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-16578 | 2021/05/29 17:26:36 | 1-44 | 2 | [5 empty] | WOOD2-3040-3 | 20-212 | H2O | [17 empty] | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-16578 | 2021/05/29 17:36:34 | 1-44 | 3 | [5 empty] | WOOD2-3040-3 | 20-212 | H2O | [17 empty] | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-16578 | 2021/05/29 17:46:32 | 1-44 | 4 | [5 empty] | WOOD2-3040-3 | 20-212 | H2O | [17 empty] | 1 | 44 | 1 | GJB_Rinse | 0
202 | P-16579 | 2021/05/29 17:57:20 | 1-45 | 1 | [5 empty] | WOOD2-B-1 | 20-212 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-16579 | 2021/05/29 18:07:18 | 1-45 | 2 | [5 empty] | WOOD2-B-1 | 20-212 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-16579 | 2021/05/29 18:17:15 | 1-45 | 3 | [5 empty] | WOOD2-B-1 | 20-212 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-16579 | 2021/05/29 18:27:14 | 1-45 | 4 | [5 empty] | WOOD2-B-1 | 20-212 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
206 | P-16580 | 2021/05/29 18:38:03 | 1-46 | 1 | [5 empty] | WOOD2-D-4 | 20-212 | H2O | [17 empty] | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-16580 | 2021/05/29 18:48:01 | 1-46 | 2 | [5 empty] | WOOD2-D-4 | 20-212 | H2O | [17 empty] | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-16580 | 2021/05/29 18:57:59 | 1-46 | 3 | [5 empty] | WOOD2-D-4 | 20-212 | H2O | [17 empty] | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-16580 | 2021/05/29 19:07:57 | 1-46 | 4 | [5 empty] | WOOD2-D-4 | 20-212 | H2O | [17 empty] | 1 | 46 | 1 | GJB_Rinse | 0
210 | P-16581 | 2021/05/29 19:18:46 | 1-47 | 1 | 15.984 | -40.239 | 11363 | -1 | 1 | WOOD2-B-2 | 20-212 | H2O | 1622337833 | 0.562 | 2.459 | 68.704 | -0.004 | 0.001 | 1.709 | 2515.371 | -73.634 | 4.935 | 85.489 | 0.862 | -11.314 | 0 | 0 | 1 | 43.812 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-16581 | 2021/05/29 19:27:41 | 1-47 | 2 | [5 empty] | WOOD2-B-2 | 20-212 | H2O | [17 empty] | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-16581 | 2021/05/29 19:37:40 | 1-47 | 3 | [5 empty] | WOOD2-B-2 | 20-212 | H2O | [17 empty] | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-16581 | 2021/05/29 19:47:37 | 1-47 | 4 | [5 empty] | WOOD2-B-2 | 20-212 | [18 empty] | 1 | 47 | 1 | GJB_Rinse | 0
214 | P-16582 | 2021/05/29 19:58:26 | 1-48 | 1 | -3.314 | -62.293 | 13809 | -1 | 1 | WOOD2-D-3 | 20-212 | H2O | 1622340212 | 0.339 | 0.873 | 78.298 | 0 | -0.001 | 1.902 | 1296.325 | -0.655 | 2.093 | 85.092 | 0.858 | -0.628 | 0 | 0 | 1 | 43.875 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-16582 | 2021/05/29 20:07:21 | 1-48 | 2 | [5 empty] | WOOD2-D-3 | 20-212 | H2O | [17 empty] | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-16582 | 2021/05/29 20:17:20 | 1-48 | 3 | [5 empty] | WOOD2-D-3 | 20-212 | H2O | [17 empty] | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-16582 | 2021/05/29 20:27:17 | 1-48 | 4 | [5 empty] | WOOD2-D-3 | 20-212 | H2O | [17 empty] | 1 | 48 | 1 | GJB_Rinse | 0
218 | P-16583 | 2021/05/29 20:38:05 | 1-49 | 1 | -7.154 | -62.178 | 13615 | -1 | 1 | EV | standard | H2O | 1622342591 | 0.288 | 0.81 | 77.938 | 0 | -0.001 | 1.904 | 1268.564 | 1.894 | 2.132 | 84.175 | 0.858 | 0.175 | 0 | 0 | 1 | 43.938 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-16583 | 2021/05/29 20:47:00 | 1-49 | 2 | -7.783 | -65.113 | 10622 | -1 | 1 | EV | standard | H2O | 1622343126 | 0.357 | 1.018 | 74.241 | 0.003 | 0.009 | 1.901 | 1248.116 | 1.446 | 1.643 | 66.283 | 0.855 | 0.107 | 0 | 0 | 1 | 43.875 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-16583 | 2021/05/29 20:55:56 | 1-49 | 3 | -7.768 | -65.298 | 8532 | -1 | 1 | EV | standard | H2O | 1622343665 | 0.346 | 1.307 | 66.835 | 0 | 0 | 1.757 | 1234.185 | 1.603 | 1.281 | 52.349 | 0.855 | -0.033 | 0 | 0 | 1 | 43.812 | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-16583 | 2021/05/29 21:04:52 | 1-49 | 4 | -8.016 | -66.568 | 7871 | 0 | 1 | EV | standard | H2O | 1622344204 | 0.34 | 1.222 | 62.296 | 0.001 | 0.005 | 1.692 | 1229.853 | 1.338 | 1.2 | 49.173 | 0.855 | 0.092 | 0 | 0 | 1 | 43.812 | 1 | 49 | 1 | GJB_Rinse | 0
222 | P-16584 | 2021/05/29 21:14:37 | 1-50 | 1 | 11.596 | -61.21 | 8633 | -1 | 1 | WOOD2-B-4 | 20-212 | H2O | 1622344786 | 0.588 | 2.473 | 64.471 | -0.004 | -0.007 | 1.684 | 2363.818 | -59.163 | 3.374 | 60.87 | 0.859 | -7.801 | 0 | 0 | 1 | 43.875 | 1 | 50 | 1 | GJB_Rinse | 0
223 | P-16584 | 2021/05/29 21:23:32 | 1-50 | 2 | [5 empty] | WOOD2-B-4 | 20-212 | H2O | [17 empty] | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-16584 | 2021/05/29 21:33:30 | 1-50 | 3 | [5 empty] | WOOD2-B-4 | 20-212 | H2O | [17 empty] | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-16584 | 2021/05/29 21:43:28 | 1-50 | 4 | 11.596 | -61.21 | 8633 | -1 | 1 | WOOD2-B-4 | 20-212 | H2O | 1622344786 | 0.588 | 2.473 | 64.471 | -0.004 | -0.007 | 1.684 | 2363.818 | -59.163 | 3.374 | 60.87 | 0.859 | -7.801 | 0 | 0 | 1 | 43.875 | 1 | 50 | 1 | GJB_Rinse | 0
226 | P-16585 | 2021/05/29 21:54:16 | 1-51 | 1 | 3.085 | -91.393 | 8976 | -1 | 1 | WOOD2-B-5 | 20-212 | H2O | 1622347164 | 0.463 | 1.806 | 62.157 | -0.002 | 0.007 | 1.615 | 1969.628 | -39.65 | 2.64 | 62.748 | 0.858 | -5.653 | 0 | 0 | 1 | 44 | 1 | 51 | 1 | GJB_Rinse | 0
227 | P-16585 | 2021/05/29 22:03:11 | 1-51 | 2 | [5 empty] | WOOD2-B-5 | 20-212 | H2O | [17 empty] | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-16585 | 2021/05/29 22:13:08 | 1-51 | 3 | [5 empty] | WOOD2-B-5 | 20-212 | H2O | [17 empty] | 1 | 51 | 1 | GJB_Rinse | 0
229 | P-16585 | 2021/05/29 22:23:03 | 1-51 | 4 | 3.085 | -91.393 | 8976 | -1 | 1 | WOOD2-B-5 | 20-212 | H2O | 1622347164 | 0.463 | 1.806 | 62.157 | -0.002 | 0.007 | 1.615 | 1969.628 | -39.65 | 2.64 | 62.748 | 0.858 | -5.653 | 0 | 0 | 1 | 44 | 1 | 51 | 1 | GJB_Rinse | 0
230 | P-16586 | 2021/05/29 22:33:51 | 1-52 | 1 | -6.728 | -84.7 | 11051 | -1 | 1 | WOOD2-D-2 | 20-212 | H2O | 1622349539 | 0.304 | 0.876 | 71.045 | -0.002 | 0.001 | 1.773 | 1268.271 | 0.357 | 1.67 | 67.954 | 0.856 | -0.324 | 0 | 0 | 1 | 43.875 | 1 | 52 | 1 | GJB_Rinse | 0
231 | P-16586 | 2021/05/29 22:42:47 | 1-52 | 2 | [5 empty] | WOOD2-D-2 | 20-212 | H2O | [17 empty] | 1 | 52 | 1 | GJB_Rinse | 0
232 | P-16586 | 2021/05/29 22:52:45 | 1-52 | 3 | [5 empty] | WOOD2-D-2 | 20-212 | H2O | [17 empty] | 1 | 52 | 1 | GJB_Rinse | 0
233 | P-16586 | 2021/05/29 23:02:43 | 1-52 | 4 | [5 empty] | WOOD2-D-2 | 20-212 | [18 empty] | 1 | 52 | 1 | GJB_Rinse | 0
234 | P-16587 | 2021/05/29 23:13:32 | 1-53 | 1 | -7.605 | -67.755 | 10886 | -1 | 1 | EV | standard | H2O | 1622351919 | 0.276 | 0.932 | 70.874 | 0 | 0 | 1.769 | 1250.344 | 1.811 | 1.696 | 67.118 | 0.856 | 0.08 | 0 | 0 | 1 | 43.875 | 1 | 53 | 1 | GJB_Rinse | 0
235 | P-16587 | 2021/05/29 23:22:28 | 1-53 | 2 | -5.935 | -63.368 | 6898 | -1 | 1 | EV | standard | H2O | 1622352467 | 0.324 | 1.372 | 50.563 | -0.002 | -0.007 | 1.587 | 1224.515 | 1.336 | 1.029 | 42.974 | 0.855 | 0.021 | 0 | 0 | 1 | 43.875 | 1 | 53 | 1 | GJB_Rinse | 0
236 | P-16587 | 2021/05/29 23:31:23 | 1-53 | 3 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 53 | 1 | GJB_Rinse | 0
237 | P-16587 | 2021/05/29 23:41:21 | 1-53 | 4 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 53 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1302.17535
varray: [1306.447, 1305.021, 1304.098, 1304.758, 1304.04, 1305.636, 1303.065, 1303.233, 1304.46, 1306.753, 1301.616, 1298.393, 1302.623, 1299.355, 1298.891, 1298.056, 1299.88, 1296.795, 1296.449, 1303.938]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_shift
baseline_shift_STD | 3.25166864491
varray: [1306.447, 1305.021, 1304.098, 1304.758, 1304.04, 1305.636, 1303.065, 1303.233, 1304.46, 1306.753, 1301.616, 1298.393, 1302.623, 1299.355, 1298.891, 1298.056, 1299.88, 1296.795, 1296.449, 1303.938]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.3726
varray: [0.145, 0.408, 0.169, 0.416, 0.42, 0.215, 0.457, 0.259, 0.474, 0.602, 0.261, 0.215, 0.602, 0.334, 0.376, 0.528, 0.334, 0.345, 0.487, 0.405]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.133449537082
varray: [0.145, 0.408, 0.169, 0.416, 0.42, 0.215, 0.457, 0.259, 0.474, 0.602, 0.261, 0.215, 0.602, 0.334, 0.376, 0.528, 0.334, 0.345, 0.487, 0.405]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: ch4_ppm
residuals_MEAN | 2.917
varray: [3.027, 2.919, 2.866, 2.969, 2.953, 2.957, 2.936, 2.93, 3.02, 3.095, 2.912, 2.789, 2.932, 2.813, 2.862, 2.892, 2.847, 2.842, 2.867, 2.912]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: residuals
residuals_STD | 0.0751412704598
varray: [3.027, 2.919, 2.866, 2.969, 2.953, 2.957, 2.936, 2.93, 3.02, 3.095, 2.912, 2.789, 2.932, 2.813, 2.862, 2.892, 2.847, 2.842, 2.867, 2.912]
inst: std_standards
operation: std_standards(residuals)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: residuals
slope_shift_MEAN | 2.1159
varray: [2.011, 1.905, 1.404, 2.084, 1.657, 1.11, 1.948, 1.517, 1.58, 1.927, 2.649, 2.668, 2.388, 2.147, 2.587, 2.745, 2.408, 2.598, 2.661, 2.324]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: slope_shift
slope_shift_STD | 0.483156823184
varray: [2.011, 1.905, 1.404, 2.084, 1.657, 1.11, 1.948, 1.517, 1.58, 1.927, 2.649, 2.668, 2.388, 2.147, 2.587, 2.745, 2.408, 2.598, 2.661, 2.324]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: slope_shift
baseline_curvature_MEAN | 111.55715
varray: [114.787, 111.167, 111.211, 111.515, 111.855, 112.501, 111.647, 112.179, 114.278, 116.014, 111.718, 107.704, 113.22, 110.054, 109.384, 110.004, 110.755, 108.335, 109.106, 113.709]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_curvature
baseline_curvature_STD | 2.1526352583
varray: [114.787, 111.167, 111.211, 111.515, 111.855, 112.501, 111.647, 112.179, 114.278, 116.014, 111.718, 107.704, 113.22, 110.054, 109.384, 110.004, 110.755, 108.335, 109.106, 113.709]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.917
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1302.17535
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.3726
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 2.1159
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.11271190569
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 3.02971190569
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.15028254092
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.06728254092
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.225423811379
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.14242381138
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.2917
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.400348611246
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.772948611246
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 1.62583432246
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1303.80118432
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1305.42701864
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 9.75500593474
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1311.93035593
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 58.5300356084
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1360.70538561
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.41578411592
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.53168411592
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 6.45790577489
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 118.015055775
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -6.45858199895
actual_array: [ 1.93 -15.84]
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493, 3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964, -11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009, -7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454, -8.56, -8.744, -8.678, -8.701, -9.514, -9.664, -9.587, -9.704, -6.61, -6.511, -6.507, -6.505, -4.029, -3.729, -3.714, -3.694, 0.0, 0.0, 13.423, 13.426, 0.0, 0.0, -9.992, -9.992, -4.577, -4.1, -4.033, -3.97, -3.682, -3.714, -3.686, -3.72, -7.043, -7.317, -7.356, -7.453, -1.304, -0.806, -0.679, -0.602, -1.62, -1.697, -1.715, -1.721, -5.176, -5.549, -5.593, -5.639, -5.019, -4.929, -4.979, -4.956, 0.0, 0.0, 4.24, 4.239, 0.0, 0.0, -7.039, -7.011, -6.87, -6.783, -6.762, -6.786, -7.081, -6.908, -7.0, -7.025, -7.435, -7.464, -7.397, -7.423, 0.0, 0.0, -2.315, -2.289, 0.0, 0.0, -1.138, -1.058, 5.508, 6.024, 6.164, 6.245, 0.199, -0.272, -0.413, -0.434, -7.629, -8.247, -8.395, -8.494, -5.804, -5.51, -5.497, -5.499, -0.921, -0.607, -0.487, -0.42, -6.387, -6.98, -7.05, -7.102, 0.0, -7.386, -7.382, -7.385, -7.385, -3.795, -3.493, -3.494, -3.454, -0.004, 0.333, 0.348, 0.451, -8.981, -9.802, -9.979, -10.043, -8.022, -7.819, -7.807, -7.776, -3.989, -3.773, -3.66, -3.599, -8.119, -8.464, -8.538, -8.537, 0.0, 0.0, -6.812, -6.803, -5.489, -5.605, -5.608, 0.0, -7.39, 0.0, 0.0, -7.39, 0.0, 0.0, -9.475, -9.475, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 15.984, 0.0, 0.0, 0.0, -3.314, 0.0, 0.0, 0.0, -7.154, -7.783, -7.768, -8.016, 11.596, 0.0, 0.0, 11.596, 3.085, 0.0, 0.0, 3.085, -6.728, 0.0, 0.0, 0.0, -7.605, -5.935, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.57084388186
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0, 5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0, 16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0, 21684.0, 22221.0, 22758.0, 23293.0, 23883.0, 24421.0, 24958.0, 25495.0, 26085.0, 26620.0, 27156.0, 27692.0, 28279.0, 28815.0, 29351.0, 29889.0, 30477.0, 31014.0, 31550.0, 32086.0, 32675.0, 33275.0, 33812.0, 34348.0, 34938.0, 35475.0, 36012.0, 36548.0, 37138.0, 37675.0, 38210.0, 38747.0, 39337.0, 39873.0, 40409.0, 40946.0, 41536.0, 42073.0, 42610.0, 43146.0, 43732.0, 44269.0, 44805.0, 45343.0, 45930.0, 46466.0, 47002.0, 47537.0, 48125.0, 48661.0, 49197.0, 49733.0, 50321.0, 50857.0, 51394.0, 51930.0, 52518.0, 53055.0, 53591.0, 54129.0, 54717.0, 55253.0, 55789.0, 56325.0, 56914.0, 57450.0, 57987.0, 58523.0, 59108.0, 59645.0, 60181.0, 60718.0, 61304.0, 61839.0, 62375.0, 62912.0, 63499.0, 64035.0, 64570.0, 65106.0, 65694.0, 66231.0, 66768.0, 67303.0, 67891.0, 68426.0, 68961.0, 69498.0, 70086.0, 70623.0, 71159.0, 71696.0, 72284.0, 72820.0, 73357.0, 73893.0, 74477.0, 75013.0, 75550.0, 76086.0, 76671.0, 77206.0, 77742.0, 78278.0, 78819.0, 79417.0, 79954.0, 80490.0, 81026.0, 81612.0, 82147.0, 82684.0, 83220.0, 83806.0, 84342.0, 84878.0, 85414.0, 86000.0, 86536.0, 87071.0, 87607.0, 88195.0, 88731.0, 89267.0, 89803.0, 90386.0, 90923.0, 91458.0, 91994.0, 92578.0, 93114.0, 93649.0, 94185.0, 94769.0, 95305.0, 95840.0, 96376.0, 96960.0, 97495.0, 98030.0, 98566.0, 99212.0, 99748.0, 100346.0, 100944.0, 101591.0, 102127.0, 102725.0, 103323.0, 103971.0, 104570.0, 105169.0, 105768.0, 106414.0, 107011.0, 107610.0, 108208.0, 108856.0, 109454.0, 110052.0, 110650.0, 111298.0, 111896.0, 112493.0, 113092.0, 113741.0, 114339.0, 114937.0, 115535.0, 116184.0, 116719.0, 117318.0, 117915.0, 118564.0, 119099.0, 119698.0, 120295.0, 120943.0, 121478.0, 122014.0, 122550.0, 123135.0, 123670.0, 124268.0, 124866.0, 125514.0, 126049.0, 126646.0, 127241.0, 127889.0, 128425.0, 129023.0, 129621.0, 130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -62.2369482688
actual_array: [ 18.1 -119. ]
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578, 23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678, -85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888, -67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209, -74.081, -76.38, -76.595, -76.815, -81.142, -82.044, -81.918, -82.169, -69.679, -68.503, -68.12, -68.547, -54.984, -52.766, -52.368, -52.343, 0.0, 0.0, -79.114, -79.981, 0.0, 0.0, -87.04, -87.04, -69.212, -66.666, -66.069, -65.684, -56.606, -55.698, -55.476, -55.266, -59.512, -60.208, -59.905, -60.153, -47.051, -45.361, -44.815, -44.662, -55.398, -56.995, -57.373, -57.328, -62.331, -63.558, -63.703, -63.847, -67.412, -67.951, -67.996, -68.074, 0.0, 0.0, -56.62, -56.069, 0.0, 0.0, -73.043, -73.108, -68.495, -67.722, -67.515, -67.328, -69.13, -67.583, -68.394, -68.454, -61.438, -60.664, -60.181, -60.382, 0.0, 0.0, -85.079, -85.515, 0.0, 0.0, -52.033, -51.446, -46.58, -45.959, -45.562, -45.562, -52.832, -53.754, -54.293, -54.215, -72.894, -75.65, -76.266, -76.773, -65.628, -63.663, -63.524, -63.383, -48.92, -46.848, -46.333, -46.094, -64.698, -67.501, -68.107, -68.438, 0.0, -61.513, -60.52, -60.162, -59.981, -52.639, -51.615, -51.425, -51.287, -50.25, -50.162, -50.021, -49.974, -81.241, -85.471, -86.633, -87.163, -74.639, -72.834, -72.178, -72.057, -58.139, -56.858, -56.191, -55.862, -74.979, -77.338, -78.157, -78.176, 0.0, 0.0, -69.344, -69.634, -67.125, -67.997, -67.464, 0.0, -65.045, 0.0, 0.0, -65.045, 0.0, 0.0, -82.299, -82.299, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -40.239, 0.0, 0.0, 0.0, -62.293, 0.0, 0.0, 0.0, -62.178, -65.113, -65.298, -66.568, -61.21, 0.0, 0.0, -61.21, -91.393, 0.0, 0.0, -91.393, -84.7, 0.0, 0.0, 0.0, -67.755, -63.368, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -46.8838396624
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0, 5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0, 16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0, 21684.0, 22221.0, 22758.0, 23293.0, 23883.0, 24421.0, 24958.0, 25495.0, 26085.0, 26620.0, 27156.0, 27692.0, 28279.0, 28815.0, 29351.0, 29889.0, 30477.0, 31014.0, 31550.0, 32086.0, 32675.0, 33275.0, 33812.0, 34348.0, 34938.0, 35475.0, 36012.0, 36548.0, 37138.0, 37675.0, 38210.0, 38747.0, 39337.0, 39873.0, 40409.0, 40946.0, 41536.0, 42073.0, 42610.0, 43146.0, 43732.0, 44269.0, 44805.0, 45343.0, 45930.0, 46466.0, 47002.0, 47537.0, 48125.0, 48661.0, 49197.0, 49733.0, 50321.0, 50857.0, 51394.0, 51930.0, 52518.0, 53055.0, 53591.0, 54129.0, 54717.0, 55253.0, 55789.0, 56325.0, 56914.0, 57450.0, 57987.0, 58523.0, 59108.0, 59645.0, 60181.0, 60718.0, 61304.0, 61839.0, 62375.0, 62912.0, 63499.0, 64035.0, 64570.0, 65106.0, 65694.0, 66231.0, 66768.0, 67303.0, 67891.0, 68426.0, 68961.0, 69498.0, 70086.0, 70623.0, 71159.0, 71696.0, 72284.0, 72820.0, 73357.0, 73893.0, 74477.0, 75013.0, 75550.0, 76086.0, 76671.0, 77206.0, 77742.0, 78278.0, 78819.0, 79417.0, 79954.0, 80490.0, 81026.0, 81612.0, 82147.0, 82684.0, 83220.0, 83806.0, 84342.0, 84878.0, 85414.0, 86000.0, 86536.0, 87071.0, 87607.0, 88195.0, 88731.0, 89267.0, 89803.0, 90386.0, 90923.0, 91458.0, 91994.0, 92578.0, 93114.0, 93649.0, 94185.0, 94769.0, 95305.0, 95840.0, 96376.0, 96960.0, 97495.0, 98030.0, 98566.0, 99212.0, 99748.0, 100346.0, 100944.0, 101591.0, 102127.0, 102725.0, 103323.0, 103971.0, 104570.0, 105169.0, 105768.0, 106414.0, 107011.0, 107610.0, 108208.0, 108856.0, 109454.0, 110052.0, 110650.0, 111298.0, 111896.0, 112493.0, 113092.0, 113741.0, 114339.0, 114937.0, 115535.0, 116184.0, 116719.0, 117318.0, 117915.0, 118564.0, 119099.0, 119698.0, 120295.0, 120943.0, 121478.0, 122014.0, 122550.0, 123135.0, 123670.0, 124268.0, 124866.0, 125514.0, 126049.0, 126646.0, 127241.0, 127889.0, 128425.0, 129023.0, 129621.0, 130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9743
varray: [3.592, 2.896, 2.922, 2.946, 2.917, 2.964, 2.912, 2.863, 2.86, 2.871]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: residuals
SAMP_residuals_STD | 0.219750894323
varray: [3.592, 2.896, 2.922, 2.946, 2.917, 2.964, 2.912, 2.863, 2.86, 2.871]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3691
varray: [0.304, 0.286, 0.447, 0.255, 0.287, 0.401, 0.598, 0.193, 0.493, 0.427]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.124765958676
varray: [0.304, 0.286, 0.447, 0.255, 0.287, 0.401, 0.598, 0.193, 0.493, 0.427]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1329.6074
varray: [1473.246, 1326.796, 1321.334, 1319.81, 1315.777, 1314.831, 1306.933, 1308.593, 1304.161, 1304.593]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 51.037006659
varray: [1473.246, 1326.796, 1321.334, 1319.81, 1315.777, 1314.831, 1306.933, 1308.593, 1304.161, 1304.593]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.2892
varray: [-2.421, 0.892, 1.259, 1.646, 1.592, 1.53, 2.207, 1.985, 2.143, 2.059]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.36826344117
varray: [-2.421, 0.892, 1.259, 1.646, 1.592, 1.53, 2.207, 1.985, 2.143, 2.059]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: slope_shift
SAMP_d18O_STD | 0.214519126316
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.44706103997
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.5965
varray: [133.479, 111.09, 112.919, 112.586, 111.334, 112.902, 110.901, 110.159, 110.263, 110.332]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.06736640011
varray: [133.479, 111.09, 112.919, 112.586, 111.334, 112.902, 110.901, 110.159, 110.263, 110.332]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9743
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3691
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1329.6074
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.2892
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0573
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0573
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.96434693178
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.03935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.5907
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
dD_mean | -61.9992
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
d18O_stdev | 0.214519126316
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
dD_stdev | 2.44706103997
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.5907
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.9992
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.214519126316
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.44706103997
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.5607586686
actual_array: [ 1.93 -15.84]
varray: [-8.16, -7.714, -7.58, -7.447, -7.498, -7.55, -7.532, -7.465, -7.468, -7.493]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.5907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.7482632142
actual_array: [ 18.1 -119. ]
varray: [-68.464, -63.583, -61.977, -61.101, -61.094, -61.128, -60.837, -60.652, -60.578, -60.578]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.9992
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [1.0, 539.0, 1075.0, 1611.0, 2147.0, 2685.0, 3222.0, 3759.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.20105
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 27.43205
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9672
varray: [3.027, 2.919, 2.866, 2.969, 2.953, 2.957, 2.936, 2.93, 3.02, 3.095]
inst: mean_all
operation: mean_all(residuals)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: residuals
SAMP_residuals_STD | 0.0649166987591
varray: [3.027, 2.919, 2.866, 2.969, 2.953, 2.957, 2.936, 2.93, 3.02, 3.095]
inst: std_all
operation: std_all(residuals)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3565
varray: [0.145, 0.408, 0.169, 0.416, 0.42, 0.215, 0.457, 0.259, 0.474, 0.602]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.150487541013
varray: [0.145, 0.408, 0.169, 0.416, 0.42, 0.215, 0.457, 0.259, 0.474, 0.602]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.7511
varray: [1306.447, 1305.021, 1304.098, 1304.758, 1304.04, 1305.636, 1303.065, 1303.233, 1304.46, 1306.753]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.24294117587
varray: [1306.447, 1305.021, 1304.098, 1304.758, 1304.04, 1305.636, 1303.065, 1303.233, 1304.46, 1306.753]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7143
varray: [2.011, 1.905, 1.404, 2.084, 1.657, 1.11, 1.948, 1.517, 1.58, 1.927]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.313163801654
varray: [2.011, 1.905, 1.404, 2.084, 1.657, 1.11, 1.948, 1.517, 1.58, 1.927]
inst: std_all
operation: std_all(slope_shift)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: slope_shift
SAMP_d18O_STD | 0.389390178499
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.03952563464
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.7154
varray: [114.787, 111.167, 111.211, 111.515, 111.855, 112.501, 111.647, 112.179, 114.278, 116.014]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.69742250355
varray: [114.787, 111.167, 111.211, 111.515, 111.855, 112.501, 111.647, 112.179, 114.278, 116.014]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9672
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.7511
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7143
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0502
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0502
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.72094617758
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.15825
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.6494
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 36.7972
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.6494
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
dD_mean | 36.7972
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
d18O_stdev | 0.389390178499
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
dD_stdev | 5.03952563464
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.6494
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
measured_dD_mean | 36.7972
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.389390178499
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.03952563464
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.628, 4.381, 4.707, 4.755, 4.782, 4.789, 4.804, 4.823, 4.861, 4.964]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.6494
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [23.67, 33.473, 36.657, 37.837, 38.308, 38.691, 39.136, 39.507, 40.015, 40.678]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 36.7972
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [5421.0, 5957.0, 6494.0, 7032.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.20815
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.57575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8668
varray: [2.912, 2.789, 2.932, 2.813, 2.862, 2.892, 2.847, 2.842, 2.867, 2.912]
inst: mean_all
operation: mean_all(residuals)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: residuals
SAMP_residuals_STD | 0.0458858487021
varray: [2.912, 2.789, 2.932, 2.813, 2.862, 2.892, 2.847, 2.842, 2.867, 2.912]
inst: std_all
operation: std_all(residuals)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3887
varray: [0.261, 0.215, 0.602, 0.334, 0.376, 0.528, 0.334, 0.345, 0.487, 0.405]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.119890737294
varray: [0.261, 0.215, 0.602, 0.334, 0.376, 0.528, 0.334, 0.345, 0.487, 0.405]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.5996
varray: [1301.616, 1298.393, 1302.623, 1299.355, 1298.891, 1298.056, 1299.88, 1296.795, 1296.449, 1303.938]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.45627867583
varray: [1301.616, 1298.393, 1302.623, 1299.355, 1298.891, 1298.056, 1299.88, 1296.795, 1296.449, 1303.938]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.5175
varray: [2.649, 2.668, 2.388, 2.147, 2.587, 2.745, 2.408, 2.598, 2.661, 2.324]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.190634758402
varray: [2.649, 2.668, 2.388, 2.147, 2.587, 2.745, 2.408, 2.598, 2.661, 2.324]
inst: std_all
operation: std_all(slope_shift)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: slope_shift
SAMP_d18O_STD | 0.592709597245
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 7.72338329792
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.3989
varray: [111.718, 107.704, 113.22, 110.054, 109.384, 110.004, 110.755, 108.335, 109.106, 113.709]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.97992376116
varray: [111.718, 107.704, 113.22, 110.054, 109.384, 110.004, 110.755, 108.335, 109.106, 113.709]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8668
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3887
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.5996
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.5175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0502
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0502
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.72094617758
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.15825
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.764
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
dD_mean | -106.0097
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
d18O_stdev | 0.592709597245
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
dD_stdev | 7.72338329792
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.764
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
measured_dD_mean | -106.0097
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.592709597245
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
measured_dD_stdev | 7.72338329792
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-11.155, -12.556, -12.751, -12.877, -12.973, -12.996, -13.055, -13.116, -13.152, -13.009]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -12.764
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-85.367, -102.233, -105.566, -107.638, -108.64, -109.484, -109.673, -110.486, -111.122, -109.888]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -106.0097
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [10841.0, 11378.0, 11915.0, 12453.0, 12989.0, 13526.0, 14062.0, 14600.0, 15137.0, 15674.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.30855
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.57575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9365
varray: [3.057, 2.833, 2.949, 2.925, 2.89, 2.893, 2.965, 2.837, 3.018, 2.998]
inst: mean_all
operation: mean_all(residuals)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: residuals
SAMP_residuals_STD | 0.0750647868326
varray: [3.057, 2.833, 2.949, 2.925, 2.89, 2.893, 2.965, 2.837, 3.018, 2.998]
inst: std_all
operation: std_all(residuals)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4498
varray: [0.459, 0.472, 0.356, 0.45, 0.254, 0.406, 0.566, 0.305, 0.653, 0.577]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.125420005493
varray: [0.459, 0.472, 0.356, 0.45, 0.254, 0.406, 0.566, 0.305, 0.653, 0.577]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.3392
varray: [1305.039, 1296.686, 1301.129, 1300.832, 1298.954, 1298.563, 1300.911, 1298.801, 1301.011, 1301.466]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.25556840828
varray: [1305.039, 1296.686, 1301.129, 1300.832, 1298.954, 1298.563, 1300.911, 1298.801, 1301.011, 1301.466]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.4863
varray: [2.506, 2.914, 2.549, 2.191, 2.247, 2.549, 2.547, 2.568, 2.55, 2.242]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.213370856075
varray: [2.506, 2.914, 2.549, 2.191, 2.247, 2.549, 2.547, 2.568, 2.55, 2.242]
inst: std_all
operation: std_all(slope_shift)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: slope_shift
SAMP_d18O_STD | 0.168814329302
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.39269713225
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.4015
varray: [116.386, 108.521, 112.599, 112.798, 110.448, 111.066, 113.138, 109.815, 114.374, 114.87]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.44829020838
varray: [116.386, 108.521, 112.599, 112.798, 110.448, 111.066, 113.138, 109.815, 114.374, 114.87]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9365
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4498
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.3392
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.4863
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0195
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0195
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.66849502914
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.84435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.5145
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
dD_mean | -61.3303
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
d18O_stdev | 0.168814329302
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
dD_stdev | 2.39269713225
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.5145
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.3303
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.168814329302
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.39269713225
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4829982083
actual_array: [ 1.93 -15.84]
varray: [-7.94, -7.654, -7.511, -7.398, -7.511, -7.455, -7.426, -7.408, -7.388, -7.454]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.5145
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.1060940333
actual_array: [ 18.1 -119. ]
varray: [-67.516, -63.208, -61.198, -60.314, -60.835, -60.319, -60.011, -59.755, -59.938, -60.209]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.3303
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [16264.0, 16801.0, 17337.0, 17874.0, 18410.0, 18948.0, 19485.0, 20021.0, 20558.0, 21095.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.23885
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.83615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.969
varray: [3.019, 2.947, 2.949, 2.961]
inst: mean_all
operation: mean_all(residuals)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: residuals
SAMP_residuals_STD | 0.0339018190269
varray: [3.019, 2.947, 2.949, 2.961]
inst: std_all
operation: std_all(residuals)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41825
varray: [0.588, 0.202, 0.321, 0.562]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.18770610894
varray: [0.588, 0.202, 0.321, 0.562]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.82725
varray: [1302.435, 1302.06, 1301.702, 1301.112]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.562969137698
varray: [1302.435, 1302.06, 1301.702, 1301.112]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.663
varray: [2.532, 2.575, 2.574, 2.971]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.206308829994
varray: [2.532, 2.575, 2.574, 2.971]
inst: std_all
operation: std_all(slope_shift)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0787374328597
varray: [-8.56, -8.744, -8.678, -8.701]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.27030845992
varray: [-74.081, -76.38, -76.595, -76.815]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.5445
varray: [115.268, 112.379, 113.462, 113.069]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23311921565
varray: [115.268, 112.379, 113.462, 113.069]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.969
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.82725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.663
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.052
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.052
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.78265341104
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.98735
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-8.56, -8.744, -8.678, -8.701]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-74.081, -76.38, -76.595, -76.815]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.67075
varray: [-8.56, -8.744, -8.678, -8.701]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
dD_mean | -75.96775
varray: [-74.081, -76.38, -76.595, -76.815]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0787374328597
varray: [-8.56, -8.744, -8.678, -8.701]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
dD_stdev | 1.27030845992
varray: [-74.081, -76.38, -76.595, -76.815]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.67075
varray: [-8.56, -8.744, -8.678, -8.701]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
measured_dD_mean | -75.96775
varray: [-74.081, -76.38, -76.595, -76.815]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0787374328597
varray: [-8.56, -8.744, -8.678, -8.701]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.27030845992
varray: [-74.081, -76.38, -76.595, -76.815]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.6629264532
actual_array: [ 1.93 -15.84]
varray: [-8.56, -8.744, -8.678, -8.701]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.67075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.1585970636
actual_array: [ 18.1 -119. ]
varray: [-74.081, -76.38, -76.595, -76.815]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -75.96775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [21684.0, 22221.0, 22758.0, 23293.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.20635
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.3481
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9055
varray: [2.874, 2.913, 2.945, 2.89]
inst: mean_all
operation: mean_all(residuals)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: residuals
SAMP_residuals_STD | 0.0308166621597
varray: [2.874, 2.913, 2.945, 2.89]
inst: std_all
operation: std_all(residuals)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3165
varray: [0.275, 0.168, 0.605, 0.218]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.197238434388
varray: [0.275, 0.168, 0.605, 0.218]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.68125
varray: [1299.002, 1300.272, 1299.65, 1299.801]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.524621371404
varray: [1299.002, 1300.272, 1299.65, 1299.801]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.63275
varray: [2.381, 2.526, 2.754, 2.87]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.22042591348
varray: [2.381, 2.526, 2.754, 2.87]
inst: std_all
operation: std_all(slope_shift)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0842352855598
varray: [-9.514, -9.664, -9.587, -9.704]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.462332041575
varray: [-81.142, -82.044, -81.918, -82.169]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.29475
varray: [110.948, 111.323, 111.99, 110.918]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.498780763462
varray: [110.948, 111.323, 111.99, 110.918]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9055
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3165
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.68125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.63275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.394240658211
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.2624
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-9.514, -9.664, -9.587, -9.704]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-81.142, -82.044, -81.918, -82.169]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.61725
varray: [-9.514, -9.664, -9.587, -9.704]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
dD_mean | -81.81825
varray: [-81.142, -82.044, -81.918, -82.169]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0842352855598
varray: [-9.514, -9.664, -9.587, -9.704]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
dD_stdev | 0.462332041575
varray: [-81.142, -82.044, -81.918, -82.169]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.61725
varray: [-9.514, -9.664, -9.587, -9.704]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
measured_dD_mean | -81.81825
varray: [-81.142, -82.044, -81.918, -82.169]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0842352855598
varray: [-9.514, -9.664, -9.587, -9.704]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.462332041575
varray: [-81.142, -82.044, -81.918, -82.169]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.6288093365
actual_array: [ 1.93 -15.84]
varray: [-9.514, -9.664, -9.587, -9.704]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.61725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -95.7752973071
actual_array: [ 18.1 -119. ]
varray: [-81.142, -82.044, -81.918, -82.169]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -81.81825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [23883.0, 24421.0, 24958.0, 25495.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.26985
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.4941
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.81075
varray: [2.881, 2.892, 2.829, 2.641]
inst: mean_all
operation: mean_all(residuals)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: residuals
SAMP_residuals_STD | 0.116454211316
varray: [2.881, 2.892, 2.829, 2.641]
inst: std_all
operation: std_all(residuals)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.432
varray: [0.541, 0.421, 0.404, 0.362]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0767810740569
varray: [0.541, 0.421, 0.404, 0.362]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.1245
varray: [1297.767, 1297.004, 1297.06, 1288.667]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.31898109126
varray: [1297.767, 1297.004, 1297.06, 1288.667]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.45675
varray: [2.75, 2.374, 2.597, 2.106]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.280201564354
varray: [2.75, 2.374, 2.597, 2.106]
inst: std_all
operation: std_all(slope_shift)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0512274340564
varray: [-6.61, -6.511, -6.507, -6.505]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.672423167854
varray: [-69.679, -68.503, -68.12, -68.547]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.89725
varray: [110.243, 109.852, 109.203, 102.291]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.76203015893
varray: [110.243, 109.852, 109.203, 102.291]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.81075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.432
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.1245
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.45675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10625
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10625
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.6424408639
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.6599
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-6.61, -6.511, -6.507, -6.505]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-69.679, -68.503, -68.12, -68.547]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.53325
varray: [-6.61, -6.511, -6.507, -6.505]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
dD_mean | -68.71225
varray: [-69.679, -68.503, -68.12, -68.547]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0512274340564
varray: [-6.61, -6.511, -6.507, -6.505]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
dD_stdev | 0.672423167854
varray: [-69.679, -68.503, -68.12, -68.547]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.53325
varray: [-6.61, -6.511, -6.507, -6.505]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.71225
varray: [-69.679, -68.503, -68.12, -68.547]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0512274340564
varray: [-6.61, -6.511, -6.507, -6.505]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.672423167854
varray: [-69.679, -68.503, -68.12, -68.547]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.48165369773
actual_array: [ 1.93 -15.84]
varray: [-6.61, -6.511, -6.507, -6.505]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.53325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.1930439285
actual_array: [ 18.1 -119. ]
varray: [-69.679, -68.503, -68.12, -68.547]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.71225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [26085.0, 26620.0, 27156.0, 27692.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3646
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.05085
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9185
varray: [2.859, 3.016, 2.922, 2.877]
inst: mean_all
operation: mean_all(residuals)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: residuals
SAMP_residuals_STD | 0.0701925922017
varray: [2.859, 3.016, 2.922, 2.877]
inst: std_all
operation: std_all(residuals)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.48275
varray: [0.496, 0.496, 0.482, 0.457]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0183915741577
varray: [0.496, 0.496, 0.482, 0.457]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1305.421
varray: [1304.307, 1309.008, 1304.71, 1303.659]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.43020781004
varray: [1304.307, 1309.008, 1304.71, 1303.659]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.051
varray: [2.098, 2.095, 2.288, 1.723]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.236571342305
varray: [2.098, 2.095, 2.288, 1.723]
inst: std_all
operation: std_all(slope_shift)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: slope_shift
SAMP_d18O_STD | 0.158981130956
varray: [-4.029, -3.729, -3.714, -3.694]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.26081385753
varray: [-54.984, -52.766, -52.368, -52.343]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.4695
varray: [110.458, 114.416, 110.771, 110.233]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.97668316463
varray: [110.458, 114.416, 110.771, 110.233]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9185
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.48275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1305.421
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.051
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0015
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0015
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0514226945492
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.08765
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-4.029, -3.729, -3.714, -3.694]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-54.984, -52.766, -52.368, -52.343]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7915
varray: [-4.029, -3.729, -3.714, -3.694]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
dD_mean | -53.11525
varray: [-54.984, -52.766, -52.368, -52.343]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
d18O_stdev | 0.158981130956
varray: [-4.029, -3.729, -3.714, -3.694]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
dD_stdev | 1.26081385753
varray: [-54.984, -52.766, -52.368, -52.343]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7915
varray: [-4.029, -3.729, -3.714, -3.694]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.11525
varray: [-54.984, -52.766, -52.368, -52.343]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.158981130956
varray: [-4.029, -3.729, -3.714, -3.694]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.26081385753
varray: [-54.984, -52.766, -52.368, -52.343]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.68375681946
actual_array: [ 1.93 -15.84]
varray: [-4.029, -3.729, -3.714, -3.694]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.7915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.2193367757
actual_array: [ 18.1 -119. ]
varray: [-54.984, -52.766, -52.368, -52.343]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -53.11525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [28279.0, 28815.0, 29351.0, 29889.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.25685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.24565
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.8845
varray: [0.0, 0.0, 7.767, 7.771]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: residuals
SAMP_residuals_STD | 4.48543453859
varray: [0.0, 0.0, 7.767, 7.771]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -8.7845
varray: [0.0, 0.0, -17.822, -17.316]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 10.1455701171
varray: [0.0, 0.0, -17.822, -17.316]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1499.93975
varray: [0.0, 0.0, 2997.737, 3002.022]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1731.9821203
varray: [0.0, 0.0, 2997.737, 3002.022]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -53.7905
varray: [0.0, 0.0, -107.027, -108.135]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: slope_shift
SAMP_slope_shift_STD | 62.1135664006
varray: [0.0, 0.0, -107.027, -108.135]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: slope_shift
SAMP_d18O_STD | 7.7506387855
varray: [0.0, 0.0, 13.423, 13.426]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 45.9281344448
varray: [0.0, 0.0, -79.114, -79.981]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.18175
varray: [0.0, 0.0, 128.568, 132.159]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 75.2796780297
varray: [0.0, 0.0, 128.568, 132.159]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.8845
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 8.7845
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1499.93975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 53.7905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.9675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.9675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 33.1676379842
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -46.3754
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 13.423, 13.426]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -79.114, -79.981]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 6.71225
varray: [0.0, 0.0, 13.423, 13.426]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
dD_mean | -39.77375
varray: [0.0, 0.0, -79.114, -79.981]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
d18O_stdev | 7.7506387855
varray: [0.0, 0.0, 13.423, 13.426]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
dD_stdev | 45.9281344448
varray: [0.0, 0.0, -79.114, -79.981]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
measured_d18O_mean | 6.71225
varray: [0.0, 0.0, 13.423, 13.426]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.77375
varray: [0.0, 0.0, -79.114, -79.981]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 7.7506387855
varray: [0.0, 0.0, 13.423, 13.426]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
measured_dD_stdev | 45.9281344448
varray: [0.0, 0.0, -79.114, -79.981]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.03509403678
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 13.423, 13.426]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 6.71225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.4109945318
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -79.114, -79.981]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -39.77375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [30477.0, 31014.0, 31550.0, 32086.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.29085
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 197.7644
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.427
varray: [0.0, 0.0, 2.854, 2.854]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: residuals
SAMP_residuals_STD | 1.64775766827
varray: [0.0, 0.0, 2.854, 2.854]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.091
varray: [0.0, 0.0, 0.182, 0.182]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105077748993
varray: [0.0, 0.0, 0.182, 0.182]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 653.8655
varray: [0.0, 0.0, 1307.731, 1307.731]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 755.018844878
varray: [0.0, 0.0, 1307.731, 1307.731]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.688
varray: [0.0, 0.0, 1.376, 1.376]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.794433970405
varray: [0.0, 0.0, 1.376, 1.376]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: slope_shift
SAMP_d18O_STD | 5.76888388974
varray: [0.0, 0.0, -9.992, -9.992]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 50.2525674303
varray: [0.0, 0.0, -87.04, -87.04]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.445
varray: [0.0, 0.0, 110.89, 110.89]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 64.0223713504
varray: [0.0, 0.0, 110.89, 110.89]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.427
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.091
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 653.8655
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.688
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.49
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.49
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.0798765855
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -56.11215
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, -9.992, -9.992]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -87.04, -87.04]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.996
varray: [0.0, 0.0, -9.992, -9.992]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
dD_mean | -43.52
varray: [0.0, 0.0, -87.04, -87.04]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
d18O_stdev | 5.76888388974
varray: [0.0, 0.0, -9.992, -9.992]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
dD_stdev | 50.2525674303
varray: [0.0, 0.0, -87.04, -87.04]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.996
varray: [0.0, 0.0, -9.992, -9.992]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.52
varray: [0.0, 0.0, -87.04, -87.04]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.76888388974
varray: [0.0, 0.0, -9.992, -9.992]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
measured_dD_stdev | 50.2525674303
varray: [0.0, 0.0, -87.04, -87.04]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.91292315114
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -9.992, -9.992]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.0075355603
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -87.04, -87.04]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -43.52
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [32675.0, 33275.0, 33812.0, 34348.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.74835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -648.30985
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.92525
varray: [2.961, 2.952, 2.872, 2.916]
inst: mean_all
operation: mean_all(residuals)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: residuals
SAMP_residuals_STD | 0.0404753011107
varray: [2.961, 2.952, 2.872, 2.916]
inst: std_all
operation: std_all(residuals)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.317
varray: [0.401, 0.332, 0.094, 0.441]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.155334048639
varray: [0.401, 0.332, 0.094, 0.441]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.01475
varray: [1301.521, 1303.477, 1302.046, 1301.015]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.06182872913
varray: [1301.521, 1303.477, 1302.046, 1301.015]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.173
varray: [1.978, 1.911, 2.184, 2.619]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.319220926632
varray: [1.978, 1.911, 2.184, 2.619]
inst: std_all
operation: std_all(slope_shift)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: slope_shift
SAMP_d18O_STD | 0.276476641569
varray: [-4.577, -4.1, -4.033, -3.97]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.58840346995
varray: [-69.212, -66.666, -66.069, -65.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.01
varray: [112.406, 112.925, 111.356, 111.353]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.786003816785
varray: [112.406, 112.925, 111.356, 111.353]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.92525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.317
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.01475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.173
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.282824820021
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.45285
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-4.577, -4.1, -4.033, -3.97]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-69.212, -66.666, -66.069, -65.684]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.17
varray: [-4.577, -4.1, -4.033, -3.97]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
dD_mean | -66.90775
varray: [-69.212, -66.666, -66.069, -65.684]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
d18O_stdev | 0.276476641569
varray: [-4.577, -4.1, -4.033, -3.97]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
dD_stdev | 1.58840346995
varray: [-69.212, -66.666, -66.069, -65.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.17
varray: [-4.577, -4.1, -4.033, -3.97]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.90775
varray: [-69.212, -66.666, -66.069, -65.684]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.276476641569
varray: [-4.577, -4.1, -4.033, -3.97]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.58840346995
varray: [-69.212, -66.666, -66.069, -65.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.07000792493
actual_array: [ 1.93 -15.84]
varray: [-4.577, -4.1, -4.033, -3.97]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.4606559977
actual_array: [ 18.1 -119. ]
varray: [-69.212, -66.666, -66.069, -65.684]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -66.90775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [34938.0, 35475.0, 36012.0, 36548.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.2501
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.1606
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.916
varray: [2.956, 2.871, 2.889, 2.948]
inst: mean_all
operation: mean_all(residuals)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: residuals
SAMP_residuals_STD | 0.0423398945047
varray: [2.956, 2.871, 2.889, 2.948]
inst: std_all
operation: std_all(residuals)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3115
varray: [0.22, 0.293, 0.254, 0.479]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11558114033
varray: [0.22, 0.293, 0.254, 0.479]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1306.16125
varray: [1308.918, 1305.78, 1304.774, 1305.173]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.88380136515
varray: [1308.918, 1305.78, 1304.774, 1305.173]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.77475
varray: [2.008, 1.872, 1.892, 1.327]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.304461683413
varray: [2.008, 1.872, 1.892, 1.327]
inst: std_all
operation: std_all(slope_shift)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0192786583212
varray: [-3.682, -3.714, -3.686, -3.72]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.589983898085
varray: [-56.606, -55.698, -55.476, -55.266]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.04825
varray: [113.551, 111.712, 110.876, 112.054]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.11736665871
varray: [113.551, 111.712, 110.876, 112.054]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.916
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3115
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1306.16125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.77475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.001
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.001
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0342817963661
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.4911
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-3.682, -3.714, -3.686, -3.72]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-56.606, -55.698, -55.476, -55.266]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7005
varray: [-3.682, -3.714, -3.686, -3.72]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
dD_mean | -55.7615
varray: [-56.606, -55.698, -55.476, -55.266]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0192786583212
varray: [-3.682, -3.714, -3.686, -3.72]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
dD_stdev | 0.589983898085
varray: [-56.606, -55.698, -55.476, -55.266]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7005
varray: [-3.682, -3.714, -3.686, -3.72]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.7615
varray: [-56.606, -55.698, -55.476, -55.266]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0192786583212
varray: [-3.682, -3.714, -3.686, -3.72]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.589983898085
varray: [-56.606, -55.698, -55.476, -55.266]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.59089327759
actual_array: [ 1.93 -15.84]
varray: [-3.682, -3.714, -3.686, -3.72]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.7005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.7598363945
actual_array: [ 18.1 -119. ]
varray: [-56.606, -55.698, -55.476, -55.266]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -55.7615
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [37138.0, 37675.0, 38210.0, 38747.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.25935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.9859
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86675
varray: [2.863, 2.795, 2.935, 2.874]
inst: mean_all
operation: mean_all(residuals)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: residuals
SAMP_residuals_STD | 0.0573665117178
varray: [2.863, 2.795, 2.935, 2.874]
inst: std_all
operation: std_all(residuals)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4845
varray: [0.498, 0.322, 0.677, 0.441]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.147802796545
varray: [0.498, 0.322, 0.677, 0.441]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.025
varray: [1296.617, 1295.722, 1299.09, 1296.671]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.44381600859
varray: [1296.617, 1295.722, 1299.09, 1296.671]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.436
varray: [2.267, 2.732, 2.809, 1.936]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.410408739348
varray: [2.267, 2.732, 2.809, 1.936]
inst: std_all
operation: std_all(slope_shift)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: slope_shift
SAMP_d18O_STD | 0.175729669284
varray: [-7.043, -7.317, -7.356, -7.453]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.317028389896
varray: [-59.512, -60.208, -59.905, -60.153]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.417
varray: [109.788, 108.202, 112.723, 110.955]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.90691251329
varray: [109.788, 108.202, 112.723, 110.955]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4845
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.436
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05025
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05025
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.7226602674
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.14015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.043, -7.317, -7.356, -7.453]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-59.512, -60.208, -59.905, -60.153]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.29225
varray: [-7.043, -7.317, -7.356, -7.453]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
dD_mean | -59.9445
varray: [-59.512, -60.208, -59.905, -60.153]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
d18O_stdev | 0.175729669284
varray: [-7.043, -7.317, -7.356, -7.453]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
dD_stdev | 0.317028389896
varray: [-59.512, -60.208, -59.905, -60.153]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.29225
varray: [-7.043, -7.317, -7.356, -7.453]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.9445
varray: [-59.512, -60.208, -59.905, -60.153]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.175729669284
varray: [-7.043, -7.317, -7.356, -7.453]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.317028389896
varray: [-59.512, -60.208, -59.905, -60.153]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2561968656
actual_array: [ 1.93 -15.84]
varray: [-7.043, -7.317, -7.356, -7.453]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.29225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.7756738645
actual_array: [ 18.1 -119. ]
varray: [-59.512, -60.208, -59.905, -60.153]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -59.9445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [39337.0, 39873.0, 40409.0, 40946.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3086
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.15035
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.88425
varray: [2.907, 2.911, 2.85, 2.869]
inst: mean_all
operation: mean_all(residuals)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: residuals
SAMP_residuals_STD | 0.0296577702016
varray: [2.907, 2.911, 2.85, 2.869]
inst: std_all
operation: std_all(residuals)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.24275
varray: [0.105, 0.24, 0.199, 0.427]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.135209405491
varray: [0.105, 0.24, 0.199, 0.427]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.4245
varray: [1303.098, 1302.912, 1302.155, 1301.533]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.720800712171
varray: [1303.098, 1302.912, 1302.155, 1301.533]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.92925
varray: [1.724, 1.983, 1.846, 2.164]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.188904517327
varray: [1.724, 1.983, 1.846, 2.164]
inst: std_all
operation: std_all(slope_shift)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: slope_shift
SAMP_d18O_STD | 0.315582398115
varray: [-1.304, -0.806, -0.679, -0.602]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.09442690482
varray: [-47.051, -45.361, -44.815, -44.662]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.869
varray: [111.683, 110.511, 110.157, 111.125]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.674101871629
varray: [111.683, 110.511, 110.157, 111.125]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.88425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.24275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.4245
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.92925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.12272883099
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.68815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-1.304, -0.806, -0.679, -0.602]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-47.051, -45.361, -44.815, -44.662]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.84775
varray: [-1.304, -0.806, -0.679, -0.602]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
dD_mean | -45.47225
varray: [-47.051, -45.361, -44.815, -44.662]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
d18O_stdev | 0.315582398115
varray: [-1.304, -0.806, -0.679, -0.602]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
dD_stdev | 1.09442690482
varray: [-47.051, -45.361, -44.815, -44.662]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.84775
varray: [-1.304, -0.806, -0.679, -0.602]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.47225
varray: [-47.051, -45.361, -44.815, -44.662]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.315582398115
varray: [-1.304, -0.806, -0.679, -0.602]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.09442690482
varray: [-47.051, -45.361, -44.815, -44.662]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.67972328781
actual_array: [ 1.93 -15.84]
varray: [-1.304, -0.806, -0.679, -0.602]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.84775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.8817690532
actual_array: [ 18.1 -119. ]
varray: [-47.051, -45.361, -44.815, -44.662]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.47225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [41536.0, 42073.0, 42610.0, 43146.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.2911
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.24915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.905
varray: [2.911, 2.852, 2.941, 2.916]
inst: mean_all
operation: mean_all(residuals)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: residuals
SAMP_residuals_STD | 0.0376917320731
varray: [2.911, 2.852, 2.941, 2.916]
inst: std_all
operation: std_all(residuals)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.36225
varray: [0.163, 0.275, 0.546, 0.465]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.174771040698
varray: [0.163, 0.275, 0.546, 0.465]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.0325
varray: [1299.824, 1299.142, 1298.285, 1298.879]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.63790830062
varray: [1299.824, 1299.142, 1298.285, 1298.879]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.0715
varray: [1.857, 2.326, 2.012, 2.091]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.195534310715
varray: [1.857, 2.326, 2.012, 2.091]
inst: std_all
operation: std_all(slope_shift)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0466288537281
varray: [-1.62, -1.697, -1.715, -1.721]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.932368489386
varray: [-55.398, -56.995, -57.373, -57.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.09525
varray: [111.046, 110.377, 111.852, 111.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.603074000434
varray: [111.046, 110.377, 111.852, 111.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.905
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.0325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.0715
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.012
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.012
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.411381556394
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.4619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-1.62, -1.697, -1.715, -1.721]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-55.398, -56.995, -57.373, -57.328]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.68825
varray: [-1.62, -1.697, -1.715, -1.721]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
dD_mean | -56.7735
varray: [-55.398, -56.995, -57.373, -57.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0466288537281
varray: [-1.62, -1.697, -1.715, -1.721]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
dD_stdev | 0.932368489386
varray: [-55.398, -56.995, -57.373, -57.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.68825
varray: [-1.62, -1.697, -1.715, -1.721]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.7735
varray: [-55.398, -56.995, -57.373, -57.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0466288537281
varray: [-1.62, -1.697, -1.715, -1.721]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.932368489386
varray: [-55.398, -56.995, -57.373, -57.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.53743545201
actual_array: [ 1.93 -15.84]
varray: [-1.62, -1.697, -1.715, -1.721]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.68825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.7313944914
actual_array: [ 18.1 -119. ]
varray: [-55.398, -56.995, -57.373, -57.328]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.7735
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [43732.0, 44269.0, 44805.0, 45343.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.27035
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.14285
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94875
varray: [2.978, 2.944, 2.955, 2.918]
inst: mean_all
operation: mean_all(residuals)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: residuals
SAMP_residuals_STD | 0.0249181995069
varray: [2.978, 2.944, 2.955, 2.918]
inst: std_all
operation: std_all(residuals)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.36775
varray: [0.364, 0.447, 0.385, 0.275]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0711682279298
varray: [0.364, 0.447, 0.385, 0.275]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.19825
varray: [1303.044, 1300.633, 1300.373, 1300.743]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.24024066885
varray: [1303.044, 1300.633, 1300.373, 1300.743]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3735
varray: [2.646, 2.103, 2.128, 2.617]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.298322532393
varray: [2.646, 2.103, 2.128, 2.617]
inst: std_all
operation: std_all(slope_shift)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: slope_shift
SAMP_d18O_STD | 0.212041466071
varray: [-5.176, -5.549, -5.593, -5.639]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.695907740533
varray: [-62.331, -63.558, -63.703, -63.847]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.535
varray: [113.839, 113.181, 112.163, 110.957]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.25780231621
varray: [113.839, 113.181, 112.163, 110.957]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.19825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3735
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.03175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.08844703462
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.97785
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-5.176, -5.549, -5.593, -5.639]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-62.331, -63.558, -63.703, -63.847]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.48925
varray: [-5.176, -5.549, -5.593, -5.639]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
dD_mean | -63.35975
varray: [-62.331, -63.558, -63.703, -63.847]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
d18O_stdev | 0.212041466071
varray: [-5.176, -5.549, -5.593, -5.639]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
dD_stdev | 0.695907740533
varray: [-62.331, -63.558, -63.703, -63.847]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.48925
varray: [-5.176, -5.549, -5.593, -5.639]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.35975
varray: [-62.331, -63.558, -63.703, -63.847]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.212041466071
varray: [-5.176, -5.549, -5.593, -5.639]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.695907740533
varray: [-62.331, -63.558, -63.703, -63.847]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.41627416243
actual_array: [ 1.93 -15.84]
varray: [-5.176, -5.549, -5.593, -5.639]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.48925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.0544424324
actual_array: [ 18.1 -119. ]
varray: [-62.331, -63.558, -63.703, -63.847]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.35975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [45930.0, 46466.0, 47002.0, 47537.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.2266
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9771
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8955
varray: [2.882, 2.882, 2.875, 2.943]
inst: mean_all
operation: mean_all(residuals)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: residuals
SAMP_residuals_STD | 0.0318381322735
varray: [2.882, 2.882, 2.875, 2.943]
inst: std_all
operation: std_all(residuals)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.43625
varray: [0.318, 0.581, 0.327, 0.519]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.133814236911
varray: [0.318, 0.581, 0.327, 0.519]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.24475
varray: [1298.801, 1297.895, 1298.116, 1298.167]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.389170378283
varray: [1298.801, 1297.895, 1298.116, 1298.167]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.22625
varray: [2.343, 2.493, 2.247, 1.822]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.287885828064
varray: [2.343, 2.493, 2.247, 1.822]
inst: std_all
operation: std_all(slope_shift)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0381083980246
varray: [-5.019, -4.929, -4.979, -4.956]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.301808300085
varray: [-67.412, -67.951, -67.996, -68.074]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.047
varray: [110.854, 110.946, 110.641, 111.747]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.483833993569
varray: [110.854, 110.946, 110.641, 111.747]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8955
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.43625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.24475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.22625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0215
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0215
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.737058621872
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.51015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-5.019, -4.929, -4.979, -4.956]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-67.412, -67.951, -67.996, -68.074]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.97075
varray: [-5.019, -4.929, -4.979, -4.956]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
dD_mean | -67.85825
varray: [-67.412, -67.951, -67.996, -68.074]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0381083980246
varray: [-5.019, -4.929, -4.979, -4.956]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
dD_stdev | 0.301808300085
varray: [-67.412, -67.951, -67.996, -68.074]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.97075
varray: [-5.019, -4.929, -4.979, -4.956]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.85825
varray: [-67.412, -67.951, -67.996, -68.074]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0381083980246
varray: [-5.019, -4.929, -4.979, -4.956]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.301808300085
varray: [-67.412, -67.951, -67.996, -68.074]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.88715606946
actual_array: [ 1.93 -15.84]
varray: [-5.019, -4.929, -4.979, -4.956]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.97075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.3731717795
actual_array: [ 18.1 -119. ]
varray: [-67.412, -67.951, -67.996, -68.074]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.85825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [48125.0, 48661.0, 49197.0, 49733.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.27985
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.9306
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.85775
varray: [0.0, 0.0, 3.595, 3.836]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: residuals
SAMP_residuals_STD | 2.14740003648
varray: [0.0, 0.0, 3.595, 3.836]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.7625
varray: [0.0, 0.0, -7.377, -7.673]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 4.34624101955
varray: [0.0, 0.0, -7.377, -7.673]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1178.16875
varray: [0.0, 0.0, 2316.093, 2396.582]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1360.82887162
varray: [0.0, 0.0, 2316.093, 2396.582]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -27.98725
varray: [0.0, 0.0, -53.957, -57.992]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: slope_shift
SAMP_slope_shift_STD | 32.3588486752
varray: [0.0, 0.0, -53.957, -57.992]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: slope_shift
SAMP_d18O_STD | 2.44767650028
varray: [0.0, 0.0, 4.24, 4.239]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 32.531289967
varray: [0.0, 0.0, -56.62, -56.069]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 49.4945
varray: [0.0, 0.0, 95.098, 102.88]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 57.2395605708
varray: [0.0, 0.0, 95.098, 102.88]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.85775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.7625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1178.16875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.98725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.05925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.05925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -36.3129928008
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -62.06265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 4.24, 4.239]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -56.62, -56.069]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.11975
varray: [0.0, 0.0, 4.24, 4.239]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
dD_mean | -28.17225
varray: [0.0, 0.0, -56.62, -56.069]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
d18O_stdev | 2.44767650028
varray: [0.0, 0.0, 4.24, 4.239]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
dD_stdev | 32.531289967
varray: [0.0, 0.0, -56.62, -56.069]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.11975
varray: [0.0, 0.0, 4.24, 4.239]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
measured_dD_mean | -28.17225
varray: [0.0, 0.0, -56.62, -56.069]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.44767650028
varray: [0.0, 0.0, 4.24, 4.239]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
measured_dD_stdev | 32.531289967
varray: [0.0, 0.0, -56.62, -56.069]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.651453392215
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 4.24, 4.239]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 2.11975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -44.2731177905
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -56.62, -56.069]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -28.17225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [50321.0, 50857.0, 51394.0, 51930.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.3176
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -124.0066
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.4285
varray: [0.0, 0.0, 2.836, 2.878]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: residuals
SAMP_residuals_STD | 1.64957883514
varray: [0.0, 0.0, 2.836, 2.878]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2565
varray: [0.0, 0.0, 0.427, 0.599]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.304390648126
varray: [0.0, 0.0, 0.427, 0.599]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 650.418
varray: [0.0, 0.0, 1302.207, 1299.465]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 751.038849012
varray: [0.0, 0.0, 1302.207, 1299.465]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.22375
varray: [0.0, 0.0, 2.521, 2.374]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.41433856743
varray: [0.0, 0.0, 2.521, 2.374]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: slope_shift
SAMP_d18O_STD | 4.0559017493
varray: [0.0, 0.0, -7.039, -7.011]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 42.1901679413
varray: [0.0, 0.0, -73.043, -73.108]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.0295
varray: [0.0, 0.0, 109.74, 110.378]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 63.5431270949
varray: [0.0, 0.0, 109.74, 110.378]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.4285
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 650.418
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.22375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.4885
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.4885
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.028453891
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -56.52765
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, -7.039, -7.011]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -73.043, -73.108]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.5125
varray: [0.0, 0.0, -7.039, -7.011]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
dD_mean | -36.53775
varray: [0.0, 0.0, -73.043, -73.108]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
d18O_stdev | 4.0559017493
varray: [0.0, 0.0, -7.039, -7.011]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
dD_stdev | 42.1901679413
varray: [0.0, 0.0, -73.043, -73.108]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.5125
varray: [0.0, 0.0, -7.039, -7.011]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.53775
varray: [0.0, 0.0, -73.043, -73.108]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.0559017493
varray: [0.0, 0.0, -7.039, -7.011]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
measured_dD_stdev | 42.1901679413
varray: [0.0, 0.0, -73.043, -73.108]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.39904332296
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -7.039, -7.011]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.5125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.3043127118
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -73.043, -73.108]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -36.53775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [52518.0, 53055.0, 53591.0, 54129.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.74685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -651.75735
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8045
varray: [2.821, 2.756, 2.811, 2.83]
inst: mean_all
operation: mean_all(residuals)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: residuals
SAMP_residuals_STD | 0.0332515663791
varray: [2.821, 2.756, 2.811, 2.83]
inst: std_all
operation: std_all(residuals)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.27825
varray: [0.337, 0.241, 0.261, 0.274]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0414517792139
varray: [0.337, 0.241, 0.261, 0.274]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.22575
varray: [1295.905, 1297.287, 1297.444, 1298.267]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.979790241157
varray: [1295.905, 1297.287, 1297.444, 1298.267]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.67125
varray: [2.311, 3.056, 2.748, 2.57]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.313025424952
varray: [2.311, 3.056, 2.748, 2.57]
inst: std_all
operation: std_all(slope_shift)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0477100618319
varray: [-6.87, -6.783, -6.762, -6.786]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.512581050502
varray: [-68.495, -67.722, -67.515, -67.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.79875
varray: [108.092, 106.064, 107.716, 109.323]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.34480962097
varray: [108.092, 106.064, 107.716, 109.323]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8045
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.27825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.22575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.67125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1125
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1125
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.85670209119
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.7584
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-6.87, -6.783, -6.762, -6.786]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-68.495, -67.722, -67.515, -67.328]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.80025
varray: [-6.87, -6.783, -6.762, -6.786]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
dD_mean | -67.765
varray: [-68.495, -67.722, -67.515, -67.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0477100618319
varray: [-6.87, -6.783, -6.762, -6.786]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
dD_stdev | 0.512581050502
varray: [-68.495, -67.722, -67.515, -67.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.80025
varray: [-6.87, -6.783, -6.762, -6.786]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.765
varray: [-68.495, -67.722, -67.515, -67.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0477100618319
varray: [-6.87, -6.783, -6.762, -6.786]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.512581050502
varray: [-68.495, -67.722, -67.515, -67.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.75412145244
actual_array: [ 1.93 -15.84]
varray: [-6.87, -6.783, -6.762, -6.786]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.80025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.2836482691
actual_array: [ 18.1 -119. ]
varray: [-68.495, -67.722, -67.515, -67.328]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.765
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [54717.0, 55253.0, 55789.0, 56325.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.37085
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.9496
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.693
varray: [2.453, 2.94, 2.643, 2.736]
inst: mean_all
operation: mean_all(residuals)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: residuals
SAMP_residuals_STD | 0.202449993826
varray: [2.453, 2.94, 2.643, 2.736]
inst: std_all
operation: std_all(residuals)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3705
varray: [0.393, 0.46, 0.226, 0.403]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.100752171192
varray: [0.393, 0.46, 0.226, 0.403]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.40075
varray: [1281.943, 1301.573, 1290.409, 1291.678]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.04123002096
varray: [1281.943, 1301.573, 1290.409, 1291.678]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.49225
varray: [2.739, 2.271, 2.74, 2.219]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.286288170672
varray: [2.739, 2.271, 2.74, 2.219]
inst: std_all
operation: std_all(slope_shift)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0721133367231
varray: [-7.081, -6.908, -7.0, -7.025]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.633235277497
varray: [-69.13, -67.583, -68.394, -68.454]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.30675
varray: [96.033, 113.195, 102.591, 105.408]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.10897265785
varray: [96.033, 113.195, 102.591, 105.408]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.693
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3705
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.40075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.49225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.224
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.224
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.67912238601
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.2504
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.081, -6.908, -7.0, -7.025]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-69.13, -67.583, -68.394, -68.454]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.0035
varray: [-7.081, -6.908, -7.0, -7.025]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
dD_mean | -68.39025
varray: [-69.13, -67.583, -68.394, -68.454]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0721133367231
varray: [-7.081, -6.908, -7.0, -7.025]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
dD_stdev | 0.633235277497
varray: [-69.13, -67.583, -68.394, -68.454]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.0035
varray: [-7.081, -6.908, -7.0, -7.025]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.39025
varray: [-69.13, -67.583, -68.394, -68.454]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0721133367231
varray: [-7.081, -6.908, -7.0, -7.025]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.633235277497
varray: [-69.13, -67.583, -68.394, -68.454]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.96153370393
actual_array: [ 1.93 -15.84]
varray: [-7.081, -6.908, -7.0, -7.025]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.0035
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.8839118068
actual_array: [ 18.1 -119. ]
varray: [-69.13, -67.583, -68.394, -68.454]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.39025
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [56914.0, 57450.0, 57987.0, 58523.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.48235
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.7746
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.831
varray: [2.835, 2.841, 2.874, 2.774]
inst: mean_all
operation: mean_all(residuals)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: residuals
SAMP_residuals_STD | 0.0416893271714
varray: [2.835, 2.841, 2.874, 2.774]
inst: std_all
operation: std_all(residuals)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34525
varray: [0.438, 0.218, 0.258, 0.467]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.125473436764
varray: [0.438, 0.218, 0.258, 0.467]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.41775
varray: [1296.706, 1296.059, 1299.623, 1293.283]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.60205526139
varray: [1296.706, 1296.059, 1299.623, 1293.283]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.61
varray: [2.267, 2.69, 3.103, 2.38]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.374164847806
varray: [2.267, 2.69, 3.103, 2.38]
inst: std_all
operation: std_all(slope_shift)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0278013788627
varray: [-7.435, -7.464, -7.397, -7.423]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.551322274174
varray: [-61.438, -60.664, -60.181, -60.382]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.66325
varray: [109.749, 108.244, 110.047, 106.613]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.57828100054
varray: [109.749, 108.244, 110.047, 106.613]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.831
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.41775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.61
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.086
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.086
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.94823448749
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.8939
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.435, -7.464, -7.397, -7.423]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-61.438, -60.664, -60.181, -60.382]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.42975
varray: [-7.435, -7.464, -7.397, -7.423]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
dD_mean | -60.66625
varray: [-61.438, -60.664, -60.181, -60.382]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0278013788627
varray: [-7.435, -7.464, -7.397, -7.423]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
dD_stdev | 0.551322274174
varray: [-61.438, -60.664, -60.181, -60.382]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.42975
varray: [-7.435, -7.464, -7.397, -7.423]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.66625
varray: [-61.438, -60.664, -60.181, -60.382]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0278013788627
varray: [-7.435, -7.464, -7.397, -7.423]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.551322274174
varray: [-61.438, -60.664, -60.181, -60.382]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3965126569
actual_array: [ 1.93 -15.84]
varray: [-7.435, -7.464, -7.397, -7.423]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.42975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.468581035
actual_array: [ 18.1 -119. ]
varray: [-61.438, -60.664, -60.181, -60.382]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.66625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [59108.0, 59645.0, 60181.0, 60718.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.34435
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.7576
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.75025
varray: [0.0, 0.0, 3.503, 3.498]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: residuals
SAMP_residuals_STD | 2.02101564813
varray: [0.0, 0.0, 3.503, 3.498]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.73475
varray: [0.0, 0.0, -5.478, -5.461]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 3.1578249239
varray: [0.0, 0.0, -5.478, -5.461]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 912.05625
varray: [0.0, 0.0, 1825.968, 1822.257]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1053.15293261
varray: [0.0, 0.0, 1825.968, 1822.257]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -16.029
varray: [0.0, 0.0, -32.306, -31.81]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: slope_shift
SAMP_slope_shift_STD | 18.5098025561
varray: [0.0, 0.0, -32.306, -31.81]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: slope_shift
SAMP_d18O_STD | 1.32910270484
varray: [0.0, 0.0, -2.315, -2.289]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 49.246567586
varray: [0.0, 0.0, -85.079, -85.515]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.21575
varray: [0.0, 0.0, 116.814, 116.049]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 67.2224833538
varray: [0.0, 0.0, 116.814, 116.049]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.75025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.73475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 912.05625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.029
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.16675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.16675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -39.9982859102
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -53.3414
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, -2.315, -2.289]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -85.079, -85.515]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.151
varray: [0.0, 0.0, -2.315, -2.289]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
dD_mean | -42.6485
varray: [0.0, 0.0, -85.079, -85.515]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
d18O_stdev | 1.32910270484
varray: [0.0, 0.0, -2.315, -2.289]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
dD_stdev | 49.246567586
varray: [0.0, 0.0, -85.079, -85.515]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.151
varray: [0.0, 0.0, -2.315, -2.289]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.6485
varray: [0.0, 0.0, -85.079, -85.515]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.32910270484
varray: [0.0, 0.0, -2.315, -2.289]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
measured_dD_stdev | 49.246567586
varray: [0.0, 0.0, -85.079, -85.515]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.98918338751
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -2.315, -2.289]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.151
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.1708627524
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -85.079, -85.515]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -42.6485
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [61304.0, 61839.0, 62375.0, 62912.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4251
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -390.1191
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.43
varray: [0.0, 0.0, 2.841, 2.879]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: residuals
SAMP_residuals_STD | 1.65129464361
varray: [0.0, 0.0, 2.841, 2.879]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.19875
varray: [0.0, 0.0, 0.367, 0.428]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.230843922741
varray: [0.0, 0.0, 0.367, 0.428]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 690.971
varray: [0.0, 0.0, 1382.561, 1381.323]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 797.864745782
varray: [0.0, 0.0, 1382.561, 1381.323]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.66825
varray: [0.0, 0.0, -1.233, -1.44]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.776242391267
varray: [0.0, 0.0, -1.233, -1.44]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: slope_shift
SAMP_d18O_STD | 0.63477134991
varray: [0.0, 0.0, -1.138, -1.058]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 29.8727754806
varray: [0.0, 0.0, -52.033, -51.446]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.036
varray: [0.0, 0.0, 109.592, 110.552]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 63.5513073141
varray: [0.0, 0.0, 109.592, 110.552]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.43
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.19875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 690.971
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.66825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.487
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.487
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.9770311964
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -56.52115
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, -1.138, -1.058]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -52.033, -51.446]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.549
varray: [0.0, 0.0, -1.138, -1.058]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
dD_mean | -25.86975
varray: [0.0, 0.0, -52.033, -51.446]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
d18O_stdev | 0.63477134991
varray: [0.0, 0.0, -1.138, -1.058]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
dD_stdev | 29.8727754806
varray: [0.0, 0.0, -52.033, -51.446]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.549
varray: [0.0, 0.0, -1.138, -1.058]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.86975
varray: [0.0, 0.0, -52.033, -51.446]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.63477134991
varray: [0.0, 0.0, -1.138, -1.058]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
measured_dD_stdev | 29.8727754806
varray: [0.0, 0.0, -52.033, -51.446]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.37485534129
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -1.138, -1.058]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.549
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.0626311124
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -52.033, -51.446]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -25.86975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [63499.0, 64035.0, 64570.0, 65106.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.74535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -611.20435
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84275
varray: [2.869, 2.847, 2.844, 2.811]
inst: mean_all
operation: mean_all(residuals)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: residuals
SAMP_residuals_STD | 0.0239217474278
varray: [2.869, 2.847, 2.844, 2.811]
inst: std_all
operation: std_all(residuals)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.153
varray: [0.094, 0.128, 0.367, 0.023]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149222429056
varray: [0.094, 0.128, 0.367, 0.023]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1321.4915
varray: [1323.541, 1320.623, 1320.615, 1321.187]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.39232599152
varray: [1323.541, 1320.623, 1320.615, 1321.187]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.1445
varray: [0.881, 1.417, 1.096, 1.184]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.221811481518
varray: [0.881, 1.417, 1.096, 1.184]
inst: std_all
operation: std_all(slope_shift)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: slope_shift
SAMP_d18O_STD | 0.331003902293
varray: [5.508, 6.024, 6.164, 6.245]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.480755221847
varray: [-46.58, -45.959, -45.562, -45.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.16775
varray: [111.101, 108.098, 109.56, 107.912]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.48465357912
varray: [111.101, 108.098, 109.56, 107.912]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.153
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1321.4915
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.1445
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.54542338019
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.3894
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [5.508, 6.024, 6.164, 6.245]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-46.58, -45.959, -45.562, -45.562]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.98525
varray: [5.508, 6.024, 6.164, 6.245]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
dD_mean | -45.91575
varray: [-46.58, -45.959, -45.562, -45.562]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
d18O_stdev | 0.331003902293
varray: [5.508, 6.024, 6.164, 6.245]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
dD_stdev | 0.480755221847
varray: [-46.58, -45.959, -45.562, -45.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.98525
varray: [5.508, 6.024, 6.164, 6.245]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.91575
varray: [-46.58, -45.959, -45.562, -45.562]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.331003902293
varray: [5.508, 6.024, 6.164, 6.245]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.480755221847
varray: [-46.58, -45.959, -45.562, -45.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.2932061803
actual_array: [ 1.93 -15.84]
varray: [5.508, 6.024, 6.164, 6.245]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 5.98525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.3075457488
actual_array: [ 18.1 -119. ]
varray: [-46.58, -45.959, -45.562, -45.562]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.91575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [65694.0, 66231.0, 66768.0, 67303.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3326
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 19.31615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82475
varray: [2.868, 2.806, 2.84, 2.785]
inst: mean_all
operation: mean_all(residuals)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: residuals
SAMP_residuals_STD | 0.0366731054953
varray: [2.868, 2.806, 2.84, 2.785]
inst: std_all
operation: std_all(residuals)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.27875
varray: [0.315, 0.318, 0.373, 0.109]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.116265142383
varray: [0.315, 0.318, 0.373, 0.109]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1357.41375
varray: [1359.552, 1357.135, 1356.133, 1356.835]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.48605593771
varray: [1359.552, 1357.135, 1356.133, 1356.835]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.00025
varray: [0.168, -0.297, 0.444, -0.314]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.370659210057
varray: [0.168, -0.297, 0.444, -0.314]
inst: std_all
operation: std_all(slope_shift)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: slope_shift
SAMP_d18O_STD | 0.294906764927
varray: [0.199, -0.272, -0.413, -0.434]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.671219040254
varray: [-52.832, -53.754, -54.293, -54.215]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.86125
varray: [109.958, 109.088, 108.213, 108.186]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.842707293984
varray: [109.958, 109.088, 108.213, 108.186]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.27875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1357.41375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.00025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.16249571478
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6959
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.199, -0.272, -0.413, -0.434]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-52.832, -53.754, -54.293, -54.215]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.23
varray: [0.199, -0.272, -0.413, -0.434]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
dD_mean | -53.7735
varray: [-52.832, -53.754, -54.293, -54.215]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
d18O_stdev | 0.294906764927
varray: [0.199, -0.272, -0.413, -0.434]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
dD_stdev | 0.671219040254
varray: [-52.832, -53.754, -54.293, -54.215]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.23
varray: [0.199, -0.272, -0.413, -0.434]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.7735
varray: [-52.832, -53.754, -54.293, -54.215]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.294906764927
varray: [0.199, -0.272, -0.413, -0.434]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.671219040254
varray: [-52.832, -53.754, -54.293, -54.215]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.0493227055
actual_array: [ 1.93 -15.84]
varray: [0.199, -0.272, -0.413, -0.434]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.23
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.8512815557
actual_array: [ 18.1 -119. ]
varray: [-52.832, -53.754, -54.293, -54.215]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -53.7735
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [67891.0, 68426.0, 68961.0, 69498.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3506
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 55.2384
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79425
varray: [2.828, 2.798, 2.794, 2.757]
inst: mean_all
operation: mean_all(residuals)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: residuals
SAMP_residuals_STD | 0.0291018327029
varray: [2.828, 2.798, 2.794, 2.757]
inst: std_all
operation: std_all(residuals)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.23975
varray: [0.211, 0.332, 0.241, 0.175]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0671583948587
varray: [0.211, 0.332, 0.241, 0.175]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.77725
varray: [1298.063, 1295.921, 1295.544, 1293.581]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.83695733483
varray: [1298.063, 1295.921, 1295.544, 1293.581]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3385
varray: [1.77, 2.526, 2.829, 2.229]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.451268212929
varray: [1.77, 2.526, 2.829, 2.229]
inst: std_all
operation: std_all(slope_shift)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: slope_shift
SAMP_d18O_STD | 0.388331795419
varray: [-7.629, -8.247, -8.395, -8.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.72988908604
varray: [-72.894, -75.65, -76.266, -76.773]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.5485
varray: [109.416, 107.686, 107.436, 105.656]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.53847056087
varray: [109.416, 107.686, 107.436, 105.656]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.23975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.77725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3385
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.20809050394
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.00865
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.629, -8.247, -8.395, -8.494]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-72.894, -75.65, -76.266, -76.773]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.19125
varray: [-7.629, -8.247, -8.395, -8.494]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
dD_mean | -75.39575
varray: [-72.894, -75.65, -76.266, -76.773]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
d18O_stdev | 0.388331795419
varray: [-7.629, -8.247, -8.395, -8.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
dD_stdev | 1.72988908604
varray: [-72.894, -75.65, -76.266, -76.773]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.19125
varray: [-7.629, -8.247, -8.395, -8.494]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
measured_dD_mean | -75.39575
varray: [-72.894, -75.65, -76.266, -76.773]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.388331795419
varray: [-7.629, -8.247, -8.395, -8.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.72988908604
varray: [-72.894, -75.65, -76.266, -76.773]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.173607021
actual_array: [ 1.93 -15.84]
varray: [-7.629, -8.247, -8.395, -8.494]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.19125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -89.6094555305
actual_array: [ 18.1 -119. ]
varray: [-72.894, -75.65, -76.266, -76.773]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -75.39575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [70086.0, 70623.0, 71159.0, 71696.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3811
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.3981
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.823
varray: [2.768, 2.843, 2.831, 2.85]
inst: mean_all
operation: mean_all(residuals)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: residuals
SAMP_residuals_STD | 0.0374966665185
varray: [2.768, 2.843, 2.831, 2.85]
inst: std_all
operation: std_all(residuals)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.439
varray: [0.298, 0.3, 0.497, 0.661]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.174976188856
varray: [0.298, 0.3, 0.497, 0.661]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.68875
varray: [1293.822, 1297.565, 1295.932, 1295.436]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.54144118171
varray: [1293.822, 1297.565, 1295.932, 1295.436]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.201
varray: [2.188, 2.246, 2.205, 2.165]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0341857670188
varray: [2.188, 2.246, 2.205, 2.165]
inst: std_all
operation: std_all(slope_shift)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: slope_shift
SAMP_d18O_STD | 0.151108129056
varray: [-5.804, -5.51, -5.497, -5.499]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.05852365743
varray: [-65.628, -63.663, -63.524, -63.383]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.986
varray: [107.613, 109.55, 109.184, 109.597]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.933764781231
varray: [107.613, 109.55, 109.184, 109.597]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.823
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.439
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.68875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.201
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.094
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.094
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.22248885842
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.57115
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-5.804, -5.51, -5.497, -5.499]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-65.628, -63.663, -63.524, -63.383]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.5775
varray: [-5.804, -5.51, -5.497, -5.499]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
dD_mean | -64.0495
varray: [-65.628, -63.663, -63.524, -63.383]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
d18O_stdev | 0.151108129056
varray: [-5.804, -5.51, -5.497, -5.499]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
dD_stdev | 1.05852365743
varray: [-65.628, -63.663, -63.524, -63.383]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.5775
varray: [-5.804, -5.51, -5.497, -5.499]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.0495
varray: [-65.628, -63.663, -63.524, -63.383]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.151108129056
varray: [-5.804, -5.51, -5.497, -5.499]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.05852365743
varray: [-65.628, -63.663, -63.524, -63.383]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.50633138847
actual_array: [ 1.93 -15.84]
varray: [-5.804, -5.51, -5.497, -5.499]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.5775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.7166283982
actual_array: [ 18.1 -119. ]
varray: [-65.628, -63.663, -63.524, -63.383]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -64.0495
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [72284.0, 72820.0, 73357.0, 73893.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.35235
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.4866
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82375
varray: [2.849, 2.821, 2.806, 2.819]
inst: mean_all
operation: mean_all(residuals)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: residuals
SAMP_residuals_STD | 0.0180992633368
varray: [2.849, 2.821, 2.806, 2.819]
inst: std_all
operation: std_all(residuals)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31675
varray: [0.352, 0.416, 0.283, 0.216]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0863765979109
varray: [0.352, 0.416, 0.283, 0.216]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.45075
varray: [1297.674, 1297.177, 1297.177, 1297.775]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.31877722106
varray: [1297.674, 1297.177, 1297.177, 1297.775]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.9265
varray: [1.937, 2.358, 1.803, 1.608]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.317803398346
varray: [1.937, 2.358, 1.803, 1.608]
inst: std_all
operation: std_all(slope_shift)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: slope_shift
SAMP_d18O_STD | 0.222075625257
varray: [-0.921, -0.607, -0.487, -0.42]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.28656166454
varray: [-48.92, -46.848, -46.333, -46.094]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.4645
varray: [110.045, 108.876, 109.243, 109.694]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.511548303356
varray: [110.045, 108.876, 109.243, 109.694]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.45075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.9265
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09325
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09325
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.19677751114
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.09265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-0.921, -0.607, -0.487, -0.42]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-48.92, -46.848, -46.333, -46.094]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.60875
varray: [-0.921, -0.607, -0.487, -0.42]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
dD_mean | -47.04875
varray: [-48.92, -46.848, -46.333, -46.094]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
d18O_stdev | 0.222075625257
varray: [-0.921, -0.607, -0.487, -0.42]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
dD_stdev | 1.28656166454
varray: [-48.92, -46.848, -46.333, -46.094]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.60875
varray: [-0.921, -0.607, -0.487, -0.42]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
measured_dD_mean | -47.04875
varray: [-48.92, -46.848, -46.333, -46.094]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.222075625257
varray: [-0.921, -0.607, -0.487, -0.42]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.28656166454
varray: [-48.92, -46.848, -46.333, -46.094]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.43582893059
actual_array: [ 1.93 -15.84]
varray: [-0.921, -0.607, -0.487, -0.42]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.60875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -62.3952684009
actual_array: [ 18.1 -119. ]
varray: [-48.92, -46.848, -46.333, -46.094]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -47.04875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [74477.0, 75013.0, 75550.0, 76086.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3516
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.7246
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.80625
varray: [2.824, 2.815, 2.776, 2.81]
inst: mean_all
operation: mean_all(residuals)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: residuals
SAMP_residuals_STD | 0.0209821352584
varray: [2.824, 2.815, 2.776, 2.81]
inst: std_all
operation: std_all(residuals)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34275
varray: [0.354, 0.393, 0.139, 0.485]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.146515926324
varray: [0.354, 0.393, 0.139, 0.485]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.68975
varray: [1296.627, 1295.779, 1295.829, 1294.524]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.8688626186
varray: [1296.627, 1295.779, 1295.829, 1294.524]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.55975
varray: [2.596, 2.82, 2.483, 2.34]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.202669476406
varray: [2.596, 2.82, 2.483, 2.34]
inst: std_all
operation: std_all(slope_shift)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: slope_shift
SAMP_d18O_STD | 0.332281381764
varray: [-6.387, -6.98, -7.05, -7.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.70343907043
varray: [-64.698, -67.501, -68.107, -68.438]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.08975
varray: [108.99, 108.446, 107.23, 107.693]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.781866303575
varray: [108.99, 108.446, 107.23, 107.693]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.80625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.68975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.55975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.79670894755
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4674
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-6.387, -6.98, -7.05, -7.102]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-64.698, -67.501, -68.107, -68.438]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.87975
varray: [-6.387, -6.98, -7.05, -7.102]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
dD_mean | -67.186
varray: [-64.698, -67.501, -68.107, -68.438]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
d18O_stdev | 0.332281381764
varray: [-6.387, -6.98, -7.05, -7.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
dD_stdev | 1.70343907043
varray: [-64.698, -67.501, -68.107, -68.438]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.87975
varray: [-6.387, -6.98, -7.05, -7.102]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.186
varray: [-64.698, -67.501, -68.107, -68.438]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.332281381764
varray: [-6.387, -6.98, -7.05, -7.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.70343907043
varray: [-64.698, -67.501, -68.107, -68.438]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.83524949177
actual_array: [ 1.93 -15.84]
varray: [-6.387, -6.98, -7.05, -7.102]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.87975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.7277864725
actual_array: [ 18.1 -119. ]
varray: [-64.698, -67.501, -68.107, -68.438]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.186
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [76671.0, 77206.0, 77742.0, 78278.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3691
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.4856
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2232
varray: [0.0, 2.744, 2.81, 2.759, 2.803]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: residuals
SAMP_residuals_STD | 1.24312457139
varray: [0.0, 2.744, 2.81, 2.759, 2.803]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.274
varray: [0.0, 0.275, 0.388, 0.185, 0.522]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.198430088444
varray: [0.0, 0.275, 0.388, 0.185, 0.522]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1035.72
varray: [0.0, 1293.754, 1294.751, 1295.07, 1295.025]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 578.985325384
varray: [0.0, 1293.754, 1294.751, 1295.07, 1295.025]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.0156
varray: [0.0, 2.688, 2.486, 2.332, 2.572]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.13421020979
varray: [0.0, 2.688, 2.486, 2.332, 2.572]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: slope_shift
SAMP_d18O_STD | 3.30244913663
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 27.0825736351
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 85.681
varray: [0.0, 105.151, 107.788, 107.612, 107.854]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 47.9104497996
varray: [0.0, 105.151, 107.788, 107.612, 107.854]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2232
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.274
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1035.72
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.0156
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.6938
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.6938
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -23.7847103188
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -25.87615
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.9076
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
dD_mean | -48.4352
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
d18O_stdev | 3.30244913663
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
dD_stdev | 27.0825736351
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.9076
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.4352
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.30244913663
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
measured_dD_stdev | 27.0825736351
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.84319133541
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.386, -7.382, -7.385, -7.385]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.9076
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.7263125941
actual_array: [ 18.1 -119. ]
varray: [0.0, -61.513, -60.52, -60.162, -59.981]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -48.4352
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [78819.0, 79417.0, 79954.0, 80490.0, 81026.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.95215
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -266.45535
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7395
varray: [2.78, 2.753, 2.717, 2.708]
inst: mean_all
operation: mean_all(residuals)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: residuals
SAMP_residuals_STD | 0.0332716095192
varray: [2.78, 2.753, 2.717, 2.708]
inst: std_all
operation: std_all(residuals)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.60875
varray: [-0.49, -0.658, -0.482, -0.805]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15395535067
varray: [-0.49, -0.658, -0.482, -0.805]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2081.415
varray: [2080.669, 2084.666, 2081.129, 2079.196]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.31885560856
varray: [2080.669, 2084.666, 2081.129, 2079.196]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -27.06725
varray: [-26.872, -27.352, -27.194, -26.851]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.246329555677
varray: [-26.872, -27.352, -27.194, -26.851]
inst: std_all
operation: std_all(slope_shift)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: slope_shift
SAMP_d18O_STD | 0.158431899145
varray: [-3.795, -3.493, -3.494, -3.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.61325660752
varray: [-52.639, -51.615, -51.425, -51.287]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 101.11275
varray: [101.836, 101.962, 101.57, 99.083]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.36299654561
varray: [101.836, 101.962, 101.57, 99.083]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7395
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.60875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2081.415
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.06725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.08501885499
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.4444
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-3.795, -3.493, -3.494, -3.454]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-52.639, -51.615, -51.425, -51.287]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.559
varray: [-3.795, -3.493, -3.494, -3.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
dD_mean | -51.7415
varray: [-52.639, -51.615, -51.425, -51.287]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
d18O_stdev | 0.158431899145
varray: [-3.795, -3.493, -3.494, -3.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
dD_stdev | 0.61325660752
varray: [-52.639, -51.615, -51.425, -51.287]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.559
varray: [-3.795, -3.493, -3.494, -3.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.7415
varray: [-52.639, -51.615, -51.425, -51.287]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.158431899145
varray: [-3.795, -3.493, -3.494, -3.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.61325660752
varray: [-52.639, -51.615, -51.425, -51.287]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.44649557238
actual_array: [ 1.93 -15.84]
varray: [-3.795, -3.493, -3.494, -3.454]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.9004850606
actual_array: [ 18.1 -119. ]
varray: [-52.639, -51.615, -51.425, -51.287]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -51.7415
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [81612.0, 82147.0, 82684.0, 83220.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.43585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 779.23965
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8455
varray: [2.845, 2.854, 2.855, 2.828]
inst: mean_all
operation: mean_all(residuals)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: residuals
SAMP_residuals_STD | 0.012503332889
varray: [2.845, 2.854, 2.855, 2.828]
inst: std_all
operation: std_all(residuals)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.32875
varray: [0.306, 0.378, 0.293, 0.338]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0378890925377
varray: [0.306, 0.378, 0.293, 0.338]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1317.16025
varray: [1322.598, 1317.238, 1315.406, 1313.399]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.94966516513
varray: [1322.598, 1317.238, 1315.406, 1313.399]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.20525
varray: [1.033, 1.034, 0.94, 1.814]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.408220018291
varray: [1.033, 1.034, 0.94, 1.814]
inst: std_all
operation: std_all(slope_shift)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: slope_shift
SAMP_d18O_STD | 0.197748999155
varray: [-0.004, 0.333, 0.348, 0.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.127081010908
varray: [-50.25, -50.162, -50.021, -49.974]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.55975
varray: [109.089, 110.366, 110.152, 108.632]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.833291615622
varray: [109.089, 110.366, 110.152, 108.632]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8455
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.32875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1317.16025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.20525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0715
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0715
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.45114844018
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.9974
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-0.004, 0.333, 0.348, 0.451]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-50.25, -50.162, -50.021, -49.974]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.282
varray: [-0.004, 0.333, 0.348, 0.451]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
dD_mean | -50.10175
varray: [-50.25, -50.162, -50.021, -49.974]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
d18O_stdev | 0.197748999155
varray: [-0.004, 0.333, 0.348, 0.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
dD_stdev | 0.127081010908
varray: [-50.25, -50.162, -50.021, -49.974]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.282
varray: [-0.004, 0.333, 0.348, 0.451]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.10175
varray: [-50.25, -50.162, -50.021, -49.974]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.197748999155
varray: [-0.004, 0.333, 0.348, 0.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.127081010908
varray: [-50.25, -50.162, -50.021, -49.974]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.52683772267
actual_array: [ 1.93 -15.84]
varray: [-0.004, 0.333, 0.348, 0.451]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.3262633318
actual_array: [ 18.1 -119. ]
varray: [-50.25, -50.162, -50.021, -49.974]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -50.10175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [83806.0, 84342.0, 84878.0, 85414.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.32985
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 14.9849
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.73175
varray: [2.64, 2.733, 2.783, 2.771]
inst: mean_all
operation: mean_all(residuals)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: residuals
SAMP_residuals_STD | 0.064773322698
varray: [2.64, 2.733, 2.783, 2.771]
inst: std_all
operation: std_all(residuals)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29075
varray: [0.097, 0.228, 0.184, 0.654]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.24820874951
varray: [0.097, 0.228, 0.184, 0.654]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.92675
varray: [1293.078, 1294.244, 1294.749, 1293.636]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.72609796171
varray: [1293.078, 1294.244, 1294.749, 1293.636]
inst: std_all
operation: std_all(baseline_shift)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.6605
varray: [2.526, 3.282, 2.552, 2.282]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431812845262
varray: [2.526, 3.282, 2.552, 2.282]
inst: std_all
operation: std_all(slope_shift)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: slope_shift
SAMP_d18O_STD | 0.490866156231
varray: [-8.981, -9.802, -9.979, -10.043]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.68530966557
varray: [-81.241, -85.471, -86.633, -87.163]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.62125
varray: [104.073, 104.389, 105.983, 108.04]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.8162904274
varray: [104.073, 104.389, 105.983, 108.04]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.73175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.92675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.6605
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.18525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.18525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.35070277683
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.9359
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-8.981, -9.802, -9.979, -10.043]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-81.241, -85.471, -86.633, -87.163]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.70125
varray: [-8.981, -9.802, -9.979, -10.043]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
dD_mean | -85.127
varray: [-81.241, -85.471, -86.633, -87.163]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
d18O_stdev | 0.490866156231
varray: [-8.981, -9.802, -9.979, -10.043]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
dD_stdev | 2.68530966557
varray: [-81.241, -85.471, -86.633, -87.163]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.70125
varray: [-8.981, -9.802, -9.979, -10.043]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.127
varray: [-81.241, -85.471, -86.633, -87.163]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.490866156231
varray: [-8.981, -9.802, -9.979, -10.043]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.68530966557
varray: [-81.241, -85.471, -86.633, -87.163]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.714529529
actual_array: [ 1.93 -15.84]
varray: [-8.981, -9.802, -9.979, -10.043]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.70125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -98.9518218658
actual_array: [ 18.1 -119. ]
varray: [-81.241, -85.471, -86.633, -87.163]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -85.127
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [86000.0, 86536.0, 87071.0, 87607.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.4436
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.2486
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.68875
varray: [2.637, 2.641, 2.757, 2.72]
inst: mean_all
operation: mean_all(residuals)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: residuals
SAMP_residuals_STD | 0.0594215168661
varray: [2.637, 2.641, 2.757, 2.72]
inst: std_all
operation: std_all(residuals)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.40125
varray: [0.395, 0.496, 0.356, 0.358]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0656626479312
varray: [0.395, 0.496, 0.356, 0.358]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.99775
varray: [1289.9, 1288.733, 1293.392, 1291.966]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.08199333572
varray: [1289.9, 1288.733, 1293.392, 1291.966]
inst: std_all
operation: std_all(baseline_shift)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.12875
varray: [1.912, 2.582, 2.123, 1.898]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.319215052903
varray: [1.912, 2.582, 2.123, 1.898]
inst: std_all
operation: std_all(slope_shift)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: slope_shift
SAMP_d18O_STD | 0.11213979371
varray: [-8.022, -7.819, -7.807, -7.776]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.19128697914
varray: [-74.639, -72.834, -72.178, -72.057]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.427
varray: [104.119, 102.118, 105.934, 105.537]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.72525302492
varray: [104.119, 102.118, 105.934, 105.537]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.68875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.40125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.99775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.12875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.22825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.22825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.82482002057
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.13015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-8.022, -7.819, -7.807, -7.776]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-74.639, -72.834, -72.178, -72.057]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.856
varray: [-8.022, -7.819, -7.807, -7.776]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
dD_mean | -72.927
varray: [-74.639, -72.834, -72.178, -72.057]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
d18O_stdev | 0.11213979371
varray: [-8.022, -7.819, -7.807, -7.776]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
dD_stdev | 1.19128697914
varray: [-74.639, -72.834, -72.178, -72.057]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.856
varray: [-8.022, -7.819, -7.807, -7.776]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.927
varray: [-74.639, -72.834, -72.178, -72.057]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.11213979371
varray: [-8.022, -7.819, -7.807, -7.776]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.19128697914
varray: [-74.639, -72.834, -72.178, -72.057]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.8314916099
actual_array: [ 1.93 -15.84]
varray: [-8.022, -7.819, -7.807, -7.776]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.856
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.2393625938
actual_array: [ 18.1 -119. ]
varray: [-74.639, -72.834, -72.178, -72.057]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -72.927
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [88195.0, 88731.0, 89267.0, 89803.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.4866
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.1776
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.75425
varray: [2.827, 2.745, 2.692, 2.753]
inst: mean_all
operation: mean_all(residuals)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: residuals
SAMP_residuals_STD | 0.0555420261304
varray: [2.827, 2.745, 2.692, 2.753]
inst: std_all
operation: std_all(residuals)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.1135
varray: [0.13, 0.184, -0.019, 0.159]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0910476065949
varray: [0.13, 0.184, -0.019, 0.159]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.8505
varray: [1300.751, 1296.713, 1296.841, 1297.097]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.94024560301
varray: [1300.751, 1296.713, 1296.841, 1297.097]
inst: std_all
operation: std_all(baseline_shift)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.972
varray: [1.884, 2.112, 2.226, 1.666]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.248660411003
varray: [1.884, 2.112, 2.226, 1.666]
inst: std_all
operation: std_all(slope_shift)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: slope_shift
SAMP_d18O_STD | 0.171698136274
varray: [-3.989, -3.773, -3.66, -3.599]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.00687387492
varray: [-58.139, -56.858, -56.191, -55.862]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.44125
varray: [108.888, 105.709, 104.654, 106.514]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.80021579725
varray: [108.888, 105.709, 104.654, 106.514]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.75425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.1135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.8505
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.972
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.57936235859
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.1159
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-3.989, -3.773, -3.66, -3.599]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-58.139, -56.858, -56.191, -55.862]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.75525
varray: [-3.989, -3.773, -3.66, -3.599]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
dD_mean | -56.7625
varray: [-58.139, -56.858, -56.191, -55.862]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
d18O_stdev | 0.171698136274
varray: [-3.989, -3.773, -3.66, -3.599]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
dD_stdev | 1.00687387492
varray: [-58.139, -56.858, -56.191, -55.862]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.75525
varray: [-3.989, -3.773, -3.66, -3.599]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.7625
varray: [-58.139, -56.858, -56.191, -55.862]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.171698136274
varray: [-3.989, -3.773, -3.66, -3.599]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.00687387492
varray: [-58.139, -56.858, -56.191, -55.862]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.64676447449
actual_array: [ 1.93 -15.84]
varray: [-3.989, -3.773, -3.66, -3.599]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.75525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.7208340773
actual_array: [ 18.1 -119. ]
varray: [-58.139, -56.858, -56.191, -55.862]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.7625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [90386.0, 90923.0, 91458.0, 91994.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.4211
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.32485
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.71775
varray: [2.669, 2.746, 2.719, 2.737]
inst: mean_all
operation: mean_all(residuals)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: residuals
SAMP_residuals_STD | 0.0343838624939
varray: [2.669, 2.746, 2.719, 2.737]
inst: std_all
operation: std_all(residuals)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.36475
varray: [0.253, 0.582, 0.349, 0.275]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.150542076067
varray: [0.253, 0.582, 0.349, 0.275]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.5505
varray: [1289.986, 1292.112, 1291.37, 1292.734]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.18267986652
varray: [1289.986, 1292.112, 1291.37, 1292.734]
inst: std_all
operation: std_all(baseline_shift)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3325
varray: [1.987, 2.528, 2.327, 2.488]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.246171349538
varray: [1.987, 2.528, 2.327, 2.488]
inst: std_all
operation: std_all(slope_shift)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: slope_shift
SAMP_d18O_STD | 0.200024165207
varray: [-8.119, -8.464, -8.538, -8.537]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.50717008551
varray: [-74.979, -77.338, -78.157, -78.176]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.27625
varray: [103.686, 106.433, 105.087, 105.899]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.19590952138
varray: [103.686, 106.433, 105.087, 105.899]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.71775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.5505
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.19925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.19925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.83064792595
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.2809
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-8.119, -8.464, -8.538, -8.537]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-74.979, -77.338, -78.157, -78.176]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.4145
varray: [-8.119, -8.464, -8.538, -8.537]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
dD_mean | -77.1625
varray: [-74.979, -77.338, -78.157, -78.176]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
d18O_stdev | 0.200024165207
varray: [-8.119, -8.464, -8.538, -8.537]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
dD_stdev | 1.50717008551
varray: [-74.979, -77.338, -78.157, -78.176]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.4145
varray: [-8.119, -8.464, -8.538, -8.537]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.1625
varray: [-74.979, -77.338, -78.157, -78.176]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.200024165207
varray: [-8.119, -8.464, -8.538, -8.537]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.50717008551
varray: [-74.979, -77.338, -78.157, -78.176]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.4014288422
actual_array: [ 1.93 -15.84]
varray: [-8.119, -8.464, -8.538, -8.537]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.4145
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.3056020402
actual_array: [ 18.1 -119. ]
varray: [-74.979, -77.338, -78.157, -78.176]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -77.1625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [92578.0, 93114.0, 93649.0, 94185.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.4576
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.62485
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.364
varray: [0.0, 0.0, 2.791, 2.665]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: residuals
SAMP_residuals_STD | 1.57585130432
varray: [0.0, 0.0, 2.791, 2.665]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.12675
varray: [0.0, 0.0, 0.179, 0.328]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.158495793845
varray: [0.0, 0.0, 0.179, 0.328]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 646.31325
varray: [0.0, 0.0, 1295.471, 1289.782]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 746.301871645
varray: [0.0, 0.0, 1295.471, 1289.782]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.19825
varray: [0.0, 0.0, 2.276, 2.517]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.3871136399
varray: [0.0, 0.0, 2.276, 2.517]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: slope_shift
SAMP_d18O_STD | 3.93031367493
varray: [0.0, 0.0, -6.812, -6.803]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 40.1196675418
varray: [0.0, 0.0, -69.344, -69.634]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 52.9355
varray: [0.0, 0.0, 108.227, 103.515]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 61.1549128934
varray: [0.0, 0.0, 108.227, 103.515]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.364
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.12675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 646.31325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.19825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.553
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.553
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -53.2396297566
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -58.62165
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, -6.812, -6.803]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -69.344, -69.634]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.40375
varray: [0.0, 0.0, -6.812, -6.803]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
dD_mean | -34.7445
varray: [0.0, 0.0, -69.344, -69.634]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
d18O_stdev | 3.93031367493
varray: [0.0, 0.0, -6.812, -6.803]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
dD_stdev | 40.1196675418
varray: [0.0, 0.0, -69.344, -69.634]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.40375
varray: [0.0, 0.0, -6.812, -6.803]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
measured_dD_mean | -34.7445
varray: [0.0, 0.0, -69.344, -69.634]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.93031367493
varray: [0.0, 0.0, -6.812, -6.803]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
measured_dD_stdev | 40.1196675418
varray: [0.0, 0.0, -69.344, -69.634]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.28806628803
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -6.812, -6.803]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.40375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -50.5827252045
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -69.344, -69.634]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -34.7445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [94769.0, 95305.0, 95840.0, 96376.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.81135
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -655.8621
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.92275
varray: [1.491, 1.054, 1.146, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: residuals
SAMP_residuals_STD | 0.643283957518
varray: [1.491, 1.054, 1.146, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.06425
varray: [0.01, 0.191, 0.056, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0879483750087
varray: [0.01, 0.191, 0.056, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 922.61675
varray: [1240.884, 1222.773, 1226.81, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 615.126820952
varray: [1240.884, 1222.773, 1226.81, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.11825
varray: [1.844, 1.234, 1.395, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.788920095236
varray: [1.844, 1.234, 1.395, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: slope_shift
SAMP_d18O_STD | 2.78421796321
varray: [-5.489, -5.605, -5.608, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 33.7662409368
varray: [-67.125, -67.997, -67.464, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 37.78525
varray: [59.759, 43.853, 47.529, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 26.0916912749
varray: [59.759, 43.853, 47.529, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.92275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.06425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 922.61675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.11825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.99425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.99425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -68.3664724032
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -73.7719
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-5.489, -5.605, -5.608, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-67.125, -67.997, -67.464, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.1755
varray: [-5.489, -5.605, -5.608, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
dD_mean | -50.6465
varray: [-67.125, -67.997, -67.464, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
d18O_stdev | 2.78421796321
varray: [-5.489, -5.605, -5.608, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
dD_stdev | 33.7662409368
varray: [-67.125, -67.997, -67.464, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.1755
varray: [-5.489, -5.605, -5.608, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.6465
varray: [-67.125, -67.997, -67.464, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.78421796321
varray: [-5.489, -5.605, -5.608, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
measured_dD_stdev | 33.7662409368
varray: [-67.125, -67.997, -67.464, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.07562055658
actual_array: [ 1.93 -15.84]
varray: [-5.489, -5.605, -5.608, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.1755
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.8492438391
actual_array: [ 18.1 -119. ]
varray: [-67.125, -67.997, -67.464, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -50.6465
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [96960.0, 97495.0, 98030.0, 98566.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.2526
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -379.5586
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.672
varray: [1.344, 0.0, 0.0, 1.344]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: residuals
SAMP_residuals_STD | 0.775958761791
varray: [1.344, 0.0, 0.0, 1.344]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0405
varray: [0.081, 0.0, 0.0, 0.081]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0467653718044
varray: [0.081, 0.0, 0.0, 0.081]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 617.8225
varray: [1235.645, 0.0, 0.0, 1235.645]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 713.399973373
varray: [1235.645, 0.0, 0.0, 1235.645]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.8445
varray: [1.689, 0.0, 0.0, 1.689]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.975144604661
varray: [1.689, 0.0, 0.0, 1.689]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: slope_shift
SAMP_d18O_STD | 4.26661848931
varray: [-7.39, 0.0, 0.0, -7.39]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 37.5537482594
varray: [-65.045, 0.0, 0.0, -65.045]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 27.3975
varray: [54.795, 0.0, 0.0, 54.795]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 31.6359080002
varray: [54.795, 0.0, 0.0, 54.795]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.672
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0405
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 617.8225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.8445
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.245
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.245
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -76.962632842
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -84.15965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.39, 0.0, 0.0, -7.39]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-65.045, 0.0, 0.0, -65.045]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.695
varray: [-7.39, 0.0, 0.0, -7.39]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
dD_mean | -32.5225
varray: [-65.045, 0.0, 0.0, -65.045]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
d18O_stdev | 4.26661848931
varray: [-7.39, 0.0, 0.0, -7.39]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
dD_stdev | 37.5537482594
varray: [-65.045, 0.0, 0.0, -65.045]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.695
varray: [-7.39, 0.0, 0.0, -7.39]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.5225
varray: [-65.045, 0.0, 0.0, -65.045]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.26661848931
varray: [-7.39, 0.0, 0.0, -7.39]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
measured_dD_stdev | 37.5537482594
varray: [-65.045, 0.0, 0.0, -65.045]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.58528064594
actual_array: [ 1.93 -15.84]
varray: [-7.39, 0.0, 0.0, -7.39]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.4495215567
actual_array: [ 18.1 -119. ]
varray: [-65.045, 0.0, 0.0, -65.045]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -32.5225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [99212.0, 99748.0, 100346.0, 100944.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.50335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -684.35285
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.642
varray: [0.0, 0.0, 1.284, 1.284]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: residuals
SAMP_residuals_STD | 0.741317745639
varray: [0.0, 0.0, 1.284, 1.284]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0285
varray: [0.0, 0.0, -0.057, -0.057]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0329089653438
varray: [0.0, 0.0, -0.057, -0.057]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 616.8655
varray: [0.0, 0.0, 1233.731, 1233.731]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 712.294924958
varray: [0.0, 0.0, 1233.731, 1233.731]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.9845
varray: [0.0, 0.0, 1.969, 1.969]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.13680268003
varray: [0.0, 0.0, 1.969, 1.969]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: slope_shift
SAMP_d18O_STD | 5.47039380057
varray: [0.0, 0.0, -9.475, -9.475]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 47.515349804
varray: [0.0, 0.0, -82.299, -82.299]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 26.2075
varray: [0.0, 0.0, 52.415, 52.415]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 30.2618143596
varray: [0.0, 0.0, 52.415, 52.415]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.642
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0285
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 616.8655
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.9845
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -77.9910867329
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -85.34965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, -9.475, -9.475]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, -82.299, -82.299]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.7375
varray: [0.0, 0.0, -9.475, -9.475]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
dD_mean | -41.1495
varray: [0.0, 0.0, -82.299, -82.299]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
d18O_stdev | 5.47039380057
varray: [0.0, 0.0, -9.475, -9.475]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
dD_stdev | 47.515349804
varray: [0.0, 0.0, -82.299, -82.299]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.7375
varray: [0.0, 0.0, -9.475, -9.475]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.1495
varray: [0.0, 0.0, -82.299, -82.299]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.47039380057
varray: [0.0, 0.0, -9.475, -9.475]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
measured_dD_stdev | 47.515349804
varray: [0.0, 0.0, -82.299, -82.299]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.64912946352
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -9.475, -9.475]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.7375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.7317663222
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -82.299, -82.299]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -41.1495
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [101591.0, 102127.0, 102725.0, 103323.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.53335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -685.30985
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.917
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.917
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -111.55715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [103971.0, 104570.0, 105169.0, 105768.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1302.17535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1302.17535
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.917
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.917
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -111.55715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [106414.0, 107011.0, 107610.0, 108208.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1302.17535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1302.17535
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.917
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.917
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -111.55715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [108856.0, 109454.0, 110052.0, 110650.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1302.17535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1302.17535
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.917
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.917
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -111.55715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [111298.0, 111896.0, 112493.0, 113092.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1302.17535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1302.17535
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.917
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.917
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -111.55715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [113741.0, 114339.0, 114937.0, 115535.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1302.17535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1302.17535
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.23375
varray: [4.935, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: residuals
SAMP_residuals_STD | 2.4675
varray: [4.935, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.8285
varray: [-11.314, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 5.657
varray: [-11.314, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 628.84275
varray: [2515.371, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1257.6855
varray: [2515.371, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -18.4085
varray: [-73.634, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: slope_shift
SAMP_slope_shift_STD | 36.817
varray: [-73.634, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: slope_shift
SAMP_d18O_STD | 7.992
varray: [15.984, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 20.1195
varray: [-40.239, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 21.37225
varray: [85.489, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 42.7445
varray: [85.489, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.23375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.8285
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 628.84275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 18.4085
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.68325
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.68325
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -57.7048337333
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -90.1849
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [15.984, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-40.239, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.996
varray: [15.984, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
dD_mean | -10.05975
varray: [-40.239, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
d18O_stdev | 7.992
varray: [15.984, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
dD_stdev | 20.1195
varray: [-40.239, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.996
varray: [15.984, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
measured_dD_mean | -10.05975
varray: [-40.239, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 7.992
varray: [15.984, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
measured_dD_stdev | 20.1195
varray: [-40.239, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.2632193598
actual_array: [ 1.93 -15.84]
varray: [15.984, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 3.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.8844359411
actual_array: [ 18.1 -119. ]
varray: [-40.239, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -10.05975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [116184.0, 116719.0, 117318.0, 117915.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.9416
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -673.3326
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.52325
varray: [2.093, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: residuals
SAMP_residuals_STD | 1.0465
varray: [2.093, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.157
varray: [-0.628, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.314
varray: [-0.628, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 324.08125
varray: [1296.325, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 648.1625
varray: [1296.325, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.16375
varray: [-0.655, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.3275
varray: [-0.655, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: slope_shift
SAMP_d18O_STD | 1.657
varray: [-3.314, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 31.1465
varray: [-62.293, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 21.273
varray: [85.092, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 42.546
varray: [85.092, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.52325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.157
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 324.08125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.16375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.39375
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.39375
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -82.0620500514
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -90.28415
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-3.314, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-62.293, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.8285
varray: [-3.314, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
dD_mean | -15.57325
varray: [-62.293, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
d18O_stdev | 1.657
varray: [-3.314, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
dD_stdev | 31.1465
varray: [-62.293, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.8285
varray: [-3.314, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
measured_dD_mean | -15.57325
varray: [-62.293, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.657
varray: [-3.314, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
measured_dD_stdev | 31.1465
varray: [-62.293, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.66007907703
actual_array: [ 1.93 -15.84]
varray: [-3.314, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.8285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.1776034982
actual_array: [ 18.1 -119. ]
varray: [-62.293, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -15.57325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [118564.0, 119099.0, 119698.0, 120295.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.6521
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -978.0941
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.564
varray: [2.132, 1.643, 1.281, 1.2]
inst: mean_all
operation: mean_all(residuals)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: residuals
SAMP_residuals_STD | 0.424833300641
varray: [2.132, 1.643, 1.281, 1.2]
inst: std_all
operation: std_all(residuals)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.08525
varray: [0.175, 0.107, -0.033, 0.092]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0867116870247
varray: [0.175, 0.107, -0.033, 0.092]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1245.1795
varray: [1268.564, 1248.116, 1234.185, 1229.853]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 17.4283151318
varray: [1268.564, 1248.116, 1234.185, 1229.853]
inst: std_all
operation: std_all(baseline_shift)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.57025
varray: [1.894, 1.446, 1.603, 1.338]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.241705571581
varray: [1.894, 1.446, 1.603, 1.338]
inst: std_all
operation: std_all(slope_shift)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: slope_shift
SAMP_d18O_STD | 0.368747768355
varray: [-7.154, -7.783, -7.768, -8.016]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.85707823834
varray: [-62.178, -65.113, -65.298, -66.568]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.995
varray: [84.175, 66.283, 52.349, 49.173]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 15.9560682709
varray: [84.175, 66.283, 52.349, 49.173]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.564
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.08525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1245.1795
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.57025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.353
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.353
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -46.3832704834
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -48.56215
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.154, -7.783, -7.768, -8.016]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-62.178, -65.113, -65.298, -66.568]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.68025
varray: [-7.154, -7.783, -7.768, -8.016]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
dD_mean | -64.78925
varray: [-62.178, -65.113, -65.298, -66.568]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
d18O_stdev | 0.368747768355
varray: [-7.154, -7.783, -7.768, -8.016]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
dD_stdev | 1.85707823834
varray: [-62.178, -65.113, -65.298, -66.568]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.68025
varray: [-7.154, -7.783, -7.768, -8.016]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.78925
varray: [-62.178, -65.113, -65.298, -66.568]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.368747768355
varray: [-7.154, -7.783, -7.768, -8.016]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.85707823834
varray: [-62.178, -65.113, -65.298, -66.568]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.6521425167
actual_array: [ 1.93 -15.84]
varray: [-7.154, -7.783, -7.768, -8.016]
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.68025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.4268162463
actual_array: [ 18.1 -119. ]
varray: [-62.178, -65.113, -65.298, -66.568]
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -64.78925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [120943.0, 121478.0, 122014.0, 122550.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.61135
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -56.99585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.687
varray: [3.374, 0.0, 0.0, 3.374]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: residuals
SAMP_residuals_STD | 1.94797980825
varray: [3.374, 0.0, 0.0, 3.374]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.9005
varray: [-7.801, 0.0, 0.0, -7.801]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 4.50390944995
varray: [-7.801, 0.0, 0.0, -7.801]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1181.909
varray: [2363.818, 0.0, 0.0, 2363.818]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1364.75095862
varray: [2363.818, 0.0, 0.0, 2363.818]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -29.5815
varray: [-59.163, 0.0, 0.0, -59.163]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: slope_shift
SAMP_slope_shift_STD | 34.1577739761
varray: [-59.163, 0.0, 0.0, -59.163]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: slope_shift
SAMP_d18O_STD | 6.69495372152
varray: [11.596, 0.0, 0.0, 11.596]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 35.3396099771
varray: [-61.21, 0.0, 0.0, -61.21]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 30.435
varray: [60.87, 0.0, 0.0, 60.87]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 35.1433108856
varray: [60.87, 0.0, 0.0, 60.87]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.687
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.9005
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1181.909
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.5815
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.23
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.23
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -42.1666095303
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -81.12215
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [11.596, 0.0, 0.0, 11.596]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-61.21, 0.0, 0.0, -61.21]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.798
varray: [11.596, 0.0, 0.0, 11.596]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
dD_mean | -30.605
varray: [-61.21, 0.0, 0.0, -61.21]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
d18O_stdev | 6.69495372152
varray: [11.596, 0.0, 0.0, 11.596]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
dD_stdev | 35.3396099771
varray: [-61.21, 0.0, 0.0, -61.21]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.798
varray: [11.596, 0.0, 0.0, 11.596]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
measured_dD_mean | -30.605
varray: [-61.21, 0.0, 0.0, -61.21]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 6.69495372152
varray: [11.596, 0.0, 0.0, 11.596]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
measured_dD_stdev | 35.3396099771
varray: [-61.21, 0.0, 0.0, -61.21]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.10212158453
actual_array: [ 1.93 -15.84]
varray: [11.596, 0.0, 0.0, 11.596]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 5.798
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -46.608649372
actual_array: [ 18.1 -119. ]
varray: [-61.21, 0.0, 0.0, -61.21]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -30.605
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [123135.0, 123670.0, 124268.0, 124866.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.48835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -120.26635
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.32
varray: [2.64, 0.0, 0.0, 2.64]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: residuals
SAMP_residuals_STD | 1.52420471066
varray: [2.64, 0.0, 0.0, 2.64]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.8265
varray: [-5.653, 0.0, 0.0, -5.653]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 3.26376107173
varray: [-5.653, 0.0, 0.0, -5.653]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 984.814
varray: [1969.628, 0.0, 0.0, 1969.628]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1137.165256
varray: [1969.628, 0.0, 0.0, 1969.628]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -19.825
varray: [-39.65, 0.0, 0.0, -39.65]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: slope_shift
SAMP_slope_shift_STD | 22.8919381734
varray: [-39.65, 0.0, 0.0, -39.65]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: slope_shift
SAMP_d18O_STD | 1.78112558045
varray: [3.085, 0.0, 0.0, 3.085]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 52.765773152
varray: [-91.393, 0.0, 0.0, -91.393]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 31.374
varray: [62.748, 0.0, 0.0, 62.748]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 36.2275746911
varray: [62.748, 0.0, 0.0, 62.748]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.32
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.8265
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 984.814
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 19.825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.597
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.597
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -54.7480287967
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -80.18315
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [3.085, 0.0, 0.0, 3.085]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-91.393, 0.0, 0.0, -91.393]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.5425
varray: [3.085, 0.0, 0.0, 3.085]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
dD_mean | -45.6965
varray: [-91.393, 0.0, 0.0, -91.393]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
d18O_stdev | 1.78112558045
varray: [3.085, 0.0, 0.0, 3.085]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
dD_stdev | 52.765773152
varray: [-91.393, 0.0, 0.0, -91.393]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.5425
varray: [3.085, 0.0, 0.0, 3.085]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.6965
varray: [-91.393, 0.0, 0.0, -91.393]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.78112558045
varray: [3.085, 0.0, 0.0, 3.085]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
measured_dD_stdev | 52.765773152
varray: [-91.393, 0.0, 0.0, -91.393]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.240524596
actual_array: [ 1.93 -15.84]
varray: [3.085, 0.0, 0.0, 3.085]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 1.5425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.0970574951
actual_array: [ 18.1 -119. ]
varray: [-91.393, 0.0, 0.0, -91.393]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.6965
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [125514.0, 126049.0, 126646.0, 127241.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.85535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -317.36135
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 52 - Sample 52 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 52
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.4175
varray: [1.67, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: residuals
SAMP_residuals_STD | 0.835
varray: [1.67, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.081
varray: [-0.324, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.162
varray: [-0.324, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 317.06775
varray: [1268.271, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 634.1355
varray: [1268.271, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.08925
varray: [0.357, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.1785
varray: [0.357, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: slope_shift
SAMP_d18O_STD | 3.364
varray: [-6.728, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 42.35
varray: [-84.7, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 16.9885
varray: [67.954, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.977
varray: [67.954, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.4175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.081
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 317.06775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.08925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.4995
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.4995
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -85.6873500171
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -94.56865
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-6.728, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-84.7, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.682
varray: [-6.728, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
dD_mean | -21.175
varray: [-84.7, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
d18O_stdev | 3.364
varray: [-6.728, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
dD_stdev | 42.35
varray: [-84.7, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.682
varray: [-6.728, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.175
varray: [-84.7, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.364
varray: [-6.728, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
measured_dD_stdev | 42.35
varray: [-84.7, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.5310574615
actual_array: [ 1.93 -15.84]
varray: [-6.728, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.682
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.5554943774
actual_array: [ 18.1 -119. ]
varray: [-84.7, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -21.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [127889.0, 128425.0, 129023.0, 129621.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.75785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -985.1076
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 53 - Sample 53 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Sample Variables
Sample 53
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.68125
varray: [1.696, 1.029, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: residuals
SAMP_residuals_STD | 0.832436333902
varray: [1.696, 1.029, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.02525
varray: [0.08, 0.021, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0378186461947
varray: [0.08, 0.021, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 618.71475
varray: [1250.344, 1224.515, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 714.508067622
varray: [1250.344, 1224.515, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.78675
varray: [1.811, 1.336, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.928926755275
varray: [1.811, 1.336, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: slope_shift
SAMP_d18O_STD | 3.96767564198
varray: [-7.605, -5.935, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 37.8942966982
varray: [-67.755, -63.368, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 27.523
varray: [67.118, 42.974, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.274256736
varray: [67.118, 42.974, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.68125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.02525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 618.71475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.78675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.23575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.23575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -76.6455262256
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -84.03415
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.764
varray: [-7.605, -5.935, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -106.0097
varray: [-67.755, -63.368, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.385
varray: [-7.605, -5.935, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
dD_mean | -32.78075
varray: [-67.755, -63.368, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
d18O_stdev | 3.96767564198
varray: [-7.605, -5.935, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
dD_stdev | 37.8942966982
varray: [-67.755, -63.368, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.385
varray: [-7.605, -5.935, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.78075
varray: [-67.755, -63.368, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.96767564198
varray: [-7.605, -5.935, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
measured_dD_stdev | 37.8942966982
varray: [-67.755, -63.368, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.26893231649
actual_array: [ 1.93 -15.84]
varray: [-7.605, -5.935, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.75813033202
except: True
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.385
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.6974512786
actual_array: [ 18.1 -119. ]
varray: [-67.755, -63.368, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.9437726477
except: True
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -32.78075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: [130270.0, 130806.0, 131341.0, 131939.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.4941
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -683.4606
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.16
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.464
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1417170685
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.16
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9547145831
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.464
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.714
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.583
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6865836655
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.714
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2687708367
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.583
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.977
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5498395489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.7269503784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.977
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.447
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.101
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4141159107
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.447
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.8859574012
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.101
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.094
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4661603133
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.8792371377
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.094
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.55
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.128
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5192251944
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.55
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9118784176
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.128
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.837
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5008565817
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6325074629
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.837
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.465
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.652
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4324845234
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4549004985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.652
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.468
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.578
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4355459589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.468
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3838577128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.578
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.493
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.578
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4610579209
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.493
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3838577128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.578
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.628
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.887683278395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 3.628
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.49739382341
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 23.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.381
operation: col(d(18_16)Mean)
inst: col
dD_actual | 33.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.65610357541
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.381
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 14.9086428597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 33.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.707
operation: col(d(18_16)Mean)
inst: col
dD_actual | 36.657
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.98877956057
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.707
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.9654027221
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 36.657
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.755
operation: col(d(18_16)Mean)
inst: col
dD_actual | 37.837
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03776252771
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.755
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.0982471435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 37.837
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.782
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.308
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06531544672
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.782
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5504248744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 38.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.789
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0724587961
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.789
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.9181192926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 38.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.804
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.136
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08776597333
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.804
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3453360447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 39.136
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.823
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.507
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10715506449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.823
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.7015100111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 39.507
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.861
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.14593324681
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.861
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.1892091349
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 40.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.964
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.678
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6494
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.7972
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.678
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.25104253047
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.8257140936
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 40.678
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.155
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.367
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.1980501223
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -11.155
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.1822309006
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -85.367
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.556
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.233
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6277404757
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -12.556
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.374225825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -102.233
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.751
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8267337797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -12.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.574031297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -105.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.638
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9553140685
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -12.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.563229298
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -107.638
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.973
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.64
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0532800028
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -12.973
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.525187018
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -108.64
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.996
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.484
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0767510078
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -12.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.335458791
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -109.484
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.055
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1369592383
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -13.055
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.516905906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -109.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.486
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1992084257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -13.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.297416511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -110.486
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.152
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.122
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.2359456511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -13.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.908000454
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -111.122
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.009
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.888
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0900172281
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -13.009
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.723314
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -109.888
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.94
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.516
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9172118024
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.94
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0445988954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.516
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.654
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6253549565
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.654
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.9087567197
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.198
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4794265336
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9790810528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.198
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.398
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.314
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3641124651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.398
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1304077744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.835
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4794265336
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6305873876
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.835
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.455
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.319
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4222797386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1352079626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.319
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.011
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3926858626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8395163679
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.011
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.408
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.755
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3743172499
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.408
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5937467307
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -59.755
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.388
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.938
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3539076803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.388
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7694336198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -59.938
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.209
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4212592601
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0296038217
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.209
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5499084613
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.3472460364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -74.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.744
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.38
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.737676502
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.744
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.5543725828
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -76.38
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.6703249222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7607806766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -76.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.701
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.815
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.6937959273
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.701
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9719889585
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -76.815
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.514
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.142
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.5234449332
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.514
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.1260718495
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -81.142
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.664
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.044
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.6765167055
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.664
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.9920258055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -82.044
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.5979398624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.8710610622
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -81.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.704
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.169
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.7173358448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.704
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.1120305111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -82.169
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.679
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.55997542123
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.1211603221
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -69.679
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.503
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.4589480515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9921560513
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.503
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.507
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.12
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.45486613757
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.507
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.6244616332
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.12
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.505
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.547
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.45282518061
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.505
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.0343977077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.547
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.029
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.984
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.92612045896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.029
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0134071253
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -54.984
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.729
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.766
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.61997691433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.729
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.8840436281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -52.766
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.714
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.368
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6046697371
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.714
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5019486453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -52.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.694
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.343
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58426016746
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.694
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.4779477042
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -52.343
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.423
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.114
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.8832700105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 13.423
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.1791155049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -79.114
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.981
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.8863314459
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 13.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.0114681433
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -79.981
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.992
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.0112336477
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.992
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.788373881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -87.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.992
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.0112336477
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.992
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.788373881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -87.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.577
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.212
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.48534266714
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.577
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.6728227418
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -69.212
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.1
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.666
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.99857443119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.1
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.228566897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -66.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.033
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.069
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.93020237289
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.033
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6554244228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -66.069
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.97
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.684
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.86591222851
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.97
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.2858099294
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -65.684
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.682
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.606
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.57201442567
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.682
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5705881859
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.606
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.714
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.698
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6046697371
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.714
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.698874004
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -55.698
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.686
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.476
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5760963396
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.686
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.4857456467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -55.476
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.72
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.266
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.61079260799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.72
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2841377412
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -55.266
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.043
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.512
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.001842604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.043
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3604575829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -59.512
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.317
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2814537081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.317
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.028643784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.356
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.905
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3212523689
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.356
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7377523775
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -59.905
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.453
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.153
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4202387816
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.453
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9758417135
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.153
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.304
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.051
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.14531659527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.304
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.3974284856
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -47.051
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.806
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.361
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.63711831119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.806
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.7749648651
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.361
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.679
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.815
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.50751754396
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.679
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.2507843108
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -44.815
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.662
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.42894070084
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.1038985511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -44.662
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.62
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.398
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.46778779561
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.62
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.4108627104
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -55.398
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.697
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.995
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.54636463873
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.697
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9440428299
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.995
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.715
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.373
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.5647332514
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.715
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3069370598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -57.373
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.721
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.5708561223
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.721
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.2637353657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -57.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.176
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.331
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.09660927791
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.176
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0668037049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -62.331
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.549
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.558
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.47724775173
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.549
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2447698956
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.558
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.703
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.52214880494
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.3839753541
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.703
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.639
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.847
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.56909081512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.639
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.522220775
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.847
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.93639415622
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.9447549803
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.929
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.951
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.84455109284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.929
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.4622152711
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.951
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.979
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.996
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.89557501694
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.979
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5054169651
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.996
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.956
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.074
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.87210401185
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -4.956
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5802999015
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.074
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.24
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.62
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.51221610943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.24
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5840287129
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.62
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.069
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.51119563095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 4.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0550479704
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.069
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.043
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99776069004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.3507269607
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -73.043
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.011
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.108
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.96918729254
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.011
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.4131294076
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -73.108
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.87
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.495
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.82529982657
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.87
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9844757501
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.495
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.783
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.722
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73651819863
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.783
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.2423666504
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.722
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.762
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.515
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7150881505
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.762
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0436388578
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.515
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.786
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73957963407
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.786
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8641118181
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.13
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0406207863
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5940996548
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -69.13
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.908
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.583
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.86407800889
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1089214177
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.583
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.394
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.95796202924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.887511948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.394
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.025
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.98347399129
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9451142067
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.435
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.438
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.401870169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2094900877
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.438
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.464
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.664
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4314640449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.464
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4664209502
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.664
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.397
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.181
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3630919866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.397
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0027227676
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.181
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.423
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.382
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3896244272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.423
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1956903343
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.382
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.315
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17702034066
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -2.315
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.9057400588
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -85.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.289
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.515
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.15048790012
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -2.289
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.3243164721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -85.515
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.138
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.033
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.97591716724
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.138
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.1803360342
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -52.033
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.058
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.446
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.89427888867
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -1.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.6167939364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -51.446
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.508
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.80618282472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 5.508
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.9452507547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -46.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.024
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.959
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.33274972148
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 6.024
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.349067377
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.959
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.164
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.562
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.47561670897
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 6.164
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9679324318
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.562
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.562
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.55827546602
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 6.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9679324318
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -45.562
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.199
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.832
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.61153743669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.199
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.9474061127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -52.832
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.272
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.09218280175
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.272
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.8325608216
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -53.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.413
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.293
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.23607026772
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.413
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3500211124
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -54.293
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.434
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.215
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.25750031585
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.434
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.2751381761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -54.215
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.629
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5998429945
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.629
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.2076813515
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -72.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.247
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.65
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2304986964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.247
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.8535451018
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -75.65
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.395
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.266
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.3815295118
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.395
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.4449282913
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -76.266
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.773
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.4825568815
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9316673774
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -76.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.804
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.628
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.73746976466
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.804
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.2320478212
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -65.628
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.663
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.43744909093
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.3455738483
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.663
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.497
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.524
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.42418287066
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.497
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2121286156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.524
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.499
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.42622382763
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.499
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.0767633077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.921
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.92
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.75447333663
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.921
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.1917388446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -48.92
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.607
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.848
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.43404309325
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.607
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.2025408436
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -46.848
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.487
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.333
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.3115856754
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.487
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.7081214563
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -46.333
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.094
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.2432136171
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.4786724591
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -46.094
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.387
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.698
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.33240871972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.387
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.3392128111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -64.698
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.98
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.501
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.9375524596
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.98
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0301983308
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.501
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.05
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.107
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0089859533
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.05
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.6119811438
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.107
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.438
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0620508344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9297536043
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -68.438
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.386
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.513
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3518667233
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.386
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2814929111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.513
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.52
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3477848094
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3281755293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.52
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.385
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.162
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3508462448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.385
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9844820523
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -60.162
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 31 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.385
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.981
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3508462448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.385
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8107152385
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -59.981
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.795
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.68732849415
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.795
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7621188472
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -52.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.493
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.615
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.37914399256
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.493
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.7790402985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -51.615
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.425
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.38016447104
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.5966331459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -51.425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.287
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.33934533176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.4641479508
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -51.287
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.004
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.25
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.81869456855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -0.004
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.4685889127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -50.25
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.333
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.162
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.47479332009
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.3841055999
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -50.162
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.348
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.021
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.45948614286
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.348
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.248740292
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -50.021
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.451
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.974
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.3543768592
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.451
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.2036185226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -49.974
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.9795299023
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.2211155763
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -81.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.471
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.8173427361
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.2820748157
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -85.471
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.979
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.633
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.9979674274
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.979
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.397638559
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -86.633
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.043
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.163
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.0632780502
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -10.043
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.906458511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -87.163
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.022
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0008910379
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.022
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.8829470425
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -74.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.819
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.834
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7937339061
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.819
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.1500790928
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -72.834
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.807
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7814881643
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.807
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.5202943975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -72.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.776
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.057
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7498533313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.776
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4041298425
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -72.057
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.989
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.139
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.88530131967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.989
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.042325896
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -58.139
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.773
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.858
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.66487796754
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.773
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8125176725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.858
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.66
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.191
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.54956389907
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.66
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1721725631
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -56.191
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.599
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.862
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.48731471166
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.599
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8563201778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -55.862
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.119
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.979
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0998774507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.119
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.2093598419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -74.979
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.464
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.338
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.451942527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.464
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.474088647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -77.338
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.538
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.157
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5274579347
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.538
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.2603594784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -78.157
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.537
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.176
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5264374562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.537
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.2786001937
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -78.176
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.812
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.344
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.76611207461
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.812
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.7995477109
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -69.344
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.634
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75692776827
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.077958628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -69.634
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.489
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.125
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.41601904281
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.489
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.6692241761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.997
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.53439454673
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5063770028
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.997
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.608
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.464
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.53745598217
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.608
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.9946769379
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.464
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.045
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3559486373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.672345874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -65.045
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.045
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3559486373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.672345874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -65.045
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.299
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.4836462724
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.236835405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -82.299
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.299
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.4836462724
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -9.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.236835405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -82.299
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 15.984
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.239
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 13.4967154031
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 15.984
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.8576520462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -40.239
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.314
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.293
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.19647834426
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -3.314
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0303222743
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -62.293
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.154
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1151157155
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.154
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.9199179451
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -62.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.783
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.113
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7569966807
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.783
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7376284339
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -65.113
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.768
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.298
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7416895035
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.768
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9152353983
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -65.298
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.016
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.568
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.994768167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -8.016
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.1344832077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -66.568
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 50 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 11.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.21
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.01885582368
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 11.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9906015045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.21
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 11.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.21
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.01885582368
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 11.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9906015045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -61.21
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 51 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.085
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.393
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.333563462621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 3.085
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.967417751
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -91.393
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.085
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.393
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.333563462621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 3.085
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.967417751
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -91.393
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 52 - Sample 52 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 52 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.728
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.68039188211
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -6.728
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.5418857912
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -84.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 52 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 52 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 52 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean

## 53 - Sample 53 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:29:14
Detail Variables
Detail 53 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.755
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5753515109
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -7.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.2740478926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -67.755
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 53 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.935
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.368
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.87115244582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: -5.935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.062362743
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: -63.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 53 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean
Detail 53 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6494
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.7972
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.009
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -109.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -106.0097
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.81461265462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.75813033202
measure_array: [ 4.6494 -12.764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.979932470456
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2266972394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.9437726477
measure_array: [ 36.7972 -106.0097]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.04162582057
tarray: []
source_col: d(D_H)Mean