## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
210601_20-212_Precip_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.5642 | -77.465 |  | 25.6175 |  | 0.33554 | 3.0168 |  | 0.2109
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.32709 | 4.8102
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.43922 | 7.0925
4 |  | EV | -10.4591 | -76.5161 |  |  |  | 0.13985 | 2.2978
5 |  | WOOD2-D-2 | -8.8942 | -85.6028 |  |  |  | 0.028359 | 0.88149
6 | * | WOOD2-B-5 | -0.013555 | -53.5798 | -4.7115 | 355.044 |  | 1.0841 | 55.5171 |  | -35.4881
7 | * | WOOD2-B-4 | 4.7205 | -42.9208 | -6.4693 | 606.7223 | 37.648 | 6.3258 | 44.4304 |  | -36.8936
8 | * | WOOD2-D-3 | -3.3103 | -36.4718 |  |  |  | 2.5669 | 37.7218
9 | * | WOOD2-B-2 | 6.5328 | -28.0917 | -6.472 | 465.9908 | 41.017 | 8.3326 | 29.0055 |  | -34.4776
10 | * | WOOD2-D-4 | -0.31622 | -36.7574 | -3.0972 | 135.0776 |  | 0.74885 | 38.0189 |  | -32.7043
11 | * | WOOD2-B-1 | 3.6393 | -42.5583 | -6.0475 | 374.1838 | 30.677 | 5.1287 | 44.0534 |  | -34.0066
12 | * | WOOD2-3040-3 | -6.5386 | -47.9598 |  |  |  | 6.1416 | 49.6706
13 |  | EV | -10.3844 | -76.6502 |  | 1.7548 |  | 0.20817 | 1.7781
14 | * | WOOD2-C-4 | -3.1013 | -32.8793 |  |  |  | 2.3354 | 33.9851
15 |  | WOOD2-A-4 | -7.3209 | -69.9747 |  | 9.7208 |  | 0.07941 | 0.22383 |  | -2.6838
16 |  | WOOD2-5060-1 | -12.7635 | -96.9635 |  |  |  | 0.10755 | 1.5828
17 | * | SLCP-21-014 | -9.8481 | -71.1096 |  |  |  | 9.8063 | 73.749
18 |  | EV | -10.6902 | -80.2052 |  |  |  | 0.38622 | 4.5844

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.995 | -77.053 | -11.4149 | -85.7466 | -1.359 | 195.615 | 2021/06/01 12:51:32
1 |  | EV | -9.509 | -72.416 | -10.9081 | -80.5988 | -1.065 | 168.692 | 2021/06/01 13:00:29
1 |  | EV | -9.211 | -69.991 | -10.5974 | -77.9066 | -1.203 | 121.85 | 2021/06/01 13:09:25
1 |  | EV | -9.023 | -69.014 | -10.4013 | -76.822 | -1.156 | 109.646 | 2021/06/01 13:18:22
1 |  | EV | -9.131 | -68.707 | -10.5139 | -76.4811 | -1.099 | 97.255 | 2021/06/01 13:27:19
1 |  | EV | -8.903 | -67.895 | -10.2762 | -75.5797 | -1.028 | 94.433 | 2021/06/01 13:36:16
1 |  | EV | -8.932 | -67.68 | -10.3064 | -75.341 | -0.935 | 92.164 | 2021/06/01 13:45:12
1 |  | EV | -8.964 | -67.771 | -10.3398 | -75.442 | -1.185 | 91.14 | 2021/06/01 13:54:09
1 |  | EV | -9.057 | -67.716 | -10.4368 | -75.381 | -1.104 | 90.498 | 2021/06/01 14:03:05
1 |  | EV | -9.067 | -67.689 | -10.4472 | -75.351 | -0.935 | 90.312 | 2021/06/01 14:12:02
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.5642 | -77.465 |  | 25.6175 |  | 0.33554 | 3.0168 |  | 0.2109
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 1.917 | 3.743 | 1.0068 | 3.9505 | -1.14 | 90.368 | 2021/06/01 14:21:50
2 |  | PZ | 2.795 | 13.99 | 1.9224 | 15.3264 | -1.326 | 90.403 | 2021/06/01 14:30:47
2 |  | PZ | 2.724 | 15.809 | 1.8483 | 17.3458 | -1.128 | 90.215 | 2021/06/01 14:39:44
2 |  | PZ | 2.966 | 17.836 | 2.1007 | 19.5961 | -1.131 | 90.278 | 2021/06/01 14:48:39
2 |  | PZ | 2.886 | 18.215 | 2.0173 | 20.0168 | -1.103 | 90.016 | 2021/06/01 14:57:36
2 |  | PZ | 2.864 | 18.066 | 1.9943 | 19.8514 | -1.165 | 89.849 | 2021/06/01 15:06:32
2 |  | PZ | 3.083 | 19.414 | 2.2227 | 21.3479 | -1.136 | 89.911 | 2021/06/01 15:15:29
2 |  | PZ | 3.002 | 19.305 | 2.1382 | 21.2269 | -1.014 | 89.507 | 2021/06/01 15:24:26
2 |  | PZ | 2.883 | 19.14 | 2.0142 | 21.0437 | -1.208 | 89.272 | 2021/06/01 15:33:22
2 |  | PZ | 2.903 | 19.366 | 2.035 | 21.2946 | -1.262 | 89.539 | 2021/06/01 15:42:19
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.32709 | 4.8102
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.051 | -88.213 | -14.6017 | -98.1361 | -0.91 | 89.452 | 2021/06/01 15:52:08
3 |  | UT2 | -14.116 | -102.883 | -15.7123 | -114.4222 | -0.888 | 89.183 | 2021/06/01 16:01:04
3 |  | UT2 | -14.301 | -106.783 | -15.9052 | -118.7519 | -0.891 | 89.211 | 2021/06/01 16:10:00
3 |  | UT2 | -14.332 | -108.584 | -15.9375 | -120.7513 | -0.907 | 89.032 | 2021/06/01 16:18:56
3 |  | UT2 | -14.273 | -109.289 | -15.876 | -121.534 | -0.958 | 89.24 | 2021/06/01 16:27:52
3 |  | UT2 | -14.607 | -110.865 | -16.2243 | -123.2836 | -1.147 | 89.05 | 2021/06/01 16:36:48
3 |  | UT2 | -14.321 | -110.439 | -15.926 | -122.8106 | -0.89 | 89.193 | 2021/06/01 16:45:45
3 |  | UT2 | -14.416 | -110.675 | -16.0251 | -123.0726 | -1.011 | 89.23 | 2021/06/01 16:54:41
3 |  | UT2 | -14.448 | -111.032 | -16.0585 | -123.469 | -0.915 | 88.966 | 2021/06/01 17:03:38
3 |  | UT2 | -14.52 | -111.302 | -16.1335 | -123.7687 | -0.84 | 88.944 | 2021/06/01 17:12:35
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.43922 | 7.0925
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.388 | -74.827 | -10.7819 | -83.2754 | -1.024 | 88.834 | 2021/06/01 17:22:24
4 |  | EV | -9.168 | -70.174 | -10.5525 | -78.1098 | -0.907 | 88.914 | 2021/06/01 17:31:21
4 |  | EV | -8.985 | -68.332 | -10.3617 | -76.0648 | -1.124 | 88.896 | 2021/06/01 17:40:17
4 |  | EV | -9.131 | -68.351 | -10.5139 | -76.0859 | -1.11 | 88.791 | 2021/06/01 17:49:14
4 |  | EV | -9.031 | -67.826 | -10.4097 | -75.5031 | -1.22 | 89.025 | 2021/06/01 17:58:11
4 |  | EV | -9.06 | -67.743 | -10.4399 | -75.4109 | -0.967 | 88.843 | 2021/06/01 18:07:09
4 |  | EV | -8.879 | -67.109 | -10.2512 | -74.7071 | -1.034 | 88.826 | 2021/06/01 18:16:05
4 |  | EV | -9.139 | -68.022 | -10.5223 | -75.7207 | -0.87 | 88.496 | 2021/06/01 18:25:02
4 |  | EV | -8.964 | -67.416 | -10.3398 | -75.0479 | -0.853 | 88.938 | 2021/06/01 18:33:58
4 |  | EV | -9.039 | -67.585 | -10.418 | -75.2355 | -1.084 | 88.682 | 2021/06/01 18:42:54
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.4591 | -76.5161 |  |  |  | 0.13985 | 2.2978
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | WOOD2-D-2 | -7.59 | -75.633 | -8.907 | -84.1702 | -1.203 | 90.54 | 2021/06/01 18:52:44
5 |  | WOOD2-D-2 | -7.586 | -77.083 | -8.9028 | -85.7799 | -1.314 | 90.254 | 2021/06/01 19:01:40
5 |  | WOOD2-D-2 | -7.599 | -77.507 | -8.9164 | -86.2506 | -1.038 | 89.734 | 2021/06/01 19:10:37
5 |  | WOOD2-D-2 | -7.536 | -77.471 | -8.8507 | -86.2107 | -1.222 | 89.811 | 2021/06/01 19:19:34
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | WOOD2-D-2 | -8.8942 | -85.6028 |  |  |  | 0.028359 | 0.88149
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | WOOD2-B-5 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 19:29:24
6 |  | WOOD2-B-5 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 19:38:20
6 |  | WOOD2-B-5 | 1.914 | -95.722 | 1.0037 | -106.4723 | -9.422 | 889.166 | 2021/06/01 19:47:16
6 |  | WOOD2-B-5 | 1.84 | -96.591 | 0.92652 | -107.4371 | -9.424 | 889.182 | 2021/06/01 19:56:12
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 | * | WOOD2-B-5 | -0.013555 | -53.5798 | -4.7115 | 355.044 |  | 1.0841 | 55.5171 |  | -35.4881
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | WOOD2-B-4 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 20:06:02
7 |  | WOOD2-B-4 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 20:14:59
7 |  | WOOD2-B-4 | 10.984 | -76.587 | 10.4618 | -85.2293 | -13.125 | 1390.809 | 2021/06/01 20:23:56
7 |  | WOOD2-B-4 | 10.929 | -77.321 | 10.4045 | -86.0441 | -12.752 | 1394.252 | 2021/06/01 20:32:52
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 | * | WOOD2-B-4 | 4.7205 | -42.9208 | -6.4693 | 606.7223 | 37.648 | 6.3258 | 44.4304 |  | -36.8936
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | WOOD2-D-3 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 20:42:39
8 |  | WOOD2-D-3 | -4.446 | -65.336 | -5.6285 | -72.7388 | -1.543 | 95.845 | 2021/06/01 20:51:36
8 |  | WOOD2-D-3 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 21:00:31
8 |  | WOOD2-D-3 | -4.446 | -65.336 | -5.6285 | -72.7388 | -1.543 | 95.845 | 2021/06/01 21:09:28
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 | * | WOOD2-D-3 | -3.3103 | -36.4718 |  |  |  | 2.5669 | 37.7218
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | WOOD2-B-2 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 21:19:15
9 |  | WOOD2-B-2 | 14.419 | -50.256 | 14.0438 | -55.9975 | -13.168 | 1110.175 | 2021/06/01 21:28:12
9 |  | WOOD2-B-2 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 21:37:08
9 |  | WOOD2-B-2 | 14.446 | -50.222 | 14.072 | -55.9597 | -12.72 | 1111.96 | 2021/06/01 21:46:06
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 | * | WOOD2-B-2 | 6.5328 | -28.0917 | -6.472 | 465.9908 | 41.017 | 8.3326 | 29.0055 |  | -34.4776
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | WOOD2-D-4 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 21:55:53
10 |  | WOOD2-D-4 | 1.27 | -65.925 | 0.33213 | -73.3927 | -6.226 | 450.308 | 2021/06/01 22:04:50
10 |  | WOOD2-D-4 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 22:13:46
10 |  | WOOD2-D-4 | 1.323 | -65.776 | 0.3874 | -73.2272 | -6.163 | 448.174 | 2021/06/01 22:22:43
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 | * | WOOD2-D-4 | -0.31622 | -36.7574 | -3.0972 | 135.0776 |  | 0.74885 | 38.0189 |  | -32.7043
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | WOOD2-B-1 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 22:32:31
11 |  | WOOD2-B-1 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 22:41:26
11 |  | WOOD2-B-1 | 8.928 | -75.932 | 8.3178 | -84.5021 | -11.986 | 927.567 | 2021/06/01 22:50:23
11 |  | WOOD2-B-1 | 8.838 | -76.67 | 8.224 | -85.3214 | -12.204 | 927.34 | 2021/06/01 22:59:19
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 | * | WOOD2-B-1 | 3.6393 | -42.5583 | -6.0475 | 374.1838 | 30.677 | 5.1287 | 44.0534 |  | -34.0066
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | WOOD2-3040-3 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 23:09:09
12 |  | WOOD2-3040-3 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/01 23:18:05
12 |  | WOOD2-3040-3 | -10.632 | -86.089 | -12.0792 | -95.7781 | -1.166 | 92.695 | 2021/06/01 23:27:01
12 |  | WOOD2-3040-3 | -10.643 | -85.975 | -12.0906 | -95.6515 | -1.204 | 92.314 | 2021/06/01 23:35:58
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 | * | WOOD2-3040-3 | -6.5386 | -47.9598 |  |  |  | 6.1416 | 49.6706
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -9.317 | -71.469 | -10.7079 | -79.5474 | -0.861 | 91.208 | 2021/06/01 23:45:47
13 |  | EV | -8.87 | -68.439 | -10.2418 | -76.1836 | -0.94 | 91.704 | 2021/06/01 23:54:43
13 |  | EV | -8.919 | -67.989 | -10.2929 | -75.684 | -1 | 91.563 | 2021/06/02 00:03:39
13 |  | EV | -8.921 | -67.54 | -10.295 | -75.1856 | -1.069 | 90.716 | 2021/06/02 00:12:36
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.3844 | -76.6502 |  | 1.7548 |  | 0.20817 | 1.7781
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD2-C-4 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/02 00:22:26
14 |  | WOOD2-C-4 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/02 00:31:21
14 |  | WOOD2-C-4 | -4.045 | -58.864 | -5.2103 | -65.5538 | -1.196 | 99.964 | 2021/06/02 00:40:17
14 |  | WOOD2-C-4 | -4.045 | -58.864 | -5.2103 | -65.5538 | -1.196 | 99.964 | 2021/06/02 00:49:13
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 | * | WOOD2-C-4 | -3.1013 | -32.8793 |  |  |  | 2.3354 | 33.9851
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD2-A-4 | -6.154 | -62.993 | -7.4096 | -70.1377 | -1.171 | 99.256 | 2021/06/02 00:59:00
15 |  | WOOD2-A-4 | -5.962 | -62.524 | -7.2093 | -69.617 | -1.164 | 99.342 | 2021/06/02 01:07:57
15 |  | WOOD2-A-4 | -6.081 | -63.003 | -7.3334 | -70.1488 | -0.8 | 99.055 | 2021/06/02 01:16:53
15 |  | WOOD2-A-4 | -6.079 | -62.865 | -7.3313 | -69.9956 | -1.242 | 99.402 | 2021/06/02 01:25:50
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | WOOD2-A-4 | -7.3209 | -69.9747 |  | 9.7208 |  | 0.07941 | 0.22383 |  | -2.6838
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | WOOD2-5060-1 | -11.182 | -84.958 | -12.6527 | -94.5225 | -0.986 | 89.515 | 2021/06/02 01:35:37
16 |  | WOOD2-5060-1 | -11.291 | -87.246 | -12.7664 | -97.0625 | -0.912 | 89.538 | 2021/06/02 01:44:33
16 |  | WOOD2-5060-1 | -11.245 | -87.738 | -12.7184 | -97.6087 | -1.086 | 89.644 | 2021/06/02 01:53:29
16 |  | WOOD2-5060-1 | -11.435 | -88.685 | -12.9165 | -98.6601 | -0.872 | 89.287 | 2021/06/02 02:02:25
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | WOOD2-5060-1 | -12.7635 | -96.9635 |  |  |  | 0.10755 | 1.5828
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | SLCP-21-014 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/02 02:12:12
17 |  | SLCP-21-014 |  |  | -0.99221 | -0.20488 |  |  | 2021/06/02 02:21:09
17 |  | SLCP-21-014 | -16.985 | -127.737 | -18.704 | -142.0143 | -0.949 | 89.145 | 2021/06/02 02:30:05
17 |  | SLCP-21-014 | -16.985 | -127.737 | -18.704 | -142.0143 | -0.949 | 89.145 | 2021/06/02 02:39:01
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 | * | SLCP-21-014 | -9.8481 | -71.1096 |  |  |  | 9.8063 | 73.749
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | EV | -9.862 | -78.796 | -11.2762 | -87.6816 | -0.855 | 89.6 | 2021/06/02 02:48:49
18 |  | EV | -9.039 | -70.901 | -10.418 | -78.9169 | -1.027 | 89.282 | 2021/06/02 02:57:45
18 |  | EV | -9.245 | -69.913 | -10.6328 | -77.82 | -1.123 | 89.236 | 2021/06/02 03:06:42
18 |  | EV | -9.054 | -68.636 | -10.4336 | -76.4023 | -1.006 | 89.338 | 2021/06/02 03:15:38
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | EV | -10.6902 | -80.2052 |  |  |  | 0.38622 | 4.5844

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-16387 | 2021/06/01 12:51:32 | 1-01 | 1 | -9.995 | -77.053 | 14429 | -1 | 1 | EV | standard | H2O | 1622573798 | 0.236 | 0.888 | 90.826 | 0.002 | 0.008 | 2.229 | 195.615 | 31.07 | 2.465 | 68.606 | 0.857 | -1.359 | 0 | 0 | 1 | 44.312 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-16387 | 2021/06/01 13:00:29 | 1-01 | 2 | -9.509 | -72.416 | 13132 | -1 | 1 | EV | standard | H2O | 1622574336 | 0.21 | 0.809 | 83.068 | 0.001 | 0.002 | 2.055 | 168.692 | 28.18 | 2.247 | 64.592 | 0.857 | -1.065 | 0 | 0 | 1 | 43.625 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-16387 | 2021/06/01 13:09:25 | 1-01 | 3 | -9.211 | -69.991 | 12852 | -1 | 1 | EV | standard | H2O | 1622574871 | 0.218 | 0.858 | 75.863 | 0 | 0.003 | 1.889 | 121.85 | 29.185 | 2.191 | 63.279 | 0.856 | -1.203 | 0 | 0 | 1 | 43.375 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-16387 | 2021/06/01 13:18:22 | 1-01 | 4 | -9.023 | -69.014 | 12686 | 0 | 1 | EV | standard | H2O | 1622575409 | 0.217 | 0.864 | 75.056 | 0.001 | -0.002 | 1.854 | 109.646 | 29.36 | 2.16 | 61.922 | 0.856 | -1.156 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-16387 | 2021/06/01 13:27:19 | 1-01 | 5 | -9.131 | -68.707 | 12229 | 0 | 1 | EV | standard | H2O | 1622575945 | 0.214 | 0.926 | 75.616 | 0.002 | 0.003 | 1.883 | 97.255 | 27.963 | 2.094 | 61.532 | 0.856 | -1.099 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-16387 | 2021/06/01 13:36:16 | 1-01 | 6 | -8.903 | -67.895 | 12372 | 0 | 1 | EV | standard | H2O | 1622576482 | 0.189 | 0.942 | 71.116 | 0.001 | 0 | 1.77 | 94.433 | 28.616 | 2.122 | 62.747 | 0.856 | -1.028 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-16387 | 2021/06/01 13:45:12 | 1-01 | 7 | -8.932 | -67.68 | 12284 | 0 | 1 | EV | standard | H2O | 1622577018 | 0.203 | 0.851 | 70.415 | 0 | 0.003 | 1.765 | 92.164 | 28.333 | 2.102 | 62.466 | 0.856 | -0.935 | 0 | 0 | 1 | 43.062 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-16387 | 2021/06/01 13:54:09 | 1-01 | 8 | -8.964 | -67.771 | 12278 | 0 | 1 | EV | standard | H2O | 1622577555 | 0.184 | 0.926 | 69.366 | 0 | 0 | 1.728 | 91.14 | 28.725 | 2.099 | 60.746 | 0.856 | -1.185 | 0 | 0 | 1 | 43.062 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-16387 | 2021/06/01 14:03:05 | 1-01 | 9 | -9.057 | -67.716 | 12195 | 0 | 1 | EV | standard | H2O | 1622578091 | 0.213 | 0.867 | 67.853 | 0.001 | 0.004 | 1.688 | 90.498 | 28.106 | 2.091 | 61.606 | 0.856 | -1.104 | 0 | 0 | 1 | 43.062 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-16387 | 2021/06/01 14:12:02 | 1-01 | 10 | -9.067 | -67.689 | 12280 | 0 | 1 | EV | standard | H2O | 1622578628 | 0.216 | 0.852 | 71.118 | 0.002 | 0.003 | 1.77 | 90.312 | 28.315 | 2.094 | 62.101 | 0.856 | -0.935 | 0 | 0 | 1 | 43.188 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-16388 | 2021/06/01 14:21:50 | 1-02 | 1 | 1.917 | 3.743 | 12749 | -1 | 1 | PZ | standard | H2O | 1622579217 | 0.211 | 0.946 | 74.155 | 0.002 | 0.009 | 1.835 | 90.368 | 30.546 | 2.289 | 64.726 | 0.856 | -1.14 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-16388 | 2021/06/01 14:30:47 | 1-02 | 2 | 2.795 | 13.99 | 12573 | -1 | 1 | PZ | standard | H2O | 1622579754 | 0.192 | 0.976 | 68.712 | -0.001 | -0.007 | 1.706 | 90.403 | 29.996 | 2.269 | 63.251 | 0.856 | -1.326 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-16388 | 2021/06/01 14:39:44 | 1-02 | 3 | 2.724 | 15.809 | 12457 | -1 | 1 | PZ | standard | H2O | 1622580290 | 0.237 | 0.944 | 77.179 | 0.001 | 0.003 | 1.91 | 90.215 | 29.102 | 2.249 | 64.525 | 0.856 | -1.128 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-16388 | 2021/06/01 14:48:39 | 1-02 | 4 | 2.966 | 17.836 | 12521 | 0 | 1 | PZ | standard | H2O | 1622580826 | 0.193 | 0.826 | 70.438 | 0 | -0.002 | 1.751 | 90.278 | 29.093 | 2.271 | 65.374 | 0.856 | -1.131 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-16388 | 2021/06/01 14:57:36 | 1-02 | 5 | 2.886 | 18.215 | 12454 | 0 | 1 | PZ | standard | H2O | 1622581361 | 0.233 | 0.902 | 72.533 | 0.001 | -0.005 | 1.794 | 90.016 | 29.087 | 2.251 | 65.397 | 0.857 | -1.103 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-16388 | 2021/06/01 15:06:32 | 1-02 | 6 | 2.864 | 18.066 | 12417 | 0 | 1 | PZ | standard | H2O | 1622581898 | 0.206 | 0.99 | 71.732 | 0.002 | 0.007 | 1.794 | 89.849 | 28.925 | 2.251 | 64.707 | 0.857 | -1.165 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-16388 | 2021/06/01 15:15:29 | 1-02 | 7 | 3.083 | 19.414 | 12498 | 0 | 1 | PZ | standard | H2O | 1622582435 | 0.205 | 0.909 | 70.059 | 0 | 0 | 1.728 | 89.911 | 29.496 | 2.262 | 64.579 | 0.857 | -1.136 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-16388 | 2021/06/01 15:24:26 | 1-02 | 8 | 3.002 | 19.305 | 12389 | 0 | 1 | PZ | standard | H2O | 1622582972 | 0.188 | 0.881 | 71.172 | 0 | -0.004 | 1.773 | 89.507 | 29.158 | 2.244 | 64.649 | 0.856 | -1.014 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-16388 | 2021/06/01 15:33:22 | 1-02 | 9 | 2.883 | 19.14 | 12243 | 0 | 1 | PZ | standard | H2O | 1622583509 | 0.258 | 0.95 | 72.681 | 0.002 | 0.004 | 1.806 | 89.272 | 28.76 | 2.232 | 63.272 | 0.857 | -1.208 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-16388 | 2021/06/01 15:42:19 | 1-02 | 10 | 2.903 | 19.366 | 12420 | 0 | 1 | PZ | standard | H2O | 1622584044 | 0.224 | 0.88 | 66.968 | 0.001 | 0.002 | 1.662 | 89.539 | 29.529 | 2.251 | 64.459 | 0.857 | -1.262 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-16389 | 2021/06/01 15:52:08 | 1-03 | 1 | -13.051 | -88.213 | 12353 | -1 | 1 | UT2 | standard | H2O | 1622584634 | 0.247 | 0.804 | 72.476 | 0.004 | -0.004 | 1.804 | 89.452 | 28.694 | 2.084 | 62.008 | 0.856 | -0.91 | 0 | 0 | 1 | 43.062 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-16389 | 2021/06/01 16:01:04 | 1-03 | 2 | -14.116 | -102.883 | 12231 | -1 | 1 | UT2 | standard | H2O | 1622585170 | 0.264 | 0.871 | 75.584 | 0.004 | 0.011 | 1.882 | 89.183 | 28.293 | 2.047 | 61.439 | 0.856 | -0.888 | 0 | 0 | 1 | 43.125 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-16389 | 2021/06/01 16:10:00 | 1-03 | 3 | -14.301 | -106.783 | 12252 | -1 | 1 | UT2 | standard | H2O | 1622585706 | 0.25 | 0.891 | 70.019 | 0.004 | 0.009 | 1.737 | 89.211 | 28.336 | 2.048 | 61.15 | 0.856 | -0.891 | 0 | 0 | 1 | 43.062 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-16389 | 2021/06/01 16:18:56 | 1-03 | 4 | -14.332 | -108.584 | 12163 | 0 | 1 | UT2 | standard | H2O | 1622586243 | 0.228 | 0.914 | 67.348 | 0.002 | 0.007 | 1.685 | 89.032 | 28.169 | 2.008 | 60.656 | 0.857 | -0.907 | 0 | 0 | 1 | 43.062 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-16389 | 2021/06/01 16:27:52 | 1-03 | 5 | -14.273 | -109.289 | 12224 | 0 | 1 | UT2 | standard | H2O | 1622586779 | 0.204 | 0.868 | 64.93 | 0 | 0.007 | 1.616 | 89.24 | 28.327 | 2.041 | 61.233 | 0.856 | -0.958 | 0 | 0 | 1 | 43.125 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-16389 | 2021/06/01 16:36:48 | 1-03 | 6 | -14.607 | -110.865 | 12156 | 0 | 1 | UT2 | standard | H2O | 1622587315 | 0.219 | 0.91 | 72.463 | 0.003 | 0.01 | 1.803 | 89.05 | 28.39 | 2.011 | 58.975 | 0.856 | -1.147 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-16389 | 2021/06/01 16:45:45 | 1-03 | 7 | -14.321 | -110.439 | 12243 | 0 | 1 | UT2 | standard | H2O | 1622587851 | 0.206 | 0.836 | 64.205 | 0 | 0.005 | 1.599 | 89.193 | 28.181 | 2.034 | 61.559 | 0.856 | -0.89 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-16389 | 2021/06/01 16:54:41 | 1-03 | 8 | -14.416 | -110.675 | 12199 | 0 | 1 | UT2 | standard | H2O | 1622588387 | 0.216 | 0.691 | 62.414 | 0 | 0 | 1.549 | 89.23 | 28.077 | 2.023 | 60.64 | 0.856 | -1.011 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-16389 | 2021/06/01 17:03:38 | 1-03 | 9 | -14.448 | -111.032 | 12128 | 0 | 1 | UT2 | standard | H2O | 1622588924 | 0.212 | 0.881 | 67.991 | 0.001 | 0.001 | 1.702 | 88.966 | 27.817 | 2.011 | 61.272 | 0.856 | -0.915 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-16389 | 2021/06/01 17:12:35 | 1-03 | 10 | -14.52 | -111.302 | 12111 | 0 | 1 | UT2 | standard | H2O | 1622589461 | 0.208 | 0.843 | 64.855 | 0.001 | 0.001 | 1.613 | 88.944 | 27.809 | 2.016 | 61.105 | 0.856 | -0.84 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-16390 | 2021/06/01 17:22:24 | 1-04 | 1 | -9.388 | -74.827 | 12049 | -1 | 1 | EV | standard | H2O | 1622590050 | 0.199 | 0.906 | 66.549 | 0 | 0.009 | 1.656 | 88.834 | 27.729 | 2.058 | 60.946 | 0.856 | -1.024 | 0 | 0 | 1 | 43.25 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-16390 | 2021/06/01 17:31:21 | 1-04 | 2 | -9.168 | -70.174 | 12165 | -1 | 1 | EV | standard | H2O | 1622590587 | 0.22 | 0.957 | 66.024 | 0.001 | 0 | 1.642 | 88.914 | 27.998 | 2.079 | 61.957 | 0.857 | -0.907 | 0 | 0 | 1 | 43.375 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-16390 | 2021/06/01 17:40:17 | 1-04 | 3 | -8.985 | -68.332 | 12117 | -1 | 1 | EV | standard | H2O | 1622591123 | 0.217 | 0.726 | 62.751 | 0 | -0.001 | 1.558 | 88.896 | 28.318 | 2.065 | 60.495 | 0.856 | -1.124 | 0 | 0 | 1 | 43.188 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-16390 | 2021/06/01 17:49:14 | 1-04 | 4 | -9.131 | -68.351 | 12130 | 0 | 1 | EV | standard | H2O | 1622591660 | 0.252 | 0.868 | 63.536 | 0.002 | -0.001 | 1.588 | 88.791 | 28.674 | 2.064 | 60.664 | 0.856 | -1.11 | 0 | 0 | 1 | 43.188 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-16390 | 2021/06/01 17:58:11 | 1-04 | 5 | -9.031 | -67.826 | 12087 | 0 | 1 | EV | standard | H2O | 1622592198 | 0.196 | 0.821 | 67.8 | 0.002 | 0.005 | 1.687 | 89.025 | 27.972 | 2.063 | 60.379 | 0.856 | -1.22 | 0 | 0 | 1 | 43.125 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-16390 | 2021/06/01 18:07:09 | 1-04 | 6 | -9.06 | -67.743 | 12126 | 0 | 1 | EV | standard | H2O | 1622592735 | 0.2 | 0.82 | 68.445 | 0.001 | 0.007 | 1.702 | 88.843 | 28.176 | 2.072 | 61.385 | 0.857 | -0.967 | 0 | 0 | 1 | 43.125 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-16390 | 2021/06/01 18:16:05 | 1-04 | 7 | -8.879 | -67.109 | 12119 | 0 | 1 | EV | standard | H2O | 1622593271 | 0.191 | 0.808 | 62.771 | 0 | 0.001 | 1.57 | 88.826 | 28.255 | 2.077 | 60.946 | 0.856 | -1.034 | 0 | 0 | 1 | 43.125 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-16390 | 2021/06/01 18:25:02 | 1-04 | 8 | -9.139 | -68.022 | 11985 | 0 | 1 | EV | standard | H2O | 1622593808 | 0.23 | 0.917 | 67.589 | 0.002 | 0.006 | 1.69 | 88.496 | 27.407 | 2.057 | 62.269 | 0.857 | -0.87 | 0 | 0 | 1 | 43.312 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-16390 | 2021/06/01 18:33:58 | 1-04 | 9 | -8.964 | -67.416 | 12189 | 0 | 1 | EV | standard | H2O | 1622594344 | 0.216 | 0.906 | 67.198 | 0.002 | 0.007 | 1.674 | 88.938 | 28.108 | 2.095 | 61.7 | 0.856 | -0.853 | 0 | 0 | 1 | 43.312 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-16390 | 2021/06/01 18:42:54 | 1-04 | 10 | -9.039 | -67.585 | 12010 | 0 | 1 | EV | standard | H2O | 1622594881 | 0.227 | 0.754 | 69.442 | 0.002 | 0.006 | 1.73 | 88.682 | 27.812 | 2.056 | 60.767 | 0.856 | -1.084 | 0 | 0 | 1 | 43.312 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-16391 | 2021/06/01 18:52:44 | 1-05 | 1 | -7.59 | -75.633 | 12607 | -1 | 1 | WOOD2-D-2 | 20-212 | H2O | 1622595470 | 0.184 | 0.808 | 65.171 | 0.001 | 0 | 1.611 | 90.54 | 29.235 | 2.149 | 63.092 | 0.856 | -1.203 | 0 | 0 | 1 | 43.312 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-16391 | 2021/06/01 19:01:40 | 1-05 | 2 | -7.586 | -77.083 | 12444 | -1 | 1 | WOOD2-D-2 | 20-212 | H2O | 1622596007 | 0.181 | 0.824 | 69.132 | 0.001 | -0.002 | 1.719 | 90.254 | 28.842 | 2.126 | 62.321 | 0.856 | -1.314 | 0 | 0 | 1 | 43.312 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-16391 | 2021/06/01 19:10:37 | 1-05 | 3 | -7.599 | -77.507 | 12254 | -1 | 1 | WOOD2-D-2 | 20-212 | H2O | 1622596543 | 0.204 | 0.861 | 66.854 | 0.001 | -0.001 | 1.658 | 89.734 | 28.113 | 2.093 | 62.186 | 0.856 | -1.038 | 0 | 0 | 1 | 43.312 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-16391 | 2021/06/01 19:19:34 | 1-05 | 4 | -7.536 | -77.471 | 12234 | 0 | 1 | WOOD2-D-2 | 20-212 | H2O | 1622597080 | 0.222 | 0.823 | 67.396 | 0 | 0.001 | 1.676 | 89.811 | 28.346 | 2.083 | 61.005 | 0.857 | -1.222 | 0 | 0 | 1 | 43.188 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-16392 | 2021/06/01 19:29:24 | 1-06 | 1 | [5 empty] | WOOD2-B-5 | 20-212 | [18 empty] | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-16392 | 2021/06/01 19:38:20 | 1-06 | 2 | [5 empty] | WOOD2-B-5 | 20-212 | [18 empty] | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-16392 | 2021/06/01 19:47:16 | 1-06 | 3 | 1.914 | -95.722 | 11690 | -1 | 1 | WOOD2-B-5 | 20-212 | H2O | 1622598743 | 0.295 | 1.421 | 67.463 | -0.002 | 0.006 | 1.692 | 889.166 | -15.123 | 4.537 | 53.948 | 0.859 | -9.422 | 0 | 0 | 1 | 43.312 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-16392 | 2021/06/01 19:56:12 | 1-06 | 4 | 1.84 | -96.591 | 11676 | 0 | 1 | WOOD2-B-5 | 20-212 | H2O | 1622599279 | 0.306 | 1.575 | 66.034 | -0.002 | 0.001 | 1.656 | 889.182 | -15.358 | 4.558 | 55.095 | 0.859 | -9.424 | 0 | 0 | 1 | 43.438 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-16393 | 2021/06/01 20:06:02 | 1-07 | 1 | [5 empty] | WOOD2-B-4 | 20-212 | [18 empty] | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-16393 | 2021/06/01 20:14:59 | 1-07 | 2 | [5 empty] | WOOD2-B-4 | 20-212 | [18 empty] | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-16393 | 2021/06/01 20:23:56 | 1-07 | 3 | 10.984 | -76.587 | 11282 | -1 | 1 | WOOD2-B-4 | 20-212 | H2O | 1622600943 | 0.35 | 2.46 | 63.396 | -0.003 | 0.003 | 1.614 | 1390.809 | -40.182 | 5.873 | 50.175 | 0.863 | -13.125 | 0 | 0 | 1 | 43.5 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-16393 | 2021/06/01 20:32:52 | 1-07 | 4 | 10.929 | -77.321 | 11290 | 0 | 1 | WOOD2-B-4 | 20-212 | H2O | 1622601479 | 0.361 | 1.91 | 63.686 | -0.003 | 0.006 | 1.597 | 1394.252 | -40.592 | 5.935 | 53.246 | 0.864 | -12.752 | 0 | 0 | 1 | 43.438 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-16394 | 2021/06/01 20:42:39 | 1-08 | 1 | [5 empty] | WOOD2-D-3 | 20-212 | [18 empty] | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-16394 | 2021/06/01 20:51:36 | 1-08 | 2 | -4.446 | -65.336 | 12117 | 0 | 1 | WOOD2-D-3 | 20-212 | H2O | 1622603674 | 0.236 | 0.856 | 67.328 | 0.002 | 0.002 | 1.676 | 95.845 | 27.253 | 2.137 | 61.187 | 0.856 | -1.543 | 0 | 0 | 1 | 43.375 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-16394 | 2021/06/01 21:00:31 | 1-08 | 3 | [5 empty] | WOOD2-D-3 | 20-212 | [18 empty] | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-16394 | 2021/06/01 21:09:28 | 1-08 | 4 | -4.446 | -65.336 | 12117 | 0 | 1 | WOOD2-D-3 | 20-212 | H2O | 1622603674 | 0.236 | 0.856 | 67.328 | 0.002 | 0.002 | 1.676 | 95.845 | 27.253 | 2.137 | 61.187 | 0.856 | -1.543 | 0 | 0 | 1 | 43.375 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-16395 | 2021/06/01 21:19:15 | 1-09 | 1 | [5 empty] | WOOD2-B-2 | 20-212 | [18 empty] | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-16395 | 2021/06/01 21:28:12 | 1-09 | 2 | 14.419 | -50.256 | 11335 | -1 | 1 | WOOD2-B-2 | 20-212 | H2O | 1622604799 | 0.348 | 2.459 | 66.842 | -0.002 | 0.008 | 1.686 | 1110.175 | -31.174 | 6.027 | 55.673 | 0.86 | -13.168 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-16395 | 2021/06/01 21:37:08 | 1-09 | 3 | [5 empty] | WOOD2-B-2 | 20-212 | [18 empty] | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-16395 | 2021/06/01 21:46:06 | 1-09 | 4 | 14.446 | -50.222 | 11379 | 0 | 1 | WOOD2-B-2 | 20-212 | H2O | 1622605872 | 0.353 | 2.391 | 67.982 | -0.002 | -0.007 | 1.714 | 1111.96 | -31.294 | 6.07 | 57.412 | 0.861 | -12.72 | -1e-04 | -1e-04 | 1 | 43.25 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-16396 | 2021/06/01 21:55:53 | 1-10 | 1 | [5 empty] | WOOD2-D-4 | 20-212 | [18 empty] | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-16396 | 2021/06/01 22:04:50 | 1-10 | 2 | 1.27 | -65.925 | 11951 | -1 | 1 | WOOD2-D-4 | 20-212 | H2O | 1622606996 | 0.244 | 1.345 | 68.338 | -0.001 | 0 | 1.709 | 450.308 | 6.052 | 3.483 | 61.067 | 0.858 | -6.226 | 0 | 0 | 1 | 43.125 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-16396 | 2021/06/01 22:13:46 | 1-10 | 3 | [5 empty] | WOOD2-D-4 | 20-212 | [18 empty] | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-16396 | 2021/06/01 22:22:43 | 1-10 | 4 | 1.323 | -65.776 | 11874 | 0 | 1 | WOOD2-D-4 | 20-212 | H2O | 1622608070 | 0.255 | 1.122 | 67.624 | -0.001 | 0.002 | 1.681 | 448.174 | 5.981 | 3.456 | 59.111 | 0.857 | -6.163 | 0 | 0 | 1 | 43.188 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-16397 | 2021/06/01 22:32:31 | 1-11 | 1 | [5 empty] | WOOD2-B-1 | 20-212 | [18 empty] | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-16397 | 2021/06/01 22:41:26 | 1-11 | 2 | [5 empty] | WOOD2-B-1 | 20-212 | [18 empty] | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-16397 | 2021/06/01 22:50:23 | 1-11 | 3 | 8.928 | -75.932 | 11482 | -1 | 1 | WOOD2-B-1 | 20-212 | H2O | 1622609729 | 0.298 | 1.709 | 64.881 | -0.001 | -0.002 | 1.627 | 927.567 | -22.527 | 5.614 | 57.708 | 0.86 | -11.986 | -1e-04 | -1e-04 | 1 | 43 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-16397 | 2021/06/01 22:59:19 | 1-11 | 4 | 8.838 | -76.67 | 11461 | 0 | 1 | WOOD2-B-1 | 20-212 | H2O | 1622610267 | 0.326 | 1.821 | 64.663 | -0.002 | 0.003 | 1.621 | 927.34 | -22.201 | 5.596 | 57.261 | 0.86 | -12.204 | 0 | 0 | 1 | 42.875 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-16398 | 2021/06/01 23:09:09 | 1-12 | 1 | [5 empty] | WOOD2-3040-3 | 20-212 | [18 empty] | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-16398 | 2021/06/01 23:18:05 | 1-12 | 2 | [5 empty] | WOOD2-3040-3 | 20-212 | [18 empty] | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-16398 | 2021/06/01 23:27:01 | 1-12 | 3 | -10.632 | -86.089 | 11896 | -1 | 1 | WOOD2-3040-3 | 20-212 | H2O | 1622611927 | 0.202 | 0.784 | 68.397 | 0 | 0 | 1.705 | 92.695 | 27.068 | 2.016 | 58.886 | 0.856 | -1.166 | 0 | 0 | 1 | 43 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-16398 | 2021/06/01 23:35:58 | 1-12 | 4 | -10.643 | -85.975 | 11937 | 0 | 1 | WOOD2-3040-3 | 20-212 | H2O | 1622612464 | 0.203 | 0.803 | 67.508 | 0 | -0.001 | 1.68 | 92.314 | 27.796 | 2.02 | 58.524 | 0.856 | -1.204 | 0 | 0 | 1 | 43 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-16399 | 2021/06/01 23:45:47 | 1-13 | 1 | -9.317 | -71.469 | 12111 | -1 | 1 | EV | standard | H2O | 1622613053 | 0.248 | 0.922 | 72.914 | 0.002 | 0.009 | 1.813 | 91.208 | 27.718 | 2.069 | 61.889 | 0.856 | -0.861 | 0 | 0 | 1 | 43 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-16399 | 2021/06/01 23:54:43 | 1-13 | 2 | -8.87 | -68.439 | 12173 | -1 | 1 | EV | standard | H2O | 1622613588 | 0.22 | 0.918 | 70.502 | 0.001 | 0.003 | 1.757 | 91.704 | 27.667 | 2.089 | 62.599 | 0.856 | -0.94 | 0 | 0 | 1 | 43 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-16399 | 2021/06/02 00:03:39 | 1-13 | 3 | -8.919 | -67.989 | 12096 | -1 | 1 | EV | standard | H2O | 1622614125 | 0.209 | 0.838 | 66.417 | 0.001 | 0.002 | 1.651 | 91.563 | 27.737 | 2.066 | 62.112 | 0.856 | -1 | 0 | 0 | 1 | 42.938 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-16399 | 2021/06/02 00:12:36 | 1-13 | 4 | -8.921 | -67.54 | 12199 | 0 | 1 | EV | standard | H2O | 1622614662 | 0.182 | 0.871 | 70.359 | 0 | 0.002 | 1.748 | 90.716 | 28.417 | 2.094 | 61.315 | 0.856 | -1.069 | 0 | 0 | 1 | 43.062 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-16400 | 2021/06/02 00:22:26 | 1-14 | 1 | [5 empty] | WOOD2-C-4 | 20-212 | [18 empty] | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-16400 | 2021/06/02 00:31:21 | 1-14 | 2 | [5 empty] | WOOD2-C-4 | 20-212 | [18 empty] | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-16400 | 2021/06/02 00:40:17 | 1-14 | 3 | -4.045 | -58.864 | 11854 | -1 | 1 | WOOD2-C-4 | 20-212 | H2O | 1622616323 | 0.202 | 0.821 | 66.579 | 0.002 | 0.002 | 1.659 | 99.964 | 27 | 2.052 | 60.101 | 0.857 | -1.196 | 0 | 0 | 1 | 43.062 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-16400 | 2021/06/02 00:49:13 | 1-14 | 4 | -4.045 | -58.864 | 11854 | -1 | 1 | WOOD2-C-4 | 20-212 | H2O | 1622616323 | 0.202 | 0.821 | 66.579 | 0.002 | 0.002 | 1.659 | 99.964 | 27 | 2.052 | 60.101 | 0.857 | -1.196 | 0 | 0 | 1 | 43.062 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-16401 | 2021/06/02 00:59:00 | 1-15 | 1 | -6.154 | -62.993 | 11822 | -1 | 1 | WOOD2-A-4 | 20-212 | H2O | 1622617447 | 0.215 | 0.807 | 70.637 | 0.002 | 0.005 | 1.759 | 99.256 | 26.942 | 2.033 | 60.059 | 0.856 | -1.171 | 0 | 0 | 1 | 43 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-16401 | 2021/06/02 01:07:57 | 1-15 | 2 | -5.962 | -62.524 | 11837 | -1 | 1 | WOOD2-A-4 | 20-212 | H2O | 1622617983 | 0.184 | 0.838 | 64.717 | 0 | -0.001 | 1.622 | 99.342 | 27.1 | 2.033 | 59.348 | 0.856 | -1.164 | 0 | 0 | 1 | 43 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-16401 | 2021/06/02 01:16:53 | 1-15 | 3 | -6.081 | -63.003 | 11781 | -1 | 1 | WOOD2-A-4 | 20-212 | H2O | 1622618520 | 0.214 | 0.927 | 68.302 | 0 | 0.004 | 1.713 | 99.055 | 26.332 | 2.051 | 61.545 | 0.856 | -0.8 | 0 | 0 | 1 | 43 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-16401 | 2021/06/02 01:25:50 | 1-15 | 4 | -6.079 | -62.865 | 11839 | 0 | 1 | WOOD2-A-4 | 20-212 | H2O | 1622619057 | 0.197 | 0.795 | 61.829 | -0.001 | 0.001 | 1.549 | 99.402 | 27.046 | 2.037 | 59.308 | 0.857 | -1.242 | 0 | 0 | 1 | 43.125 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-16402 | 2021/06/02 01:35:37 | 1-16 | 1 | -11.182 | -84.958 | 11938 | -1 | 1 | WOOD2-5060-1 | 20-212 | H2O | 1622619643 | 0.233 | 0.814 | 69.799 | 0.003 | 0.003 | 1.741 | 89.515 | 27.571 | 2.014 | 60.273 | 0.857 | -0.986 | 0 | 0 | 1 | 43.188 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-16402 | 2021/06/02 01:44:33 | 1-16 | 2 | -11.291 | -87.246 | 12054 | -1 | 1 | WOOD2-5060-1 | 20-212 | H2O | 1622620180 | 0.194 | 0.81 | 67.628 | 0.002 | 0.004 | 1.682 | 89.538 | 27.461 | 2.034 | 61.715 | 0.856 | -0.912 | 0 | 0 | 1 | 43.188 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-16402 | 2021/06/02 01:53:29 | 1-16 | 3 | -11.245 | -87.738 | 12022 | -1 | 1 | WOOD2-5060-1 | 20-212 | H2O | 1622620715 | 0.198 | 0.761 | 67.386 | 0.001 | 0.003 | 1.688 | 89.644 | 27.774 | 2.024 | 59.73 | 0.857 | -1.086 | 0 | 0 | 1 | 43.062 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-16402 | 2021/06/02 02:02:25 | 1-16 | 4 | -11.435 | -88.685 | 11983 | 0 | 1 | WOOD2-5060-1 | 20-212 | H2O | 1622621252 | 0.205 | 0.801 | 67.317 | 0.002 | 0.004 | 1.674 | 89.287 | 27.403 | 2.027 | 61.009 | 0.856 | -0.872 | 0 | 0 | 1 | 43.188 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-16403 | 2021/06/02 02:12:12 | 1-17 | 1 | [5 empty] | SLCP-21-014 | Precipitation | [18 empty] | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-16403 | 2021/06/02 02:21:09 | 1-17 | 2 | [5 empty] | SLCP-21-014 | Precipitation | [18 empty] | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-16403 | 2021/06/02 02:30:05 | 1-17 | 3 | -16.985 | -127.737 | 12019 | 0 | 1 | SLCP-21-014 | Precipitation | H2O | 1622623448 | 0.223 | 0.923 | 73.153 | 0.003 | 0.013 | 1.823 | 89.145 | 28.017 | 1.973 | 58.98 | 0.856 | -0.949 | 0 | 0 | 1 | 43.188 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-16403 | 2021/06/02 02:39:01 | 1-17 | 4 | -16.985 | -127.737 | 12019 | 0 | 1 | SLCP-21-014 | Precipitation | H2O | 1622623448 | 0.223 | 0.923 | 73.153 | 0.003 | 0.013 | 1.823 | 89.145 | 28.017 | 1.973 | 58.98 | 0.856 | -0.949 | 0 | 0 | 1 | 43.188 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-16404 | 2021/06/02 02:48:49 | 1-18 | 1 | -9.862 | -78.796 | 12339 | -1 | 1 | EV | standard | H2O | 1622624035 | 0.249 | 1.007 | 76.372 | 0.004 | 0.016 | 1.901 | 89.6 | 28.717 | 2.091 | 62.27 | 0.856 | -0.855 | 0 | 0 | 1 | 43.188 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-16404 | 2021/06/02 02:57:45 | 1-18 | 2 | -9.039 | -70.901 | 12130 | -1 | 1 | EV | standard | H2O | 1622624572 | 0.197 | 0.853 | 65.617 | 0 | 0.001 | 1.633 | 89.282 | 28.214 | 2.069 | 60.268 | 0.856 | -1.027 | 0 | 0 | 1 | 43.188 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-16404 | 2021/06/02 03:06:42 | 1-18 | 3 | -9.245 | -69.913 | 12115 | -1 | 1 | EV | standard | H2O | 1622625108 | 0.233 | 0.937 | 72.521 | 0.002 | 0.004 | 1.806 | 89.236 | 28.407 | 2.065 | 60.545 | 0.856 | -1.123 | 0 | 0 | 1 | 43.125 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-16404 | 2021/06/02 03:15:38 | 1-18 | 4 | -9.054 | -68.636 | 12174 | 0 | 1 | EV | standard | H2O | 1622625644 | 0.221 | 0.834 | 73.752 | 0.002 | 0.006 | 1.834 | 89.338 | 28.306 | 2.083 | 61.09 | 0.856 | -1.006 | 0 | 0 | 1 | 43.25 | 1 | 18 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 89.54295
varray: [90.368, 90.403, 90.215, 90.278, 90.016, 89.849, 89.911, 89.507, 89.272, 89.539, 89.452, 89.183, 89.211, 89.032, 89.24, 89.05, 89.193, 89.23, 88.966, 88.944]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_shift
baseline_shift_STD | 0.49714230463
varray: [90.368, 90.403, 90.215, 90.278, 90.016, 89.849, 89.911, 89.507, 89.272, 89.539, 89.452, 89.183, 89.211, 89.032, 89.24, 89.05, 89.193, 89.23, 88.966, 88.944]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_shift
ch4_ppm_MEAN | -1.0485
varray: [-1.14, -1.326, -1.128, -1.131, -1.103, -1.165, -1.136, -1.014, -1.208, -1.262, -0.91, -0.888, -0.891, -0.907, -0.958, -1.147, -0.89, -1.011, -0.915, -0.84]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.143344083644
varray: [-1.14, -1.326, -1.128, -1.131, -1.103, -1.165, -1.136, -1.014, -1.208, -1.262, -0.91, -0.888, -0.891, -0.907, -0.958, -1.147, -0.89, -1.011, -0.915, -0.84]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: ch4_ppm
residuals_MEAN | 2.1446
varray: [2.289, 2.269, 2.249, 2.271, 2.251, 2.251, 2.262, 2.244, 2.232, 2.251, 2.084, 2.047, 2.048, 2.008, 2.041, 2.011, 2.034, 2.023, 2.011, 2.016]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: residuals
residuals_STD | 0.116899597046
varray: [2.289, 2.269, 2.249, 2.271, 2.251, 2.251, 2.262, 2.244, 2.232, 2.251, 2.084, 2.047, 2.048, 2.008, 2.041, 2.011, 2.034, 2.023, 2.011, 2.016]
inst: std_standards
operation: std_standards(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: residuals
slope_shift_MEAN | 28.78925
varray: [30.546, 29.996, 29.102, 29.093, 29.087, 28.925, 29.496, 29.158, 28.76, 29.529, 28.694, 28.293, 28.336, 28.169, 28.327, 28.39, 28.181, 28.077, 27.817, 27.809]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: slope_shift
slope_shift_STD | 0.726386867708
varray: [30.546, 29.996, 29.102, 29.093, 29.087, 28.925, 29.496, 29.158, 28.76, 29.529, 28.694, 28.293, 28.336, 28.169, 28.327, 28.39, 28.181, 28.077, 27.817, 27.809]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: slope_shift
baseline_curvature_MEAN | 62.7488
varray: [64.726, 63.251, 64.525, 65.374, 65.397, 64.707, 64.579, 64.649, 63.272, 64.459, 62.008, 61.439, 61.15, 60.656, 61.233, 58.975, 61.559, 60.64, 61.272, 61.105]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.94259375172
varray: [64.726, 63.251, 64.525, 65.374, 65.397, 64.707, 64.579, 64.649, 63.272, 64.459, 62.008, 61.439, 61.15, 60.656, 61.233, 58.975, 61.559, 60.64, 61.272, 61.105]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.1446
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 89.54295
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 1.0485
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 28.78925
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.175349395569
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.31994939557
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.233799194092
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.37839919409
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.350698791137
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.49529879114
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.21446
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.430032250933
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.47853225093
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.248571152315
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 89.7915211523
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 90.0400923046
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 1.49142691389
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 91.0343769139
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 8.94856148334
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 98.4915114833
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 3.63193433854
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 32.4211843385
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 5.82778125515
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 68.5765812552
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -6.06757798563
actual_array: [ 1.93 -15.84]
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067, 1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903, -13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52, -9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039, -7.59, -7.586, -7.599, -7.536, 0.0, 0.0, 1.914, 1.84, 0.0, 0.0, 10.984, 10.929, 0.0, -4.446, 0.0, -4.446, 0.0, 14.419, 0.0, 14.446, 0.0, 1.27, 0.0, 1.323, 0.0, 0.0, 8.928, 8.838, 0.0, 0.0, -10.632, -10.643, -9.317, -8.87, -8.919, -8.921, 0.0, 0.0, -4.045, -4.045, -6.154, -5.962, -6.081, -6.079, -11.182, -11.291, -11.245, -11.435, 0.0, 0.0, -16.985, -16.985, -9.862, -9.039, -9.245, -9.054]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -4.86709375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0, 5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0, 16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0, 21673.0, 22209.0, 22746.0, 23283.0, 23873.0, 24409.0, 24945.0, 25481.0, 26071.0, 26608.0, 27145.0, 27681.0, 28268.0, 28805.0, 29340.0, 29877.0, 30464.0, 31001.0, 31537.0, 32075.0, 32662.0, 33199.0, 33735.0, 34272.0, 34860.0, 35395.0, 35932.0, 36468.0, 37058.0, 37594.0, 38130.0, 38667.0, 39256.0, 39792.0, 40328.0, 40865.0, 41455.0, 41990.0, 42526.0, 43062.0, 43649.0, 44186.0, 44722.0, 45259.0, 45846.0, 46382.0, 46918.0, 47454.0, 48041.0, 48578.0, 49114.0, 49650.0, 50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -59.9553241419
actual_array: [ 18.1 -119. ]
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689, 3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366, -88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302, -74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585, -75.633, -77.083, -77.507, -77.471, 0.0, 0.0, -95.722, -96.591, 0.0, 0.0, -76.587, -77.321, 0.0, -65.336, 0.0, -65.336, 0.0, -50.256, 0.0, -50.222, 0.0, -65.925, 0.0, -65.776, 0.0, 0.0, -75.932, -76.67, 0.0, 0.0, -86.089, -85.975, -71.469, -68.439, -67.989, -67.54, 0.0, 0.0, -58.864, -58.864, -62.993, -62.524, -63.003, -62.865, -84.958, -87.246, -87.738, -88.685, 0.0, 0.0, -127.737, -127.737, -78.796, -70.901, -69.913, -68.636]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -53.8211145833
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0, 5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0, 10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0, 16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0, 21673.0, 22209.0, 22746.0, 23283.0, 23873.0, 24409.0, 24945.0, 25481.0, 26071.0, 26608.0, 27145.0, 27681.0, 28268.0, 28805.0, 29340.0, 29877.0, 30464.0, 31001.0, 31537.0, 32075.0, 32662.0, 33199.0, 33735.0, 34272.0, 34860.0, 35395.0, 35932.0, 36468.0, 37058.0, 37594.0, 38130.0, 38667.0, 39256.0, 39792.0, 40328.0, 40865.0, 41455.0, 41990.0, 42526.0, 43062.0, 43649.0, 44186.0, 44722.0, 45259.0, 45846.0, 46382.0, 46918.0, 47454.0, 48041.0, 48578.0, 49114.0, 49650.0, 50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1665
varray: [2.465, 2.247, 2.191, 2.16, 2.094, 2.122, 2.102, 2.099, 2.091, 2.094]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: residuals
SAMP_residuals_STD | 0.11696841929
varray: [2.465, 2.247, 2.191, 2.16, 2.094, 2.122, 2.102, 2.099, 2.091, 2.094]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1069
varray: [-1.359, -1.065, -1.203, -1.156, -1.099, -1.028, -0.935, -1.185, -1.104, -0.935]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.12843108831
varray: [-1.359, -1.065, -1.203, -1.156, -1.099, -1.028, -0.935, -1.185, -1.104, -0.935]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 115.1605
varray: [195.615, 168.692, 121.85, 109.646, 97.255, 94.433, 92.164, 91.14, 90.498, 90.312]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 37.2610217259
varray: [195.615, 168.692, 121.85, 109.646, 97.255, 94.433, 92.164, 91.14, 90.498, 90.312]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.7853
varray: [31.07, 28.18, 29.185, 29.36, 27.963, 28.616, 28.333, 28.725, 28.106, 28.315]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.923277019943
varray: [31.07, 28.18, 29.185, 29.36, 27.963, 28.616, 28.333, 28.725, 28.106, 28.315]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: slope_shift
SAMP_d18O_STD | 0.335535657983
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.01681162819
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.9597
varray: [68.606, 64.592, 63.279, 61.922, 61.532, 62.747, 62.466, 60.746, 61.606, 62.101]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.24898807121
varray: [68.606, 64.592, 63.279, 61.922, 61.532, 62.747, 62.466, 60.746, 61.606, 62.101]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1665
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1069
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 115.1605
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.7853
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0219
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0219
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.02116944885
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.2109
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1792
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
dD_mean | -69.5932
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
d18O_stdev | 0.335535657983
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
dD_stdev | 3.01681162819
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1792
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.5932
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.335535657983
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.01681162819
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.5642053777
actual_array: [ 1.93 -15.84]
varray: [-9.995, -9.509, -9.211, -9.023, -9.131, -8.903, -8.932, -8.964, -9.057, -9.067]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.1792
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.464977663
actual_array: [ 18.1 -119. ]
varray: [-77.053, -72.416, -69.991, -69.014, -68.707, -67.895, -67.68, -67.771, -67.716, -67.689]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -69.5932
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2685.0, 3221.0, 3758.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.37645
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 25.61755
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2569
varray: [2.289, 2.269, 2.249, 2.271, 2.251, 2.251, 2.262, 2.244, 2.232, 2.251]
inst: mean_all
operation: mean_all(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: residuals
SAMP_residuals_STD | 0.0161757706325
varray: [2.289, 2.269, 2.249, 2.271, 2.251, 2.251, 2.262, 2.244, 2.232, 2.251]
inst: std_all
operation: std_all(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1613
varray: [-1.14, -1.326, -1.128, -1.131, -1.103, -1.165, -1.136, -1.014, -1.208, -1.262]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0867167163175
varray: [-1.14, -1.326, -1.128, -1.131, -1.103, -1.165, -1.136, -1.014, -1.208, -1.262]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 89.9358
varray: [90.368, 90.403, 90.215, 90.278, 90.016, 89.849, 89.911, 89.507, 89.272, 89.539]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.393934230269
varray: [90.368, 90.403, 90.215, 90.278, 90.016, 89.849, 89.911, 89.507, 89.272, 89.539]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.3692
varray: [30.546, 29.996, 29.102, 29.093, 29.087, 28.925, 29.496, 29.158, 28.76, 29.529]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.543620946371
varray: [30.546, 29.996, 29.102, 29.093, 29.087, 28.925, 29.496, 29.158, 28.76, 29.529]
inst: std_all
operation: std_all(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: slope_shift
SAMP_d18O_STD | 0.327087501035
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.81015983102
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.4939
varray: [64.726, 63.251, 64.525, 65.374, 65.397, 64.707, 64.579, 64.649, 63.272, 64.459]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.727306362164
varray: [64.726, 63.251, 64.525, 65.374, 65.397, 64.707, 64.579, 64.649, 63.272, 64.459]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2569
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1613
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 89.9358
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.3692
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1123
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1123
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.23640772172
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.7451
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.8023
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 16.4884
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.8023
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
dD_mean | 16.4884
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
d18O_stdev | 0.327087501035
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
dD_stdev | 4.81015983102
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.8023
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
measured_dD_mean | 16.4884
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.327087501035
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.81015983102
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [1.917, 2.795, 2.724, 2.966, 2.886, 2.864, 3.083, 3.002, 2.883, 2.903]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.8023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [3.743, 13.99, 15.809, 17.836, 18.215, 18.066, 19.414, 19.305, 19.14, 19.366]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 16.4884
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [5419.0, 5956.0, 6493.0, 7028.0, 7565.0, 8101.0, 8638.0, 9175.0, 9711.0, 10248.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.28605
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.39285
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0323
varray: [2.084, 2.047, 2.048, 2.008, 2.041, 2.011, 2.034, 2.023, 2.011, 2.016]
inst: mean_all
operation: mean_all(residuals)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: residuals
SAMP_residuals_STD | 0.0237208300398
varray: [2.084, 2.047, 2.048, 2.008, 2.041, 2.011, 2.034, 2.023, 2.011, 2.016]
inst: std_all
operation: std_all(residuals)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9357
varray: [-0.91, -0.888, -0.891, -0.907, -0.958, -1.147, -0.89, -1.011, -0.915, -0.84]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0870811243739
varray: [-0.91, -0.888, -0.891, -0.907, -0.958, -1.147, -0.89, -1.011, -0.915, -0.84]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 89.1501
varray: [89.452, 89.183, 89.211, 89.032, 89.24, 89.05, 89.193, 89.23, 88.966, 88.944]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.153687597996
varray: [89.452, 89.183, 89.211, 89.032, 89.24, 89.05, 89.193, 89.23, 88.966, 88.944]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.2093
varray: [28.694, 28.293, 28.336, 28.169, 28.327, 28.39, 28.181, 28.077, 27.817, 27.809]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.266367769989
varray: [28.694, 28.293, 28.336, 28.169, 28.327, 28.39, 28.181, 28.077, 27.817, 27.809]
inst: std_all
operation: std_all(slope_shift)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: slope_shift
SAMP_d18O_STD | 0.439221343845
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 7.09249290526
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.0037
varray: [62.008, 61.439, 61.15, 60.656, 61.233, 58.975, 61.559, 60.64, 61.272, 61.105]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.818636271151
varray: [62.008, 61.439, 61.15, 60.656, 61.233, 58.975, 61.559, 60.64, 61.272, 61.105]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0323
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9357
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 89.1501
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.2093
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1123
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1123
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.23640772172
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7451
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.2385
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
dD_mean | -107.0065
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
d18O_stdev | 0.439221343845
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
dD_stdev | 7.09249290526
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.2385
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.0065
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.439221343845
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
measured_dD_stdev | 7.09249290526
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.051, -14.116, -14.301, -14.332, -14.273, -14.607, -14.321, -14.416, -14.448, -14.52]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.2385
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-88.213, -102.883, -106.783, -108.584, -109.289, -110.865, -110.439, -110.675, -111.032, -111.302]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -107.0065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [10837.0, 11373.0, 11909.0, 12445.0, 12981.0, 13517.0, 14054.0, 14590.0, 15127.0, 15664.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.51065
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.39285
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0686
varray: [2.058, 2.079, 2.065, 2.064, 2.063, 2.072, 2.077, 2.057, 2.095, 2.056]
inst: mean_all
operation: mean_all(residuals)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: residuals
SAMP_residuals_STD | 0.0122854927998
varray: [2.058, 2.079, 2.065, 2.064, 2.063, 2.072, 2.077, 2.057, 2.095, 2.056]
inst: std_all
operation: std_all(residuals)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0193
varray: [-1.024, -0.907, -1.124, -1.11, -1.22, -0.967, -1.034, -0.87, -0.853, -1.084]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.119808226392
varray: [-1.024, -0.907, -1.124, -1.11, -1.22, -0.967, -1.034, -0.87, -0.853, -1.084]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 88.8245
varray: [88.834, 88.914, 88.896, 88.791, 89.025, 88.843, 88.826, 88.496, 88.938, 88.682]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.14756561629
varray: [88.834, 88.914, 88.896, 88.791, 89.025, 88.843, 88.826, 88.496, 88.938, 88.682]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.0449
varray: [27.729, 27.998, 28.318, 28.674, 27.972, 28.176, 28.255, 27.407, 28.108, 27.812]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.350258142898
varray: [27.729, 27.998, 28.318, 28.674, 27.972, 28.176, 28.255, 27.407, 28.108, 27.812]
inst: std_all
operation: std_all(slope_shift)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: slope_shift
SAMP_d18O_STD | 0.139852461942
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.29777134778
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.1508
varray: [60.946, 61.957, 60.495, 60.664, 60.379, 61.385, 60.946, 62.269, 61.7, 60.767]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.645558466239
varray: [60.946, 61.957, 60.495, 60.664, 60.379, 61.385, 60.946, 62.269, 61.7, 60.767]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0686
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0193
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 88.8245
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.0449
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.076
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.076
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.5437843887
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.598
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.0784
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
dD_mean | -68.7385
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
d18O_stdev | 0.139852461942
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
dD_stdev | 2.29777134778
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.0784
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.7385
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.139852461942
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.29777134778
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4590920027
actual_array: [ 1.93 -15.84]
varray: [-9.388, -9.168, -8.985, -9.131, -9.031, -9.06, -8.879, -9.139, -8.964, -9.039]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.0784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.5161176696
actual_array: [ 18.1 -119. ]
varray: [-74.827, -70.174, -68.332, -68.351, -67.826, -67.743, -67.109, -68.022, -67.416, -67.585]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.7385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [16253.0, 16790.0, 17326.0, 17863.0, 18400.0, 18938.0, 19474.0, 20011.0, 20547.0, 21083.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.47435
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.71845
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11275
varray: [2.149, 2.126, 2.093, 2.083]
inst: mean_all
operation: mean_all(residuals)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: residuals
SAMP_residuals_STD | 0.0303575910331
varray: [2.149, 2.126, 2.093, 2.083]
inst: std_all
operation: std_all(residuals)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.19425
varray: [-1.203, -1.314, -1.038, -1.222]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114892340911
varray: [-1.203, -1.314, -1.038, -1.222]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 90.08475
varray: [90.54, 90.254, 89.734, 89.811]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.380290568732
varray: [90.54, 90.254, 89.734, 89.811]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.634
varray: [29.235, 28.842, 28.113, 28.346]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.502941348469
varray: [29.235, 28.842, 28.113, 28.346]
inst: std_all
operation: std_all(slope_shift)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: slope_shift
SAMP_d18O_STD | 0.028359301825
varray: [-7.59, -7.586, -7.599, -7.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.881487190302
varray: [-75.633, -77.083, -77.507, -77.471]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.151
varray: [63.092, 62.321, 62.186, 61.005]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.861959395795
varray: [63.092, 62.321, 62.186, 61.005]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.19425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 90.08475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.634
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.48512543132
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.5978
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-7.59, -7.586, -7.599, -7.536]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-75.633, -77.083, -77.507, -77.471]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.57775
varray: [-7.59, -7.586, -7.599, -7.536]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
dD_mean | -76.9235
varray: [-75.633, -77.083, -77.507, -77.471]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
d18O_stdev | 0.028359301825
varray: [-7.59, -7.586, -7.599, -7.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
dD_stdev | 0.881487190302
varray: [-75.633, -77.083, -77.507, -77.471]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.57775
varray: [-7.59, -7.586, -7.599, -7.536]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
measured_dD_mean | -76.9235
varray: [-75.633, -77.083, -77.507, -77.471]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.028359301825
varray: [-7.59, -7.586, -7.599, -7.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.881487190302
varray: [-75.633, -77.083, -77.507, -77.471]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.89422706094
actual_array: [ 1.93 -15.84]
varray: [-7.59, -7.586, -7.599, -7.536]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -7.57775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -85.6028370402
actual_array: [ 18.1 -119. ]
varray: [-75.633, -77.083, -77.507, -77.471]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -76.9235
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [21673.0, 22209.0, 22746.0, 23283.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.4302
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.5418
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.27375
varray: [0.0, 0.0, 4.537, 4.558]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: residuals
SAMP_residuals_STD | 2.62551434643
varray: [0.0, 0.0, 4.537, 4.558]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.7115
varray: [0.0, 0.0, -9.422, -9.424]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 5.44037164784
varray: [0.0, 0.0, -9.422, -9.424]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 444.587
varray: [0.0, 0.0, 889.166, 889.182]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 513.364848298
varray: [0.0, 0.0, 889.166, 889.182]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -7.62025
varray: [0.0, 0.0, -15.123, -15.358]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: slope_shift
SAMP_slope_shift_STD | 8.79962977914
varray: [0.0, 0.0, -15.123, -15.358]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: slope_shift
SAMP_d18O_STD | 1.084107467
varray: [0.0, 0.0, 1.914, 1.84]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 55.5171146967
varray: [0.0, 0.0, -95.722, -96.591]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 27.26075
varray: [0.0, 0.0, 53.948, 55.095]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 31.4814853882
varray: [0.0, 0.0, 53.948, 55.095]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.27375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.7115
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 444.587
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 7.62025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.12915
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12915
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.02210202369
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.48805
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 0.0, 1.914, 1.84]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, 0.0, -95.722, -96.591]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.9385
varray: [0.0, 0.0, 1.914, 1.84]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
dD_mean | -48.07825
varray: [0.0, 0.0, -95.722, -96.591]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
d18O_stdev | 1.084107467
varray: [0.0, 0.0, 1.914, 1.84]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
dD_stdev | 55.5171146967
varray: [0.0, 0.0, -95.722, -96.591]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.9385
varray: [0.0, 0.0, 1.914, 1.84]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.07825
varray: [0.0, 0.0, -95.722, -96.591]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.084107467
varray: [0.0, 0.0, 1.914, 1.84]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
measured_dD_stdev | 55.5171146967
varray: [0.0, 0.0, -95.722, -96.591]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.0135546453218
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 1.914, 1.84]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.9385
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.5797836591
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -95.722, -96.591]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -48.07825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [23873.0, 24409.0, 24945.0, 25481.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.2692
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 355.04405
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.952
varray: [0.0, 0.0, 5.873, 5.935]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: residuals
SAMP_residuals_STD | 3.40876996388
varray: [0.0, 0.0, 5.873, 5.935]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -6.46925
varray: [0.0, 0.0, -13.125, -12.752]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 7.47159837317
varray: [0.0, 0.0, -13.125, -12.752]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 696.26525
varray: [0.0, 0.0, 1390.809, 1394.252]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 803.979087737
varray: [0.0, 0.0, 1390.809, 1394.252]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -20.1935
varray: [0.0, 0.0, -40.182, -40.592]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: slope_shift
SAMP_slope_shift_STD | 23.3180460802
varray: [0.0, 0.0, -40.182, -40.592]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: slope_shift
SAMP_d18O_STD | 6.32577807467
varray: [0.0, 0.0, 10.984, 10.929]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 44.430423113
varray: [0.0, 0.0, -76.587, -77.321]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 25.85525
varray: [0.0, 0.0, 50.175, 53.246]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 29.8813840083
varray: [0.0, 0.0, 50.175, 53.246]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.952
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 6.46925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 696.26525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.1935
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.8074
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.8074
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 37.6480462557
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -36.89355
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 0.0, 10.984, 10.929]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, 0.0, -76.587, -77.321]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.47825
varray: [0.0, 0.0, 10.984, 10.929]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
dD_mean | -38.477
varray: [0.0, 0.0, -76.587, -77.321]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
d18O_stdev | 6.32577807467
varray: [0.0, 0.0, 10.984, 10.929]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
dD_stdev | 44.430423113
varray: [0.0, 0.0, -76.587, -77.321]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.47825
varray: [0.0, 0.0, 10.984, 10.929]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.477
varray: [0.0, 0.0, -76.587, -77.321]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 6.32577807467
varray: [0.0, 0.0, 10.984, 10.929]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
measured_dD_stdev | 44.430423113
varray: [0.0, 0.0, -76.587, -77.321]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.72045769565
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 10.984, 10.929]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 5.47825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.9207898464
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -76.587, -77.321]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -38.477
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [26071.0, 26608.0, 27145.0, 27681.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 86.59095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 606.7223
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.0685
varray: [0.0, 2.137, 0.0, 2.137]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: residuals
SAMP_residuals_STD | 1.23379752526
varray: [0.0, 2.137, 0.0, 2.137]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.7715
varray: [0.0, -1.543, 0.0, -1.543]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.89085146536
varray: [0.0, -1.543, 0.0, -1.543]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 47.9225
varray: [0.0, 95.845, 0.0, 95.845]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 55.3361365505
varray: [0.0, 95.845, 0.0, 95.845]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.6265
varray: [0.0, 27.253, 0.0, 27.253]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: slope_shift
SAMP_slope_shift_STD | 15.7345268862
varray: [0.0, 27.253, 0.0, 27.253]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: slope_shift
SAMP_d18O_STD | 2.56689929682
varray: [0.0, -4.446, 0.0, -4.446]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 37.7217571878
varray: [0.0, -65.336, 0.0, -65.336]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 30.5935
varray: [0.0, 61.187, 0.0, 61.187]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 35.3263309209
varray: [0.0, 61.187, 0.0, 61.187]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.0685
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.7715
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 47.9225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.6265
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.0761
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.0761
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.1771892194
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -32.1553
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, -4.446, 0.0, -4.446]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, -65.336, 0.0, -65.336]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.223
varray: [0.0, -4.446, 0.0, -4.446]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
dD_mean | -32.668
varray: [0.0, -65.336, 0.0, -65.336]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
d18O_stdev | 2.56689929682
varray: [0.0, -4.446, 0.0, -4.446]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
dD_stdev | 37.7217571878
varray: [0.0, -65.336, 0.0, -65.336]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.223
varray: [0.0, -4.446, 0.0, -4.446]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.668
varray: [0.0, -65.336, 0.0, -65.336]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.56689929682
varray: [0.0, -4.446, 0.0, -4.446]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
measured_dD_stdev | 37.7217571878
varray: [0.0, -65.336, 0.0, -65.336]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.3103397141
actual_array: [ 1.93 -15.84]
varray: [0.0, -4.446, 0.0, -4.446]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -2.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.4718279864
actual_array: [ 18.1 -119. ]
varray: [0.0, -65.336, 0.0, -65.336]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -32.668
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [28268.0, 28805.0, 29340.0, 29877.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 88.47445
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -41.62045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.02425
varray: [0.0, 6.027, 0.0, 6.07]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: residuals
SAMP_residuals_STD | 3.49214722628
varray: [0.0, 6.027, 0.0, 6.07]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -6.472
varray: [0.0, -13.168, 0.0, -12.72]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 7.47545958453
varray: [0.0, -13.168, 0.0, -12.72]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 555.53375
varray: [0.0, 1110.175, 0.0, 1111.96]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 641.475534132
varray: [0.0, 1110.175, 0.0, 1111.96]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -15.617
varray: [0.0, -31.174, 0.0, -31.294]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: slope_shift
SAMP_slope_shift_STD | 18.0330248526
varray: [0.0, -31.174, 0.0, -31.294]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: slope_shift
SAMP_d18O_STD | 8.33261505071
varray: [0.0, 14.419, 0.0, 14.446]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 29.005503495
varray: [0.0, -50.256, 0.0, -50.222]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 28.27125
varray: [0.0, 55.673, 0.0, 57.412]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 32.6525464385
varray: [0.0, 55.673, 0.0, 57.412]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.02425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 6.472
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 555.53375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 15.617
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.87965
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.87965
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 41.0169728621
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -34.47755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 14.419, 0.0, 14.446]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, -50.256, 0.0, -50.222]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 7.21625
varray: [0.0, 14.419, 0.0, 14.446]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
dD_mean | -25.1195
varray: [0.0, -50.256, 0.0, -50.222]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
d18O_stdev | 8.33261505071
varray: [0.0, 14.419, 0.0, 14.446]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
dD_stdev | 29.005503495
varray: [0.0, -50.256, 0.0, -50.222]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
measured_d18O_mean | 7.21625
varray: [0.0, 14.419, 0.0, 14.446]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.1195
varray: [0.0, -50.256, 0.0, -50.222]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 8.33261505071
varray: [0.0, 14.419, 0.0, 14.446]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
measured_dD_stdev | 29.005503495
varray: [0.0, -50.256, 0.0, -50.222]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 6.53282918055
actual_array: [ 1.93 -15.84]
varray: [0.0, 14.419, 0.0, 14.446]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 7.21625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -28.0917301038
actual_array: [ 18.1 -119. ]
varray: [0.0, -50.256, 0.0, -50.222]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -25.1195
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [30464.0, 31001.0, 31537.0, 32075.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 86.5187
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 465.9908
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.73475
varray: [0.0, 3.483, 0.0, 3.456]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: residuals
SAMP_residuals_STD | 2.00314708646
varray: [0.0, 3.483, 0.0, 3.456]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.09725
varray: [0.0, -6.226, 0.0, -6.163]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 3.57648872266
varray: [0.0, -6.226, 0.0, -6.163]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 224.6205
varray: [0.0, 450.308, 0.0, 448.174]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 259.370875427
varray: [0.0, 450.308, 0.0, 448.174]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 3.00825
varray: [0.0, 6.052, 0.0, 5.981]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.47374882752
varray: [0.0, 6.052, 0.0, 5.981]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: slope_shift
SAMP_d18O_STD | 0.748847280826
varray: [0.0, 1.27, 0.0, 1.323]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 38.0188525636
varray: [0.0, -65.925, 0.0, -65.776]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 30.0445
varray: [0.0, 61.067, 0.0, 59.111]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 34.7015892479
varray: [0.0, 61.067, 0.0, 59.111]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.73475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.09725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 224.6205
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.00825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.40985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.40985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -19.1107898909
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -32.7043
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 1.27, 0.0, 1.323]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, -65.925, 0.0, -65.776]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.64825
varray: [0.0, 1.27, 0.0, 1.323]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
dD_mean | -32.92525
varray: [0.0, -65.925, 0.0, -65.776]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
d18O_stdev | 0.748847280826
varray: [0.0, 1.27, 0.0, 1.323]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
dD_stdev | 38.0188525636
varray: [0.0, -65.925, 0.0, -65.776]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.64825
varray: [0.0, 1.27, 0.0, 1.323]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.92525
varray: [0.0, -65.925, 0.0, -65.776]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.748847280826
varray: [0.0, 1.27, 0.0, 1.323]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
measured_dD_stdev | 38.0188525636
varray: [0.0, -65.925, 0.0, -65.776]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.316224854467
actual_array: [ 1.93 -15.84]
varray: [0.0, 1.27, 0.0, 1.323]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.64825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.757418525
actual_array: [ 18.1 -119. ]
varray: [0.0, -65.925, 0.0, -65.776]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -32.92525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [32662.0, 33199.0, 33735.0, 34272.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.8082
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 135.07755
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8025
varray: [0.0, 0.0, 5.614, 5.596]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: residuals
SAMP_residuals_STD | 3.23605660231
varray: [0.0, 0.0, 5.614, 5.596]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -6.0475
varray: [0.0, 0.0, -11.986, -12.204]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 6.98361861788
varray: [0.0, 0.0, -11.986, -12.204]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 463.72675
varray: [0.0, 0.0, 927.567, 927.34]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 535.465535905
varray: [0.0, 0.0, 927.567, 927.34]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -11.182
varray: [0.0, 0.0, -22.527, -22.201]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: slope_shift
SAMP_slope_shift_STD | 12.9125473087
varray: [0.0, 0.0, -22.527, -22.201]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: slope_shift
SAMP_d18O_STD | 5.12873405433
varray: [0.0, 0.0, 8.928, 8.838]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 44.0534331731
varray: [0.0, 0.0, -75.932, -76.67]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 28.74225
varray: [0.0, 0.0, 57.708, 57.261]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.1891932449
varray: [0.0, 0.0, 57.708, 57.261]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 6.0475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 463.72675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 11.182
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.6579
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.6579
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 30.6770493332
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -34.00655
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 0.0, 8.928, 8.838]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, 0.0, -75.932, -76.67]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.4415
varray: [0.0, 0.0, 8.928, 8.838]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
dD_mean | -38.1505
varray: [0.0, 0.0, -75.932, -76.67]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
d18O_stdev | 5.12873405433
varray: [0.0, 0.0, 8.928, 8.838]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
dD_stdev | 44.0534331731
varray: [0.0, 0.0, -75.932, -76.67]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.4415
varray: [0.0, 0.0, 8.928, 8.838]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.1505
varray: [0.0, 0.0, -75.932, -76.67]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.12873405433
varray: [0.0, 0.0, 8.928, 8.838]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
measured_dD_stdev | 44.0534331731
varray: [0.0, 0.0, -75.932, -76.67]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.63934369278
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 8.928, 8.838]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 4.4415
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.5583202221
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -75.932, -76.67]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -38.1505
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [34860.0, 35395.0, 35932.0, 36468.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 86.74045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 374.1838
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.009
varray: [0.0, 0.0, 2.016, 2.02]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: residuals
SAMP_residuals_STD | 1.16509398763
varray: [0.0, 0.0, 2.016, 2.02]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.5925
varray: [0.0, 0.0, -1.166, -1.204]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.684335931153
varray: [0.0, 0.0, -1.166, -1.204]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.25225
varray: [0.0, 0.0, 92.695, 92.314]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 53.4077244749
varray: [0.0, 0.0, 92.695, 92.314]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.716
varray: [0.0, 0.0, 27.068, 27.796]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: slope_shift
SAMP_slope_shift_STD | 15.840660929
varray: [0.0, 0.0, 27.068, 27.796]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: slope_shift
SAMP_d18O_STD | 6.14156513032
varray: [0.0, 0.0, -10.632, -10.643]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 49.6706201626
varray: [0.0, 0.0, -86.089, -85.975]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 29.3525
varray: [0.0, 0.0, 58.886, 58.524]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.8936697482
varray: [0.0, 0.0, 58.886, 58.524]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.009
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.25225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.716
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1356
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1356
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -52.9515993658
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.3963
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 0.0, -10.632, -10.643]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, 0.0, -86.089, -85.975]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.31875
varray: [0.0, 0.0, -10.632, -10.643]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
dD_mean | -43.016
varray: [0.0, 0.0, -86.089, -85.975]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
d18O_stdev | 6.14156513032
varray: [0.0, 0.0, -10.632, -10.643]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
dD_stdev | 49.6706201626
varray: [0.0, 0.0, -86.089, -85.975]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.31875
varray: [0.0, 0.0, -10.632, -10.643]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.016
varray: [0.0, 0.0, -86.089, -85.975]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 6.14156513032
varray: [0.0, 0.0, -10.632, -10.643]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
measured_dD_stdev | 49.6706201626
varray: [0.0, 0.0, -86.089, -85.975]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.53856124712
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -10.632, -10.643]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -5.31875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -47.9598392322
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -86.089, -85.975]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -43.016
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [37058.0, 37594.0, 38130.0, 38667.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 88.53395
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -43.2907
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0795
varray: [2.069, 2.089, 2.066, 2.094]
inst: mean_all
operation: mean_all(residuals)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: residuals
SAMP_residuals_STD | 0.014059397806
varray: [2.069, 2.089, 2.066, 2.094]
inst: std_all
operation: std_all(residuals)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9675
varray: [-0.861, -0.94, -1.0, -1.069]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0884251095561
varray: [-0.861, -0.94, -1.0, -1.069]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 91.29775
varray: [91.208, 91.704, 91.563, 90.716]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.440410698477
varray: [91.208, 91.704, 91.563, 90.716]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.88475
varray: [27.718, 27.667, 27.737, 28.417]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.356062143453
varray: [27.718, 27.667, 27.737, 28.417]
inst: std_all
operation: std_all(slope_shift)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: slope_shift
SAMP_d18O_STD | 0.208173605435
varray: [-9.317, -8.87, -8.919, -8.921]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.77812267575
varray: [-71.469, -68.439, -67.989, -67.54]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.97875
varray: [61.889, 62.599, 62.112, 61.315]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.532630187529
varray: [61.889, 62.599, 62.112, 61.315]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0795
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 91.29775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.88475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0651
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0651
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.03553110137
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.77005
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-9.317, -8.87, -8.919, -8.921]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-71.469, -68.439, -67.989, -67.54]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.00675
varray: [-9.317, -8.87, -8.919, -8.921]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
dD_mean | -68.85925
varray: [-71.469, -68.439, -67.989, -67.54]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
d18O_stdev | 0.208173605435
varray: [-9.317, -8.87, -8.919, -8.921]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
dD_stdev | 1.77812267575
varray: [-71.469, -68.439, -67.989, -67.54]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.00675
varray: [-9.317, -8.87, -8.919, -8.921]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.85925
varray: [-71.469, -68.439, -67.989, -67.54]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.208173605435
varray: [-9.317, -8.87, -8.919, -8.921]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.77812267575
varray: [-71.469, -68.439, -67.989, -67.54]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3843759976
actual_array: [ 1.93 -15.84]
varray: [-9.317, -8.87, -8.919, -8.921]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.00675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.6501703714
actual_array: [ 18.1 -119. ]
varray: [-71.469, -68.439, -67.989, -67.54]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.85925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [39256.0, 39792.0, 40328.0, 40865.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.46345
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.7548
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.026
varray: [0.0, 0.0, 2.052, 2.052]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: residuals
SAMP_residuals_STD | 1.18472275238
varray: [0.0, 0.0, 2.052, 2.052]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.598
varray: [0.0, 0.0, -1.196, -1.196]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.690510921951
varray: [0.0, 0.0, -1.196, -1.196]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 49.982
varray: [0.0, 0.0, 99.964, 99.964]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 57.7142423093
varray: [0.0, 0.0, 99.964, 99.964]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.5
varray: [0.0, 0.0, 27.0, 27.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: slope_shift
SAMP_slope_shift_STD | 15.5884572681
varray: [0.0, 0.0, 27.0, 27.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: slope_shift
SAMP_d18O_STD | 2.33538183887
varray: [0.0, 0.0, -4.045, -4.045]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 33.9851462456
varray: [0.0, 0.0, -58.864, -58.864]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 30.0505
varray: [0.0, 0.0, 60.101, 60.101]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 34.6993285286
varray: [0.0, 0.0, 60.101, 60.101]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.026
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.598
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 49.982
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.5
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1186
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1186
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -52.1589107526
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -32.6983
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 0.0, -4.045, -4.045]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, 0.0, -58.864, -58.864]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.0225
varray: [0.0, 0.0, -4.045, -4.045]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
dD_mean | -29.432
varray: [0.0, 0.0, -58.864, -58.864]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
d18O_stdev | 2.33538183887
varray: [0.0, 0.0, -4.045, -4.045]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
dD_stdev | 33.9851462456
varray: [0.0, 0.0, -58.864, -58.864]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.0225
varray: [0.0, 0.0, -4.045, -4.045]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
measured_dD_mean | -29.432
varray: [0.0, 0.0, -58.864, -58.864]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.33538183887
varray: [0.0, 0.0, -4.045, -4.045]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
measured_dD_stdev | 33.9851462456
varray: [0.0, 0.0, -58.864, -58.864]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.10126003474
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -4.045, -4.045]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -2.0225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.8793265957
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -58.864, -58.864]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -29.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [41455.0, 41990.0, 42526.0, 43062.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 88.51695
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -39.56095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0385
varray: [2.033, 2.033, 2.051, 2.037]
inst: mean_all
operation: mean_all(residuals)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: residuals
SAMP_residuals_STD | 0.00854400374532
varray: [2.033, 2.033, 2.051, 2.037]
inst: std_all
operation: std_all(residuals)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.09425
varray: [-1.171, -1.164, -0.8, -1.242]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.199306087882
varray: [-1.171, -1.164, -0.8, -1.242]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 99.26375
varray: [99.256, 99.342, 99.055, 99.402]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.151517600738
varray: [99.256, 99.342, 99.055, 99.402]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.855
varray: [26.942, 27.1, 26.332, 27.046]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.354778804327
varray: [26.942, 27.1, 26.332, 27.046]
inst: std_all
operation: std_all(slope_shift)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0794103267844
varray: [-6.154, -5.962, -6.081, -6.079]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.223832340529
varray: [-62.993, -62.524, -63.003, -62.865]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.065
varray: [60.059, 59.348, 61.545, 59.308]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.04523904762
varray: [60.059, 59.348, 61.545, 59.308]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0385
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.09425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 99.26375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.855
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1061
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1061
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.94730952159
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6838
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-6.154, -5.962, -6.081, -6.079]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-62.993, -62.524, -63.003, -62.865]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.069
varray: [-6.154, -5.962, -6.081, -6.079]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
dD_mean | -62.84625
varray: [-62.993, -62.524, -63.003, -62.865]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0794103267844
varray: [-6.154, -5.962, -6.081, -6.079]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
dD_stdev | 0.223832340529
varray: [-62.993, -62.524, -63.003, -62.865]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.069
varray: [-6.154, -5.962, -6.081, -6.079]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.84625
varray: [-62.993, -62.524, -63.003, -62.865]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0794103267844
varray: [-6.154, -5.962, -6.081, -6.079]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.223832340529
varray: [-62.993, -62.524, -63.003, -62.865]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.32091550866
actual_array: [ 1.93 -15.84]
varray: [-6.154, -5.962, -6.081, -6.079]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -6.069
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.9747343817
actual_array: [ 18.1 -119. ]
varray: [-62.993, -62.524, -63.003, -62.865]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -62.84625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [43649.0, 44186.0, 44722.0, 45259.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.50445
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 9.7208
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.02475
varray: [2.014, 2.034, 2.024, 2.027]
inst: mean_all
operation: mean_all(residuals)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: residuals
SAMP_residuals_STD | 0.00830160627027
varray: [2.014, 2.034, 2.024, 2.027]
inst: std_all
operation: std_all(residuals)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.964
varray: [-0.986, -0.912, -1.086, -0.872]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0940496322871
varray: [-0.986, -0.912, -1.086, -0.872]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 89.496
varray: [89.515, 89.538, 89.644, 89.287]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.150233152134
varray: [89.515, 89.538, 89.644, 89.287]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.55225
varray: [27.571, 27.461, 27.774, 27.403]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.163428628663
varray: [27.571, 27.461, 27.774, 27.403]
inst: std_all
operation: std_all(slope_shift)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: slope_shift
SAMP_d18O_STD | 0.107552700261
varray: [-11.182, -11.291, -11.245, -11.435]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.58281044875
varray: [-84.958, -87.246, -87.738, -88.685]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.68175
varray: [60.273, 61.715, 59.73, 61.009]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.865563929855
varray: [60.273, 61.715, 59.73, 61.009]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.02475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.964
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 89.496
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.55225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.58845472349
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.06705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-11.182, -11.291, -11.245, -11.435]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-84.958, -87.246, -87.738, -88.685]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.28825
varray: [-11.182, -11.291, -11.245, -11.435]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
dD_mean | -87.15675
varray: [-84.958, -87.246, -87.738, -88.685]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
d18O_stdev | 0.107552700261
varray: [-11.182, -11.291, -11.245, -11.435]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
dD_stdev | 1.58281044875
varray: [-84.958, -87.246, -87.738, -88.685]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.28825
varray: [-11.182, -11.291, -11.245, -11.435]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
measured_dD_mean | -87.15675
varray: [-84.958, -87.246, -87.738, -88.685]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.107552700261
varray: [-11.182, -11.291, -11.245, -11.435]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.58281044875
varray: [-84.958, -87.246, -87.738, -88.685]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.7635046183
actual_array: [ 1.93 -15.84]
varray: [-11.182, -11.291, -11.245, -11.435]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -11.28825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -96.9634565881
actual_array: [ 18.1 -119. ]
varray: [-84.958, -87.246, -87.738, -88.685]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -87.15675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [45846.0, 46382.0, 46918.0, 47454.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.5182
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.04695
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.9865
varray: [0.0, 0.0, 1.973, 1.973]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: residuals
SAMP_residuals_STD | 1.13911208111
varray: [0.0, 0.0, 1.973, 1.973]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.4745
varray: [0.0, 0.0, -0.949, -0.949]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.547905405461
varray: [0.0, 0.0, -0.949, -0.949]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 44.5725
varray: [0.0, 0.0, 89.145, 89.145]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 51.4678897469
varray: [0.0, 0.0, 89.145, 89.145]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 14.0085
varray: [0.0, 0.0, 28.017, 28.017]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: slope_shift
SAMP_slope_shift_STD | 16.1756224919
varray: [0.0, 0.0, 28.017, 28.017]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: slope_shift
SAMP_d18O_STD | 9.80629432219
varray: [0.0, 0.0, -16.985, -16.985]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 73.7489913355
varray: [0.0, 0.0, -127.737, -127.737]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 29.49
varray: [0.0, 0.0, 58.98, 58.98]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 34.0521188768
varray: [0.0, 0.0, 58.98, 58.98]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.9865
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4745
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 44.5725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.0085
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1581
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1581
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -54.0007460599
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.2588
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [0.0, 0.0, -16.985, -16.985]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [0.0, 0.0, -127.737, -127.737]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.4925
varray: [0.0, 0.0, -16.985, -16.985]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
dD_mean | -63.8685
varray: [0.0, 0.0, -127.737, -127.737]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
d18O_stdev | 9.80629432219
varray: [0.0, 0.0, -16.985, -16.985]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
dD_stdev | 73.7489913355
varray: [0.0, 0.0, -127.737, -127.737]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.4925
varray: [0.0, 0.0, -16.985, -16.985]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.8685
varray: [0.0, 0.0, -127.737, -127.737]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 9.80629432219
varray: [0.0, 0.0, -16.985, -16.985]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
measured_dD_stdev | 73.7489913355
varray: [0.0, 0.0, -127.737, -127.737]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.84812051077
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -16.985, -16.985]
fit_rsq: 1.0
fit_intercept: 0.951498424311
except: True
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.4925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.1096029067
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -127.737, -127.737]
fit_rsq: 1.0
fit_intercept: 0.184551057622
except: True
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -63.8685
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [48041.0, 48578.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 88.55645
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -44.97045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.077
varray: [2.091, 2.069, 2.065, 2.083]
inst: mean_all
operation: mean_all(residuals)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: residuals
SAMP_residuals_STD | 0.0121106014164
varray: [2.091, 2.069, 2.065, 2.083]
inst: std_all
operation: std_all(residuals)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00275
varray: [-0.855, -1.027, -1.123, -1.006]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11088845747
varray: [-0.855, -1.027, -1.123, -1.006]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 89.364
varray: [89.6, 89.282, 89.236, 89.338]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.162767728169
varray: [89.6, 89.282, 89.236, 89.338]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.411
varray: [28.717, 28.214, 28.407, 28.306]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.218697660405
varray: [28.717, 28.214, 28.407, 28.306]
inst: std_all
operation: std_all(slope_shift)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: slope_shift
SAMP_d18O_STD | 0.386223596724
varray: [-9.862, -9.039, -9.245, -9.054]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.58440665008
varray: [-78.796, -70.901, -69.913, -68.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.04325
varray: [62.27, 60.268, 60.545, 61.09]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.88625932059
varray: [62.27, 60.268, 60.545, 61.09]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.077
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 89.364
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.411
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0676
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0676
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.15210295626
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.70555
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2385
varray: [-9.862, -9.039, -9.245, -9.054]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.0065
varray: [-78.796, -70.901, -69.913, -68.636]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.3
varray: [-9.862, -9.039, -9.245, -9.054]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
dD_mean | -72.0615
varray: [-78.796, -70.901, -69.913, -68.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
d18O_stdev | 0.386223596724
varray: [-9.862, -9.039, -9.245, -9.054]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
dD_stdev | 4.58440665008
varray: [-78.796, -70.901, -69.913, -68.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.3
varray: [-9.862, -9.039, -9.245, -9.054]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.0615
varray: [-78.796, -70.901, -69.913, -68.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.386223596724
varray: [-9.862, -9.039, -9.245, -9.054]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.58440665008
varray: [-78.796, -70.901, -69.913, -68.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.690174581
actual_array: [ 1.93 -15.84]
varray: [-9.862, -9.039, -9.245, -9.054]
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.3
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.2052036157
actual_array: [ 18.1 -119. ]
varray: [-78.796, -70.901, -69.913, -68.636]
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -72.0615
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: [50238.0, 50774.0, 51311.0, 51847.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 87.46595
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.17895
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.995
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.4149146167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.995
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.7466037059
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -77.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.509
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.416
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9081179874
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.509
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.5987579244
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -72.416
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.211
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.991
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5973661448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.211
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.906602216
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -69.991
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.023
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4013213582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8219687615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -69.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.131
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.707
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5139428313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.131
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.4811473996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.707
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.903
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.895
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.276186388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.903
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.57969155
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.895
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.68
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3064273391
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.34100558
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.68
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3397966645
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4420308045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.057
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.716
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4367762664
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.057
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3809716029
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.716
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.067
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.689
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4472041806
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.067
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3509970857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.689
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.917
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.743
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.00681675743
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 1.917
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 3.95047366328
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 3.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.795
operation: col(d(18_16)Mean)
inst: col
dD_actual | 13.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.92238762265
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.795
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.3263580115
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 13.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.809
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.84834943195
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.3457523347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 15.809
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.966
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.836
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10070495517
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5960614568
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 17.836
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.886
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.215
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.01728164171
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.886
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.0168148644
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 18.215
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.066
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.99434023051
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.8513999364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 18.066
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.083
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.414
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.2227115511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 3.083
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.3479054601
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 19.414
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.13824544622
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 3.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.2268972241
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 19.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.883
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.14
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.01415326745
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.883
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.0437196192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 19.14
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.903
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.366
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8023
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 16.4884
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.903
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03500909582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 2.903
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.2946174295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 19.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.051
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.213
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.6016851908
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -13.051
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.1360708013
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -88.213
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.883
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7122580513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.422225128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -102.883
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.301
operation: col(d(18_16)Mean)
inst: col
dD_actual | -106.783
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9051744636
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.301
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.751877608
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -106.783
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.332
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.584
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9375009976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.332
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.75128892
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -108.584
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.273
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.289
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8759763039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.533956868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -109.289
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.607
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.2242686376
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.607
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.283580537
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -110.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.321
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.926030292
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.321
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.810649266
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -110.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.416
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.675
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0250954767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.416
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.072648749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -110.675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.032
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0584648021
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.468978476
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -111.032
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.52
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.302
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1335457842
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -14.52
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.768723648
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -111.302
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.388
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.827
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7819402258
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.388
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.2753712906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -74.827
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.168
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5525261138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.168
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.1097628323
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -70.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.985
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.332
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3616952843
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0648346612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.332
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.131
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.351
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5139428313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.131
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0859278399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.351
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.031
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.826
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4096636895
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.031
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.5030900061
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.826
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.06
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.743
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4399046406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.06
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.41094612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.879
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.109
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2511593939
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.879
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7071000503
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.109
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.139
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.022
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5222851627
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.139
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7206827974
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.022
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.416
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3397966645
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0479214121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.416
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.585
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4180060208
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2355396863
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.585
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.59
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.633
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.90700125581
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -7.59
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.1701661364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -75.633
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.586
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.083
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.90283009014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -7.586
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.779908725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -77.083
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.599
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.507
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.91638637857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -7.599
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.2506196612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -77.507
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.471
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.85069051922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -7.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.2106536383
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -77.471
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.722
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.00368838317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 1.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -106.472317075
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -95.722
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.84
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.591
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.926521818224
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 1.84
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -107.437052461
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -96.591
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 10.984
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.587
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.4618065466
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 10.984
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.229265743
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -76.587
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 10.929
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.321
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.4044530186
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 10.929
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.044128543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -77.321
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.62846503685
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -4.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.738773423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -65.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.62846503685
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -4.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.738773423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -65.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 14.419
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.256
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 14.0437950683
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 14.419
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.997450502
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -50.256
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 14.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 14.0719504366
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 14.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.9597048137
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -50.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.27
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.332130709826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 1.27
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3926619642
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -65.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.323
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.776
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.387398654993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 1.323
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2272470361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -65.776
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.928
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.932
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.31782739073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 8.928
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.5021061598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -75.932
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.838
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.22397616309
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 8.838
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.3214096291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -76.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.0791727501
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -10.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.7780754509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -86.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.975
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.0906434557
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -10.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.6515163784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -85.975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.317
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.469
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7079020351
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.317
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.547429489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -71.469
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.87
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2417742712
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.87
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1836225626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.919
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.989
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2928710507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.919
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6840472764
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.989
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.921
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.54
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2949566335
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -8.921
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1855821576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -67.54
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.864
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.21030567814
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -4.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.5537706415
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -58.864
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.864
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.21030567814
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -4.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.5537706415
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -58.864
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.154
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.40955277921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -6.154
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.1376514334
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -62.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.962
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.524
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.20933682691
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -5.962
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6169829685
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -62.524
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.003
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.33342900568
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -6.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.1487531064
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -63.003
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.079
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.33134342284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -6.079
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.9955500187
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -62.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.182
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.958
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.6527080301
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -11.182
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.5224762318
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -84.958
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.291
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.246
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.7663722947
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -11.291
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.0625390198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -87.246
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.738
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.7184038895
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -11.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.6087413326
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -87.738
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.435
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.916534259
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -11.435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.6600697681
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -88.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.992214391343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.204882549806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -16.985
operation: col(d(18_16)Mean)
inst: col
dD_actual | -127.737
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -18.7040266302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -16.985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -142.014323264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -127.737
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -16.985
operation: col(d(18_16)Mean)
inst: col
dD_actual | -127.737
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -18.7040266302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -16.985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -142.014323264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -127.737
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:32:30
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.862
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.796
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2762233581
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.6816253141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -78.796
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.901
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4180060208
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.9168544612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -70.901
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.913
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.632821053
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.8200091664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -69.913
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8023
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 16.4884
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.52
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.302
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2385
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.0065
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.054
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4336478921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.951498424311
measure_array: [ 2.8023 -14.2385]
inst: fit_type1
actual_mean: -9.054
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.958964546989
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.4023255211
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.184551057622
measure_array: [ 16.4884 -107.0065]
inst: fit_type1
actual_mean: -68.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.900765134938
tarray: []
source_col: d(D_H)Mean