## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
210730_21-030_20-213_Precip_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.7459 | -70.5747 |  |  |  | 2.4542 | 20.0104
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.37902 | 4.929
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.53952 | 7.4191
4 |  | EV | -10.4442 | -75.5283 |  |  |  | 0.16235 | 2.4448
5 | * | NIWO2-B-5 | -2.6906 | -16.993 |  |  |  | 0 | 0
6 | * | NIWO2-B-4 | -2.6906 | -16.993 |  |  |  | 0 | 0
7 |  | NIWO2-010-2 | -10.3937 | -97.195 |  |  |  | 0.21742 | 0.30308
8 |  | NIWO2-1020-2 | -6.9977 | -78.8123 |  |  |  | 0.14925 | 1.3122
9 | * | NIWO2-E-5 | -2.6906 | -16.993 |  |  |  | 0 | 0
10 | * | NIWO2-2030-3 | -2.6906 | -16.993 |  |  |  | 0 | 0
11 |  | NIWO2-D-5 | -7.2725 | -94.9471 |  |  |  | 0.31638 | 0.40786
12 | * | NIWO2-3040-3 | -2.6906 | -16.993 |  |  |  | 0 | 0
13 |  | EV | -10.4292 | -76.5781 |  |  |  | 0.13954 | 2.6481
14 | * | NIWO2-D-3 | -7.4675 | -65.7534 |  |  |  | 5.4334 | 59.198
15 |  | SLCP-21-015 | -4.9364 | -44.2942 |  |  |  | 0.32105 | 5.1417
16 |  | 262 | -11.5223 | -93.947 |  |  |  | 0.31059 | 3.9597
17 |  | 237 | -11.0162 | -97.6663 |  |  |  | 0.072844 | 0.11277
18 |  | 243 | -11.8857 | -68.1967 |  |  |  | 0.07908 | 2.0732
19 |  | 250 | -4.7009 | -69.6027 | -3.7275 | 778.6451 | 2.922 | 0.1466 | 0.53477 | -35.1573 | -3.1132
20 |  | 258 | -11.4713 | -89.9112 |  |  |  | 0.13192 | 1.1603
21 | * | 248 | -5.9744 | -51.5243 |  |  |  | 3.7352 | 41.9232
22 |  | EV | -10.3173 | -75.1541 |  |  |  | 0.10109 | 0.95628
23 |  | 255 | -5.4092 | -67.083 | -3.109 | 610.3224 |  | 0.059059 | 0.24729 | -27.7905 | -2.7365
24 |  | 259 | -8.1501 | -85.2729 | -2.6507 | 654.8551 |  | 0.073824 | 1.1835 | -28.057 | -4.3297
25 |  | 257 | -7.1495 | -82.1937 | -2.8392 | 787.7803 |  | 0.043046 | 0.31378 | -33.474 | -5.888
26 |  | 249 | -12.8531 | -119.8342 |  |  |  | 0.070967 | 2.2975
27 | * | 244 | -9.9297 | -85.8417 |  |  |  | 4.7542 | 48.2638
28 | * | 101 | -7.5923 | -61.5088 |  |  |  | 5.5755 | 54.0449
29 | * | 123 | -2.4046 | -46.5895 | -3.3872 |  |  | 0.32535 | 35.9321 | -28.975
30 |  | 245 | 2.3654 | -79.5327 | -9.0083 | 519.3388 | 59.0434 | 0.091617 | 0.27102 | -42.177 | 10.5615
31 | * | EV | -6.4711 | -45.9411 |  |  |  | 4.3001 | 35.145
32 |  | 253 | -6.6348 | -82.5145 | -3.2452 | 641.1634 |  | 0.13857 | 0.39439 | -29.7993 | -7.1275
33 | * | 241 | -7.168 | -57.4982 |  |  |  | 5.0928 | 49.1759
34 |  | 232 | -12.0583 | -102.3353 |  |  |  | 0.06868 | 0.53463
35 |  | 227 | -12.2953 | -103.3042 |  |  |  | 0.032294 | 0.083372
36 |  | 229 | -11.7184 | -100.4053 |  |  |  | 0.039808 | 0.2035
37 | * | 193 | -6.6459 | -49.192 |  |  |  | 4.4991 | 39.0917
38 |  | 233 | -12.3839 | -102.8227 |  |  |  | 0.045022 | 1.6374
39 | * | 242 | -2.6906 | -16.993 |  |  |  | 0 | 0
40 |  | EV | -10.4827 | -76.3097 |  |  |  | 0.14132 | 1.6528
41 |  | 167 | -14.0736 | -123.3638 |  |  |  | 0.075899 | 3.3179
42 | * | 265 | -2.6906 | -16.993 |  |  |  | 0 | 0
43 |  | EV | -10.6883 | -78.5683 |  |  |  | 0.19858 | 2.1683

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -8.36 | -70.41 | -11.1774 | -83.9604 | -0.011 | 1350.554 | 2021/07/30 11:42:41
1 |  | EV | -7.801 | -63.79 | -10.6099 | -77.6641 | 0.337 | 1326.736 | 2021/07/30 11:51:37
1 |  | EV | -7.664 | -62.35 | -10.4708 | -76.2945 | 0.333 | 1321.295 | 2021/07/30 12:00:33
1 |  | EV |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 12:09:29
1 |  | EV | -7.765 | -62.198 | -10.5733 | -76.1499 | 0.534 | 1325.085 | 2021/07/30 12:19:28
1 |  | EV | -7.591 | -61.15 | -10.3967 | -75.1531 | 0.278 | 1314.472 | 2021/07/30 12:28:25
1 |  | EV | -7.691 | -61.474 | -10.4982 | -75.4613 | 0.289 | 1311.73 | 2021/07/30 12:37:21
1 |  | EV | -7.523 | -60.455 | -10.3277 | -74.4921 | 0.399 | 1311.704 | 2021/07/30 12:46:18
1 |  | EV | -7.502 | -60.758 | -10.3064 | -74.7803 | 0.068 | 1308.902 | 2021/07/30 12:55:14
1 |  | EV | -7.602 | -60.777 | -10.4079 | -74.7984 | 0.348 | 1307.671 | 2021/07/30 13:04:10
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.7459 | -70.5747 |  |  |  | 2.4542 | 20.0104
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.561 | 23.967 | 0.92438 | 5.8022 | 0.085 | 1302.892 | 2021/07/30 13:13:58
2 |  | PZ | 4.312 | 33.771 | 1.6868 | 15.1268 | 0.449 | 1303.547 | 2021/07/30 13:22:54
2 |  | PZ | 4.563 | 36.931 | 1.9416 | 18.1323 | 0.309 | 1302.983 | 2021/07/30 13:31:51
2 |  | PZ | 4.574 | 37.624 | 1.9527 | 18.7915 | 0.683 | 1300.027 | 2021/07/30 13:40:47
2 |  | PZ | 4.665 | 38.399 | 2.0451 | 19.5286 | 0.321 | 1300.835 | 2021/07/30 13:49:44
2 |  | PZ | 4.713 | 39.231 | 2.0938 | 20.3199 | 0.438 | 1302.708 | 2021/07/30 13:58:41
2 |  | PZ | 4.751 | 39.293 | 2.1324 | 20.3789 | 0.36 | 1301.723 | 2021/07/30 14:07:37
2 |  | PZ | 4.802 | 39.609 | 2.1842 | 20.6794 | 0.22 | 1301.025 | 2021/07/30 14:16:33
2 |  | PZ | 4.766 | 39.757 | 2.1477 | 20.8202 | 0.317 | 1300.142 | 2021/07/30 14:25:30
2 |  | PZ | 4.809 | 40.388 | 2.1913 | 21.4203 | 0.262 | 1301.477 | 2021/07/30 14:34:27
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.37902 | 4.929
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -11.497 | -87.305 | -14.3619 | -100.0293 | 0.39 | 1298.725 | 2021/07/30 14:44:15
3 |  | UT2 | -12.736 | -103.58 | -15.6197 | -115.5086 | 0.16 | 1301.282 | 2021/07/30 14:53:12
3 |  | UT2 | -12.946 | -107.434 | -15.8329 | -119.1741 | 0.364 | 1300.901 | 2021/07/30 15:02:09
3 |  | UT2 | -13.032 | -108.894 | -15.9202 | -120.5628 | 0.407 | 1301.655 | 2021/07/30 15:11:05
3 |  | UT2 | -13.169 | -109.909 | -16.0593 | -121.5281 | 0.463 | 1300.85 | 2021/07/30 15:20:02
3 |  | UT2 | -13.153 | -110.406 | -16.043 | -122.0008 | 0.353 | 1299.701 | 2021/07/30 15:28:59
3 |  | UT2 | -13.204 | -110.832 | -16.0948 | -122.406 | 0.212 | 1299.389 | 2021/07/30 15:37:55
3 |  | UT2 | -13.227 | -111.006 | -16.1182 | -122.5715 | 0.59 | 1299.208 | 2021/07/30 15:46:52
3 |  | UT2 | -13.327 | -111.752 | -16.2197 | -123.281 | 0.551 | 1297.65 | 2021/07/30 15:55:49
3 |  | UT2 | -13.239 | -111.391 | -16.1303 | -122.9377 | 0.604 | 1297.888 | 2021/07/30 16:04:45
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.53952 | 7.4191
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -8.081 | -68.241 | -10.8941 | -81.8974 | 0.504 | 1299.645 | 2021/07/30 16:14:35
4 |  | EV | -7.675 | -62.388 | -10.482 | -76.3306 | 0.405 | 1300.072 | 2021/07/30 16:23:31
4 |  | EV | -7.617 | -61.335 | -10.4231 | -75.3291 | 0.692 | 1298.335 | 2021/07/30 16:32:28
4 |  | EV | -7.619 | -61.15 | -10.4251 | -75.1531 | 0.382 | 1298.41 | 2021/07/30 16:41:25
4 |  | EV | -7.532 | -60.435 | -10.3368 | -74.4731 | 0.273 | 1298.974 | 2021/07/30 16:50:21
4 |  | EV | -7.575 | -60.318 | -10.3805 | -74.3618 | 0.299 | 1298.107 | 2021/07/30 16:59:18
4 |  | EV | -7.578 | -60.36 | -10.3835 | -74.4018 | 0.215 | 1297.796 | 2021/07/30 17:08:14
4 |  | EV | -7.521 | -60.384 | -10.3256 | -74.4246 | 0.553 | 1297.696 | 2021/07/30 17:17:11
4 |  | EV | -7.619 | -60.343 | -10.4251 | -74.3856 | 0.422 | 1297.241 | 2021/07/30 17:26:08
4 |  | EV | -7.561 | -60.491 | -10.3662 | -74.5264 | 0.428 | 1296.744 | 2021/07/30 17:35:04
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.4442 | -75.5283 |  |  |  | 0.16235 | 2.4448
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | NIWO2-B-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 17:44:54
5 |  | NIWO2-B-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 17:54:52
5 |  | NIWO2-B-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 18:04:52
5 |  | NIWO2-B-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 18:14:50
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 | * | NIWO2-B-5 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | NIWO2-B-4 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 18:25:42
6 |  | NIWO2-B-4 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 18:35:41
6 |  | NIWO2-B-4 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 18:45:40
6 |  | NIWO2-B-4 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 18:55:39
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 | * | NIWO2-B-4 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | NIWO2-010-2 | -7.912 | -84.773 | -10.7226 | -97.6211 | 0.475 | 1293.773 | 2021/07/30 19:06:31
7 |  | NIWO2-010-2 | -7.496 | -84.147 | -10.3003 | -97.0257 | 0.132 | 1294.332 | 2021/07/30 19:15:28
7 |  | NIWO2-010-2 | -7.499 | -84.132 | -10.3033 | -97.0115 | 0.225 | 1294.314 | 2021/07/30 19:24:25
7 |  | NIWO2-010-2 | -7.445 | -84.248 | -10.2485 | -97.1218 | 0.194 | 1293.566 | 2021/07/30 19:33:22
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | NIWO2-010-2 | -10.3937 | -97.195 |  |  |  | 0.21742 | 0.30308
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | NIWO2-1020-2 | -4.465 | -66.958 | -7.2233 | -80.6772 | 0.104 | 1289.294 | 2021/07/30 19:43:08
8 |  | NIWO2-1020-2 | -4.157 | -64.358 | -6.9106 | -78.2043 | 0.087 | 1290.86 | 2021/07/30 19:52:04
8 |  | NIWO2-1020-2 | -4.155 | -64.476 | -6.9086 | -78.3165 | 0.208 | 1282.225 | 2021/07/30 20:01:01
8 |  | NIWO2-1020-2 | -4.194 | -64.197 | -6.9482 | -78.0512 | 0.148 | 1286.184 | 2021/07/30 20:09:57
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | NIWO2-1020-2 | -6.9977 | -78.8123 |  |  |  | 0.14925 | 1.3122
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | NIWO2-E-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 20:19:44
9 |  | NIWO2-E-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 20:29:42
9 |  | NIWO2-E-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 20:39:41
9 |  | NIWO2-E-5 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 20:49:40
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 | * | NIWO2-E-5 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | NIWO2-2030-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 21:00:30
10 |  | NIWO2-2030-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 21:10:29
10 |  | NIWO2-2030-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 21:20:28
10 |  | NIWO2-2030-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 21:30:26
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 | * | NIWO2-2030-3 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | NIWO2-D-5 | -4.979 | -82.5 | -7.7451 | -95.4593 | 0.092 | 1287.615 | 2021/07/30 21:41:17
11 |  | NIWO2-D-5 | -4.444 | -81.68 | -7.202 | -94.6794 | 0.392 | 1297.98 | 2021/07/30 21:50:14
11 |  | NIWO2-D-5 | -4.302 | -81.613 | -7.0578 | -94.6156 | 0.287 | 1301.587 | 2021/07/30 21:59:09
11 |  | NIWO2-D-5 | -4.329 | -82.053 | -7.0852 | -95.0341 | 0.375 | 1298.294 | 2021/07/30 22:08:06
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | NIWO2-D-5 | -7.2725 | -94.9471 |  |  |  | 0.31638 | 0.40786
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | NIWO2-3040-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 22:17:55
12 |  | NIWO2-3040-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 22:27:53
12 |  | NIWO2-3040-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 22:37:51
12 |  | NIWO2-3040-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 22:47:50
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 | * | NIWO2-3040-3 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -7.825 | -66.519 | -10.6343 | -80.2596 | 0.408 | 1297.088 | 2021/07/30 22:58:42
13 |  | EV | -7.593 | -62.195 | -10.3987 | -76.147 | 0.46 | 1297.374 | 2021/07/30 23:07:38
13 |  | EV | -7.506 | -61.01 | -10.3104 | -75.02 | 0.234 | 1297.256 | 2021/07/30 23:16:34
13 |  | EV | -7.568 | -60.869 | -10.3734 | -74.8859 | 0.427 | 1297.119 | 2021/07/30 23:25:31
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.4292 | -76.5781 |  |  |  | 0.13954 | 2.6481
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | NIWO2-D-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 23:35:20
14 |  | NIWO2-D-3 |  |  | -2.6906 | -16.993 |  |  | 2021/07/30 23:45:19
14 |  | NIWO2-D-3 | -9.411 | -102.534 | -12.2443 | -114.5137 | 0.033 | 1295.492 | 2021/07/30 23:55:17
14 |  | NIWO2-D-3 | -9.411 | -102.534 | -12.2443 | -114.5137 | 0.033 | 1295.492 | 2021/07/31 00:05:15
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 | * | NIWO2-D-3 | -7.4675 | -65.7534 |  |  |  | 5.4334 | 59.198
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | SLCP-21-015 | -2.688 | -36.276 | -5.4194 | -51.4953 | 0.617 | 1295.882 | 2021/07/31 00:15:02
15 |  | SLCP-21-015 | -2.12 | -27.551 | -4.8428 | -43.1969 | 0.149 | 1297.145 | 2021/07/31 00:23:58
15 |  | SLCP-21-015 | -2.041 | -25.677 | -4.7626 | -41.4145 | 0.284 | 1297.26 | 2021/07/31 00:32:54
15 |  | SLCP-21-015 | -2 | -25.315 | -4.7209 | -41.0702 | 0.458 | 1296.329 | 2021/07/31 00:41:51
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | SLCP-21-015 | -4.9364 | -44.2942 |  |  |  | 0.32105 | 5.1417
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | 262 | -8.244 | -75.114 | -11.0596 | -88.4344 | 0.419 | 1298.409 | 2021/07/31 00:51:37
16 |  | 262 | -8.761 | -81.622 | -11.5844 | -94.6242 | 0.381 | 1297.734 | 2021/07/31 01:00:34
16 |  | 262 | -8.905 | -83.367 | -11.7306 | -96.2839 | 0.372 | 1295.372 | 2021/07/31 01:09:30
16 |  | 262 | -8.889 | -83.537 | -11.7144 | -96.4456 | 0.274 | 1298.193 | 2021/07/31 01:18:27
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | 262 | -11.5223 | -93.947 |  |  |  | 0.31059 | 3.9597
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | 237 | -8.292 | -84.775 | -11.1083 | -97.623 | 0.564 | 1297.274 | 2021/07/31 01:28:14
17 |  | 237 | -8.223 | -84.97 | -11.0383 | -97.8085 | 0.327 | 1296.25 | 2021/07/31 01:37:11
17 |  | 237 | -8.124 | -84.704 | -10.9378 | -97.5555 | 0.205 | 1297.115 | 2021/07/31 01:46:08
17 |  | 237 | -8.166 | -84.833 | -10.9804 | -97.6782 | 0.275 | 1297.534 | 2021/07/31 01:55:04
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | 237 | -11.0162 | -97.6663 |  |  |  | 0.072844 | 0.11277
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | 243 | -8.969 | -56.85 | -11.7956 | -71.0634 | 0.59 | 1297.437 | 2021/07/31 02:04:52
18 |  | 243 | -9.064 | -53.542 | -11.892 | -67.9171 | 0.437 | 1295.075 | 2021/07/31 02:13:49
18 |  | 243 | -9.038 | -52.365 | -11.8656 | -66.7977 | 0.01 | 1297.924 | 2021/07/31 02:22:45
18 |  | 243 | -9.16 | -52.587 | -11.9895 | -67.0088 | 0.12 | 1294.242 | 2021/07/31 02:31:41
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | 243 | -11.8857 | -68.1967 |  |  |  | 0.07908 | 2.0732
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | 250 | -2.2 | -54.561 | -4.924 | -68.8863 | -3.846 | 2076.673 | 2021/07/31 02:41:29
19 |  | 250 | -1.901 | -55.326 | -4.6204 | -69.6139 | -3.704 | 2082.088 | 2021/07/31 02:50:25
19 |  | 250 | -1.914 | -55.773 | -4.6336 | -70.039 | -3.877 | 2078.85 | 2021/07/31 02:59:22
19 |  | 250 | -1.906 | -55.597 | -4.6255 | -69.8716 | -3.483 | 2079.891 | 2021/07/31 03:08:19
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | 250 | -4.7009 | -69.6027 | -3.7275 | 778.6451 | 2.922 | 0.1466 | 0.53477 | -35.1573 | -3.1132
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | 258 | -8.46 | -74.944 | -11.2789 | -88.2727 | -0.037 | 1308.904 | 2021/07/31 03:18:07
20 |  | 258 | -8.738 | -77.057 | -11.5611 | -90.2824 | 0.237 | 1302.34 | 2021/07/31 03:27:03
20 |  | 258 | -8.741 | -77.458 | -11.5641 | -90.6638 | 0.375 | 1299.022 | 2021/07/31 03:35:59
20 |  | 258 | -8.659 | -77.208 | -11.4809 | -90.426 | 0.158 | 1299.496 | 2021/07/31 03:44:55
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | 258 | -11.4713 | -89.9112 |  |  |  | 0.13192 | 1.1603
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | 248 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 03:54:44
21 |  | 248 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 04:03:41
21 |  | 248 | -6.441 | -72.554 | -9.2293 | -85.9996 | 0.197 | 1298.64 | 2021/07/31 04:12:38
21 |  | 248 | -6.498 | -72.672 | -9.2871 | -86.1118 | 0.338 | 1296.927 | 2021/07/31 04:21:34
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 | * | 248 | -5.9744 | -51.5243 |  |  |  | 3.7352 | 41.9232
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -7.516 | -62.562 | -10.3206 | -76.4961 | 0.208 | 1297.365 | 2021/07/31 04:31:19
22 |  | EV | -7.418 | -60.564 | -10.2211 | -74.5958 | 0.253 | 1298.574 | 2021/07/31 04:40:16
22 |  | EV | -7.465 | -60.554 | -10.2688 | -74.5863 | 0.571 | 1296.7 | 2021/07/31 04:49:12
22 |  | EV | -7.652 | -60.924 | -10.4586 | -74.9382 | 0.473 | 1295.014 | 2021/07/31 04:58:09
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.3173 | -75.1541 |  |  |  | 0.10109 | 0.95628
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | 255 | -2.766 | -52.983 | -5.4985 | -67.3854 | -3.029 | 1900.72 | 2021/07/31 05:07:55
23 |  | 255 | -2.642 | -52.473 | -5.3727 | -66.9004 | -3.225 | 1915.582 | 2021/07/31 05:16:51
23 |  | 255 | -2.646 | -52.739 | -5.3767 | -67.1534 | -3.256 | 1913.422 | 2021/07/31 05:25:48
23 |  | 255 | -2.658 | -52.465 | -5.3889 | -66.8928 | -2.926 | 1914.487 | 2021/07/31 05:34:44
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | 255 | -5.4092 | -67.083 | -3.109 | 610.3224 |  | 0.059059 | 0.24729 | -27.7905 | -2.7365
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | 259 | -5.27 | -70.043 | -8.0405 | -83.6113 | -2.724 | 1955.233 | 2021/07/31 05:44:30
24 |  | 259 | -5.409 | -72.167 | -8.1816 | -85.6315 | -2.709 | 1943.517 | 2021/07/31 05:53:27
24 |  | 259 | -5.436 | -72.289 | -8.209 | -85.7475 | -2.601 | 1962.286 | 2021/07/31 06:02:23
24 |  | 259 | -5.397 | -72.661 | -8.1694 | -86.1013 | -2.569 | 1961.306 | 2021/07/31 06:11:18
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | 259 | -8.1501 | -85.2729 | -2.6507 | 654.8551 |  | 0.073824 | 1.1835 | -28.057 | -4.3297
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | 257 | -4.439 | -68.734 | -7.1969 | -82.3663 | -2.918 | 2092.433 | 2021/07/31 06:21:04
25 |  | 257 | -4.402 | -68.876 | -7.1594 | -82.5014 | -2.849 | 2081.598 | 2021/07/31 06:30:00
25 |  | 257 | -4.393 | -68.423 | -7.1502 | -82.0705 | -2.71 | 2081.045 | 2021/07/31 06:38:56
25 |  | 257 | -4.335 | -68.177 | -7.0913 | -81.8366 | -2.88 | 2098.967 | 2021/07/31 06:47:52
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | 257 | -7.1495 | -82.1937 | -2.8392 | 787.7803 |  | 0.043046 | 0.31378 | -33.474 | -5.888
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | 249 | -9.925 | -104.817 | -12.7661 | -116.6851 | 0.294 | 1306.278 | 2021/07/31 06:57:39
26 |  | 249 | -9.983 | -108.365 | -12.825 | -120.0596 | 0.448 | 1300.538 | 2021/07/31 07:06:35
26 |  | 249 | -10.051 | -109.45 | -12.894 | -121.0916 | 0.402 | 1300.902 | 2021/07/31 07:15:31
26 |  | 249 | -10.084 | -109.88 | -12.9275 | -121.5006 | 0.25 | 1299.478 | 2021/07/31 07:24:27
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | 249 | -12.8531 | -119.8342 |  |  |  | 0.070967 | 2.2975
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | 244 | -9.482 | -97.491 | -12.3164 | -109.7173 | 0.187 | 1291.887 | 2021/07/31 07:34:15
27 |  | 244 | -9.48 | -95.858 | -12.3143 | -108.1641 | -0.059 | 1288.678 | 2021/07/31 07:43:11
27 |  | 244 | -9.562 | -96.203 | -12.3976 | -108.4923 | 0.061 | 1282.55 | 2021/07/31 07:52:07
27 |  | 244 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 08:01:04
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 | * | 244 | -9.9297 | -85.8417 |  |  |  | 4.7542 | 48.2638
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | 101 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 08:11:53
28 |  | 101 | -9.691 | -93.674 | -12.5285 | -106.0869 | 0.526 | 1298.697 | 2021/07/31 08:20:49
28 |  | 101 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 08:29:45
28 |  | 101 | -9.623 | -93.543 | -12.4595 | -105.9623 | 0.182 | 1298.246 | 2021/07/31 08:38:42
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 | * | 101 | -7.5923 | -61.5088 |  |  |  | 5.5755 | 54.0449
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | 123 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 08:48:26
29 |  | 123 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 08:57:22
29 |  | 123 | 0.567 | -62.354 | -2.115 | -76.2983 | -6.849 | 2453.585 | 2021/07/31 09:06:18
29 |  | 123 | 0.56 | -62.118 | -2.1221 | -76.0738 | -6.7 | 2442.222 | 2021/07/31 09:15:15
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 | * | 123 | -2.4046 | -46.5895 | -3.3872 |  |  | 0.32535 | 35.9321 | -28.975
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | 245 | 4.844 | -65.401 | 2.2268 | -79.1963 | -8.462 | 1785.716 | 2021/07/31 09:25:00
30 |  | 245 | 5.041 | -65.695 | 2.4268 | -79.4759 | -9.194 | 1834.287 | 2021/07/31 09:33:56
30 |  | 245 | 5.018 | -65.906 | 2.4035 | -79.6766 | -9.005 | 1831.931 | 2021/07/31 09:42:52
30 |  | 245 | 5.019 | -66.017 | 2.4045 | -79.7822 | -9.372 | 1828.343 | 2021/07/31 09:51:48
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | 245 | 2.3654 | -79.5327 | -9.0083 | 519.3388 | 59.0434 | 0.091617 | 0.27102 | -42.177 | 10.5615
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 10:01:34
31 |  | EV |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 10:10:29
31 |  | EV | -7.471 | -61.027 | -10.2749 | -75.0361 | 0.086 | 1298.191 | 2021/07/31 10:19:26
31 |  | EV | -7.425 | -60.718 | -10.2282 | -74.7423 | 0.367 | 1297.718 | 2021/07/31 10:28:22
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 | * | EV | -6.4711 | -45.9411 |  |  |  | 4.3001 | 35.145
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | 253 | -4.066 | -68.891 | -6.8183 | -82.5157 | -2.91 | 1842.489 | 2021/07/31 10:38:07
32 |  | 253 | -3.751 | -68.355 | -6.4985 | -82.0059 | -3.308 | 1977.119 | 2021/07/31 10:47:04
32 |  | 253 | -3.916 | -69.295 | -6.666 | -82.8999 | -3.478 | 1972.197 | 2021/07/31 10:56:00
32 |  | 253 | -3.808 | -69.018 | -6.5563 | -82.6364 | -3.285 | 1975.77 | 2021/07/31 11:04:56
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | 253 | -6.6348 | -82.5145 | -3.2452 | 641.1634 |  | 0.13857 | 0.39439 | -29.7993 | -7.1275
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | 241 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 11:14:43
33 |  | 241 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 11:23:38
33 |  | 241 | -8.803 | -85.091 | -11.6271 | -97.9236 | 0.247 | 1300.992 | 2021/07/31 11:32:34
33 |  | 241 | -8.839 | -85.259 | -11.6636 | -98.0834 | 0.302 | 1297.012 | 2021/07/31 11:41:31
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 | * | 241 | -7.168 | -57.4982 |  |  |  | 5.0928 | 49.1759
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | 232 | -9.191 | -89.124 | -12.021 | -101.7594 | 0.265 | 1297.583 | 2021/07/31 11:51:17
34 |  | 232 | -9.225 | -89.691 | -12.0555 | -102.2987 | 0.44 | 1298.173 | 2021/07/31 12:00:13
34 |  | 232 | -9.17 | -89.676 | -11.9996 | -102.2844 | 0.32 | 1296.701 | 2021/07/31 12:09:09
34 |  | 232 | -9.325 | -90.427 | -12.157 | -102.9987 | 0.281 | 1294.153 | 2021/07/31 12:18:05
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | 232 | -12.0583 | -102.3353 |  |  |  | 0.06868 | 0.53463
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | 227 | -9.507 | -90.767 | -12.3418 | -103.3221 | 0.061 | 1289.131 | 2021/07/31 12:27:52
35 |  | 227 | -9.459 | -90.674 | -12.293 | -103.2336 | 0.27 | 1292.023 | 2021/07/31 12:36:49
35 |  | 227 | -9.446 | -90.858 | -12.2798 | -103.4086 | 0.251 | 1290.795 | 2021/07/31 12:45:44
35 |  | 227 | -9.433 | -90.694 | -12.2666 | -103.2526 | 0.292 | 1291.007 | 2021/07/31 12:54:40
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | 227 | -12.2953 | -103.3042 |  |  |  | 0.032294 | 0.083372
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | 229 | -8.926 | -87.89 | -11.7519 | -100.5857 | 0.544 | 1297.011 | 2021/07/31 13:04:22
36 |  | 229 | -8.913 | -87.771 | -11.7387 | -100.4725 | 0.339 | 1295.884 | 2021/07/31 13:13:17
36 |  | 229 | -8.836 | -87.413 | -11.6606 | -100.132 | 0.438 | 1296.948 | 2021/07/31 13:22:13
36 |  | 229 | -8.897 | -87.727 | -11.7225 | -100.4307 | 0.266 | 1296.075 | 2021/07/31 13:31:09
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | 229 | -11.7184 | -100.4053 |  |  |  | 0.039808 | 0.2035
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | 193 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 13:40:53
37 |  | 193 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 13:49:49
37 |  | 193 | -7.761 | -67.566 | -10.5693 | -81.2554 | -0.005 | 1287.516 | 2021/07/31 13:58:44
37 |  | 193 | -7.824 | -67.851 | -10.6332 | -81.5265 | 0.24 | 1278.572 | 2021/07/31 14:07:40
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 | * | 193 | -6.6459 | -49.192 |  |  |  | 4.4991 | 39.0917
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | 233 | -9.489 | -87.803 | -12.3235 | -100.503 | 0.413 | 1295.489 | 2021/07/31 14:17:22
38 |  | 233 | -9.562 | -90.783 | -12.3976 | -103.3373 | 0.23 | 1292.627 | 2021/07/31 14:26:17
38 |  | 233 | -9.597 | -91.204 | -12.4331 | -103.7377 | 0.23 | 1295.376 | 2021/07/31 14:35:13
38 |  | 233 | -9.546 | -91.178 | -12.3813 | -103.713 | 0.065 | 1297.451 | 2021/07/31 14:44:09
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | 233 | -12.3839 | -102.8227 |  |  |  | 0.045022 | 1.6374
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | 242 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 14:53:53
39 |  | 242 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 15:03:51
39 |  | 242 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 15:13:49
39 |  | 242 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 15:23:47
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 | * | 242 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.875 | -64.747 | -10.685 | -78.5743 | 0.361 | 1294.363 | 2021/07/31 15:34:34
40 |  | EV | -7.677 | -62.211 | -10.484 | -76.1623 | 0.078 | 1291.654 | 2021/07/31 15:43:29
40 |  | EV | -7.568 | -61.363 | -10.3734 | -75.3557 | 0.339 | 1291.095 | 2021/07/31 15:52:24
40 |  | EV | -7.583 | -61.143 | -10.3886 | -75.1465 | 0.319 | 1291.533 | 2021/07/31 16:01:19
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.4827 | -76.3097 |  |  |  | 0.14132 | 1.6528
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | 167 | -11.114 | -107.072 | -13.9731 | -118.8298 | 0.286 | 1289.991 | 2021/07/31 16:11:03
41 |  | 167 | -11.219 | -112.149 | -14.0797 | -123.6586 | 0.038 | 1297.465 | 2021/07/31 16:20:00
41 |  | 167 | -11.22 | -113.699 | -14.0807 | -125.1328 | 0.035 | 1298.39 | 2021/07/31 16:28:55
41 |  | 167 | -11.299 | -114.436 | -14.1609 | -125.8338 | 0.47 | 1297.184 | 2021/07/31 16:37:51
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | 167 | -14.0736 | -123.3638 |  |  |  | 0.075899 | 3.3179
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | 265 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 16:47:36
42 |  | 265 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 16:57:34
42 |  | 265 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 17:07:32
42 |  | 265 |  |  | -2.6906 | -16.993 |  |  | 2021/07/31 17:17:30
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 | * | 265 | -2.6906 | -16.993 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | EV | -8.039 | -67.292 | -10.8515 | -80.9948 | 0.167 | 1294.244 | 2021/07/31 17:28:16
43 |  | EV | -7.741 | -63.976 | -10.549 | -77.841 | 0.279 | 1281.655 | 2021/07/31 17:37:12
43 |  | EV | -8.058 | -65.493 | -10.8708 | -79.2838 | 0.051 | 1257.446 | 2021/07/31 17:46:07
43 |  | EV | -7.675 | -62.202 | -10.482 | -76.1537 | 0.178 | 1283.518 | 2021/07/31 17:55:02
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | EV | -10.6883 | -78.5683 |  |  |  | 0.19858 | 2.1683

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-17265 | 2021/07/30 11:42:41 | 1-01 | 1 | -8.36 | -70.41 | 20127 | -1 | 1 | EV | standard | H2O | 1627667267 | 0.339 | 0.697 | 137.146 | 0.001 | 0.001 | 3.226 | 1350.554 | 1.068 | 3.164 | 121.858 | 0.876 | -0.011 | 0 | 0 | 1 | 45.125 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-17265 | 2021/07/30 11:51:37 | 1-01 | 2 | -7.801 | -63.79 | 18700 | -1 | 1 | EV | standard | H2O | 1627667803 | 0.313 | 0.749 | 121.876 | 0 | -0.001 | 2.881 | 1326.736 | 1.129 | 3.023 | 114.888 | 0.876 | 0.337 | 0 | 0 | 1 | 44.625 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-17265 | 2021/07/30 12:00:33 | 1-01 | 3 | -7.664 | -62.35 | 18664 | -1 | 1 | EV | standard | H2O | 1627668339 | 0.297 | 0.685 | 118.41 | -0.001 | -0.002 | 2.813 | 1321.295 | 1.706 | 2.971 | 114.557 | 0.877 | 0.333 | 0 | 0 | 1 | 44.562 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-17265 | 2021/07/30 12:09:29 | 1-01 | 4 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-17265 | 2021/07/30 12:19:28 | 1-01 | 5 | -7.765 | -62.198 | 18395 | 0 | 1 | EV | standard | H2O | 1627669474 | 0.325 | 0.748 | 106.521 | -0.001 | -0.003 | 2.536 | 1325.085 | 1.189 | 2.947 | 113.578 | 0.875 | 0.534 | 0 | 0 | 1 | 44.25 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-17265 | 2021/07/30 12:28:25 | 1-01 | 6 | -7.591 | -61.15 | 18572 | 0 | 1 | EV | standard | H2O | 1627670011 | 0.323 | 0.715 | 110.053 | -0.001 | 0.002 | 2.602 | 1314.472 | 1.888 | 2.963 | 113.042 | 0.875 | 0.278 | 0 | 0 | 1 | 44.25 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-17265 | 2021/07/30 12:37:21 | 1-01 | 7 | -7.691 | -61.474 | 18393 | 0 | 1 | EV | standard | H2O | 1627670547 | 0.307 | 0.733 | 108.621 | 0 | 0.006 | 2.562 | 1311.73 | 1.561 | 2.897 | 112.398 | 0.873 | 0.289 | 0 | 0 | 1 | 44.25 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-17265 | 2021/07/30 12:46:18 | 1-01 | 8 | -7.523 | -60.455 | 18625 | 0 | 1 | EV | standard | H2O | 1627671084 | 0.298 | 0.716 | 108.796 | 0 | 0 | 2.561 | 1311.704 | 2.023 | 2.997 | 114.596 | 0.869 | 0.399 | 0 | 0 | 1 | 44.312 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-17265 | 2021/07/30 12:55:14 | 1-01 | 9 | -7.502 | -60.758 | 18406 | 0 | 1 | EV | standard | H2O | 1627671621 | 0.289 | 0.788 | 105.711 | -0.001 | 0.002 | 2.518 | 1308.902 | 1.946 | 2.899 | 111.481 | 0.87 | 0.068 | 0 | 0 | 1 | 44.25 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-17265 | 2021/07/30 13:04:10 | 1-01 | 10 | -7.602 | -60.777 | 18477 | 0 | 1 | EV | standard | H2O | 1627672157 | 0.278 | 0.723 | 105.456 | -0.001 | 0.001 | 2.487 | 1307.671 | 1.639 | 2.915 | 113.568 | 0.868 | 0.348 | 0 | 0 | 1 | 44.375 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-17266 | 2021/07/30 13:13:58 | 1-02 | 1 | 3.561 | 23.967 | 18579 | -1 | 1 | PZ | standard | H2O | 1627672744 | 0.319 | 0.981 | 103.05 | -0.002 | -0.015 | 2.448 | 1302.892 | 1.848 | 2.97 | 113.497 | 0.877 | 0.085 | 0 | 0 | 1 | 44.312 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-17266 | 2021/07/30 13:22:54 | 1-02 | 2 | 4.312 | 33.771 | 18409 | -1 | 1 | PZ | standard | H2O | 1627673281 | 0.291 | 0.804 | 105.541 | -0.001 | -0.01 | 2.498 | 1303.547 | 1.682 | 2.956 | 113.877 | 0.875 | 0.449 | 0 | 0 | 1 | 44.25 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-17266 | 2021/07/30 13:31:51 | 1-02 | 3 | 4.563 | 36.931 | 18397 | -1 | 1 | PZ | standard | H2O | 1627673817 | 0.32 | 0.825 | 104.467 | 0 | -0.009 | 2.477 | 1302.983 | 1.41 | 2.928 | 113.54 | 0.868 | 0.309 | 0 | 0 | 1 | 44.25 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-17266 | 2021/07/30 13:40:47 | 1-02 | 4 | 4.574 | 37.624 | 18123 | 0 | 1 | PZ | standard | H2O | 1627674354 | 0.329 | 0.804 | 102.323 | -0.001 | -0.007 | 2.426 | 1300.027 | 1.869 | 2.911 | 112.573 | 0.873 | 0.683 | 0 | 0 | 1 | 44.25 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-17266 | 2021/07/30 13:49:44 | 1-02 | 5 | 4.665 | 38.399 | 18225 | 0 | 1 | PZ | standard | H2O | 1627674890 | 0.309 | 0.806 | 101.104 | -0.001 | -0.006 | 2.408 | 1300.835 | 1.517 | 2.874 | 112.248 | 0.874 | 0.321 | 0 | 0 | 1 | 44.25 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-17266 | 2021/07/30 13:58:41 | 1-02 | 6 | 4.713 | 39.231 | 18573 | 0 | 1 | PZ | standard | H2O | 1627675427 | 0.328 | 0.766 | 106.164 | -0.001 | -0.004 | 2.503 | 1302.708 | 1.956 | 2.97 | 114.218 | 0.872 | 0.438 | 0 | 0 | 1 | 44.188 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-17266 | 2021/07/30 14:07:37 | 1-02 | 7 | 4.751 | 39.293 | 18514 | 0 | 1 | PZ | standard | H2O | 1627675964 | 0.33 | 0.721 | 105.042 | -0.001 | -0.003 | 2.49 | 1301.723 | 2.17 | 2.951 | 113.838 | 0.874 | 0.36 | 0 | 0 | 1 | 44.188 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-17266 | 2021/07/30 14:16:33 | 1-02 | 8 | 4.802 | 39.609 | 18428 | 0 | 1 | PZ | standard | H2O | 1627676500 | 0.31 | 0.71 | 104.473 | -0.001 | -0.004 | 2.48 | 1301.025 | 1.847 | 2.947 | 112.672 | 0.868 | 0.22 | 0 | 0 | 1 | 44.25 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-17266 | 2021/07/30 14:25:30 | 1-02 | 9 | 4.766 | 39.757 | 18384 | 0 | 1 | PZ | standard | H2O | 1627677036 | 0.33 | 0.774 | 104.219 | -0.001 | -0.003 | 2.466 | 1300.142 | 2.449 | 2.97 | 112.301 | 0.875 | 0.317 | 0 | 0 | 1 | 44.188 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-17266 | 2021/07/30 14:34:27 | 1-02 | 10 | 4.809 | 40.388 | 18580 | 0 | 1 | PZ | standard | H2O | 1627677573 | 0.288 | 0.697 | 102.534 | -0.001 | -0.001 | 2.419 | 1301.477 | 2.018 | 2.992 | 113.572 | 0.873 | 0.262 | 0 | 0 | 1 | 44.188 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-17267 | 2021/07/30 14:44:15 | 1-03 | 1 | -11.497 | -87.305 | 18331 | -1 | 1 | UT2 | standard | H2O | 1627678162 | 0.327 | 0.96 | 102.479 | 0.002 | 0.017 | 2.429 | 1298.725 | 3.254 | 2.876 | 110.97 | 0.873 | 0.39 | 0 | 0 | 1 | 44.188 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-17267 | 2021/07/30 14:53:12 | 1-03 | 2 | -12.736 | -103.58 | 18602 | -1 | 1 | UT2 | standard | H2O | 1627678698 | 0.313 | 0.977 | 105.967 | 0 | 0.016 | 2.504 | 1301.282 | 3.347 | 2.901 | 111.351 | 0.876 | 0.16 | 0 | 0 | 1 | 44.188 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-17267 | 2021/07/30 15:02:09 | 1-03 | 3 | -12.946 | -107.434 | 18575 | -1 | 1 | UT2 | standard | H2O | 1627679235 | 0.295 | 0.845 | 103.227 | 0 | 0.011 | 2.444 | 1300.901 | 3.085 | 2.922 | 112.779 | 0.865 | 0.364 | 0 | 0 | 1 | 44.188 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-17267 | 2021/07/30 15:11:05 | 1-03 | 4 | -13.032 | -108.894 | 18682 | 0 | 1 | UT2 | standard | H2O | 1627679772 | 0.317 | 0.804 | 104.457 | 0 | 0.009 | 2.483 | 1301.655 | 3.172 | 2.928 | 113.125 | 0.871 | 0.407 | 0 | 0 | 1 | 44.25 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-17267 | 2021/07/30 15:20:02 | 1-03 | 5 | -13.169 | -109.909 | 18587 | 0 | 1 | UT2 | standard | H2O | 1627680308 | 0.307 | 0.815 | 101.958 | 0.001 | 0.008 | 2.424 | 1300.85 | 2.825 | 2.924 | 113.353 | 0.871 | 0.463 | 0 | 0 | 1 | 44.25 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-17267 | 2021/07/30 15:28:59 | 1-03 | 6 | -13.153 | -110.406 | 18428 | 0 | 1 | UT2 | standard | H2O | 1627680845 | 0.318 | 0.752 | 105.668 | 0 | 0.006 | 2.512 | 1299.701 | 3.024 | 2.903 | 111.797 | 0.871 | 0.353 | 0 | 0 | 1 | 44.25 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-17267 | 2021/07/30 15:37:55 | 1-03 | 7 | -13.204 | -110.832 | 18372 | 0 | 1 | UT2 | standard | H2O | 1627681381 | 0.3 | 0.818 | 105.066 | 0 | 0.007 | 2.486 | 1299.389 | 2.976 | 2.853 | 111.107 | 0.875 | 0.212 | 0 | 0 | 1 | 44.312 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-17267 | 2021/07/30 15:46:52 | 1-03 | 8 | -13.227 | -111.006 | 18400 | 0 | 1 | UT2 | standard | H2O | 1627681918 | 0.301 | 0.721 | 105.156 | -0.001 | 0.006 | 2.504 | 1299.208 | 3.111 | 2.921 | 112.8 | 0.875 | 0.59 | 0 | 0 | 1 | 44.25 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-17267 | 2021/07/30 15:55:49 | 1-03 | 9 | -13.327 | -111.752 | 18178 | 0 | 1 | UT2 | standard | H2O | 1627682455 | 0.297 | 0.863 | 103.933 | 0 | 0.01 | 2.455 | 1297.65 | 3.161 | 2.86 | 110.948 | 0.867 | 0.551 | 0 | 0 | 1 | 44.25 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-17267 | 2021/07/30 16:04:45 | 1-03 | 10 | -13.239 | -111.391 | 18206 | 0 | 1 | UT2 | standard | H2O | 1627682992 | 0.316 | 0.787 | 100.247 | 0.001 | 0.004 | 2.382 | 1297.888 | 3.016 | 2.861 | 111.553 | 0.872 | 0.604 | 0 | 0 | 1 | 44.25 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-17268 | 2021/07/30 16:14:35 | 1-04 | 1 | -8.081 | -68.241 | 18446 | -1 | 1 | EV | standard | H2O | 1627683581 | 0.307 | 0.742 | 103.025 | -0.001 | -0.001 | 2.447 | 1299.645 | 2.292 | 2.945 | 113.698 | 0.872 | 0.504 | 0 | 0 | 1 | 44.25 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-17268 | 2021/07/30 16:23:31 | 1-04 | 2 | -7.675 | -62.388 | 18513 | -1 | 1 | EV | standard | H2O | 1627684117 | 0.281 | 0.762 | 104.03 | 0.001 | 0 | 2.467 | 1300.072 | 2.73 | 2.908 | 113.151 | 0.873 | 0.405 | 0 | 0 | 1 | 44.25 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-17268 | 2021/07/30 16:32:28 | 1-04 | 3 | -7.617 | -61.335 | 18313 | -1 | 1 | EV | standard | H2O | 1627684655 | 0.314 | 0.737 | 102.371 | 0 | 0.002 | 2.422 | 1298.335 | 2.989 | 2.963 | 112.86 | 0.871 | 0.692 | 0 | 0 | 1 | 44.25 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-17268 | 2021/07/30 16:41:25 | 1-04 | 4 | -7.619 | -61.15 | 18290 | 0 | 1 | EV | standard | H2O | 1627685191 | 0.334 | 0.786 | 103.716 | 0.001 | 0.007 | 2.462 | 1298.41 | 3.17 | 2.91 | 111.024 | 0.868 | 0.382 | 0 | 0 | 1 | 44.188 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-17268 | 2021/07/30 16:50:21 | 1-04 | 5 | -7.532 | -60.435 | 18328 | 0 | 1 | EV | standard | H2O | 1627685727 | 0.294 | 0.714 | 104.001 | 0 | 0.001 | 2.491 | 1298.974 | 2.977 | 2.844 | 110.334 | 0.871 | 0.273 | 0 | 0 | 1 | 44.125 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-17268 | 2021/07/30 16:59:18 | 1-04 | 6 | -7.575 | -60.318 | 18222 | 0 | 1 | EV | standard | H2O | 1627686264 | 0.319 | 0.767 | 103.459 | 0.001 | 0.004 | 2.453 | 1298.107 | 2.953 | 2.878 | 110.293 | 0.874 | 0.299 | 0 | 0 | 1 | 44.188 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-17268 | 2021/07/30 17:08:14 | 1-04 | 7 | -7.578 | -60.36 | 18171 | 0 | 1 | EV | standard | H2O | 1627686801 | 0.309 | 0.738 | 102.611 | 0 | 0.001 | 2.443 | 1297.796 | 2.661 | 2.824 | 109.864 | 0.87 | 0.215 | 0 | 0 | 1 | 44.188 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-17268 | 2021/07/30 17:17:11 | 1-04 | 8 | -7.521 | -60.384 | 18171 | 0 | 1 | EV | standard | H2O | 1627687337 | 0.301 | 0.758 | 104.646 | 0 | 0.003 | 2.488 | 1297.696 | 2.157 | 2.911 | 112.944 | 0.867 | 0.553 | 0 | 0 | 1 | 44.188 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-17268 | 2021/07/30 17:26:08 | 1-04 | 9 | -7.619 | -60.343 | 18121 | 0 | 1 | EV | standard | H2O | 1627687874 | 0.363 | 0.685 | 103.061 | 0 | 0.003 | 2.443 | 1297.241 | 2.525 | 2.895 | 111.178 | 0.868 | 0.422 | 0 | 0 | 1 | 44.188 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-17268 | 2021/07/30 17:35:04 | 1-04 | 10 | -7.561 | -60.491 | 18045 | 0 | 1 | EV | standard | H2O | 1627688410 | 0.304 | 0.757 | 102.449 | 0 | 0.002 | 2.438 | 1296.744 | 2.723 | 2.836 | 110.182 | 0.87 | 0.428 | 0 | 0 | 1 | 44.25 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-17269 | 2021/07/30 17:44:54 | 1-05 | 1 | [5 empty] | NIWO2-B-5 | 20-213 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-17269 | 2021/07/30 17:54:52 | 1-05 | 2 | [5 empty] | NIWO2-B-5 | 20-213 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-17269 | 2021/07/30 18:04:52 | 1-05 | 3 | [5 empty] | NIWO2-B-5 | 20-213 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-17269 | 2021/07/30 18:14:50 | 1-05 | 4 | [5 empty] | NIWO2-B-5 | 20-213 | H2O | [17 empty] | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-17270 | 2021/07/30 18:25:42 | 1-06 | 1 | [5 empty] | NIWO2-B-4 | 20-213 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-17270 | 2021/07/30 18:35:41 | 1-06 | 2 | [5 empty] | NIWO2-B-4 | 20-213 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-17270 | 2021/07/30 18:45:40 | 1-06 | 3 | [5 empty] | NIWO2-B-4 | 20-213 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-17270 | 2021/07/30 18:55:39 | 1-06 | 4 | [5 empty] | NIWO2-B-4 | 20-213 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-17271 | 2021/07/30 19:06:31 | 1-07 | 1 | -7.912 | -84.773 | 17597 | -1 | 1 | NIWO2-010-2 | 20-213 | H2O | 1627693898 | 0.303 | 0.781 | 83.8 | 0 | 0.003 | 1.997 | 1293.773 | 2.501 | 2.803 | 107.943 | 0.868 | 0.475 | 0 | 0 | 1 | 44.5 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-17271 | 2021/07/30 19:15:28 | 1-07 | 2 | -7.496 | -84.147 | 17635 | -1 | 1 | NIWO2-010-2 | 20-213 | H2O | 1627694435 | 0.299 | 0.733 | 96.883 | -0.001 | 0.002 | 2.314 | 1294.332 | 2.492 | 2.703 | 106.577 | 0.867 | 0.132 | 0 | 0 | 1 | 44.25 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-17271 | 2021/07/30 19:24:25 | 1-07 | 3 | -7.499 | -84.132 | 17665 | -1 | 1 | NIWO2-010-2 | 20-213 | H2O | 1627694971 | 0.322 | 0.691 | 95.015 | -0.001 | 0.002 | 2.26 | 1294.314 | 2.76 | 2.768 | 106.9 | 0.876 | 0.225 | 0 | 0 | 1 | 44.125 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-17271 | 2021/07/30 19:33:22 | 1-07 | 4 | -7.445 | -84.248 | 17540 | 0 | 1 | NIWO2-010-2 | 20-213 | H2O | 1627695508 | 0.266 | 0.755 | 96.61 | -0.001 | 0.003 | 2.302 | 1293.566 | 2.32 | 2.704 | 106.568 | 0.87 | 0.194 | 0 | 0 | 1 | 44.062 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-17272 | 2021/07/30 19:43:08 | 1-08 | 1 | -4.465 | -66.958 | 17000 | -1 | 1 | NIWO2-1020-2 | 20-213 | H2O | 1627696094 | 0.283 | 0.711 | 95.104 | -0.001 | 0 | 2.276 | 1289.294 | 3.099 | 2.637 | 102.532 | 0.871 | 0.104 | 0 | 0 | 1 | 44.062 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-17272 | 2021/07/30 19:52:04 | 1-08 | 2 | -4.157 | -64.358 | 17211 | -1 | 1 | NIWO2-1020-2 | 20-213 | H2O | 1627696630 | 0.276 | 0.655 | 97.27 | 0 | 0 | 2.327 | 1290.86 | 3.051 | 2.664 | 103.631 | 0.865 | 0.087 | 0 | 0 | 1 | 44.062 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-17272 | 2021/07/30 20:01:01 | 1-08 | 3 | -4.155 | -64.476 | 15955 | -1 | 1 | NIWO2-1020-2 | 20-213 | H2O | 1627697167 | 0.282 | 0.783 | 94.694 | -0.001 | 0 | 2.268 | 1282.225 | 2.173 | 2.51 | 97.006 | 0.865 | 0.208 | 0 | 0 | 1 | 44.125 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-17272 | 2021/07/30 20:09:57 | 1-08 | 4 | -4.194 | -64.197 | 16552 | 0 | 1 | NIWO2-1020-2 | 20-213 | H2O | 1627697703 | 0.301 | 0.748 | 97.105 | 0.001 | 0.005 | 2.328 | 1286.184 | 2.437 | 2.572 | 101.081 | 0.869 | 0.148 | 0 | 0 | 1 | 44.125 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-17273 | 2021/07/30 20:19:44 | 1-09 | 1 | [5 empty] | NIWO2-E-5 | 20-213 | H2O | [17 empty] | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-17273 | 2021/07/30 20:29:42 | 1-09 | 2 | [5 empty] | NIWO2-E-5 | 20-213 | H2O | [17 empty] | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-17273 | 2021/07/30 20:39:41 | 1-09 | 3 | [5 empty] | NIWO2-E-5 | 20-213 | H2O | [17 empty] | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-17273 | 2021/07/30 20:49:40 | 1-09 | 4 | [5 empty] | NIWO2-E-5 | 20-213 | H2O | [17 empty] | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-17274 | 2021/07/30 21:00:30 | 1-10 | 1 | [5 empty] | NIWO2-2030-3 | 20-213 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-17274 | 2021/07/30 21:10:29 | 1-10 | 2 | [5 empty] | NIWO2-2030-3 | 20-213 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-17274 | 2021/07/30 21:20:28 | 1-10 | 3 | [5 empty] | NIWO2-2030-3 | 20-213 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-17274 | 2021/07/30 21:30:26 | 1-10 | 4 | [5 empty] | NIWO2-2030-3 | 20-213 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-17275 | 2021/07/30 21:41:17 | 1-11 | 1 | -4.979 | -82.5 | 16642 | -1 | 1 | NIWO2-D-5 | 20-213 | H2O | 1627703183 | 0.322 | 0.725 | 84.527 | -0.001 | 0.001 | 2.032 | 1287.615 | 2.803 | 2.54 | 100.63 | 0.867 | 0.092 | 0 | 0 | 1 | 44.5 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-17275 | 2021/07/30 21:50:14 | 1-11 | 2 | -4.444 | -81.68 | 18129 | -1 | 1 | NIWO2-D-5 | 20-213 | H2O | 1627703720 | 0.315 | 0.787 | 92.271 | 0 | 0.004 | 2.203 | 1297.98 | 2.576 | 2.884 | 111.123 | 0.873 | 0.392 | 0 | 0 | 1 | 44.25 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-17275 | 2021/07/30 21:59:09 | 1-11 | 3 | -4.302 | -81.613 | 18627 | -1 | 1 | NIWO2-D-5 | 20-213 | H2O | 1627704256 | 0.316 | 0.743 | 99.515 | -0.001 | 0.002 | 2.366 | 1301.587 | 2.62 | 2.897 | 113.625 | 0.87 | 0.287 | 0 | 0 | 1 | 44.125 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-17275 | 2021/07/30 22:08:06 | 1-11 | 4 | -4.329 | -82.053 | 18186 | 0 | 1 | NIWO2-D-5 | 20-213 | H2O | 1627704793 | 0.338 | 0.744 | 101.575 | -0.001 | 0.003 | 2.414 | 1298.294 | 2.705 | 2.903 | 111.427 | 0.871 | 0.375 | 0 | 0 | 1 | 44.125 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-17276 | 2021/07/30 22:17:55 | 1-12 | 1 | [5 empty] | NIWO2-3040-3 | 20-213 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-17276 | 2021/07/30 22:27:53 | 1-12 | 2 | [5 empty] | NIWO2-3040-3 | 20-213 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-17276 | 2021/07/30 22:37:51 | 1-12 | 3 | [5 empty] | NIWO2-3040-3 | 20-213 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-17276 | 2021/07/30 22:47:50 | 1-12 | 4 | [5 empty] | NIWO2-3040-3 | 20-213 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-17277 | 2021/07/30 22:58:42 | 1-13 | 1 | -7.825 | -66.519 | 18092 | -1 | 1 | EV | standard | H2O | 1627707828 | 0.311 | 0.756 | 88.493 | 0 | 0.001 | 2.117 | 1297.088 | 2.76 | 2.859 | 110.597 | 0.865 | 0.408 | 0 | 0 | 1 | 44.125 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-17277 | 2021/07/30 23:07:38 | 1-13 | 2 | -7.593 | -62.195 | 18139 | -1 | 1 | EV | standard | H2O | 1627708365 | 0.314 | 0.705 | 96.876 | 0.001 | 0.002 | 2.29 | 1297.374 | 2.817 | 2.882 | 110.859 | 0.87 | 0.46 | 0 | 0 | 1 | 44.125 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-17277 | 2021/07/30 23:16:34 | 1-13 | 3 | -7.506 | -61.01 | 18128 | -1 | 1 | EV | standard | H2O | 1627708901 | 0.282 | 0.658 | 99.496 | 0 | 0.002 | 2.368 | 1297.256 | 3.084 | 2.864 | 109.843 | 0.87 | 0.234 | 0 | 0 | 1 | 44.125 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-17277 | 2021/07/30 23:25:31 | 1-13 | 4 | -7.568 | -60.869 | 18105 | 0 | 1 | EV | standard | H2O | 1627709437 | 0.333 | 0.737 | 99.711 | 0 | 0.006 | 2.372 | 1297.119 | 2.552 | 2.862 | 111.38 | 0.865 | 0.427 | 0 | 0 | 1 | 44.062 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-17278 | 2021/07/30 23:35:20 | 1-14 | 1 | [5 empty] | NIWO2-D-3 | 20-213 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-17278 | 2021/07/30 23:45:19 | 1-14 | 2 | [5 empty] | NIWO2-D-3 | 20-213 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-17278 | 2021/07/30 23:55:17 | 1-14 | 3 | -9.411 | -102.534 | 17654 | 0 | 1 | NIWO2-D-3 | 20-213 | H2O | 1627711821 | 0.284 | 0.855 | 90.369 | 0 | 0.009 | 2.14 | 1295.492 | 2.052 | 2.72 | 106.917 | 0.864 | 0.033 | 0 | 0 | 1 | 44.438 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-17278 | 2021/07/31 00:05:15 | 1-14 | 4 | -9.411 | -102.534 | 17654 | 0 | 1 | NIWO2-D-3 | 20-213 | H2O | 1627711821 | 0.284 | 0.855 | 90.369 | 0 | 0.009 | 2.14 | 1295.492 | 2.052 | 2.72 | 106.917 | 0.864 | 0.033 | 0 | 0 | 1 | 44.438 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-17279 | 2021/07/31 00:15:02 | 1-15 | 1 | -2.688 | -36.276 | 17953 | -1 | 1 | SLCP-21-015 | Precipitation | H2O | 1627712408 | 0.283 | 0.801 | 94.268 | -0.001 | -0.006 | 2.241 | 1295.882 | 2.404 | 2.901 | 111.15 | 0.871 | 0.617 | 0 | 0 | 1 | 44.438 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-17279 | 2021/07/31 00:23:58 | 1-15 | 2 | -2.12 | -27.551 | 18087 | -1 | 1 | SLCP-21-015 | Precipitation | H2O | 1627712944 | 0.307 | 0.752 | 96.692 | -0.001 | -0.005 | 2.299 | 1297.145 | 2.5 | 2.826 | 109.318 | 0.866 | 0.149 | 0 | 0 | 1 | 44.438 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-17279 | 2021/07/31 00:32:54 | 1-15 | 3 | -2.041 | -25.677 | 18134 | -1 | 1 | SLCP-21-015 | Precipitation | H2O | 1627713480 | 0.289 | 0.686 | 98.944 | 0 | -0.001 | 2.347 | 1297.26 | 2.425 | 2.896 | 110.888 | 0.873 | 0.284 | 0 | 0 | 1 | 44.375 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-17279 | 2021/07/31 00:41:51 | 1-15 | 4 | -2 | -25.315 | 18030 | 0 | 1 | SLCP-21-015 | Precipitation | H2O | 1627714017 | 0.318 | 0.75 | 100.182 | 0 | 0 | 2.376 | 1296.329 | 2.734 | 2.836 | 110.7 | 0.869 | 0.458 | 0 | 0 | 1 | 44.375 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-17280 | 2021/07/31 00:51:37 | 1-16 | 1 | -8.244 | -75.114 | 18251 | -1 | 1 | 262 | 21-030 | H2O | 1627714604 | 0.308 | 0.741 | 101.001 | 0 | 0.004 | 2.403 | 1298.409 | 2.442 | 2.914 | 112.137 | 0.871 | 0.419 | 0 | 0 | 1 | 44.438 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-17280 | 2021/07/31 01:00:34 | 1-16 | 2 | -8.761 | -81.622 | 18170 | -1 | 1 | 262 | 21-030 | H2O | 1627715140 | 0.317 | 0.703 | 99.791 | 0 | 0.005 | 2.37 | 1297.734 | 2.714 | 2.878 | 111.328 | 0.867 | 0.381 | 0 | 0 | 1 | 44.375 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-17280 | 2021/07/31 01:09:30 | 1-16 | 3 | -8.905 | -83.367 | 17837 | -1 | 1 | 262 | 21-030 | H2O | 1627715676 | 0.303 | 0.721 | 98.178 | 0 | 0.003 | 2.339 | 1295.372 | 2.661 | 2.809 | 109.046 | 0.865 | 0.372 | 0 | 0 | 1 | 44.375 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-17280 | 2021/07/31 01:18:27 | 1-16 | 4 | -8.889 | -83.537 | 18212 | 0 | 1 | 262 | 21-030 | H2O | 1627716213 | 0.322 | 0.805 | 103.637 | 0 | 0.005 | 2.445 | 1298.193 | 2.327 | 2.873 | 111.328 | 0.873 | 0.274 | 0 | 0 | 1 | 44.375 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-17281 | 2021/07/31 01:28:14 | 1-17 | 1 | -8.292 | -84.775 | 18103 | -1 | 1 | 237 | 21-030 | H2O | 1627716801 | 0.329 | 0.804 | 101.205 | 0.001 | 0.007 | 2.411 | 1297.274 | 2.059 | 2.896 | 112.914 | 0.869 | 0.564 | 0 | 0 | 1 | 44.375 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-17281 | 2021/07/31 01:37:11 | 1-17 | 2 | -8.223 | -84.97 | 17944 | -1 | 1 | 237 | 21-030 | H2O | 1627717338 | 0.307 | 0.782 | 102.201 | 0.001 | 0.007 | 2.421 | 1296.25 | 2.031 | 2.84 | 111.034 | 0.869 | 0.327 | 0 | 0 | 1 | 44.312 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-17281 | 2021/07/31 01:46:08 | 1-17 | 3 | -8.124 | -84.704 | 18051 | -1 | 1 | 237 | 21-030 | H2O | 1627717874 | 0.292 | 0.669 | 105.353 | 0 | 0.004 | 2.512 | 1297.115 | 2.99 | 2.785 | 108.742 | 0.869 | 0.205 | 0 | 0 | 1 | 44.375 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-17281 | 2021/07/31 01:55:04 | 1-17 | 4 | -8.166 | -84.833 | 18091 | 0 | 1 | 237 | 21-030 | H2O | 1627718411 | 0.312 | 0.734 | 99.58 | 0 | 0.002 | 2.373 | 1297.534 | 2.301 | 2.814 | 110.575 | 0.869 | 0.275 | 0 | 0 | 1 | 44.312 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-17282 | 2021/07/31 02:04:52 | 1-18 | 1 | -8.969 | -56.85 | 18121 | -1 | 1 | 243 | 21-030 | H2O | 1627718999 | 0.306 | 0.75 | 101.65 | 0 | 0 | 2.421 | 1297.437 | 2.358 | 2.935 | 112.158 | 0.871 | 0.59 | 0 | 0 | 1 | 44.375 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-17282 | 2021/07/31 02:13:49 | 1-18 | 2 | -9.064 | -53.542 | 17751 | -1 | 1 | 243 | 21-030 | H2O | 1627719536 | 0.294 | 0.689 | 97.746 | 0.001 | -0.002 | 2.343 | 1295.075 | 2.133 | 2.833 | 109.494 | 0.874 | 0.437 | 0 | 0 | 1 | 44.438 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-17282 | 2021/07/31 02:22:45 | 1-18 | 3 | -9.038 | -52.365 | 18111 | -1 | 1 | 243 | 21-030 | H2O | 1627720072 | 0.343 | 0.767 | 98.264 | 0 | 0.001 | 2.334 | 1297.924 | 2.389 | 2.775 | 109.368 | 0.871 | 0.01 | 0 | 0 | 1 | 44.5 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-17282 | 2021/07/31 02:31:41 | 1-18 | 4 | -9.16 | -52.587 | 17639 | 0 | 1 | 243 | 21-030 | H2O | 1627720608 | 0.333 | 0.766 | 99.984 | 0 | 0.002 | 2.379 | 1294.242 | 2.331 | 2.729 | 107.797 | 0.87 | 0.12 | 0 | 0 | 1 | 44.5 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-17283 | 2021/07/31 02:41:29 | 1-19 | 1 | -2.2 | -54.561 | 17683 | -1 | 1 | 250 | 21-030 | H2O | 1627721196 | 0.423 | 1.168 | 100.317 | -0.001 | 0.004 | 2.387 | 2076.673 | -34.66 | 3.02 | 109.114 | 0.886 | -3.846 | 0 | 0 | 1 | 44.5 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-17283 | 2021/07/31 02:50:25 | 1-19 | 2 | -1.901 | -55.326 | 17716 | -1 | 1 | 250 | 21-030 | H2O | 1627721732 | 0.458 | 1.239 | 99.742 | 0 | 0.002 | 2.378 | 2082.088 | -35.503 | 3.042 | 110.504 | 0.888 | -3.704 | 0 | 0 | 1 | 44.5 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-17283 | 2021/07/31 02:59:22 | 1-19 | 3 | -1.914 | -55.773 | 17633 | -1 | 1 | 250 | 21-030 | H2O | 1627722269 | 0.419 | 1.144 | 103.864 | -0.002 | 0.005 | 2.459 | 2078.85 | -35.249 | 2.945 | 108.581 | 0.89 | -3.877 | 0 | 0 | 1 | 44.562 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-17283 | 2021/07/31 03:08:19 | 1-19 | 4 | -1.906 | -55.597 | 17630 | 0 | 1 | 250 | 21-030 | H2O | 1627722805 | 0.429 | 1.384 | 96.703 | 0 | 0.005 | 2.31 | 2079.891 | -35.217 | 3.018 | 109.772 | 0.888 | -3.483 | 0 | 0 | 1 | 44.5 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-17284 | 2021/07/31 03:18:07 | 1-20 | 1 | -8.46 | -74.944 | 18298 | -1 | 1 | 258 | 21-030 | H2O | 1627723394 | 0.316 | 0.758 | 99.031 | 0.001 | 0.004 | 2.35 | 1308.904 | 1.933 | 2.871 | 111.061 | 0.867 | -0.037 | 0 | 0 | 1 | 44.5 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-17284 | 2021/07/31 03:27:03 | 1-20 | 2 | -8.738 | -77.057 | 18248 | -1 | 1 | 258 | 21-030 | H2O | 1627723929 | 0.307 | 0.77 | 104.472 | 0 | 0.006 | 2.484 | 1302.34 | 2.147 | 2.853 | 111.816 | 0.876 | 0.237 | 0 | 0 | 1 | 44.5 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-17284 | 2021/07/31 03:35:59 | 1-20 | 3 | -8.741 | -77.458 | 18000 | -1 | 1 | 258 | 21-030 | H2O | 1627724465 | 0.287 | 0.757 | 103.927 | 0 | 0.007 | 2.465 | 1299.022 | 2.224 | 2.897 | 110.583 | 0.869 | 0.375 | 0 | 0 | 1 | 44.5 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-17284 | 2021/07/31 03:44:55 | 1-20 | 4 | -8.659 | -77.208 | 18175 | 0 | 1 | 258 | 21-030 | H2O | 1627725002 | 0.321 | 0.811 | 103.743 | 0.001 | 0.008 | 2.464 | 1299.496 | 2.922 | 2.814 | 110.302 | 0.871 | 0.158 | 0 | 0 | 1 | 44.562 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-17285 | 2021/07/31 03:54:44 | 1-21 | 1 | [5 empty] | 248 | 21-030 | [18 empty] | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-17285 | 2021/07/31 04:03:41 | 1-21 | 2 | [5 empty] | 248 | 21-030 | [18 empty] | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-17285 | 2021/07/31 04:12:38 | 1-21 | 3 | -6.441 | -72.554 | 18139 | -1 | 1 | 248 | 21-030 | H2O | 1627726664 | 0.324 | 0.721 | 101.62 | 0 | 0.004 | 2.412 | 1298.64 | 2.35 | 2.848 | 110.206 | 0.867 | 0.197 | 0 | 0 | 1 | 44.5 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-17285 | 2021/07/31 04:21:34 | 1-21 | 4 | -6.498 | -72.672 | 17954 | 0 | 1 | 248 | 21-030 | H2O | 1627727201 | 0.331 | 0.79 | 103.16 | -0.001 | 0.002 | 2.449 | 1296.927 | 2.656 | 2.852 | 109.586 | 0.868 | 0.338 | 0 | 0 | 1 | 44.438 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-17286 | 2021/07/31 04:31:19 | 1-22 | 1 | -7.516 | -62.562 | 18075 | -1 | 1 | EV | standard | H2O | 1627727786 | 0.298 | 0.768 | 102.59 | 0.001 | 0.007 | 2.443 | 1297.365 | 2.256 | 2.86 | 110.593 | 0.87 | 0.208 | 0 | 0 | 1 | 44.5 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-17286 | 2021/07/31 04:40:16 | 1-22 | 2 | -7.418 | -60.564 | 18229 | -1 | 1 | EV | standard | H2O | 1627728323 | 0.314 | 0.682 | 102.51 | 0 | -0.001 | 2.443 | 1298.574 | 2.919 | 2.843 | 110.013 | 0.869 | 0.253 | 0 | 0 | 1 | 44.5 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-17286 | 2021/07/31 04:49:12 | 1-22 | 3 | -7.465 | -60.554 | 18000 | -1 | 1 | EV | standard | H2O | 1627728859 | 0.316 | 0.705 | 100.13 | -0.001 | 0 | 2.387 | 1296.7 | 2.483 | 2.88 | 110.817 | 0.87 | 0.571 | 0 | 0 | 1 | 44.5 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-17286 | 2021/07/31 04:58:09 | 1-22 | 4 | -7.652 | -60.924 | 17772 | 0 | 1 | EV | standard | H2O | 1627729395 | 0.323 | 0.721 | 101.167 | 0.002 | 0.003 | 2.406 | 1295.014 | 2.17 | 2.801 | 109.991 | 0.871 | 0.473 | 0 | 0 | 1 | 44.5 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-17287 | 2021/07/31 05:07:55 | 1-23 | 1 | -2.766 | -52.983 | 17614 | -1 | 1 | 255 | 21-030 | H2O | 1627729981 | 0.398 | 1.118 | 98.377 | 0 | 0.002 | 2.348 | 1900.72 | -27.157 | 2.917 | 109.399 | 0.884 | -3.029 | 0 | 0 | 1 | 44.562 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-17287 | 2021/07/31 05:16:51 | 1-23 | 2 | -2.642 | -52.473 | 17862 | -1 | 1 | 255 | 21-030 | H2O | 1627730517 | 0.381 | 1.08 | 102.096 | -0.001 | -0.001 | 2.435 | 1915.582 | -27.852 | 2.955 | 109.565 | 0.889 | -3.225 | 0 | 0 | 1 | 44.5 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-17287 | 2021/07/31 05:25:48 | 1-23 | 3 | -2.646 | -52.739 | 17754 | -1 | 1 | 255 | 21-030 | H2O | 1627731054 | 0.372 | 1.01 | 100.388 | 0 | 0.008 | 2.403 | 1913.422 | -27.96 | 2.95 | 109.732 | 0.886 | -3.256 | 0 | 0 | 1 | 44.5 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-17287 | 2021/07/31 05:34:44 | 1-23 | 4 | -2.658 | -52.465 | 17808 | 0 | 1 | 255 | 21-030 | H2O | 1627731590 | 0.379 | 0.952 | 101.398 | -0.001 | 0.004 | 2.403 | 1914.487 | -28.193 | 2.961 | 110.782 | 0.874 | -2.926 | 0 | 0 | 1 | 44.5 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-17288 | 2021/07/31 05:44:30 | 1-24 | 1 | -5.27 | -70.043 | 17685 | -1 | 1 | 259 | 21-030 | H2O | 1627732177 | 0.39 | 1.167 | 102.521 | -0.001 | 0.006 | 2.452 | 1955.233 | -27.744 | 2.856 | 107.523 | 0.882 | -2.724 | 0 | 0 | 1 | 44.5 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-17288 | 2021/07/31 05:53:27 | 1-24 | 2 | -5.409 | -72.167 | 17441 | -1 | 1 | 259 | 21-030 | H2O | 1627732713 | 0.388 | 1.031 | 97.608 | 0 | 0.002 | 2.356 | 1943.517 | -27.432 | 2.811 | 105.948 | 0.878 | -2.709 | 0 | 0 | 1 | 44.5 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-17288 | 2021/07/31 06:02:23 | 1-24 | 3 | -5.436 | -72.289 | 17836 | -1 | 1 | 259 | 21-030 | H2O | 1627733249 | 0.373 | 1.048 | 100.678 | 0 | 0.001 | 2.4 | 1962.286 | -28.357 | 2.89 | 109.093 | 0.874 | -2.601 | 0 | 0 | 1 | 44.5 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-17288 | 2021/07/31 06:11:18 | 1-24 | 4 | -5.397 | -72.661 | 17859 | 0 | 1 | 259 | 21-030 | H2O | 1627733784 | 0.419 | 1.049 | 100.316 | 0 | 0.002 | 2.384 | 1961.306 | -28.695 | 2.838 | 110.541 | 0.882 | -2.569 | 0 | 0 | 1 | 44.438 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-17289 | 2021/07/31 06:21:04 | 1-25 | 1 | -4.439 | -68.734 | 17584 | -1 | 1 | 257 | 21-030 | H2O | 1627734371 | 0.398 | 1.084 | 97.598 | 0 | 0.001 | 2.322 | 2092.433 | -34.04 | 2.869 | 107.327 | 0.88 | -2.918 | 0 | 0 | 1 | 44.438 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-17289 | 2021/07/31 06:30:00 | 1-25 | 2 | -4.402 | -68.876 | 17348 | -1 | 1 | 257 | 21-030 | H2O | 1627734907 | 0.454 | 1.257 | 98.662 | 0 | 0.006 | 2.358 | 2081.598 | -33.207 | 2.919 | 106.546 | 0.876 | -2.849 | 0 | 0 | 1 | 44.438 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-17289 | 2021/07/31 06:38:56 | 1-25 | 3 | -4.393 | -68.423 | 17341 | -1 | 1 | 257 | 21-030 | H2O | 1627735442 | 0.417 | 1.157 | 96.756 | 0.002 | 0.001 | 2.309 | 2081.045 | -33.212 | 2.861 | 105.375 | 0.878 | -2.71 | 0 | 0 | 1 | 44.5 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-17289 | 2021/07/31 06:47:52 | 1-25 | 4 | -4.335 | -68.177 | 17700 | 0 | 1 | 257 | 21-030 | H2O | 1627735979 | 0.426 | 1.239 | 100.013 | 0.001 | 0.006 | 2.389 | 2098.967 | -33.437 | 2.901 | 107.624 | 0.882 | -2.88 | 0 | 0 | 1 | 44.438 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-17290 | 2021/07/31 06:57:39 | 1-26 | 1 | -9.925 | -104.817 | 17734 | -1 | 1 | 249 | 21-030 | H2O | 1627736565 | 0.309 | 0.825 | 100.984 | 0.001 | 0.012 | 2.402 | 1306.278 | 1.851 | 2.794 | 109.62 | 0.871 | 0.294 | 0 | 0 | 1 | 44.438 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-17290 | 2021/07/31 07:06:35 | 1-26 | 2 | -9.983 | -108.365 | 17824 | -1 | 1 | 249 | 21-030 | H2O | 1627737102 | 0.314 | 0.771 | 98.199 | 0 | 0.007 | 2.331 | 1300.538 | 2.269 | 2.786 | 109.719 | 0.868 | 0.448 | 0 | 0 | 1 | 44.438 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-17290 | 2021/07/31 07:15:31 | 1-26 | 3 | -10.051 | -109.45 | 18069 | -1 | 1 | 249 | 21-030 | H2O | 1627737638 | 0.292 | 0.697 | 100.941 | -0.001 | 0.006 | 2.381 | 1300.902 | 2.084 | 2.864 | 110.814 | 0.868 | 0.402 | 0 | 0 | 1 | 44.438 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-17290 | 2021/07/31 07:24:27 | 1-26 | 4 | -10.084 | -109.88 | 18039 | 0 | 1 | 249 | 21-030 | H2O | 1627738174 | 0.314 | 0.715 | 99.82 | 0 | 0.004 | 2.374 | 1299.478 | 2.851 | 2.819 | 109.722 | 0.871 | 0.25 | 0 | 0 | 1 | 44.5 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-17291 | 2021/07/31 07:34:15 | 1-27 | 1 | -9.482 | -97.491 | 17031 | -1 | 1 | 244 | 21-030 | H2O | 1627738761 | 0.31 | 0.688 | 97.156 | -0.001 | 0.003 | 2.313 | 1291.887 | 1.791 | 2.683 | 105.185 | 0.865 | 0.187 | 0 | 0 | 1 | 44.5 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-17291 | 2021/07/31 07:43:11 | 1-27 | 2 | -9.48 | -95.858 | 16520 | -1 | 1 | 244 | 21-030 | H2O | 1627739298 | 0.327 | 0.79 | 95.813 | -0.002 | 0.003 | 2.289 | 1288.678 | 1.908 | 2.603 | 101.415 | 0.867 | -0.059 | 0 | 0 | 1 | 44.5 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-17291 | 2021/07/31 07:52:07 | 1-27 | 3 | -9.562 | -96.203 | 15694 | -1 | 1 | 244 | 21-030 | H2O | 1627739834 | 0.321 | 0.813 | 94.094 | -0.001 | 0.003 | 2.267 | 1282.55 | 2.096 | 2.445 | 96.92 | 0.865 | 0.061 | 0 | 0 | 1 | 44.5 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-17291 | 2021/07/31 08:01:04 | 1-27 | 4 | [5 empty] | 244 | 21-030 | H2O | [17 empty] | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-17292 | 2021/07/31 08:11:53 | 1-28 | 1 | [5 empty] | 101 | 21-030 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-17292 | 2021/07/31 08:20:49 | 1-28 | 2 | -9.691 | -93.674 | 18188 | -1 | 1 | 101 | 21-030 | H2O | 1627741556 | 0.305 | 0.702 | 96.875 | 0 | 0.003 | 2.306 | 1298.697 | 2.507 | 2.873 | 112.307 | 0.87 | 0.526 | 0 | 0 | 1 | 44.562 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-17292 | 2021/07/31 08:29:45 | 1-28 | 3 | [5 empty] | 101 | 21-030 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-17292 | 2021/07/31 08:38:42 | 1-28 | 4 | -9.623 | -93.543 | 18103 | 0 | 1 | 101 | 21-030 | H2O | 1627742628 | 0.353 | 0.76 | 100.823 | 0 | 0.003 | 2.396 | 1298.246 | 2.761 | 2.831 | 109.563 | 0.871 | 0.182 | 0 | 0 | 1 | 44.438 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-17293 | 2021/07/31 08:48:26 | 1-29 | 1 | [5 empty] | 123 | 21-030 | [18 empty] | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-17293 | 2021/07/31 08:57:22 | 1-29 | 2 | [5 empty] | 123 | 21-030 | [18 empty] | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-17293 | 2021/07/31 09:06:18 | 1-29 | 3 | 0.567 | -62.354 | 17459 | -1 | 1 | 123 | 21-030 | H2O | 1627743749 | 0.461 | 1.479 | 99.357 | -0.002 | 0 | 2.362 | 2453.585 | -58.217 | 3.804 | 115.025 | 0.897 | -6.849 | 0 | 0 | 1 | 44.438 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-17293 | 2021/07/31 09:15:15 | 1-29 | 4 | 0.56 | -62.118 | 17251 | -1 | 1 | 123 | 21-030 | H2O | 1627744285 | 0.508 | 1.331 | 99.145 | -0.001 | -0.001 | 2.365 | 2442.222 | -57.683 | 3.768 | 113.694 | 0.893 | -6.7 | 0 | 0 | 1 | 44.438 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-17294 | 2021/07/31 09:25:00 | 1-30 | 1 | 4.844 | -65.401 | 16356 | -1 | 1 | 245 | 21-030 | H2O | 1627745406 | 0.359 | 1.286 | 98.619 | 0 | 0.008 | 2.358 | 1785.716 | -39.44 | 4.375 | 115.677 | 0.876 | -8.462 | 0 | 0 | 1 | 44.438 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-17294 | 2021/07/31 09:33:56 | 1-30 | 2 | 5.041 | -65.695 | 17769 | -1 | 1 | 245 | 21-030 | H2O | 1627745943 | 0.379 | 1.258 | 100.859 | -0.002 | 0.008 | 2.398 | 1834.287 | -43.676 | 4.738 | 126.304 | 0.877 | -9.194 | 0 | 0 | 1 | 44.438 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-17294 | 2021/07/31 09:42:52 | 1-30 | 3 | 5.018 | -65.906 | 17743 | -1 | 1 | 245 | 21-030 | H2O | 1627746478 | 0.373 | 1.244 | 101.194 | -0.003 | 0.005 | 2.409 | 1831.931 | -42.842 | 4.752 | 125.654 | 0.875 | -9.005 | 0 | 0 | 1 | 44.375 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-17294 | 2021/07/31 09:51:48 | 1-30 | 4 | 5.019 | -66.017 | 17675 | 0 | 1 | 245 | 21-030 | H2O | 1627747015 | 0.405 | 1.138 | 100.309 | -0.002 | 0.005 | 2.384 | 1828.343 | -42.75 | 4.717 | 125.035 | 0.887 | -9.372 | 0 | 0 | 1 | 44.438 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-17295 | 2021/07/31 10:01:34 | 1-31 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-17295 | 2021/07/31 10:10:29 | 1-31 | 2 | [5 empty] | EV | standard | [18 empty] | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-17295 | 2021/07/31 10:19:26 | 1-31 | 3 | -7.471 | -61.027 | 17843 | -1 | 1 | EV | standard | H2O | 1627748672 | 0.306 | 0.705 | 100.343 | 0 | 0.001 | 2.38 | 1298.191 | 2.749 | 2.736 | 107.535 | 0.873 | 0.086 | 0 | 0 | 1 | 44.438 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-17295 | 2021/07/31 10:28:22 | 1-31 | 4 | -7.425 | -60.718 | 17853 | 0 | 1 | EV | standard | H2O | 1627749209 | 0.318 | 0.75 | 102.206 | 0 | 0.001 | 2.441 | 1297.718 | 2.475 | 2.772 | 108.935 | 0.868 | 0.367 | 0 | 0 | 1 | 44.438 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-17296 | 2021/07/31 10:38:07 | 1-32 | 1 | -4.066 | -68.891 | 14856 | -1 | 1 | 253 | 21-030 | H2O | 1627749794 | 0.354 | 1.085 | 92.599 | -0.001 | 0.005 | 2.24 | 1842.489 | -25.7 | 2.531 | 92.883 | 0.877 | -2.91 | 0 | 0 | 1 | 44.5 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-17296 | 2021/07/31 10:47:04 | 1-32 | 2 | -3.751 | -68.355 | 17646 | -1 | 1 | 253 | 21-030 | H2O | 1627750330 | 0.399 | 1.201 | 98.65 | 0 | 0.006 | 2.349 | 1977.119 | -31.26 | 2.97 | 109.839 | 0.875 | -3.308 | 0 | 0 | 1 | 44.438 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-17296 | 2021/07/31 10:56:00 | 1-32 | 3 | -3.916 | -69.295 | 17493 | -1 | 1 | 253 | 21-030 | H2O | 1627750867 | 0.391 | 1.046 | 101.082 | 0.001 | 0.008 | 2.411 | 1972.197 | -31.042 | 2.956 | 108.704 | 0.874 | -3.478 | 0 | 0 | 1 | 44.438 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-17296 | 2021/07/31 11:04:56 | 1-32 | 4 | -3.808 | -69.018 | 17597 | 0 | 1 | 253 | 21-030 | H2O | 1627751403 | 0.376 | 1.172 | 97.919 | -0.001 | 0.005 | 2.336 | 1975.77 | -31.195 | 2.973 | 110.488 | 0.874 | -3.285 | 0 | 0 | 1 | 44.5 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-17297 | 2021/07/31 11:14:43 | 1-33 | 1 | [5 empty] | 241 | 21-030 | [18 empty] | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-17297 | 2021/07/31 11:23:38 | 1-33 | 2 | [5 empty] | 241 | 21-030 | [18 empty] | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-17297 | 2021/07/31 11:32:34 | 1-33 | 3 | -8.803 | -85.091 | 18213 | -1 | 1 | 241 | 21-030 | H2O | 1627753061 | 0.306 | 0.744 | 102.892 | 0 | 0.006 | 2.448 | 1300.992 | 2.525 | 2.834 | 110.68 | 0.87 | 0.247 | 0 | 0 | 1 | 44.375 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-17297 | 2021/07/31 11:41:31 | 1-33 | 4 | -8.839 | -85.259 | 17797 | 0 | 1 | 241 | 21-030 | H2O | 1627753596 | 0.31 | 0.733 | 101.238 | -0.001 | 0.006 | 2.405 | 1297.012 | 2.33 | 2.785 | 109.551 | 0.872 | 0.302 | 0 | 0 | 1 | 44.375 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-17298 | 2021/07/31 11:51:17 | 1-34 | 1 | -9.191 | -89.124 | 17963 | -1 | 1 | 232 | 21-030 | H2O | 1627754183 | 0.3 | 0.75 | 101.434 | 0 | 0.002 | 2.41 | 1297.583 | 2.355 | 2.837 | 110.339 | 0.872 | 0.265 | 0 | 0 | 1 | 44.375 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-17298 | 2021/07/31 12:00:13 | 1-34 | 2 | -9.225 | -89.691 | 18013 | -1 | 1 | 232 | 21-030 | H2O | 1627754720 | 0.319 | 0.823 | 99.787 | 0 | 0.007 | 2.376 | 1298.173 | 2.451 | 2.845 | 110.939 | 0.87 | 0.44 | 0 | 0 | 1 | 44.312 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-17298 | 2021/07/31 12:09:09 | 1-34 | 3 | -9.17 | -89.676 | 17799 | -1 | 1 | 232 | 21-030 | H2O | 1627755256 | 0.309 | 0.734 | 101.359 | 0 | 0.003 | 2.399 | 1296.701 | 2.336 | 2.779 | 108.963 | 0.865 | 0.32 | 0 | 0 | 1 | 44.375 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-17298 | 2021/07/31 12:18:05 | 1-34 | 4 | -9.325 | -90.427 | 17485 | 0 | 1 | 232 | 21-030 | H2O | 1627755791 | 0.311 | 0.784 | 99.786 | 0.001 | 0.007 | 2.377 | 1294.153 | 2.246 | 2.721 | 107.336 | 0.87 | 0.281 | 0 | 0 | 1 | 44.375 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-17299 | 2021/07/31 12:27:52 | 1-35 | 1 | -9.507 | -90.767 | 16865 | -1 | 1 | 227 | 21-030 | H2O | 1627756379 | 0.294 | 0.811 | 96.713 | 0.001 | 0.002 | 2.322 | 1289.131 | 2.514 | 2.609 | 102.482 | 0.868 | 0.061 | 0 | 0 | 1 | 44.375 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-17299 | 2021/07/31 12:36:49 | 1-35 | 2 | -9.459 | -90.674 | 17228 | -1 | 1 | 227 | 21-030 | H2O | 1627756915 | 0.282 | 0.788 | 99.652 | -0.001 | 0.002 | 2.377 | 1292.023 | 2.057 | 2.712 | 105.506 | 0.868 | 0.27 | 0 | 0 | 1 | 44.375 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-17299 | 2021/07/31 12:45:44 | 1-35 | 3 | -9.446 | -90.858 | 17087 | -1 | 1 | 227 | 21-030 | H2O | 1627757450 | 0.284 | 0.751 | 97.817 | -0.001 | 0.002 | 2.326 | 1290.795 | 2.559 | 2.671 | 104.804 | 0.866 | 0.251 | 0 | 0 | 1 | 44.312 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-17299 | 2021/07/31 12:54:40 | 1-35 | 4 | -9.433 | -90.694 | 17121 | 0 | 1 | 227 | 21-030 | H2O | 1627757986 | 0.299 | 0.75 | 98.051 | 0 | 0.002 | 2.324 | 1291.007 | 2.399 | 2.689 | 104.681 | 0.874 | 0.292 | 0 | 0 | 1 | 44.312 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-17300 | 2021/07/31 13:04:22 | 1-36 | 1 | -8.926 | -87.89 | 18033 | -1 | 1 | 229 | 21-030 | H2O | 1627758569 | 0.324 | 0.661 | 98.557 | -0.001 | 0.002 | 2.342 | 1297.011 | 2.549 | 2.821 | 110.959 | 0.871 | 0.544 | 0 | 0 | 1 | 44.312 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-17300 | 2021/07/31 13:13:17 | 1-36 | 2 | -8.913 | -87.771 | 17851 | -1 | 1 | 229 | 21-030 | H2O | 1627759104 | 0.324 | 0.748 | 100.725 | 0.001 | 0.004 | 2.408 | 1295.884 | 2.932 | 2.802 | 107.899 | 0.871 | 0.339 | 0 | 0 | 1 | 44.25 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-17300 | 2021/07/31 13:22:13 | 1-36 | 3 | -8.836 | -87.413 | 17991 | -1 | 1 | 229 | 21-030 | H2O | 1627759640 | 0.329 | 0.644 | 101.229 | -0.001 | 0.003 | 2.388 | 1296.948 | 2.082 | 2.839 | 110.819 | 0.868 | 0.438 | 0 | 0 | 1 | 44.312 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-17300 | 2021/07/31 13:31:09 | 1-36 | 4 | -8.897 | -87.727 | 17856 | 0 | 1 | 229 | 21-030 | H2O | 1627760175 | 0.334 | 0.758 | 99.55 | 0.001 | 0.005 | 2.368 | 1296.075 | 2.632 | 2.757 | 108.611 | 0.87 | 0.266 | 0 | 0 | 1 | 44.312 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-17301 | 2021/07/31 13:40:53 | 1-37 | 1 | [5 empty] | 193 | 21-030 | [18 empty] | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-17301 | 2021/07/31 13:49:49 | 1-37 | 2 | [5 empty] | 193 | 21-030 | [18 empty] | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-17301 | 2021/07/31 13:58:44 | 1-37 | 3 | -7.761 | -67.566 | 16336 | -1 | 1 | 193 | 21-030 | H2O | 1627761295 | 0.307 | 0.708 | 96.037 | 0.001 | 0.003 | 2.314 | 1287.516 | 2.298 | 2.55 | 99.089 | 0.871 | -0.005 | 0 | 0 | 1 | 44.25 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-17301 | 2021/07/31 14:07:40 | 1-37 | 4 | -7.824 | -67.851 | 15099 | -1 | 1 | 193 | 21-030 | H2O | 1627761830 | 0.293 | 0.734 | 93.359 | 0 | 0 | 2.253 | 1278.572 | 2.192 | 2.315 | 92.563 | 0.863 | 0.24 | 0 | 0 | 1 | 44.25 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-17302 | 2021/07/31 14:17:22 | 1-38 | 1 | -9.489 | -87.803 | 17828 | -1 | 1 | 233 | 21-030 | H2O | 1627762949 | 0.328 | 0.684 | 94.777 | 0 | 0.003 | 2.251 | 1295.489 | 2.56 | 2.824 | 109.615 | 0.871 | 0.413 | 0 | 0 | 1 | 44.312 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-17302 | 2021/07/31 14:26:17 | 1-38 | 2 | -9.562 | -90.783 | 17403 | -1 | 1 | 233 | 21-030 | H2O | 1627763484 | 0.319 | 0.731 | 98.329 | 0 | 0.001 | 2.349 | 1292.627 | 2.422 | 2.739 | 106.742 | 0.872 | 0.23 | 0 | 0 | 1 | 44.25 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-17302 | 2021/07/31 14:35:13 | 1-38 | 3 | -9.597 | -91.204 | 17797 | -1 | 1 | 233 | 21-030 | H2O | 1627764019 | 0.293 | 0.681 | 99.509 | -0.001 | 0.002 | 2.363 | 1295.376 | 2.785 | 2.725 | 107.998 | 0.871 | 0.23 | 0 | 0 | 1 | 44.25 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-17302 | 2021/07/31 14:44:09 | 1-38 | 4 | -9.546 | -91.178 | 18054 | 0 | 1 | 233 | 21-030 | H2O | 1627764555 | 0.294 | 0.662 | 104.172 | 0 | 0.002 | 2.486 | 1297.451 | 2.568 | 2.784 | 109.088 | 0.868 | 0.065 | 0 | 0 | 1 | 44.25 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-17303 | 2021/07/31 14:53:53 | 1-39 | 1 | [5 empty] | 242 | 21-030 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-17303 | 2021/07/31 15:03:51 | 1-39 | 2 | [5 empty] | 242 | 21-030 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-17303 | 2021/07/31 15:13:49 | 1-39 | 3 | [5 empty] | 242 | 21-030 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-17303 | 2021/07/31 15:23:47 | 1-39 | 4 | [5 empty] | 242 | 21-030 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-17304 | 2021/07/31 15:34:34 | 1-40 | 1 | -7.875 | -64.747 | 17704 | -1 | 1 | EV | standard | H2O | 1627767580 | 0.304 | 0.831 | 86.096 | 0 | -0.002 | 2.052 | 1294.363 | 2.909 | 2.811 | 108.11 | 0.869 | 0.361 | 0 | 0 | 1 | 44.688 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-17304 | 2021/07/31 15:43:29 | 1-40 | 2 | -7.677 | -62.211 | 17297 | -1 | 1 | EV | standard | H2O | 1627768115 | 0.319 | 0.68 | 94.853 | 0 | 0 | 2.258 | 1291.654 | 2.817 | 2.679 | 103.924 | 0.869 | 0.078 | 0 | 0 | 1 | 44.5 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-17304 | 2021/07/31 15:52:24 | 1-40 | 3 | -7.568 | -61.363 | 17272 | -1 | 1 | EV | standard | H2O | 1627768650 | 0.301 | 0.756 | 93.265 | 0 | 0.002 | 2.222 | 1291.095 | 3.003 | 2.702 | 105.47 | 0.864 | 0.339 | 0 | 0 | 1 | 44.312 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-17304 | 2021/07/31 16:01:19 | 1-40 | 4 | -7.583 | -61.143 | 17292 | 0 | 1 | EV | standard | H2O | 1627769185 | 0.332 | 0.775 | 96.483 | -0.001 | 0.001 | 2.305 | 1291.533 | 2.63 | 2.71 | 105.117 | 0.867 | 0.319 | 0 | 0 | 1 | 44.312 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-17305 | 2021/07/31 16:11:03 | 1-41 | 1 | -11.114 | -107.072 | 16911 | -1 | 1 | 167 | 21-030 | H2O | 1627769770 | 0.319 | 0.901 | 98.293 | 0.002 | 0.012 | 2.346 | 1289.991 | 2.936 | 2.622 | 102.563 | 0.868 | 0.286 | 0 | 0 | 1 | 44.25 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-17305 | 2021/07/31 16:20:00 | 1-41 | 2 | -11.219 | -112.149 | 17899 | -1 | 1 | 167 | 21-030 | H2O | 1627770306 | 0.291 | 0.722 | 98.966 | 0 | 0.006 | 2.367 | 1297.465 | 2.277 | 2.763 | 108.098 | 0.871 | 0.038 | 0 | 0 | 1 | 44.25 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-17305 | 2021/07/31 16:28:55 | 1-41 | 3 | -11.22 | -113.699 | 18032 | -1 | 1 | 167 | 21-030 | H2O | 1627770842 | 0.325 | 0.813 | 100.225 | 0 | 0.004 | 2.387 | 1298.39 | 2.757 | 2.769 | 108.679 | 0.869 | 0.035 | 0 | 0 | 1 | 44.312 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-17305 | 2021/07/31 16:37:51 | 1-41 | 4 | -11.299 | -114.436 | 17907 | 0 | 1 | 167 | 21-030 | H2O | 1627771378 | 0.314 | 0.858 | 100.677 | 0 | 0.008 | 2.389 | 1297.184 | 2.376 | 2.862 | 110.386 | 0.873 | 0.47 | 0 | 0 | 1 | 44.312 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-17306 | 2021/07/31 16:47:36 | 1-42 | 1 | [5 empty] | 265 | 21-030 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-17306 | 2021/07/31 16:57:34 | 1-42 | 2 | [5 empty] | 265 | 21-030 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-17306 | 2021/07/31 17:07:32 | 1-42 | 3 | [5 empty] | 265 | 21-030 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-17306 | 2021/07/31 17:17:30 | 1-42 | 4 | [5 empty] | 265 | 21-030 | H2O | [17 empty] | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-17307 | 2021/07/31 17:28:16 | 1-43 | 1 | -8.039 | -67.292 | 17680 | -1 | 1 | EV | standard | H2O | 1627774403 | 0.324 | 0.716 | 89.913 | -0.001 | 0 | 2.135 | 1294.244 | 2.575 | 2.742 | 107.336 | 0.867 | 0.167 | 0 | 0 | 1 | 44.375 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-17307 | 2021/07/31 17:37:12 | 1-43 | 2 | -7.741 | -63.976 | 15894 | -1 | 1 | EV | standard | H2O | 1627774938 | 0.3 | 0.719 | 90.691 | 0 | -0.001 | 2.167 | 1281.655 | 2.425 | 2.492 | 97.44 | 0.867 | 0.279 | 0 | 0 | 1 | 44.25 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-17307 | 2021/07/31 17:46:07 | 1-43 | 3 | -8.058 | -65.493 | 12344 | -1 | 1 | EV | standard | H2O | 1627775474 | 0.336 | 0.799 | 81.829 | -0.001 | 0.001 | 2.021 | 1257.446 | 2.004 | 1.86 | 75.225 | 0.859 | 0.051 | 0 | 0 | 1 | 44.25 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-17307 | 2021/07/31 17:55:02 | 1-43 | 4 | -7.675 | -62.202 | 16168 | 0 | 1 | EV | standard | H2O | 1627776008 | 0.297 | 0.802 | 92.488 | 0 | 0.001 | 2.223 | 1283.518 | 2.432 | 2.462 | 98.206 | 0.866 | 0.178 | 0 | 0 | 1 | 44.188 | 1 | 43 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1300.7304
varray: [1302.892, 1303.547, 1302.983, 1300.027, 1300.835, 1302.708, 1301.723, 1301.025, 1300.142, 1301.477, 1298.725, 1301.282, 1300.901, 1301.655, 1300.85, 1299.701, 1299.389, 1299.208, 1297.65, 1297.888]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_shift
baseline_shift_STD | 1.65499640324
varray: [1302.892, 1303.547, 1302.983, 1300.027, 1300.835, 1302.708, 1301.723, 1301.025, 1300.142, 1301.477, 1298.725, 1301.282, 1300.901, 1301.655, 1300.85, 1299.701, 1299.389, 1299.208, 1297.65, 1297.888]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.3769
varray: [0.085, 0.449, 0.309, 0.683, 0.321, 0.438, 0.36, 0.22, 0.317, 0.262, 0.39, 0.16, 0.364, 0.407, 0.463, 0.353, 0.212, 0.59, 0.551, 0.604]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.15355880067
varray: [0.085, 0.449, 0.309, 0.683, 0.321, 0.438, 0.36, 0.22, 0.317, 0.262, 0.39, 0.16, 0.364, 0.407, 0.463, 0.353, 0.212, 0.59, 0.551, 0.604]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: ch4_ppm
residuals_MEAN | 2.9209
varray: [2.97, 2.956, 2.928, 2.911, 2.874, 2.97, 2.951, 2.947, 2.97, 2.992, 2.876, 2.901, 2.922, 2.928, 2.924, 2.903, 2.853, 2.921, 2.86, 2.861]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: residuals
residuals_STD | 0.0411389174179
varray: [2.97, 2.956, 2.928, 2.911, 2.874, 2.97, 2.951, 2.947, 2.97, 2.992, 2.876, 2.901, 2.922, 2.928, 2.924, 2.903, 2.853, 2.921, 2.86, 2.861]
inst: std_standards
operation: std_standards(residuals)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: residuals
slope_shift_MEAN | 2.48685
varray: [1.848, 1.682, 1.41, 1.869, 1.517, 1.956, 2.17, 1.847, 2.449, 2.018, 3.254, 3.347, 3.085, 3.172, 2.825, 3.024, 2.976, 3.111, 3.161, 3.016]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: slope_shift
slope_shift_STD | 0.667773147346
varray: [1.848, 1.682, 1.41, 1.869, 1.517, 1.956, 2.17, 1.847, 2.449, 2.018, 3.254, 3.347, 3.085, 3.172, 2.825, 3.024, 2.976, 3.111, 3.161, 3.016]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: slope_shift
baseline_curvature_MEAN | 112.60595
varray: [113.497, 113.877, 113.54, 112.573, 112.248, 114.218, 113.838, 112.672, 112.301, 113.572, 110.97, 111.351, 112.779, 113.125, 113.353, 111.797, 111.107, 112.8, 110.948, 111.553]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.03779997439
varray: [113.497, 113.877, 113.54, 112.573, 112.248, 114.218, 113.838, 112.672, 112.301, 113.572, 110.97, 111.351, 112.779, 113.125, 113.353, 111.797, 111.107, 112.8, 110.948, 111.553]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.9209
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1300.7304
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.3769
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 2.48685
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0617083761268
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.98260837613
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.0822778348358
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.00317783484
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.123416752254
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.04431675225
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.29209
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.46067640201
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.83757640201
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.82749820162
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1301.5578982
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1302.3853964
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 4.96498920972
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1305.69538921
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 29.7899352583
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1330.52033526
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 3.33886573673
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 5.82571573673
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 3.11339992317
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 115.719349923
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -7.75944730213
actual_array: [ 1.93 -15.84]
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602, 3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809, -11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239, -8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -7.912, -7.496, -7.499, -7.445, -4.465, -4.157, -4.155, -4.194, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -4.979, -4.444, -4.302, -4.329, 0.0, 0.0, 0.0, 0.0, -7.825, -7.593, -7.506, -7.568, 0.0, 0.0, -9.411, -9.411, -2.688, -2.12, -2.041, -2.0, -8.244, -8.761, -8.905, -8.889, -8.292, -8.223, -8.124, -8.166, -8.969, -9.064, -9.038, -9.16, -2.2, -1.901, -1.914, -1.906, -8.46, -8.738, -8.741, -8.659, 0.0, 0.0, -6.441, -6.498, -7.516, -7.418, -7.465, -7.652, -2.766, -2.642, -2.646, -2.658, -5.27, -5.409, -5.436, -5.397, -4.439, -4.402, -4.393, -4.335, -9.925, -9.983, -10.051, -10.084, -9.482, -9.48, -9.562, 0.0, 0.0, -9.691, 0.0, -9.623, 0.0, 0.0, 0.567, 0.56, 4.844, 5.041, 5.018, 5.019, 0.0, 0.0, -7.471, -7.425, -4.066, -3.751, -3.916, -3.808, 0.0, 0.0, -8.803, -8.839, -9.191, -9.225, -9.17, -9.325, -9.507, -9.459, -9.446, -9.433, -8.926, -8.913, -8.836, -8.897, 0.0, 0.0, -7.761, -7.824, -9.489, -9.562, -9.597, -9.546, 0.0, 0.0, 0.0, 0.0, -7.875, -7.677, -7.568, -7.583, -11.114, -11.219, -11.22, -11.299, 0.0, 0.0, 0.0, 0.0, -8.039, -7.741, -8.058, -7.675]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.99313265306
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0, 5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0, 16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0, 21734.0, 22332.0, 22932.0, 23530.0, 24182.0, 24781.0, 25380.0, 25979.0, 26631.0, 27168.0, 27705.0, 28242.0, 28828.0, 29364.0, 29901.0, 30437.0, 31024.0, 31622.0, 32221.0, 32820.0, 33470.0, 34069.0, 34668.0, 35266.0, 35917.0, 36454.0, 36989.0, 37526.0, 38115.0, 38713.0, 39311.0, 39910.0, 40562.0, 41098.0, 41634.0, 42171.0, 42760.0, 43359.0, 43957.0, 44555.0, 45142.0, 45678.0, 46214.0, 46751.0, 47337.0, 47874.0, 48410.0, 48947.0, 49534.0, 50071.0, 50608.0, 51144.0, 51732.0, 52269.0, 52805.0, 53341.0, 53929.0, 54465.0, 55002.0, 55539.0, 56127.0, 56663.0, 57199.0, 57735.0, 58324.0, 58861.0, 59398.0, 59934.0, 60519.0, 61056.0, 61592.0, 62129.0, 62715.0, 63251.0, 63788.0, 64324.0, 64910.0, 65447.0, 65983.0, 66518.0, 67104.0, 67640.0, 68176.0, 68712.0, 69299.0, 69835.0, 70371.0, 70907.0, 71495.0, 72031.0, 72567.0, 73104.0, 73753.0, 74289.0, 74825.0, 75362.0, 75946.0, 76482.0, 77018.0, 77555.0, 78140.0, 78676.0, 79212.0, 79748.0, 80334.0, 80869.0, 81406.0, 81942.0, 82527.0, 83064.0, 83600.0, 84136.0, 84723.0, 85258.0, 85794.0, 86331.0, 86917.0, 87453.0, 87989.0, 88525.0, 89112.0, 89649.0, 90184.0, 90720.0, 91302.0, 91837.0, 92373.0, 92909.0, 93493.0, 94029.0, 94564.0, 95100.0, 95682.0, 96217.0, 96753.0, 97289.0, 97873.0, 98471.0, 99069.0, 99667.0, 100314.0, 100849.0, 101384.0, 101919.0, 102503.0, 103040.0, 103575.0, 104111.0, 104696.0, 105294.0, 105892.0, 106490.0, 107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -67.1459694458
actual_array: [ 18.1 -119. ]
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777, 23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388, -87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391, -68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -84.773, -84.147, -84.132, -84.248, -66.958, -64.358, -64.476, -64.197, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -82.5, -81.68, -81.613, -82.053, 0.0, 0.0, 0.0, 0.0, -66.519, -62.195, -61.01, -60.869, 0.0, 0.0, -102.534, -102.534, -36.276, -27.551, -25.677, -25.315, -75.114, -81.622, -83.367, -83.537, -84.775, -84.97, -84.704, -84.833, -56.85, -53.542, -52.365, -52.587, -54.561, -55.326, -55.773, -55.597, -74.944, -77.057, -77.458, -77.208, 0.0, 0.0, -72.554, -72.672, -62.562, -60.564, -60.554, -60.924, -52.983, -52.473, -52.739, -52.465, -70.043, -72.167, -72.289, -72.661, -68.734, -68.876, -68.423, -68.177, -104.817, -108.365, -109.45, -109.88, -97.491, -95.858, -96.203, 0.0, 0.0, -93.674, 0.0, -93.543, 0.0, 0.0, -62.354, -62.118, -65.401, -65.695, -65.906, -66.017, 0.0, 0.0, -61.027, -60.718, -68.891, -68.355, -69.295, -69.018, 0.0, 0.0, -85.091, -85.259, -89.124, -89.691, -89.676, -90.427, -90.767, -90.674, -90.858, -90.694, -87.89, -87.771, -87.413, -87.727, 0.0, 0.0, -67.566, -67.851, -87.803, -90.783, -91.204, -91.178, 0.0, 0.0, 0.0, 0.0, -64.747, -62.211, -61.363, -61.143, -107.072, -112.149, -113.699, -114.436, 0.0, 0.0, 0.0, 0.0, -67.292, -63.976, -65.493, -62.202]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.7312091837
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0, 5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0, 10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0, 16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0, 21734.0, 22332.0, 22932.0, 23530.0, 24182.0, 24781.0, 25380.0, 25979.0, 26631.0, 27168.0, 27705.0, 28242.0, 28828.0, 29364.0, 29901.0, 30437.0, 31024.0, 31622.0, 32221.0, 32820.0, 33470.0, 34069.0, 34668.0, 35266.0, 35917.0, 36454.0, 36989.0, 37526.0, 38115.0, 38713.0, 39311.0, 39910.0, 40562.0, 41098.0, 41634.0, 42171.0, 42760.0, 43359.0, 43957.0, 44555.0, 45142.0, 45678.0, 46214.0, 46751.0, 47337.0, 47874.0, 48410.0, 48947.0, 49534.0, 50071.0, 50608.0, 51144.0, 51732.0, 52269.0, 52805.0, 53341.0, 53929.0, 54465.0, 55002.0, 55539.0, 56127.0, 56663.0, 57199.0, 57735.0, 58324.0, 58861.0, 59398.0, 59934.0, 60519.0, 61056.0, 61592.0, 62129.0, 62715.0, 63251.0, 63788.0, 64324.0, 64910.0, 65447.0, 65983.0, 66518.0, 67104.0, 67640.0, 68176.0, 68712.0, 69299.0, 69835.0, 70371.0, 70907.0, 71495.0, 72031.0, 72567.0, 73104.0, 73753.0, 74289.0, 74825.0, 75362.0, 75946.0, 76482.0, 77018.0, 77555.0, 78140.0, 78676.0, 79212.0, 79748.0, 80334.0, 80869.0, 81406.0, 81942.0, 82527.0, 83064.0, 83600.0, 84136.0, 84723.0, 85258.0, 85794.0, 86331.0, 86917.0, 87453.0, 87989.0, 88525.0, 89112.0, 89649.0, 90184.0, 90720.0, 91302.0, 91837.0, 92373.0, 92909.0, 93493.0, 94029.0, 94564.0, 95100.0, 95682.0, 96217.0, 96753.0, 97289.0, 97873.0, 98471.0, 99069.0, 99667.0, 100314.0, 100849.0, 101384.0, 101919.0, 102503.0, 103040.0, 103575.0, 104111.0, 104696.0, 105294.0, 105892.0, 106490.0, 107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6776
varray: [3.164, 3.023, 2.971, 0.0, 2.947, 2.963, 2.897, 2.997, 2.899, 2.915]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: residuals
SAMP_residuals_STD | 0.944049810362
varray: [3.164, 3.023, 2.971, 0.0, 2.947, 2.963, 2.897, 2.997, 2.899, 2.915]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2575
varray: [-0.011, 0.337, 0.333, 0.0, 0.534, 0.278, 0.289, 0.399, 0.068, 0.348]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.180248747877
varray: [-0.011, 0.337, 0.333, 0.0, 0.534, 0.278, 0.289, 0.399, 0.068, 0.348]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1187.8149
varray: [1350.554, 1326.736, 1321.295, 0.0, 1325.085, 1314.472, 1311.73, 1311.704, 1308.902, 1307.671]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 417.548452729
varray: [1350.554, 1326.736, 1321.295, 0.0, 1325.085, 1314.472, 1311.73, 1311.704, 1308.902, 1307.671]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.4149
varray: [1.068, 1.129, 1.706, 0.0, 1.189, 1.888, 1.561, 2.023, 1.946, 1.639]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.604339207546
varray: [1.068, 1.129, 1.706, 0.0, 1.189, 1.888, 1.561, 2.023, 1.946, 1.639]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: slope_shift
SAMP_d18O_STD | 2.45417831689
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 20.0103642412
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.9966
varray: [121.858, 114.888, 114.557, 0.0, 113.578, 113.042, 112.398, 114.596, 111.481, 113.568]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 36.2991403241
varray: [121.858, 114.888, 114.557, 0.0, 113.578, 113.042, 112.398, 114.596, 111.481, 113.568]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6776
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1187.8149
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.4149
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2433
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2433
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.32962443083
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.60935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.9499
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
dD_mean | -56.3362
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
d18O_stdev | 2.45417831689
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
dD_stdev | 20.0103642412
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.9499
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.3362
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.45417831689
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
measured_dD_stdev | 20.0103642412
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.74588262514
actual_array: [ 1.93 -15.84]
varray: [-8.36, -7.801, -7.664, 0.0, -7.765, -7.591, -7.691, -7.523, -7.502, -7.602]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -6.9499
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.5746995274
actual_array: [ 18.1 -119. ]
varray: [-70.41, -63.79, -62.35, 0.0, -62.198, -61.15, -61.474, -60.455, -60.758, -60.777]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -56.3362
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [1.0, 537.0, 1073.0, 1609.0, 2208.0, 2745.0, 3281.0, 3818.0, 4354.0, 4890.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.0528
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -112.9155
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9469
varray: [2.97, 2.956, 2.928, 2.911, 2.874, 2.97, 2.951, 2.947, 2.97, 2.992]
inst: mean_all
operation: mean_all(residuals)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: residuals
SAMP_residuals_STD | 0.0344720498697
varray: [2.97, 2.956, 2.928, 2.911, 2.874, 2.97, 2.951, 2.947, 2.97, 2.992]
inst: std_all
operation: std_all(residuals)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3444
varray: [0.085, 0.449, 0.309, 0.683, 0.321, 0.438, 0.36, 0.22, 0.317, 0.262]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.158710218113
varray: [0.085, 0.449, 0.309, 0.683, 0.321, 0.438, 0.36, 0.22, 0.317, 0.262]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.7359
varray: [1302.892, 1303.547, 1302.983, 1300.027, 1300.835, 1302.708, 1301.723, 1301.025, 1300.142, 1301.477]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.24552259893
varray: [1302.892, 1303.547, 1302.983, 1300.027, 1300.835, 1302.708, 1301.723, 1301.025, 1300.142, 1301.477]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.8766
varray: [1.848, 1.682, 1.41, 1.869, 1.517, 1.956, 2.17, 1.847, 2.449, 2.018]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.303071755347
varray: [1.848, 1.682, 1.41, 1.869, 1.517, 1.956, 2.17, 1.847, 2.449, 2.018]
inst: std_all
operation: std_all(slope_shift)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: slope_shift
SAMP_d18O_STD | 0.379022192256
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.92896723462
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.2336
varray: [113.497, 113.877, 113.54, 112.573, 112.248, 114.218, 113.838, 112.672, 112.301, 113.572]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.716396724355
varray: [113.497, 113.877, 113.54, 112.573, 112.248, 114.218, 113.838, 112.672, 112.301, 113.572]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9469
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3444
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.7359
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.8766
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.026
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.026
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.890136601732
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.62765
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.5516
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 36.897
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.5516
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
dD_mean | 36.897
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
d18O_stdev | 0.379022192256
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
dD_stdev | 4.92896723462
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.5516
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
measured_dD_mean | 36.897
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.379022192256
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.92896723462
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.561, 4.312, 4.563, 4.574, 4.665, 4.713, 4.751, 4.802, 4.766, 4.809]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.5516
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [23.967, 33.771, 36.931, 37.624, 38.399, 39.231, 39.293, 39.609, 39.757, 40.388]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 36.897
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [5478.0, 6014.0, 6551.0, 7087.0, 7624.0, 8161.0, 8697.0, 9233.0, 9770.0, 10307.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.7835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.0055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8949
varray: [2.876, 2.901, 2.922, 2.928, 2.924, 2.903, 2.853, 2.921, 2.86, 2.861]
inst: mean_all
operation: mean_all(residuals)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: residuals
SAMP_residuals_STD | 0.0297039092077
varray: [2.876, 2.901, 2.922, 2.928, 2.924, 2.903, 2.853, 2.921, 2.86, 2.861]
inst: std_all
operation: std_all(residuals)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4094
varray: [0.39, 0.16, 0.364, 0.407, 0.463, 0.353, 0.212, 0.59, 0.551, 0.604]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149145864471
varray: [0.39, 0.16, 0.364, 0.407, 0.463, 0.353, 0.212, 0.59, 0.551, 0.604]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.7249
varray: [1298.725, 1301.282, 1300.901, 1301.655, 1300.85, 1299.701, 1299.389, 1299.208, 1297.65, 1297.888]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.408652031
varray: [1298.725, 1301.282, 1300.901, 1301.655, 1300.85, 1299.701, 1299.389, 1299.208, 1297.65, 1297.888]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 3.0971
varray: [3.254, 3.347, 3.085, 3.172, 2.825, 3.024, 2.976, 3.111, 3.161, 3.016]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.148220893714
varray: [3.254, 3.347, 3.085, 3.172, 2.825, 3.024, 2.976, 3.111, 3.161, 3.016]
inst: std_all
operation: std_all(slope_shift)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: slope_shift
SAMP_d18O_STD | 0.539518303675
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 7.41914393759
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.9783
varray: [110.97, 111.351, 112.779, 113.125, 113.353, 111.797, 111.107, 112.8, 110.948, 111.553]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.940782068281
varray: [110.97, 111.351, 112.779, 113.125, 113.353, 111.797, 111.107, 112.8, 110.948, 111.553]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8949
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4094
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.7249
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.0971
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.026
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.026
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.890136601732
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.62765
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.953
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
dD_mean | -107.2509
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
d18O_stdev | 0.539518303675
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
dD_stdev | 7.41914393759
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.953
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.2509
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.539518303675
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
measured_dD_stdev | 7.41914393759
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-11.497, -12.736, -12.946, -13.032, -13.169, -13.153, -13.204, -13.227, -13.327, -13.239]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -12.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-87.305, -103.58, -107.434, -108.894, -109.909, -110.406, -110.832, -111.006, -111.752, -111.391]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -107.2509
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [10895.0, 11432.0, 11969.0, 12505.0, 13042.0, 13579.0, 14115.0, 14652.0, 15189.0, 15725.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.8355
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8914
varray: [2.945, 2.908, 2.963, 2.91, 2.844, 2.878, 2.824, 2.911, 2.895, 2.836]
inst: mean_all
operation: mean_all(residuals)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: residuals
SAMP_residuals_STD | 0.0460149734084
varray: [2.945, 2.908, 2.963, 2.91, 2.844, 2.878, 2.824, 2.911, 2.895, 2.836]
inst: std_all
operation: std_all(residuals)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4173
varray: [0.504, 0.405, 0.692, 0.382, 0.273, 0.299, 0.215, 0.553, 0.422, 0.428]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.140897165022
varray: [0.504, 0.405, 0.692, 0.382, 0.273, 0.299, 0.215, 0.553, 0.422, 0.428]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.302
varray: [1299.645, 1300.072, 1298.335, 1298.41, 1298.974, 1298.107, 1297.796, 1297.696, 1297.241, 1296.744]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.03361652034
varray: [1299.645, 1300.072, 1298.335, 1298.41, 1298.974, 1298.107, 1297.796, 1297.696, 1297.241, 1296.744]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.7177
varray: [2.292, 2.73, 2.989, 3.17, 2.977, 2.953, 2.661, 2.157, 2.525, 2.723]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.322647743935
varray: [2.292, 2.73, 2.989, 3.17, 2.977, 2.953, 2.661, 2.157, 2.525, 2.723]
inst: std_all
operation: std_all(slope_shift)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: slope_shift
SAMP_d18O_STD | 0.162345174106
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.44478316466
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.5528
varray: [113.698, 113.151, 112.86, 111.024, 110.334, 110.293, 109.864, 112.944, 111.178, 110.182]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.4538893279
varray: [113.698, 113.151, 112.86, 111.024, 110.334, 110.293, 109.864, 112.944, 111.178, 110.182]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8914
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4173
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.302
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.7177
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0295
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0295
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.00996268273
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.05315
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6378
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
dD_mean | -61.5445
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
d18O_stdev | 0.162345174106
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
dD_stdev | 2.44478316466
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6378
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.5445
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.162345174106
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.44478316466
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4442123785
actual_array: [ 1.93 -15.84]
varray: [-8.081, -7.675, -7.617, -7.619, -7.532, -7.575, -7.578, -7.521, -7.619, -7.561]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.6378
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.5283473432
actual_array: [ 18.1 -119. ]
varray: [-68.241, -62.388, -61.335, -61.15, -60.435, -60.318, -60.36, -60.384, -60.343, -60.491]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.5445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [16315.0, 16851.0, 17388.0, 17925.0, 18461.0, 18998.0, 19534.0, 20071.0, 20608.0, 21144.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.839
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.4284
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [21734.0, 22332.0, 22932.0, 23530.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [24182.0, 24781.0, 25380.0, 25979.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7445
varray: [2.803, 2.703, 2.768, 2.704]
inst: mean_all
operation: mean_all(residuals)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: residuals
SAMP_residuals_STD | 0.0494536820335
varray: [2.803, 2.703, 2.768, 2.704]
inst: std_all
operation: std_all(residuals)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2565
varray: [0.475, 0.132, 0.225, 0.194]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.150710539335
varray: [0.475, 0.132, 0.225, 0.194]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.99625
varray: [1293.773, 1294.332, 1294.314, 1293.566]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.38671641203
varray: [1293.773, 1294.332, 1294.314, 1293.566]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.51825
varray: [2.501, 2.492, 2.76, 2.32]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.181413661742
varray: [2.501, 2.492, 2.76, 2.32]
inst: std_all
operation: std_all(slope_shift)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: slope_shift
SAMP_d18O_STD | 0.217416650696
varray: [-7.912, -7.496, -7.499, -7.445]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.303076448002
varray: [-84.773, -84.147, -84.132, -84.248]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.997
varray: [107.943, 106.577, 106.9, 106.568]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.649298595922
varray: [107.943, 106.577, 106.9, 106.568]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7445
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.99625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.51825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1764
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1764
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.03923448252
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.60895
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-7.912, -7.496, -7.499, -7.445]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-84.773, -84.147, -84.132, -84.248]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.588
varray: [-7.912, -7.496, -7.499, -7.445]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
dD_mean | -84.325
varray: [-84.773, -84.147, -84.132, -84.248]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
d18O_stdev | 0.217416650696
varray: [-7.912, -7.496, -7.499, -7.445]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
dD_stdev | 0.303076448002
varray: [-84.773, -84.147, -84.132, -84.248]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.588
varray: [-7.912, -7.496, -7.499, -7.445]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
measured_dD_mean | -84.325
varray: [-84.773, -84.147, -84.132, -84.248]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.217416650696
varray: [-7.912, -7.496, -7.499, -7.445]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.303076448002
varray: [-84.773, -84.147, -84.132, -84.248]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3936573244
actual_array: [ 1.93 -15.84]
varray: [-7.912, -7.496, -7.499, -7.445]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -97.1950282314
actual_array: [ 18.1 -119. ]
varray: [-84.773, -84.147, -84.132, -84.248]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [26631.0, 27168.0, 27705.0, 28242.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.9859
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.73415
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.59575
varray: [2.637, 2.664, 2.51, 2.572]
inst: mean_all
operation: mean_all(residuals)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: residuals
SAMP_residuals_STD | 0.0689849017298
varray: [2.637, 2.664, 2.51, 2.572]
inst: std_all
operation: std_all(residuals)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.13675
varray: [0.104, 0.087, 0.208, 0.148]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0540084869874
varray: [0.104, 0.087, 0.208, 0.148]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.14075
varray: [1289.294, 1290.86, 1282.225, 1286.184]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.81004263974
varray: [1289.294, 1290.86, 1282.225, 1286.184]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.69
varray: [3.099, 3.051, 2.173, 2.437]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.457857328579
varray: [3.099, 3.051, 2.173, 2.437]
inst: std_all
operation: std_all(slope_shift)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: slope_shift
SAMP_d18O_STD | 0.149247836389
varray: [-4.465, -4.157, -4.155, -4.194]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.31215887631
varray: [-66.958, -64.358, -64.476, -64.197]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 101.0625
varray: [102.532, 103.631, 97.006, 101.081]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.89897435886
varray: [102.532, 103.631, 97.006, 101.081]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.59575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.13675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.14075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.69
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.32515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.32515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.1318429251
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -11.54345
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-4.465, -4.157, -4.155, -4.194]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-66.958, -64.358, -64.476, -64.197]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.24275
varray: [-4.465, -4.157, -4.155, -4.194]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
dD_mean | -64.99725
varray: [-66.958, -64.358, -64.476, -64.197]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
d18O_stdev | 0.149247836389
varray: [-4.465, -4.157, -4.155, -4.194]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
dD_stdev | 1.31215887631
varray: [-66.958, -64.358, -64.476, -64.197]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.24275
varray: [-4.465, -4.157, -4.155, -4.194]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.99725
varray: [-66.958, -64.358, -64.476, -64.197]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.149247836389
varray: [-4.465, -4.157, -4.155, -4.194]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.31215887631
varray: [-66.958, -64.358, -64.476, -64.197]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.9976875507
actual_array: [ 1.93 -15.84]
varray: [-4.465, -4.157, -4.155, -4.194]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.24275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.8122801997
actual_array: [ 18.1 -119. ]
varray: [-66.958, -64.358, -64.476, -64.197]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -64.99725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [28828.0, 29364.0, 29901.0, 30437.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.13465
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -13.58965
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [31024.0, 31622.0, 32221.0, 32820.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [33470.0, 34069.0, 34668.0, 35266.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.806
varray: [2.54, 2.884, 2.897, 2.903]
inst: mean_all
operation: mean_all(residuals)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: residuals
SAMP_residuals_STD | 0.177510563066
varray: [2.54, 2.884, 2.897, 2.903]
inst: std_all
operation: std_all(residuals)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2865
varray: [0.092, 0.392, 0.287, 0.375]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137589970565
varray: [0.092, 0.392, 0.287, 0.375]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.369
varray: [1287.615, 1297.98, 1301.587, 1298.294]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.059730082
varray: [1287.615, 1297.98, 1301.587, 1298.294]
inst: std_all
operation: std_all(baseline_shift)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.676
varray: [2.803, 2.576, 2.62, 2.705]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.100176510886
varray: [2.803, 2.576, 2.62, 2.705]
inst: std_all
operation: std_all(slope_shift)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: slope_shift
SAMP_d18O_STD | 0.316382152889
varray: [-4.979, -4.444, -4.302, -4.329]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.407859861554
varray: [-82.5, -81.68, -81.613, -82.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.20125
varray: [100.63, 111.123, 113.625, 111.427]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.82188333646
varray: [100.63, 111.123, 113.625, 111.427]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.806
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2865
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.369
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.676
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1149
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1149
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.93371905919
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4047
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-4.979, -4.444, -4.302, -4.329]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-82.5, -81.68, -81.613, -82.053]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.5135
varray: [-4.979, -4.444, -4.302, -4.329]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
dD_mean | -81.9615
varray: [-82.5, -81.68, -81.613, -82.053]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
d18O_stdev | 0.316382152889
varray: [-4.979, -4.444, -4.302, -4.329]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
dD_stdev | 0.407859861554
varray: [-82.5, -81.68, -81.613, -82.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.5135
varray: [-4.979, -4.444, -4.302, -4.329]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
measured_dD_mean | -81.9615
varray: [-82.5, -81.68, -81.613, -82.053]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.316382152889
varray: [-4.979, -4.444, -4.302, -4.329]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.407859861554
varray: [-82.5, -81.68, -81.613, -82.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.27254258881
actual_array: [ 1.93 -15.84]
varray: [-4.979, -4.444, -4.302, -4.329]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.5135
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -94.9470880949
actual_array: [ 18.1 -119. ]
varray: [-82.5, -81.68, -81.613, -82.053]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -81.9615
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [35917.0, 36454.0, 36989.0, 37526.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.9244
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.3614
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [38115.0, 38713.0, 39311.0, 39910.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86675
varray: [2.859, 2.882, 2.864, 2.862]
inst: mean_all
operation: mean_all(residuals)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: residuals
SAMP_residuals_STD | 0.0103722385883
varray: [2.859, 2.882, 2.864, 2.862]
inst: std_all
operation: std_all(residuals)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.38225
varray: [0.408, 0.46, 0.234, 0.427]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101141402666
varray: [0.408, 0.46, 0.234, 0.427]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.20925
varray: [1297.088, 1297.374, 1297.256, 1297.119]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.131877152431
varray: [1297.088, 1297.374, 1297.256, 1297.119]
inst: std_all
operation: std_all(baseline_shift)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.80325
varray: [2.76, 2.817, 3.084, 2.552]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.219094157841
varray: [2.76, 2.817, 3.084, 2.552]
inst: std_all
operation: std_all(slope_shift)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: slope_shift
SAMP_d18O_STD | 0.139544497085
varray: [-7.825, -7.593, -7.506, -7.568]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.64812731001
varray: [-66.519, -62.195, -61.01, -60.869]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.66975
varray: [110.597, 110.859, 109.843, 111.38]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.64007258703
varray: [110.597, 110.859, 109.843, 111.38]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.38225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.20925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.80325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05415
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05415
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.85388065322
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.9362
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-7.825, -7.593, -7.506, -7.568]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-66.519, -62.195, -61.01, -60.869]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.623
varray: [-7.825, -7.593, -7.506, -7.568]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
dD_mean | -62.64825
varray: [-66.519, -62.195, -61.01, -60.869]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
d18O_stdev | 0.139544497085
varray: [-7.825, -7.593, -7.506, -7.568]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
dD_stdev | 2.64812731001
varray: [-66.519, -62.195, -61.01, -60.869]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.623
varray: [-7.825, -7.593, -7.506, -7.568]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.64825
varray: [-66.519, -62.195, -61.01, -60.869]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.139544497085
varray: [-7.825, -7.593, -7.506, -7.568]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.64812731001
varray: [-66.519, -62.195, -61.01, -60.869]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4291879849
actual_array: [ 1.93 -15.84]
varray: [-7.825, -7.593, -7.506, -7.568]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.623
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.578131107
actual_array: [ 18.1 -119. ]
varray: [-66.519, -62.195, -61.01, -60.869]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.64825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [40562.0, 41098.0, 41634.0, 42171.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.86365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.52115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.36
varray: [0.0, 0.0, 2.72, 2.72]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: residuals
SAMP_residuals_STD | 1.5703927322
varray: [0.0, 0.0, 2.72, 2.72]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0165
varray: [0.0, 0.0, 0.033, 0.033]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0190525588833
varray: [0.0, 0.0, 0.033, 0.033]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 647.746
varray: [0.0, 0.0, 1295.492, 1295.492]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 747.952654933
varray: [0.0, 0.0, 1295.492, 1295.492]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.026
varray: [0.0, 0.0, 2.052, 2.052]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.18472275238
varray: [0.0, 0.0, 2.052, 2.052]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: slope_shift
SAMP_d18O_STD | 5.43344338334
varray: [0.0, 0.0, -9.411, -9.411]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 59.1980325011
varray: [0.0, 0.0, -102.534, -102.534]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 53.4585
varray: [0.0, 0.0, 106.917, 106.917]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 61.7285587309
varray: [0.0, 0.0, 106.917, 106.917]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.36
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0165
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 647.746
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.026
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.5609
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.5609
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -53.4390085248
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -59.14745
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, -9.411, -9.411]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, -102.534, -102.534]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.7055
varray: [0.0, 0.0, -9.411, -9.411]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
dD_mean | -51.267
varray: [0.0, 0.0, -102.534, -102.534]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
d18O_stdev | 5.43344338334
varray: [0.0, 0.0, -9.411, -9.411]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
dD_stdev | 59.1980325011
varray: [0.0, 0.0, -102.534, -102.534]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.7055
varray: [0.0, 0.0, -9.411, -9.411]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.267
varray: [0.0, 0.0, -102.534, -102.534]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.43344338334
varray: [0.0, 0.0, -9.411, -9.411]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
measured_dD_stdev | 59.1980325011
varray: [0.0, 0.0, -102.534, -102.534]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.46745364076
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -9.411, -9.411]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.7055
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.7533506211
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -102.534, -102.534]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -51.267
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [42760.0, 43359.0, 43957.0, 44555.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3704
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -652.9844
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86475
varray: [2.901, 2.826, 2.896, 2.836]
inst: mean_all
operation: mean_all(residuals)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: residuals
SAMP_residuals_STD | 0.0392375245566
varray: [2.901, 2.826, 2.896, 2.836]
inst: std_all
operation: std_all(residuals)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.377
varray: [0.617, 0.149, 0.284, 0.458]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.203955877581
varray: [0.617, 0.149, 0.284, 0.458]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.654
varray: [1295.882, 1297.145, 1297.26, 1296.329]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.660788922425
varray: [1295.882, 1297.145, 1297.26, 1296.329]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.51575
varray: [2.404, 2.5, 2.425, 2.734]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.151222518164
varray: [2.404, 2.5, 2.425, 2.734]
inst: std_all
operation: std_all(slope_shift)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: slope_shift
SAMP_d18O_STD | 0.321052825352
varray: [-2.688, -2.12, -2.041, -2.0]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.14174551263
varray: [-36.276, -27.551, -25.677, -25.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.514
varray: [111.15, 109.318, 110.888, 110.7]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.81840984028
varray: [111.15, 109.318, 110.888, 110.7]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.377
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.654
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.51575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05615
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05615
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.92235269951
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.09195
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-2.688, -2.12, -2.041, -2.0]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-36.276, -27.551, -25.677, -25.315]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.21225
varray: [-2.688, -2.12, -2.041, -2.0]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
dD_mean | -28.70475
varray: [-36.276, -27.551, -25.677, -25.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
d18O_stdev | 0.321052825352
varray: [-2.688, -2.12, -2.041, -2.0]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
dD_stdev | 5.14174551263
varray: [-36.276, -27.551, -25.677, -25.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.21225
varray: [-2.688, -2.12, -2.041, -2.0]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
measured_dD_mean | -28.70475
varray: [-36.276, -27.551, -25.677, -25.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.321052825352
varray: [-2.688, -2.12, -2.041, -2.0]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.14174551263
varray: [-36.276, -27.551, -25.677, -25.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.93640166014
actual_array: [ 1.93 -15.84]
varray: [-2.688, -2.12, -2.041, -2.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.21225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -44.2942487195
actual_array: [ 18.1 -119. ]
varray: [-36.276, -27.551, -25.677, -25.315]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -28.70475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [45142.0, 45678.0, 46214.0, 46751.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.86565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.0764
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8685
varray: [2.914, 2.878, 2.809, 2.873]
inst: mean_all
operation: mean_all(residuals)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: residuals
SAMP_residuals_STD | 0.0436692111218
varray: [2.914, 2.878, 2.809, 2.873]
inst: std_all
operation: std_all(residuals)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3615
varray: [0.419, 0.381, 0.372, 0.274]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0617872694547
varray: [0.419, 0.381, 0.372, 0.274]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.427
varray: [1298.409, 1297.734, 1295.372, 1298.193]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.39861288425
varray: [1298.409, 1297.734, 1295.372, 1298.193]
inst: std_all
operation: std_all(baseline_shift)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.536
varray: [2.442, 2.714, 2.661, 2.327]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.182415277138
varray: [2.442, 2.714, 2.661, 2.327]
inst: std_all
operation: std_all(slope_shift)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: slope_shift
SAMP_d18O_STD | 0.310592310486
varray: [-8.244, -8.761, -8.905, -8.889]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.95973643904
varray: [-75.114, -81.622, -83.367, -83.537]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.95975
varray: [112.137, 111.328, 109.046, 111.328]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.33161214949
varray: [112.137, 111.328, 109.046, 111.328]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8685
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3615
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.427
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.536
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0524
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0524
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.79396761272
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.6462
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.244, -8.761, -8.905, -8.889]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-75.114, -81.622, -83.367, -83.537]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.69975
varray: [-8.244, -8.761, -8.905, -8.889]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
dD_mean | -80.91
varray: [-75.114, -81.622, -83.367, -83.537]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
d18O_stdev | 0.310592310486
varray: [-8.244, -8.761, -8.905, -8.889]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
dD_stdev | 3.95973643904
varray: [-75.114, -81.622, -83.367, -83.537]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.69975
varray: [-8.244, -8.761, -8.905, -8.889]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
measured_dD_mean | -80.91
varray: [-75.114, -81.622, -83.367, -83.537]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.310592310486
varray: [-8.244, -8.761, -8.905, -8.889]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.95973643904
varray: [-75.114, -81.622, -83.367, -83.537]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.5222633765
actual_array: [ 1.93 -15.84]
varray: [-8.244, -8.761, -8.905, -8.889]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.69975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -93.9469996441
actual_array: [ 18.1 -119. ]
varray: [-75.114, -81.622, -83.367, -83.537]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -80.91
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [47337.0, 47874.0, 48410.0, 48947.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.8619
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.3034
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.83375
varray: [2.896, 2.84, 2.785, 2.814]
inst: mean_all
operation: mean_all(residuals)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: residuals
SAMP_residuals_STD | 0.0471902179129
varray: [2.896, 2.84, 2.785, 2.814]
inst: std_all
operation: std_all(residuals)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34275
varray: [0.564, 0.327, 0.205, 0.275]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.155739900689
varray: [0.564, 0.327, 0.205, 0.275]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.04325
varray: [1297.274, 1296.25, 1297.115, 1297.534]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.556319677284
varray: [1297.274, 1296.25, 1297.115, 1297.534]
inst: std_all
operation: std_all(baseline_shift)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.34525
varray: [2.059, 2.031, 2.99, 2.301]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.446599279743
varray: [2.059, 2.031, 2.99, 2.301]
inst: std_all
operation: std_all(slope_shift)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0728440114217
varray: [-8.292, -8.223, -8.124, -8.166]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.112766720859
varray: [-84.775, -84.97, -84.704, -84.833]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.81625
varray: [112.914, 111.034, 108.742, 110.575]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.71354260233
varray: [112.914, 111.034, 108.742, 110.575]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.83375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.04325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.34525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08715
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08715
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.98366941696
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7897
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.292, -8.223, -8.124, -8.166]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-84.775, -84.97, -84.704, -84.833]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.20125
varray: [-8.292, -8.223, -8.124, -8.166]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
dD_mean | -84.8205
varray: [-84.775, -84.97, -84.704, -84.833]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0728440114217
varray: [-8.292, -8.223, -8.124, -8.166]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
dD_stdev | 0.112766720859
varray: [-84.775, -84.97, -84.704, -84.833]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.20125
varray: [-8.292, -8.223, -8.124, -8.166]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
measured_dD_mean | -84.8205
varray: [-84.775, -84.97, -84.704, -84.833]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0728440114217
varray: [-8.292, -8.223, -8.124, -8.166]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.112766720859
varray: [-84.775, -84.97, -84.704, -84.833]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.0162052546
actual_array: [ 1.93 -15.84]
varray: [-8.292, -8.223, -8.124, -8.166]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.20125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -97.6663014862
actual_array: [ 18.1 -119. ]
varray: [-84.775, -84.97, -84.704, -84.833]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.8205
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [49534.0, 50071.0, 50608.0, 51144.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.89665
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.68715
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.818
varray: [2.935, 2.833, 2.775, 2.729]
inst: mean_all
operation: mean_all(residuals)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: residuals
SAMP_residuals_STD | 0.0888519367637
varray: [2.935, 2.833, 2.775, 2.729]
inst: std_all
operation: std_all(residuals)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28925
varray: [0.59, 0.437, 0.01, 0.12]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.270127593309
varray: [0.59, 0.437, 0.01, 0.12]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.1695
varray: [1297.437, 1295.075, 1297.924, 1294.242]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.78866924462
varray: [1297.437, 1295.075, 1297.924, 1294.242]
inst: std_all
operation: std_all(baseline_shift)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.30275
varray: [2.358, 2.133, 2.389, 2.331]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.115621148585
varray: [2.358, 2.133, 2.389, 2.331]
inst: std_all
operation: std_all(slope_shift)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0790796012467
varray: [-8.969, -9.064, -9.038, -9.16]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.07319994212
varray: [-56.85, -53.542, -52.365, -52.587]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.70425
varray: [112.158, 109.494, 109.368, 107.797]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.80884500442
varray: [112.158, 109.494, 109.368, 107.797]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.818
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.1695
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.30275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1029
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1029
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.52288678147
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.9017
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.969, -9.064, -9.038, -9.16]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-56.85, -53.542, -52.365, -52.587]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.05775
varray: [-8.969, -9.064, -9.038, -9.16]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
dD_mean | -53.836
varray: [-56.85, -53.542, -52.365, -52.587]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0790796012467
varray: [-8.969, -9.064, -9.038, -9.16]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
dD_stdev | 2.07319994212
varray: [-56.85, -53.542, -52.365, -52.587]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.05775
varray: [-8.969, -9.064, -9.038, -9.16]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.836
varray: [-56.85, -53.542, -52.365, -52.587]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0790796012467
varray: [-8.969, -9.064, -9.038, -9.16]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.07319994212
varray: [-56.85, -53.542, -52.365, -52.587]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.8856912754
actual_array: [ 1.93 -15.84]
varray: [-8.969, -9.064, -9.038, -9.16]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.05775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.1967431367
actual_array: [ 18.1 -119. ]
varray: [-56.85, -53.542, -52.365, -52.587]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -53.836
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [51732.0, 52269.0, 52805.0, 53341.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.9124
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.5609
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.00625
varray: [3.02, 3.042, 2.945, 3.018]
inst: mean_all
operation: mean_all(residuals)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: residuals
SAMP_residuals_STD | 0.0422561632586
varray: [3.02, 3.042, 2.945, 3.018]
inst: std_all
operation: std_all(residuals)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.7275
varray: [-3.846, -3.704, -3.877, -3.483]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.179559646543
varray: [-3.846, -3.704, -3.877, -3.483]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2079.3755
varray: [2076.673, 2082.088, 2078.85, 2079.891]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.25115177928
varray: [2076.673, 2082.088, 2078.85, 2079.891]
inst: std_all
operation: std_all(baseline_shift)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -35.15725
varray: [-34.66, -35.503, -35.249, -35.217]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.355334935894
varray: [-34.66, -35.503, -35.249, -35.217]
inst: std_all
operation: std_all(slope_shift)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: slope_shift
SAMP_d18O_STD | 0.146597805804
varray: [-2.2, -1.901, -1.914, -1.906]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.534765602858
varray: [-54.561, -55.326, -55.773, -55.597]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.49275
varray: [109.114, 110.504, 108.581, 109.772]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.831734482873
varray: [109.114, 110.504, 108.581, 109.772]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.00625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.7275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2079.3755
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 35.15725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.08535
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08535
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.9220445753
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.1132
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-2.2, -1.901, -1.914, -1.906]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-54.561, -55.326, -55.773, -55.597]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.98025
varray: [-2.2, -1.901, -1.914, -1.906]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
dD_mean | -55.31425
varray: [-54.561, -55.326, -55.773, -55.597]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
d18O_stdev | 0.146597805804
varray: [-2.2, -1.901, -1.914, -1.906]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
dD_stdev | 0.534765602858
varray: [-54.561, -55.326, -55.773, -55.597]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.98025
varray: [-2.2, -1.901, -1.914, -1.906]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.31425
varray: [-54.561, -55.326, -55.773, -55.597]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.146597805804
varray: [-2.2, -1.901, -1.914, -1.906]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.534765602858
varray: [-54.561, -55.326, -55.773, -55.597]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.70088413903
actual_array: [ 1.93 -15.84]
varray: [-2.2, -1.901, -1.914, -1.906]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -1.98025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.6027162727
actual_array: [ 18.1 -119. ]
varray: [-54.561, -55.326, -55.773, -55.597]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -55.31425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [53929.0, 54465.0, 55002.0, 55539.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.72415
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 778.6451
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.85875
varray: [2.871, 2.853, 2.897, 2.814]
inst: mean_all
operation: mean_all(residuals)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: residuals
SAMP_residuals_STD | 0.0348747759849
varray: [2.871, 2.853, 2.897, 2.814]
inst: std_all
operation: std_all(residuals)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.18325
varray: [-0.037, 0.237, 0.375, 0.158]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.172051106748
varray: [-0.037, 0.237, 0.375, 0.158]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.4405
varray: [1308.904, 1302.34, 1299.022, 1299.496]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.55130585656
varray: [1308.904, 1302.34, 1299.022, 1299.496]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3065
varray: [1.933, 2.147, 2.224, 2.922]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.428403625879
varray: [1.933, 2.147, 2.224, 2.922]
inst: std_all
operation: std_all(slope_shift)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: slope_shift
SAMP_d18O_STD | 0.131915376915
varray: [-8.46, -8.738, -8.741, -8.659]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.16034344772
varray: [-74.944, -77.057, -77.458, -77.208]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.9405
varray: [111.061, 111.816, 110.583, 110.302]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.662447230603
varray: [111.061, 111.816, 110.583, 110.302]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.85875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.18325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.4405
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3065
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06215
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06215
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.12776883837
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.66545
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.46, -8.738, -8.741, -8.659]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-74.944, -77.057, -77.458, -77.208]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.6495
varray: [-8.46, -8.738, -8.741, -8.659]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
dD_mean | -76.66675
varray: [-74.944, -77.057, -77.458, -77.208]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
d18O_stdev | 0.131915376915
varray: [-8.46, -8.738, -8.741, -8.659]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
dD_stdev | 1.16034344772
varray: [-74.944, -77.057, -77.458, -77.208]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.6495
varray: [-8.46, -8.738, -8.741, -8.659]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
measured_dD_mean | -76.66675
varray: [-74.944, -77.057, -77.458, -77.208]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.131915376915
varray: [-8.46, -8.738, -8.741, -8.659]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.16034344772
varray: [-74.944, -77.057, -77.458, -77.208]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.4712514996
actual_array: [ 1.93 -15.84]
varray: [-8.46, -8.738, -8.741, -8.659]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.6495
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -89.9112171249
actual_array: [ 18.1 -119. ]
varray: [-74.944, -77.057, -77.458, -77.208]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -76.66675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [56127.0, 56663.0, 57199.0, 57735.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.87165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.7101
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.425
varray: [0.0, 0.0, 2.848, 2.852]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: residuals
SAMP_residuals_STD | 1.64544907751
varray: [0.0, 0.0, 2.848, 2.852]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.13375
varray: [0.0, 0.0, 0.197, 0.338]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.164819851151
varray: [0.0, 0.0, 0.197, 0.338]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 648.89175
varray: [0.0, 0.0, 1298.64, 1296.927]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 749.275979431
varray: [0.0, 0.0, 1298.64, 1296.927]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.2515
varray: [0.0, 0.0, 2.35, 2.656]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.45049727105
varray: [0.0, 0.0, 2.35, 2.656]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: slope_shift
SAMP_d18O_STD | 3.73524005253
varray: [0.0, 0.0, -6.441, -6.498]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 41.9231627743
varray: [0.0, 0.0, -72.554, -72.672]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.948
varray: [0.0, 0.0, 110.206, 109.586]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 63.4489900524
varray: [0.0, 0.0, 110.206, 109.586]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.13375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 648.89175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.2515
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.4959
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.4959
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.2136670204
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.65795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, -6.441, -6.498]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, -72.554, -72.672]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.23475
varray: [0.0, 0.0, -6.441, -6.498]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
dD_mean | -36.3065
varray: [0.0, 0.0, -72.554, -72.672]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
d18O_stdev | 3.73524005253
varray: [0.0, 0.0, -6.441, -6.498]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
dD_stdev | 41.9231627743
varray: [0.0, 0.0, -72.554, -72.672]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.23475
varray: [0.0, 0.0, -6.441, -6.498]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.3065
varray: [0.0, 0.0, -72.554, -72.672]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.73524005253
varray: [0.0, 0.0, -6.441, -6.498]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
measured_dD_stdev | 41.9231627743
varray: [0.0, 0.0, -72.554, -72.672]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.97440452795
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -6.441, -6.498]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.23475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -51.5243223106
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -72.554, -72.672]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -36.3065
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [58324.0, 58861.0, 59398.0, 59934.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3054
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -651.83865
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.846
varray: [2.86, 2.843, 2.88, 2.801]
inst: mean_all
operation: mean_all(residuals)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: residuals
SAMP_residuals_STD | 0.0335956346371
varray: [2.86, 2.843, 2.88, 2.801]
inst: std_all
operation: std_all(residuals)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.37625
varray: [0.208, 0.253, 0.571, 0.473]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.173960484019
varray: [0.208, 0.253, 0.571, 0.473]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.91325
varray: [1297.365, 1298.574, 1296.7, 1295.014]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.48490120322
varray: [1297.365, 1298.574, 1296.7, 1295.014]
inst: std_all
operation: std_all(baseline_shift)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.457
varray: [2.256, 2.919, 2.483, 2.17]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.335106948103
varray: [2.256, 2.919, 2.483, 2.17]
inst: std_all
operation: std_all(slope_shift)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: slope_shift
SAMP_d18O_STD | 0.10109195484
varray: [-7.516, -7.418, -7.465, -7.652]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.956282385073
varray: [-62.562, -60.564, -60.554, -60.924]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.3535
varray: [110.593, 110.013, 110.817, 109.991]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.416148611276
varray: [110.593, 110.013, 110.817, 109.991]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.846
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.37625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.91325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.457
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0749
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0749
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.56427813345
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.25245
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-7.516, -7.418, -7.465, -7.652]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-62.562, -60.564, -60.554, -60.924]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.51275
varray: [-7.516, -7.418, -7.465, -7.652]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
dD_mean | -61.151
varray: [-62.562, -60.564, -60.554, -60.924]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
d18O_stdev | 0.10109195484
varray: [-7.516, -7.418, -7.465, -7.652]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
dD_stdev | 0.956282385073
varray: [-62.562, -60.564, -60.554, -60.924]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.51275
varray: [-7.516, -7.418, -7.465, -7.652]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.151
varray: [-62.562, -60.564, -60.554, -60.924]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.10109195484
varray: [-7.516, -7.418, -7.465, -7.652]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.956282385073
varray: [-62.562, -60.564, -60.554, -60.924]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3172664043
actual_array: [ 1.93 -15.84]
varray: [-7.516, -7.418, -7.465, -7.652]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.51275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.1540869482
actual_array: [ 18.1 -119. ]
varray: [-62.562, -60.564, -60.554, -60.924]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [60519.0, 61056.0, 61592.0, 62129.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.8844
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.81715
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94575
varray: [2.917, 2.955, 2.95, 2.961]
inst: mean_all
operation: mean_all(residuals)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: residuals
SAMP_residuals_STD | 0.01968713624
varray: [2.917, 2.955, 2.95, 2.961]
inst: std_all
operation: std_all(residuals)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.109
varray: [-3.029, -3.225, -3.256, -2.926]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.158065387314
varray: [-3.029, -3.225, -3.256, -2.926]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1911.05275
varray: [1900.72, 1915.582, 1913.422, 1914.487]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.94471613891
varray: [1900.72, 1915.582, 1913.422, 1914.487]
inst: std_all
operation: std_all(baseline_shift)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -27.7905
varray: [-27.157, -27.852, -27.96, -28.193]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.445660932399
varray: [-27.157, -27.852, -27.96, -28.193]
inst: std_all
operation: std_all(slope_shift)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0590592922409
varray: [-2.766, -2.642, -2.646, -2.658]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.247294695994
varray: [-52.983, -52.473, -52.739, -52.465]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.8695
varray: [109.399, 109.565, 109.732, 110.782]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.623338591778
varray: [109.399, 109.565, 109.732, 110.782]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.109
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1911.05275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.7905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.850765175117
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.73645
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-2.766, -2.642, -2.646, -2.658]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-52.983, -52.473, -52.739, -52.465]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.678
varray: [-2.766, -2.642, -2.646, -2.658]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
dD_mean | -52.665
varray: [-52.983, -52.473, -52.739, -52.465]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0590592922409
varray: [-2.766, -2.642, -2.646, -2.658]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
dD_stdev | 0.247294695994
varray: [-52.983, -52.473, -52.739, -52.465]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.678
varray: [-2.766, -2.642, -2.646, -2.658]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.665
varray: [-52.983, -52.473, -52.739, -52.465]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0590592922409
varray: [-2.766, -2.642, -2.646, -2.658]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.247294695994
varray: [-52.983, -52.473, -52.739, -52.465]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.40921323538
actual_array: [ 1.93 -15.84]
varray: [-2.766, -2.642, -2.646, -2.658]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.082997463
actual_array: [ 18.1 -119. ]
varray: [-52.983, -52.473, -52.739, -52.465]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [62715.0, 63251.0, 63788.0, 64324.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.78465
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 610.32235
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84875
varray: [2.856, 2.811, 2.89, 2.838]
inst: mean_all
operation: mean_all(residuals)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: residuals
SAMP_residuals_STD | 0.0331398551596
varray: [2.856, 2.811, 2.89, 2.838]
inst: std_all
operation: std_all(residuals)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.65075
varray: [-2.724, -2.709, -2.601, -2.569]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0772803338502
varray: [-2.724, -2.709, -2.601, -2.569]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1955.5855
varray: [1955.233, 1943.517, 1962.286, 1961.306]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.62928674148
varray: [1955.233, 1943.517, 1962.286, 1961.306]
inst: std_all
operation: std_all(baseline_shift)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -28.057
varray: [-27.744, -27.432, -28.357, -28.695]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.573189322999
varray: [-27.744, -27.432, -28.357, -28.695]
inst: std_all
operation: std_all(slope_shift)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0738241153012
varray: [-5.27, -5.409, -5.436, -5.397]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.18346665916
varray: [-70.043, -72.167, -72.289, -72.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.27625
varray: [107.523, 105.948, 109.093, 110.541]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.98194405824
varray: [107.523, 105.948, 109.093, 110.541]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.65075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1955.5855
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.057
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07215
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07215
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.47012906981
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.3297
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-5.27, -5.409, -5.436, -5.397]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-70.043, -72.167, -72.289, -72.661]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.378
varray: [-5.27, -5.409, -5.436, -5.397]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
dD_mean | -71.79
varray: [-70.043, -72.167, -72.289, -72.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0738241153012
varray: [-5.27, -5.409, -5.436, -5.397]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
dD_stdev | 1.18346665916
varray: [-70.043, -72.167, -72.289, -72.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.378
varray: [-5.27, -5.409, -5.436, -5.397]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.79
varray: [-70.043, -72.167, -72.289, -72.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0738241153012
varray: [-5.27, -5.409, -5.436, -5.397]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.18346665916
varray: [-70.043, -72.167, -72.289, -72.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.15014990345
actual_array: [ 1.93 -15.84]
varray: [-5.27, -5.409, -5.436, -5.397]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -5.378
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -85.2729086584
actual_array: [ 18.1 -119. ]
varray: [-70.043, -72.167, -72.289, -72.661]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -71.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [64910.0, 65447.0, 65983.0, 66518.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.88165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 654.8551
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8875
varray: [2.869, 2.919, 2.861, 2.901]
inst: mean_all
operation: mean_all(residuals)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: residuals
SAMP_residuals_STD | 0.0271968135388
varray: [2.869, 2.919, 2.861, 2.901]
inst: std_all
operation: std_all(residuals)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.83925
varray: [-2.918, -2.849, -2.71, -2.88]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0906692707959
varray: [-2.918, -2.849, -2.71, -2.88]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2088.51075
varray: [2092.433, 2081.598, 2081.045, 2098.967]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.72239922556
varray: [2092.433, 2081.598, 2081.045, 2098.967]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -33.474
varray: [-34.04, -33.207, -33.212, -33.437]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.392283061067
varray: [-34.04, -33.207, -33.212, -33.437]
inst: std_all
operation: std_all(slope_shift)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0430455185434
varray: [-4.439, -4.402, -4.393, -4.335]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.313775397378
varray: [-68.734, -68.876, -68.423, -68.177]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.718
varray: [107.327, 106.546, 105.375, 107.624]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.00414972323
varray: [107.327, 106.546, 105.375, 107.624]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.83925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2088.51075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 33.474
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0334
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0334
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.14348317299
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.88795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-4.439, -4.402, -4.393, -4.335]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-68.734, -68.876, -68.423, -68.177]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.39225
varray: [-4.439, -4.402, -4.393, -4.335]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
dD_mean | -68.5525
varray: [-68.734, -68.876, -68.423, -68.177]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0430455185434
varray: [-4.439, -4.402, -4.393, -4.335]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
dD_stdev | 0.313775397378
varray: [-68.734, -68.876, -68.423, -68.177]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.39225
varray: [-4.439, -4.402, -4.393, -4.335]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.5525
varray: [-68.734, -68.876, -68.423, -68.177]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0430455185434
varray: [-4.439, -4.402, -4.393, -4.335]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.313775397378
varray: [-68.734, -68.876, -68.423, -68.177]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.14945422917
actual_array: [ 1.93 -15.84]
varray: [-4.439, -4.402, -4.393, -4.335]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.39225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.1937014691
actual_array: [ 18.1 -119. ]
varray: [-68.734, -68.876, -68.423, -68.177]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.5525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [67104.0, 67640.0, 68176.0, 68712.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.8429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 787.78035
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.81575
varray: [2.794, 2.786, 2.864, 2.819]
inst: mean_all
operation: mean_all(residuals)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: residuals
SAMP_residuals_STD | 0.0351034186369
varray: [2.794, 2.786, 2.864, 2.819]
inst: std_all
operation: std_all(residuals)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3485
varray: [0.294, 0.448, 0.402, 0.25]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0920778655994
varray: [0.294, 0.448, 0.402, 0.25]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.799
varray: [1306.278, 1300.538, 1300.902, 1299.478]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.04648496905
varray: [1306.278, 1300.538, 1300.902, 1299.478]
inst: std_all
operation: std_all(baseline_shift)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.26375
varray: [1.851, 2.269, 2.084, 2.851]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.427224667671
varray: [1.851, 2.269, 2.084, 2.851]
inst: std_all
operation: std_all(slope_shift)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0709665414121
varray: [-9.925, -9.983, -10.051, -10.084]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.2975405691
varray: [-104.817, -108.365, -109.45, -109.88]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.96875
varray: [109.62, 109.719, 110.814, 109.722]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.565489389821
varray: [109.62, 109.719, 110.814, 109.722]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.81575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3485
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.799
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.26375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.59991783354
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6372
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-9.925, -9.983, -10.051, -10.084]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-104.817, -108.365, -109.45, -109.88]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -10.01075
varray: [-9.925, -9.983, -10.051, -10.084]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
dD_mean | -108.128
varray: [-104.817, -108.365, -109.45, -109.88]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0709665414121
varray: [-9.925, -9.983, -10.051, -10.084]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
dD_stdev | 2.2975405691
varray: [-104.817, -108.365, -109.45, -109.88]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
measured_d18O_mean | -10.01075
varray: [-9.925, -9.983, -10.051, -10.084]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.128
varray: [-104.817, -108.365, -109.45, -109.88]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0709665414121
varray: [-9.925, -9.983, -10.051, -10.084]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.2975405691
varray: [-104.817, -108.365, -109.45, -109.88]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.8531404031
actual_array: [ 1.93 -15.84]
varray: [-9.925, -9.983, -10.051, -10.084]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -10.01075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.834215483
actual_array: [ 18.1 -119. ]
varray: [-104.817, -108.365, -109.45, -109.88]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -108.128
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [69299.0, 69835.0, 70371.0, 70907.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.91465
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.0686
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.93275
varray: [2.683, 2.603, 2.445, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: residuals
SAMP_residuals_STD | 1.29228902211
varray: [2.683, 2.603, 2.445, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.04725
varray: [0.187, -0.059, 0.061, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105262766447
varray: [0.187, -0.059, 0.061, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 965.77875
varray: [1291.887, 1288.678, 1282.55, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 643.864151082
varray: [1291.887, 1288.678, 1282.55, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.44875
varray: [1.791, 1.908, 2.096, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.973970353758
varray: [1.791, 1.908, 2.096, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: slope_shift
SAMP_d18O_STD | 4.75415341219
varray: [-9.482, -9.48, -9.562, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 48.2637831022
varray: [-97.491, -95.858, -96.203, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 75.88
varray: [105.185, 101.415, 96.92, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 50.6993597921
varray: [105.185, 101.415, 96.92, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.93275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.04725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 965.77875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.44875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.98815
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.98815
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -33.8303262693
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -36.72595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-9.482, -9.48, -9.562, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-97.491, -95.858, -96.203, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.131
varray: [-9.482, -9.48, -9.562, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
dD_mean | -72.388
varray: [-97.491, -95.858, -96.203, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
d18O_stdev | 4.75415341219
varray: [-9.482, -9.48, -9.562, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
dD_stdev | 48.2637831022
varray: [-97.491, -95.858, -96.203, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.131
varray: [-9.482, -9.48, -9.562, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.388
varray: [-97.491, -95.858, -96.203, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.75415341219
varray: [-9.482, -9.48, -9.562, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
measured_dD_stdev | 48.2637831022
varray: [-97.491, -95.858, -96.203, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.92972841425
actual_array: [ 1.93 -15.84]
varray: [-9.482, -9.48, -9.562, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.131
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -85.8416703261
actual_array: [ 18.1 -119. ]
varray: [-97.491, -95.858, -96.203, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -72.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [71495.0, 72031.0, 72567.0, 73104.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.79765
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -334.95165
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.426
varray: [0.0, 2.873, 0.0, 2.831]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: residuals
SAMP_residuals_STD | 1.64669224002
varray: [0.0, 2.873, 0.0, 2.831]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.177
varray: [0.0, 0.526, 0.0, 0.182]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.247981182082
varray: [0.0, 0.526, 0.0, 0.182]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 649.23575
varray: [0.0, 1298.697, 0.0, 1298.246]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 749.67289267
varray: [0.0, 1298.697, 0.0, 1298.246]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.317
varray: [0.0, 2.507, 0.0, 2.761]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.52427184802
varray: [0.0, 2.507, 0.0, 2.761]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: slope_shift
SAMP_d18O_STD | 5.57554066138
varray: [0.0, -9.691, 0.0, -9.623]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 54.0449191345
varray: [0.0, -93.674, 0.0, -93.543]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.4675
varray: [0.0, 112.307, 0.0, 109.563]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 64.0581480454
varray: [0.0, 112.307, 0.0, 109.563]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.426
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.177
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 649.23575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.317
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.4949
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.4949
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.1794309973
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.13845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, -9.691, 0.0, -9.623]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, -93.674, 0.0, -93.543]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.8285
varray: [0.0, -9.691, 0.0, -9.623]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
dD_mean | -46.80425
varray: [0.0, -93.674, 0.0, -93.543]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
d18O_stdev | 5.57554066138
varray: [0.0, -9.691, 0.0, -9.623]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
dD_stdev | 54.0449191345
varray: [0.0, -93.674, 0.0, -93.543]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.8285
varray: [0.0, -9.691, 0.0, -9.623]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.80425
varray: [0.0, -93.674, 0.0, -93.543]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.57554066138
varray: [0.0, -9.691, 0.0, -9.623]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
measured_dD_stdev | 54.0449191345
varray: [0.0, -93.674, 0.0, -93.543]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.59231853341
actual_array: [ 1.93 -15.84]
varray: [0.0, -9.691, 0.0, -9.623]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.8285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.5088002323
actual_array: [ 18.1 -119. ]
varray: [0.0, -93.674, 0.0, -93.543]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -46.80425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [73753.0, 74289.0, 74825.0, 75362.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3044
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -651.49465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.893
varray: [0.0, 0.0, 3.804, 3.768]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: residuals
SAMP_residuals_STD | 2.18589752733
varray: [0.0, 0.0, 3.804, 3.768]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.38725
varray: [0.0, 0.0, -6.849, -6.7]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 3.91173238476
varray: [0.0, 0.0, -6.849, -6.7]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1223.95175
varray: [0.0, 0.0, 2453.585, 2442.222]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1413.30535792
varray: [0.0, 0.0, 2453.585, 2442.222]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -28.975
varray: [0.0, 0.0, -58.217, -57.683]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: slope_shift
SAMP_slope_shift_STD | 33.4581583374
varray: [0.0, 0.0, -58.217, -57.683]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: slope_shift
SAMP_d18O_STD | 0.325349427539
varray: [0.0, 0.0, 0.567, 0.56]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 35.9321005231
varray: [0.0, 0.0, -62.354, -62.118]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.17975
varray: [0.0, 0.0, 115.025, 113.694]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 66.0277240275
varray: [0.0, 0.0, 115.025, 113.694]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.893
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.38725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1223.95175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.0279
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.0279
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -35.1912081893
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -55.4262
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.567, 0.56]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, -62.354, -62.118]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.28175
varray: [0.0, 0.0, 0.567, 0.56]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
dD_mean | -31.118
varray: [0.0, 0.0, -62.354, -62.118]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
d18O_stdev | 0.325349427539
varray: [0.0, 0.0, 0.567, 0.56]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
dD_stdev | 35.9321005231
varray: [0.0, 0.0, -62.354, -62.118]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.28175
varray: [0.0, 0.0, 0.567, 0.56]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
measured_dD_mean | -31.118
varray: [0.0, 0.0, -62.354, -62.118]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.325349427539
varray: [0.0, 0.0, 0.567, 0.56]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
measured_dD_stdev | 35.9321005231
varray: [0.0, 0.0, -62.354, -62.118]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.40458830822
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.567, 0.56]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.28175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -46.5895064028
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -62.354, -62.118]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -31.118
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [75946.0, 76482.0, 77018.0, 77555.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.8374
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -76.77865
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.6455
varray: [4.375, 4.738, 4.752, 4.717]
inst: mean_all
operation: mean_all(residuals)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: residuals
SAMP_residuals_STD | 0.180906052967
varray: [4.375, 4.738, 4.752, 4.717]
inst: std_all
operation: std_all(residuals)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -9.00825
varray: [-8.462, -9.194, -9.005, -9.372]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.393792140602
varray: [-8.462, -9.194, -9.005, -9.372]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1820.06925
varray: [1785.716, 1834.287, 1831.931, 1828.343]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 23.0321966875
varray: [1785.716, 1834.287, 1831.931, 1828.343]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -42.177
varray: [-39.44, -43.676, -42.842, -42.75]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.87160572771
varray: [-39.44, -43.676, -42.842, -42.75]
inst: std_all
operation: std_all(slope_shift)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0916169562181
varray: [4.844, 5.041, 5.018, 5.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.271023215488
varray: [-65.401, -65.695, -65.906, -66.017]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 123.1675
varray: [115.677, 126.304, 125.654, 125.035]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.02047345045
varray: [115.677, 126.304, 125.654, 125.035]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.6455
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 9.00825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1820.06925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 42.177
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.7246
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.7246
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 59.0434455134
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 10.56155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [4.844, 5.041, 5.018, 5.019]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-65.401, -65.695, -65.906, -66.017]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.9805
varray: [4.844, 5.041, 5.018, 5.019]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
dD_mean | -65.75475
varray: [-65.401, -65.695, -65.906, -66.017]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0916169562181
varray: [4.844, 5.041, 5.018, 5.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
dD_stdev | 0.271023215488
varray: [-65.401, -65.695, -65.906, -66.017]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.9805
varray: [4.844, 5.041, 5.018, 5.019]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.75475
varray: [-65.401, -65.695, -65.906, -66.017]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0916169562181
varray: [4.844, 5.041, 5.018, 5.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.271023215488
varray: [-65.401, -65.695, -65.906, -66.017]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.36540286553
actual_array: [ 1.93 -15.84]
varray: [4.844, 5.041, 5.018, 5.019]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.9805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.532743349
actual_array: [ 18.1 -119. ]
varray: [-65.401, -65.695, -65.906, -66.017]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -65.75475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [78140.0, 78676.0, 79212.0, 79748.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1296.0849
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 519.33885
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.377
varray: [0.0, 0.0, 2.736, 2.772]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: residuals
SAMP_residuals_STD | 1.59009056346
varray: [0.0, 0.0, 2.736, 2.772]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.11325
varray: [0.0, 0.0, 0.086, 0.367]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.173956651688
varray: [0.0, 0.0, 0.086, 0.367]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 648.97725
varray: [0.0, 0.0, 1298.191, 1297.718]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 749.37440485
varray: [0.0, 0.0, 1298.191, 1297.718]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.306
varray: [0.0, 0.0, 2.749, 2.475]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.51218186737
varray: [0.0, 0.0, 2.749, 2.475]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: slope_shift
SAMP_d18O_STD | 4.30014581148
varray: [0.0, 0.0, -7.471, -7.425]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 35.1449806599
varray: [0.0, 0.0, -61.027, -60.718]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.1175
varray: [0.0, 0.0, 107.535, 108.935]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 62.4921201033
varray: [0.0, 0.0, 107.535, 108.935]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.377
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.11325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 648.97725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.306
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.5439
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.5439
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -52.8569961313
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -58.48845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, -7.471, -7.425]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, -61.027, -60.718]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.724
varray: [0.0, 0.0, -7.471, -7.425]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
dD_mean | -30.43625
varray: [0.0, 0.0, -61.027, -60.718]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
d18O_stdev | 4.30014581148
varray: [0.0, 0.0, -7.471, -7.425]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
dD_stdev | 35.1449806599
varray: [0.0, 0.0, -61.027, -60.718]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.724
varray: [0.0, 0.0, -7.471, -7.425]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
measured_dD_mean | -30.43625
varray: [0.0, 0.0, -61.027, -60.718]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.30014581148
varray: [0.0, 0.0, -7.471, -7.425]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
measured_dD_stdev | 35.1449806599
varray: [0.0, 0.0, -61.027, -60.718]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.47107240383
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -7.471, -7.425]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -45.9410895684
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -61.027, -60.718]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -30.43625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [80334.0, 80869.0, 81406.0, 81942.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.3534
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -651.75315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8575
varray: [2.531, 2.97, 2.956, 2.973]
inst: mean_all
operation: mean_all(residuals)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: residuals
SAMP_residuals_STD | 0.217792714907
varray: [2.531, 2.97, 2.956, 2.973]
inst: std_all
operation: std_all(residuals)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.24525
varray: [-2.91, -3.308, -3.478, -3.285]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.239501391784
varray: [-2.91, -3.308, -3.478, -3.285]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1941.89375
varray: [1842.489, 1977.119, 1972.197, 1975.77]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 66.302362539
varray: [1842.489, 1977.119, 1972.197, 1975.77]
inst: std_all
operation: std_all(baseline_shift)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -29.79925
varray: [-25.7, -31.26, -31.042, -31.195]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.73436078758
varray: [-25.7, -31.26, -31.042, -31.195]
inst: std_all
operation: std_all(slope_shift)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: slope_shift
SAMP_d18O_STD | 0.138572183356
varray: [-4.066, -3.751, -3.916, -3.808]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.394391831389
varray: [-68.891, -68.355, -69.295, -69.018]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.4785
varray: [92.883, 109.839, 108.704, 110.488]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 8.42930445925
varray: [92.883, 109.839, 108.704, 110.488]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.24525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1941.89375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.79925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0634
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0634
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.1705638673
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.12745
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-4.066, -3.751, -3.916, -3.808]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-68.891, -68.355, -69.295, -69.018]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.88525
varray: [-4.066, -3.751, -3.916, -3.808]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
dD_mean | -68.88975
varray: [-68.891, -68.355, -69.295, -69.018]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
d18O_stdev | 0.138572183356
varray: [-4.066, -3.751, -3.916, -3.808]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
dD_stdev | 0.394391831389
varray: [-68.891, -68.355, -69.295, -69.018]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.88525
varray: [-4.066, -3.751, -3.916, -3.808]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.88975
varray: [-68.891, -68.355, -69.295, -69.018]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.138572183356
varray: [-4.066, -3.751, -3.916, -3.808]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.394391831389
varray: [-68.891, -68.355, -69.295, -69.018]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.63476723261
actual_array: [ 1.93 -15.84]
varray: [-4.066, -3.751, -3.916, -3.808]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.88525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.5144621254
actual_array: [ 18.1 -119. ]
varray: [-68.891, -68.355, -69.295, -69.018]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.88975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [82527.0, 83064.0, 83600.0, 84136.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.8729
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 641.16335
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.40475
varray: [0.0, 0.0, 2.834, 2.785]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: residuals
SAMP_residuals_STD | 1.62218892755
varray: [0.0, 0.0, 2.834, 2.785]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.13725
varray: [0.0, 0.0, 0.247, 0.302]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.160065351237
varray: [0.0, 0.0, 0.247, 0.302]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 649.501
varray: [0.0, 0.0, 1300.992, 1297.012]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 749.980914469
varray: [0.0, 0.0, 1300.992, 1297.012]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.21375
varray: [0.0, 0.0, 2.525, 2.33]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.40377689942
varray: [0.0, 0.0, 2.525, 2.33]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: slope_shift
SAMP_d18O_STD | 5.09282793086
varray: [0.0, 0.0, -8.803, -8.839]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 49.1758570066
varray: [0.0, 0.0, -85.091, -85.259]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.05775
varray: [0.0, 0.0, 110.68, 109.551]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 63.5768843232
varray: [0.0, 0.0, 110.68, 109.551]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.40475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.13725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 649.501
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.21375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.51615
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.51615
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.9069464891
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.5482
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, -8.803, -8.839]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, -85.091, -85.259]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.4105
varray: [0.0, 0.0, -8.803, -8.839]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
dD_mean | -42.5875
varray: [0.0, 0.0, -85.091, -85.259]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
d18O_stdev | 5.09282793086
varray: [0.0, 0.0, -8.803, -8.839]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
dD_stdev | 49.1758570066
varray: [0.0, 0.0, -85.091, -85.259]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.4105
varray: [0.0, 0.0, -8.803, -8.839]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.5875
varray: [0.0, 0.0, -85.091, -85.259]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.09282793086
varray: [0.0, 0.0, -8.803, -8.839]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
measured_dD_stdev | 49.1758570066
varray: [0.0, 0.0, -85.091, -85.259]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.16798093073
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -8.803, -8.839]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.4105
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.4982220345
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -85.091, -85.259]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -42.5875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [84723.0, 85258.0, 85794.0, 86331.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.32565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -651.2294
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7955
varray: [2.837, 2.845, 2.779, 2.721]
inst: mean_all
operation: mean_all(residuals)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: residuals
SAMP_residuals_STD | 0.0577205913576
varray: [2.837, 2.845, 2.779, 2.721]
inst: std_all
operation: std_all(residuals)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3265
varray: [0.265, 0.44, 0.32, 0.281]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0791138420253
varray: [0.265, 0.44, 0.32, 0.281]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.6525
varray: [1297.583, 1298.173, 1296.701, 1294.153]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.77271947395
varray: [1297.583, 1298.173, 1296.701, 1294.153]
inst: std_all
operation: std_all(baseline_shift)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.347
varray: [2.355, 2.451, 2.336, 2.246]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0840674332505
varray: [2.355, 2.451, 2.336, 2.246]
inst: std_all
operation: std_all(slope_shift)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0686798126575
varray: [-9.191, -9.225, -9.17, -9.325]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.534627907988
varray: [-89.124, -89.691, -89.676, -90.427]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.39425
varray: [110.339, 110.939, 108.963, 107.336]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.60220418487
varray: [110.339, 110.939, 108.963, 107.336]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7955
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3265
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.6525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.347
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1254
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1254
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.2931973022
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.2117
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-9.191, -9.225, -9.17, -9.325]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-89.124, -89.691, -89.676, -90.427]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.22775
varray: [-9.191, -9.225, -9.17, -9.325]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
dD_mean | -89.7295
varray: [-89.124, -89.691, -89.676, -90.427]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0686798126575
varray: [-9.191, -9.225, -9.17, -9.325]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
dD_stdev | 0.534627907988
varray: [-89.124, -89.691, -89.676, -90.427]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.22775
varray: [-9.191, -9.225, -9.17, -9.325]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
measured_dD_mean | -89.7295
varray: [-89.124, -89.691, -89.676, -90.427]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0686798126575
varray: [-9.191, -9.225, -9.17, -9.325]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.534627907988
varray: [-89.124, -89.691, -89.676, -90.427]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.0582687694
actual_array: [ 1.93 -15.84]
varray: [-9.191, -9.225, -9.17, -9.325]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.22775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -102.335283136
actual_array: [ 18.1 -119. ]
varray: [-89.124, -89.691, -89.676, -90.427]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -89.7295
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [86917.0, 87453.0, 87989.0, 88525.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.9349
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.0779
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.67025
varray: [2.609, 2.712, 2.671, 2.689]
inst: mean_all
operation: mean_all(residuals)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: residuals
SAMP_residuals_STD | 0.044146536293
varray: [2.609, 2.712, 2.671, 2.689]
inst: std_all
operation: std_all(residuals)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2185
varray: [0.061, 0.27, 0.251, 0.292]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.106328108545
varray: [0.061, 0.27, 0.251, 0.292]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.739
varray: [1289.131, 1292.023, 1290.795, 1291.007]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.1985101863
varray: [1289.131, 1292.023, 1290.795, 1291.007]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.38225
varray: [2.514, 2.057, 2.559, 2.399]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.227058546048
varray: [2.514, 2.057, 2.559, 2.399]
inst: std_all
operation: std_all(slope_shift)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0322942203291
varray: [-9.507, -9.459, -9.446, -9.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0833721576227
varray: [-90.767, -90.674, -90.858, -90.694]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.36825
varray: [102.482, 105.506, 104.804, 104.681]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.30895693206
varray: [102.482, 105.506, 104.804, 104.681]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.67025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2185
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.739
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.38225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.25065
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.25065
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.58125920093
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.2377
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-9.507, -9.459, -9.446, -9.433]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-90.767, -90.674, -90.858, -90.694]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.46125
varray: [-9.507, -9.459, -9.446, -9.433]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
dD_mean | -90.74825
varray: [-90.767, -90.674, -90.858, -90.694]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0322942203291
varray: [-9.507, -9.459, -9.446, -9.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
dD_stdev | 0.0833721576227
varray: [-90.767, -90.674, -90.858, -90.694]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.46125
varray: [-9.507, -9.459, -9.446, -9.433]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
measured_dD_mean | -90.74825
varray: [-90.767, -90.674, -90.858, -90.694]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0322942203291
varray: [-9.507, -9.459, -9.446, -9.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0833721576227
varray: [-90.767, -90.674, -90.858, -90.694]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.2953090331
actual_array: [ 1.93 -15.84]
varray: [-9.507, -9.459, -9.446, -9.433]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.46125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -103.30422285
actual_array: [ 18.1 -119. ]
varray: [-90.767, -90.674, -90.858, -90.694]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.74825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [89112.0, 89649.0, 90184.0, 90720.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.06015
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.9914
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.80475
varray: [2.821, 2.802, 2.839, 2.757]
inst: mean_all
operation: mean_all(residuals)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: residuals
SAMP_residuals_STD | 0.0352361083738
varray: [2.821, 2.802, 2.839, 2.757]
inst: std_all
operation: std_all(residuals)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.39675
varray: [0.544, 0.339, 0.438, 0.266]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.120850803335
varray: [0.544, 0.339, 0.438, 0.266]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.4795
varray: [1297.011, 1295.884, 1296.948, 1296.075]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.583159497908
varray: [1297.011, 1295.884, 1296.948, 1296.075]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.54875
varray: [2.549, 2.932, 2.082, 2.632]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.351978574159
varray: [2.549, 2.932, 2.082, 2.632]
inst: std_all
operation: std_all(slope_shift)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0398078719183
varray: [-8.926, -8.913, -8.836, -8.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.203501637995
varray: [-87.89, -87.771, -87.413, -87.727]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.572
varray: [110.959, 107.899, 110.819, 108.611]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.54932544462
varray: [110.959, 107.899, 110.819, 108.611]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.80475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.39675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.4795
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.54875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11615
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11615
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.97651408812
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.03395
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.926, -8.913, -8.836, -8.897]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-87.89, -87.771, -87.413, -87.727]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.893
varray: [-8.926, -8.913, -8.836, -8.897]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
dD_mean | -87.70025
varray: [-87.89, -87.771, -87.413, -87.727]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0398078719183
varray: [-8.926, -8.913, -8.836, -8.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
dD_stdev | 0.203501637995
varray: [-87.89, -87.771, -87.413, -87.727]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.893
varray: [-8.926, -8.913, -8.836, -8.897]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
measured_dD_mean | -87.70025
varray: [-87.89, -87.771, -87.413, -87.727]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0398078719183
varray: [-8.926, -8.913, -8.836, -8.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.203501637995
varray: [-87.89, -87.771, -87.413, -87.727]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.7184433806
actual_array: [ 1.93 -15.84]
varray: [-8.926, -8.913, -8.836, -8.897]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -100.405250337
actual_array: [ 18.1 -119. ]
varray: [-87.89, -87.771, -87.413, -87.727]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.70025
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [91302.0, 91837.0, 92373.0, 92909.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.92565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.2509
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.21625
varray: [0.0, 0.0, 2.55, 2.315]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: residuals
SAMP_residuals_STD | 1.40767760869
varray: [0.0, 0.0, 2.55, 2.315]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.05875
varray: [0.0, 0.0, -0.005, 0.24]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.120856319653
varray: [0.0, 0.0, -0.005, 0.24]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 641.522
varray: [0.0, 0.0, 1287.516, 1278.572]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 740.774797876
varray: [0.0, 0.0, 1287.516, 1278.572]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.1225
varray: [0.0, 0.0, 2.298, 2.192]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.29687354819
varray: [0.0, 0.0, 2.298, 2.192]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: slope_shift
SAMP_d18O_STD | 4.49907548836
varray: [0.0, 0.0, -7.761, -7.824]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 39.0916938524
varray: [0.0, 0.0, -67.566, -67.851]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 47.913
varray: [0.0, 0.0, 99.089, 92.563]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 55.3892787881
varray: [0.0, 0.0, 99.089, 92.563]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.21625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.05875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 641.522
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.1225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.70465
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.70465
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -58.3604368517
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -64.69295
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, -7.761, -7.824]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, -67.566, -67.851]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.89625
varray: [0.0, 0.0, -7.761, -7.824]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
dD_mean | -33.85425
varray: [0.0, 0.0, -67.566, -67.851]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
d18O_stdev | 4.49907548836
varray: [0.0, 0.0, -7.761, -7.824]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
dD_stdev | 39.0916938524
varray: [0.0, 0.0, -67.566, -67.851]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.89625
varray: [0.0, 0.0, -7.761, -7.824]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
measured_dD_mean | -33.85425
varray: [0.0, 0.0, -67.566, -67.851]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.49907548836
varray: [0.0, 0.0, -7.761, -7.824]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
measured_dD_stdev | 39.0916938524
varray: [0.0, 0.0, -67.566, -67.851]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.64593401163
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -7.761, -7.824]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.89625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.1919714751
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -67.566, -67.851]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -33.85425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [93493.0, 94029.0, 94564.0, 95100.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.51415
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -659.2084
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.768
varray: [2.824, 2.739, 2.725, 2.784]
inst: mean_all
operation: mean_all(residuals)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: residuals
SAMP_residuals_STD | 0.0450259184619
varray: [2.824, 2.739, 2.725, 2.784]
inst: std_all
operation: std_all(residuals)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2345
varray: [0.413, 0.23, 0.23, 0.065]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.142165396634
varray: [0.413, 0.23, 0.23, 0.065]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.23575
varray: [1295.489, 1292.627, 1295.376, 1297.451]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.98298720369
varray: [1295.489, 1292.627, 1295.376, 1297.451]
inst: std_all
operation: std_all(baseline_shift)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.58375
varray: [2.56, 2.422, 2.785, 2.568]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.149974164442
varray: [2.56, 2.422, 2.785, 2.568]
inst: std_all
operation: std_all(slope_shift)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: slope_shift
SAMP_d18O_STD | 0.045022216738
varray: [-9.489, -9.562, -9.597, -9.546]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.63737004573
varray: [-87.803, -90.783, -91.204, -91.178]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.36075
varray: [109.615, 106.742, 107.998, 109.088]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.27200298087
varray: [109.615, 106.742, 107.998, 109.088]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.768
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2345
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.23575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.58375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1529
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1529
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.23468793865
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.2452
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-9.489, -9.562, -9.597, -9.546]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-87.803, -90.783, -91.204, -91.178]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.5485
varray: [-9.489, -9.562, -9.597, -9.546]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
dD_mean | -90.242
varray: [-87.803, -90.783, -91.204, -91.178]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
d18O_stdev | 0.045022216738
varray: [-9.489, -9.562, -9.597, -9.546]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
dD_stdev | 1.63737004573
varray: [-87.803, -90.783, -91.204, -91.178]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.5485
varray: [-9.489, -9.562, -9.597, -9.546]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
measured_dD_mean | -90.242
varray: [-87.803, -90.783, -91.204, -91.178]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.045022216738
varray: [-9.489, -9.562, -9.597, -9.546]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.63737004573
varray: [-87.803, -90.783, -91.204, -91.178]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.3838818939
actual_array: [ 1.93 -15.84]
varray: [-9.489, -9.562, -9.597, -9.546]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.5485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -102.822725201
actual_array: [ 18.1 -119. ]
varray: [-87.803, -90.783, -91.204, -91.178]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [95682.0, 96217.0, 96753.0, 97289.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.9624
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.49465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [97873.0, 98471.0, 99069.0, 99667.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7255
varray: [2.811, 2.679, 2.702, 2.71]
inst: mean_all
operation: mean_all(residuals)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: residuals
SAMP_residuals_STD | 0.0584950140325
varray: [2.811, 2.679, 2.702, 2.71]
inst: std_all
operation: std_all(residuals)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.27425
varray: [0.361, 0.078, 0.339, 0.319]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.13195295879
varray: [0.361, 0.078, 0.339, 0.319]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.16125
varray: [1294.363, 1291.654, 1291.095, 1291.533]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.48734581834
varray: [1294.363, 1291.654, 1291.095, 1291.533]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.83975
varray: [2.909, 2.817, 3.003, 2.63]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.159121284979
varray: [2.909, 2.817, 3.003, 2.63]
inst: std_all
operation: std_all(slope_shift)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: slope_shift
SAMP_d18O_STD | 0.141320852436
varray: [-7.875, -7.677, -7.568, -7.583]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.65276777961
varray: [-64.747, -62.211, -61.363, -61.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.65525
varray: [108.11, 103.924, 105.47, 105.117]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7651302832
varray: [108.11, 103.924, 105.47, 105.117]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7255
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.27425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.16125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.83975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1954
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1954
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.68971892225
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.9507
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-7.875, -7.677, -7.568, -7.583]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-64.747, -62.211, -61.363, -61.143]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.67575
varray: [-7.875, -7.677, -7.568, -7.583]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
dD_mean | -62.366
varray: [-64.747, -62.211, -61.363, -61.143]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
d18O_stdev | 0.141320852436
varray: [-7.875, -7.677, -7.568, -7.583]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
dD_stdev | 1.65276777961
varray: [-64.747, -62.211, -61.363, -61.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.67575
varray: [-7.875, -7.677, -7.568, -7.583]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.366
varray: [-64.747, -62.211, -61.363, -61.143]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.141320852436
varray: [-7.875, -7.677, -7.568, -7.583]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.65276777961
varray: [-64.747, -62.211, -61.363, -61.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4827377661
actual_array: [ 1.93 -15.84]
varray: [-7.875, -7.677, -7.568, -7.583]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.67575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.3096813065
actual_array: [ 18.1 -119. ]
varray: [-64.747, -62.211, -61.363, -61.143]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [100314.0, 100849.0, 101384.0, 101919.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.0049
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.56915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.754
varray: [2.622, 2.763, 2.769, 2.862]
inst: mean_all
operation: mean_all(residuals)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: residuals
SAMP_residuals_STD | 0.0989848473252
varray: [2.622, 2.763, 2.769, 2.862]
inst: std_all
operation: std_all(residuals)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.20725
varray: [0.286, 0.038, 0.035, 0.47]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.210993483312
varray: [0.286, 0.038, 0.035, 0.47]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.7575
varray: [1289.991, 1297.465, 1298.39, 1297.184]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.87870412208
varray: [1289.991, 1297.465, 1298.39, 1297.184]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.5865
varray: [2.936, 2.277, 2.757, 2.376]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.31161996085
varray: [2.936, 2.277, 2.757, 2.376]
inst: std_all
operation: std_all(slope_shift)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0758990557693
varray: [-11.114, -11.219, -11.22, -11.299]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.31785060946
varray: [-107.072, -112.149, -113.699, -114.436]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.4315
varray: [102.563, 108.098, 108.679, 110.386]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.38781350333
varray: [102.563, 108.098, 108.679, 110.386]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.754
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.20725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.7575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.5865
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1669
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1669
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.71399226266
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.17445
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-11.114, -11.219, -11.22, -11.299]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-107.072, -112.149, -113.699, -114.436]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.213
varray: [-11.114, -11.219, -11.22, -11.299]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
dD_mean | -111.839
varray: [-107.072, -112.149, -113.699, -114.436]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0758990557693
varray: [-11.114, -11.219, -11.22, -11.299]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
dD_stdev | 3.31785060946
varray: [-107.072, -112.149, -113.699, -114.436]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.213
varray: [-11.114, -11.219, -11.22, -11.299]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
measured_dD_mean | -111.839
varray: [-107.072, -112.149, -113.699, -114.436]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0758990557693
varray: [-11.114, -11.219, -11.22, -11.299]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.31785060946
varray: [-107.072, -112.149, -113.699, -114.436]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -14.0736185917
actual_array: [ 1.93 -15.84]
varray: [-11.114, -11.219, -11.22, -11.299]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -11.213
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -123.363771585
actual_array: [ 18.1 -119. ]
varray: [-107.072, -112.149, -113.699, -114.436]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -111.839
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [102503.0, 103040.0, 103575.0, 104111.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.9764
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.9729
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.9209
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.9209
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -112.60595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.65042509848
except: True
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 17.8665332604
except: True
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [104696.0, 105294.0, 105892.0, 106490.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1300.7304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.389
varray: [2.742, 2.492, 1.86, 2.462]
inst: mean_all
operation: mean_all(residuals)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: residuals
SAMP_residuals_STD | 0.374338527717
varray: [2.742, 2.492, 1.86, 2.462]
inst: std_all
operation: std_all(residuals)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.16875
varray: [0.167, 0.279, 0.051, 0.178]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0932894241952
varray: [0.167, 0.279, 0.051, 0.178]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1279.21575
varray: [1294.244, 1281.655, 1257.446, 1283.518]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 15.5373700987
varray: [1294.244, 1281.655, 1257.446, 1283.518]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.359
varray: [2.575, 2.425, 2.004, 2.432]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.24655357768
varray: [2.575, 2.425, 2.004, 2.432]
inst: std_all
operation: std_all(slope_shift)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: slope_shift
SAMP_d18O_STD | 0.198577230988
varray: [-8.039, -7.741, -8.058, -7.675]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.16832122082
varray: [-67.292, -63.976, -65.493, -62.202]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 94.55175
varray: [107.336, 97.44, 75.225, 98.206]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 13.646194155
varray: [107.336, 97.44, 75.225, 98.206]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.389
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.16875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1279.21575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.359
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.5319
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.5319
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -18.2101407101
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -18.0542
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -12.953
varray: [-8.039, -7.741, -8.058, -7.675]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.2509
varray: [-67.292, -63.976, -65.493, -62.202]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.87825
varray: [-8.039, -7.741, -8.058, -7.675]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
dD_mean | -64.74075
varray: [-67.292, -63.976, -65.493, -62.202]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
d18O_stdev | 0.198577230988
varray: [-8.039, -7.741, -8.058, -7.675]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
dD_stdev | 2.16832122082
varray: [-67.292, -63.976, -65.493, -62.202]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.87825
varray: [-8.039, -7.741, -8.058, -7.675]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.74075
varray: [-67.292, -63.976, -65.493, -62.202]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.198577230988
varray: [-8.039, -7.741, -8.058, -7.675]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.16832122082
varray: [-67.292, -63.976, -65.493, -62.202]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.6883080162
actual_array: [ 1.93 -15.84]
varray: [-8.039, -7.741, -8.058, -7.675]
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.87825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.5683213907
actual_array: [ 18.1 -119. ]
varray: [-67.292, -63.976, -65.493, -62.202]
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -64.74075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: [107136.0, 107672.0, 108207.0, 108742.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.3414
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -21.51465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.36
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.41
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1773621791
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.36
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.9603817329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -70.41
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.801
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6098867726
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.801
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.6640569165
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -63.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.664
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.35
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4708096158
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.664
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.294463603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.35
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.765
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.198
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5733409504
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.765
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1498954199
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.198
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3967028095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1531358417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.474
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4982189824
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4612943373
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.523
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.455
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.327671812
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.523
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4921168467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.455
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.502
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.758
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3063534157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7803021064
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.758
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.777
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4078695886
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7983731293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.777
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.967
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.924380791335
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 3.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.80219337222
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 23.967
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.312
operation: col(d(18_16)Mean)
inst: col
dD_actual | 33.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.68676724975
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.312
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.1268411819
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 33.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.563
operation: col(d(18_16)Mean)
inst: col
dD_actual | 36.931
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94157284371
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.563
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.1323376199
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 36.931
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.574
operation: col(d(18_16)Mean)
inst: col
dD_actual | 37.624
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.95273962273
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.574
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.791454402
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 37.624
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.665
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04511934006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.665
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5285619145
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 38.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.713
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.09384710305
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.713
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3198824957
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 39.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.751
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.293
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.13242324875
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3788510967
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 39.293
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.609
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.18419649692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6794007405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 39.609
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.766
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.757
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.14765067468
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.766
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.8201644977
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 39.757
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.809
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5516
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 36.897
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.809
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.388
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.19130262902
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.809
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.4203126789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 40.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.497
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.3619245227
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -11.497
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.029325505
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.736
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6197099048
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -12.736
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.508583268
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -103.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.946
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.434
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8328938679
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -12.946
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.174147594
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -107.434
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.032
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9201977766
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.562763037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -108.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.909
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0592749334
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.528136102
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -109.909
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.153
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0430323458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.153
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.000836016
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -110.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.204
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.832
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.094805594
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.204
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.406007372
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -110.832
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.227
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.006
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1181543137
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.227
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.571499897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -111.006
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.327
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.2196704866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.327
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.281025322
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -111.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.391
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1303362545
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -13.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.937675887
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -111.391
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8941320567
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8974318044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.675
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.388
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4819763948
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.3306056488
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.617
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.335
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4230970145
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.617
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3290905383
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.335
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.425127338
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1531358417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.435
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3368082675
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4730947173
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.435
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.575
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.318
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3804602219
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3618152606
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.318
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.578
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.36
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3835057071
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.578
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4017617322
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.36
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.521
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.384
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3256414885
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.521
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4245882874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.384
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.343
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.425127338
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3855929223
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.343
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.491
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3662479577
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5263566795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.491
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.912
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.773
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7225697245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.912
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.621123929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.496
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3002624453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.496
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.0257312802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.499
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.132
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3033079305
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.499
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.0114646831
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.132
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.445
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.248
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2484891971
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.445
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.1217930334
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.248
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.465
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.958
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.22330724495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6771622063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -66.958
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.157
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.358
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.91063743245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.157
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2042853902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -64.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.155
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.476
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.90860710899
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.155
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.3165159534
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -64.476
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.194
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.94819841642
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.194
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.0511572489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -64.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.979
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.5
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.74510037362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.979
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.4592589278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -82.5
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.444
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.68
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.20198884865
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.444
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.6793516243
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -81.68
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.302
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.613
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.05783588314
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.302
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.6156274909
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -81.613
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.329
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.08524524982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.329
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.0341143367
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -82.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.825
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.519
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6342506541
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.259626467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -66.519
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.195
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.398733133
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1470421005
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.195
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3104140626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0199809362
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.568
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.869
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3733540898
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.568
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8858749243
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.869
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.534
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2442971562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.513725902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -102.534
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.534
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2442971562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.513725902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -102.534
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.688
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.276
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.41936485267
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.688
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.4953135634
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -36.276
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.12
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.551
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.84275299064
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.12
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.1969096324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -27.551
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.677
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.76255521406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.4145361119
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -25.677
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.315
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.72093358317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.0702355705
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -25.315
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.244
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.114
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0596034185
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.244
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.4343865571
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -75.114
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5844420324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.6241874491
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -81.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.905
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.367
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7306253213
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.905
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.2838682353
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -83.367
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.537
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7143827337
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.4455563348
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -83.537
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.292
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.775
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1083311815
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.292
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.6230261419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.223
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0382850222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.8084919031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.124
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.704
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9377840111
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.124
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.5554975827
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.704
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.833
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9804208037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.678190317
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -84.833
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.969
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.85
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.795595672
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.969
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0633780305
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -56.85
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.064
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.542
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.8920360362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.064
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.9171178352
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -53.542
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.038
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.365
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.8656418313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.038
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.7976655227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.365
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.16
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.587
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.9894915622
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.16
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.0088111585
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.587
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.2
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.92396592896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.2
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.8862953258
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -54.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.901
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.326
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.62043257201
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -1.901
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6138917737
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -55.326
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.773
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.63362967449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -1.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0390363647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -55.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.906
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.597
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.62550838065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -1.906
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8716416264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -55.597
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.46
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.944
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.278878352
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.46
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.2726984576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -74.944
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.738
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.057
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5610933126
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.738
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.2823864239
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -77.057
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.458
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5641387978
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.6637801175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -77.458
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.659
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.480895536
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.659
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.4260035006
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -77.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.441
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.22926682129
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -6.441
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.9995539997
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -72.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.28713103984
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -6.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.1117845629
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -72.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.516
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.562
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3205656799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.516
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.4960981742
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.562
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.418
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.564
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2210798304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.418
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5957874516
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.564
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.465
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2687924317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.586276387
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.652
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.924
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.458627675
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.652
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.93818578
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.924
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.766
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.49854746752
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.766
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.3854493198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.642
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.37266741314
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.642
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.9003850212
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.646
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.37672806005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.646
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.1533793416
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.658
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.465
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.3889100008
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -2.658
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.8927761695
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -52.465
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.27
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.043
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.04051243673
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -5.27
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.6113256593
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -70.043
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.409
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.18161991705
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -5.409
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.6314757967
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -72.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.436
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.289
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.20902928373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -5.436
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.7475107858
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -72.289
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.397
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.1694379763
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -5.397
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.1013223918
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -72.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.439
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.734
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.19691304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.439
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.366327293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.402
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.876
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.15935205603
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.402
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5013844114
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.876
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.393
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.423
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.15021560047
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.393
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0705331815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.423
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.335
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.177
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.09133622019
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.335
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8365609905
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.177
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.925
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.817
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.7660902848
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.685101968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -104.817
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.983
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.365
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.8249696651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.983
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.059627716
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -108.365
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.051
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.8940006627
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -10.051
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.091578233
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -109.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.084
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.88
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.9275009997
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -10.084
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.500554014
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -109.88
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.482
operation: col(d(18_16)Mean)
inst: col
dD_actual | -97.491
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3163736389
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.482
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.717295986
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -97.491
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.48
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.858
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3143433155
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.48
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -108.164139124
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -95.858
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.203
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3975865772
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -108.492270855
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -96.203
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.674
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.5285424403
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -106.086922598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -93.674
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.623
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.543
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.4595114427
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.623
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -105.962327651
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -93.543
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.567
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.354
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.11501342504
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.567
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2982680289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.354
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.118
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.12211955714
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0738069025
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.118
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.844
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.401
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.22683328954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 4.844
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1962894361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -65.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.695
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.42682015013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 5.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.4759147376
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -65.695
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.018
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.906
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.40347143037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 5.018
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.6765982023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -65.906
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.017
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.4044865921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 5.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7821710202
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -66.017
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.471
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.027
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2748834021
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.471
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0361497462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.027
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.425
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.718
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2281859625
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7422578477
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -60.718
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.066
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.891
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.81825771511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -4.066
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5156510084
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.891
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.751
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.355
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.49848177051
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0058579417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -68.355
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.916
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.295
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.66598345578
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.916
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.8998980214
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -69.295
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.808
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.55634598905
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -3.808
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.6364415298
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -69.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.091
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.627078825
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.9235757857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -85.091
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.839
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.259
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.6636246472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.839
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.0833616723
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -85.259
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.191
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.124
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.0209615758
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.191
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.75938817
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -89.124
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.225
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.0554770746
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.298665537
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -89.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.17
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.676
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.9996431795
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.28439894
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -89.676
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.325
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.427
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.1569932475
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.998679898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.427
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.507
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.767
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3417526822
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.507
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.322056097
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.767
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.459
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.674
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2930249192
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.233603195
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.674
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.858
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2798278167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.408606785
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.858
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.433
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.694
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2666307142
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.433
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.252625324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.694
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.89
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7519437177
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.585722789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.89
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.913
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7387466152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.913
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.472541119
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.836
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.413
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.660579162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.836
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.132045004
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.413
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.727
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7225040275
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.430692435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.727
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5692803035
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.2554349387
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -67.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.824
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.851
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6332354924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.824
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.526500282
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -67.851
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.489
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.803
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.323479771
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.489
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.502976526
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -87.803
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.783
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3975865772
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.3372738
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -90.783
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.4331172378
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.737689623
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -91.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.546
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3813439896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -9.546
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.712960855
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -91.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.747
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6850087406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.5742658062
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -64.747
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.677
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.211
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4840067182
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.677
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.162259804
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.211
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.568
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.363
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3733540898
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.568
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3557215194
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.363
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.583
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3885815157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.583
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1464780965
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -61.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.114
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.9731175805
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -11.114
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.829847053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -107.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.219
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.149
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.0797095621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -11.219
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.65861459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -112.149
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.22
operation: col(d(18_16)Mean)
inst: col
dD_actual | -113.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.0807247238
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -11.22
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -125.132829615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -113.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.299
operation: col(d(18_16)Mean)
inst: col
dD_actual | -114.436
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.1609225004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -11.299
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -125.833795081
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -114.436
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.69061012534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9929753399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:40:23
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.292
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8514952641
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.9948317665
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -67.292
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.976
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5489770689
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.8409627195
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -63.976
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.058
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.870783337
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -8.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.2837912311
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -65.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5516
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 36.897
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.239
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.391
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -12.953
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.2509
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.675
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.202
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4819763948
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.65042509848
measure_array: [ 4.5516 -12.953 ]
inst: fit_type1
actual_mean: -7.675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.985064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1536998458
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 17.8665332604
measure_array: [ 36.897 -107.2509]
inst: fit_type1
actual_mean: -62.202
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.05140700219
tarray: []
source_col: d(D_H)Mean