## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
220104_20-196_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.6077 | -76.949 |  | 6.76 |  | 0.51749 | 3.1651 |  | 0.3856
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.29197 | 3.9317
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.39189 | 5.6373
4 |  | EV | -10.3463 | -75.412 |  |  |  | 0.1194 | 1.8303
5 |  | WOOD1-deep-3 | -11.7629 | -87.8841 |  |  |  | 0.051409 | 0.69674
6 |  | WOOD1-2030-2 | -10.2252 | -81.1009 |  |  |  | 0.063174 | 0.4487
7 | * | SCBI1-C-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
8 |  | UNDE1-3040-3 | -6.2361 | -45.8167 |  |  |  | 0.16851 | 1.9711
9 |  | SCBI1-D-4 | 1.1227 | -27.806 |  |  |  | 0.22837 | 1.0653
10 | * | WOOD1-3040-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
11 | * | UNDE1-deep-3 | -3.3904 | -22.2277 |  |  |  | 2.7965 | 22.8817
12 | * | WOOD1-2030-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
13 |  | EV | -10.1817 | -72.673 |  |  |  | 0.039255 | 0.68353
14 | * | WOOD1-deep-2 | -0.881 | -0.56235 |  |  |  | 0 | 0
15 | * | WOOD1-3040-2 | -0.881 | -0.56235 |  |  |  | 0 | 0
16 |  | SCBI1-D-5 | -3.3725 | -47.2568 | -2.5585 | 551.4805 | 7.6659 | 0.19399 | 1.5939 |  | -11.6646
17 |  | WOOD1-010-2 | -6.8531 | -70.0325 |  | 2.3568 |  | 0.051624 | 1.1964
18 |  | WOOD1-5060-3 | -12.6162 | -95.6138 |  |  |  | 0.16906 | 1.4402
19 | * | WOOD1-3040-1 | -5.5641 | -40.6104 |  |  |  | 5.2187 | 42.2903
20 | * | SCBI1-C-2 | -2.1174 | -21.8078 |  |  |  | 1.3851 | 22.4811
21 | * | WOOD1-1020-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
22 |  | EV | -10.2027 | -73.0893 |  |  |  | 0.023022 | 0.82167
23 | * | WOOD1-deep-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
24 | * | WOOD1-010-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
25 | * | WOOD1-010-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
26 | * | WOOD1-6070-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
27 | * | WOOD1-2030-1 | -6.749 | -55.7282 |  |  |  | 3.7755 | 33.6335
28 | * | WOOD1-6070-2 | -5.8356 | -38.1599 |  |  |  | 5.5212 | 39.7026
29 | * | WOOD1-1020-2 | -0.881 | -0.56235 |  |  |  | 0 | 0
30 | * | WOOD1-1020-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
31 |  | EV | -10.2801 | -74.0754 |  |  |  | 0.080413 | 0.33816
32 |  | SCBI1-B-1 | -3.4663 | -39.1208 |  |  |  | 0.21259 | 1.9379
33 | * | SCBI1-A-4 | -0.44295 | -23.3261 | -2.6833 | 237.6585 |  | 0.28221 | 13.8791 |  | -19.2966
34 |  | SCBI1-B-3 | -3.2534 | -41.5939 |  |  |  | 0.047781 | 0.28954
35 |  | EV | -10.0136 | -71.6309 |  |  |  | 0.18922 | 1.5743

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -10.796 | -77.868 | -12.0678 | -85.7091 | -1.297 | 98.482 | 2022/01/04 11:05:10
1 |  | EV | -9.608 | -72.591 | -10.8368 | -79.9389 | -1.123 | 104.579 | 2022/01/04 11:14:06
1 |  | EV | -9.333 | -70.124 | -10.5519 | -77.2413 | -1.089 | 87.424 | 2022/01/04 11:23:02
1 |  | EV | -9.258 | -69.023 | -10.4741 | -76.0373 | -0.863 | 82.082 | 2022/01/04 11:31:59
1 |  | EV | -9.192 | -68.633 | -10.4058 | -75.6109 | -0.915 | 79.415 | 2022/01/04 11:40:56
1 |  | EV | -9.166 | -68.277 | -10.3788 | -75.2216 | -1.199 | 78.827 | 2022/01/04 11:49:53
1 |  | EV | -9.163 | -68.241 | -10.3757 | -75.1822 | -1.014 | 78.623 | 2022/01/04 11:58:51
1 |  | EV | -9.086 | -67.947 | -10.2959 | -74.8608 | -0.862 | 78.439 | 2022/01/04 12:07:46
1 |  | EV | -9.146 | -68.123 | -10.3581 | -75.0532 | -0.736 | 77.923 | 2022/01/04 12:16:43
1 |  | EV | -9.121 | -67.74 | -10.3322 | -74.6344 | -0.985 | 78.085 | 2022/01/04 12:25:40
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.6077 | -76.949 |  | 6.76 |  | 0.51749 | 3.1651 |  | 0.3856
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 1.94 | 6.748 | 1.1292 | 6.8164 | -1.117 | 77.971 | 2022/01/04 12:35:29
2 |  | PZ | 2.547 | 14.7 | 1.7582 | 15.5117 | -0.94 | 77.638 | 2022/01/04 12:44:26
2 |  | PZ | 2.715 | 16.79 | 1.9323 | 17.7971 | -1.087 | 77.75 | 2022/01/04 12:53:22
2 |  | PZ | 2.775 | 17.75 | 1.9945 | 18.8468 | -1.161 | 77.716 | 2022/01/04 13:02:19
2 |  | PZ | 2.777 | 18.305 | 1.9965 | 19.4537 | -0.924 | 77.539 | 2022/01/04 13:11:16
2 |  | PZ | 2.834 | 18.66 | 2.0556 | 19.8419 | -0.953 | 77.559 | 2022/01/04 13:20:14
2 |  | PZ | 2.855 | 19.201 | 2.0773 | 20.4335 | -0.989 | 77.605 | 2022/01/04 13:29:10
2 |  | PZ | 2.891 | 19.412 | 2.1147 | 20.6642 | -0.998 | 77.605 | 2022/01/04 13:38:07
2 |  | PZ | 2.897 | 19.44 | 2.1209 | 20.6948 | -1.191 | 77.829 | 2022/01/04 13:47:04
2 |  | PZ | 2.897 | 19.664 | 2.1209 | 20.9398 | -1.192 | 77.774 | 2022/01/04 13:56:00
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.29197 | 3.9317
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.381 | -93.373 | -14.7464 | -102.6635 | -1.031 | 78.125 | 2022/01/04 14:05:50
3 |  | UT2 | -14.276 | -105.128 | -15.6738 | -115.5173 | -0.906 | 77.647 | 2022/01/04 14:14:47
3 |  | UT2 | -14.433 | -108.108 | -15.8365 | -118.7758 | -0.908 | 77.512 | 2022/01/04 14:23:43
3 |  | UT2 | -14.491 | -109.403 | -15.8966 | -120.1919 | -0.806 | 77.516 | 2022/01/04 14:32:41
3 |  | UT2 | -14.542 | -110.162 | -15.9494 | -121.0218 | -0.564 | 77.457 | 2022/01/04 14:41:38
3 |  | UT2 | -14.614 | -110.794 | -16.024 | -121.7129 | -0.753 | 77.451 | 2022/01/04 14:50:35
3 |  | UT2 | -14.64 | -111.243 | -16.051 | -122.2039 | -0.85 | 77.54 | 2022/01/04 14:59:31
3 |  | UT2 | -14.623 | -111.467 | -16.0334 | -122.4488 | -0.688 | 77.395 | 2022/01/04 15:08:28
3 |  | UT2 | -14.683 | -111.603 | -16.0955 | -122.5975 | -0.834 | 77.493 | 2022/01/04 15:17:25
3 |  | UT2 | -14.681 | -111.849 | -16.0935 | -122.8665 | -0.823 | 77.437 | 2022/01/04 15:26:21
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.39189 | 5.6373
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.453 | -73.348 | -10.6762 | -80.7666 | -0.965 | 78.029 | 2022/01/04 15:36:10
4 |  | EV | -9.19 | -69.45 | -10.4037 | -76.5042 | -0.938 | 77.586 | 2022/01/04 15:45:07
4 |  | EV | -9.125 | -68.299 | -10.3363 | -75.2457 | -0.877 | 77.638 | 2022/01/04 15:54:03
4 |  | EV | -9.133 | -68.096 | -10.3446 | -75.0237 | -1.003 | 77.439 | 2022/01/04 16:02:59
4 |  | EV | -9.082 | -67.871 | -10.2918 | -74.7777 | -0.871 | 77.412 | 2022/01/04 16:11:56
4 |  | EV | -9.076 | -67.651 | -10.2856 | -74.5371 | -0.795 | 77.305 | 2022/01/04 16:20:53
4 |  | EV | -9.088 | -67.466 | -10.298 | -74.3348 | -0.913 | 77.248 | 2022/01/04 16:29:50
4 |  | EV | -9.092 | -67.566 | -10.3021 | -74.4441 | -0.9 | 77.181 | 2022/01/04 16:38:47
4 |  | EV | -9.068 | -67.445 | -10.2773 | -74.3118 | -0.924 | 77.039 | 2022/01/04 16:47:44
4 |  | EV | -9.039 | -67.319 | -10.2472 | -74.1741 | -1.047 | 77.096 | 2022/01/04 16:56:40
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3463 | -75.412 |  |  |  | 0.1194 | 1.8303
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | WOOD1-deep-3 | -10.45 | -78.849 | -11.7093 | -86.7818 | -1.072 | 77.497 | 2022/01/04 17:06:31
5 |  | WOOD1-deep-3 | -10.484 | -79.936 | -11.7445 | -87.9704 | -0.829 | 77.63 | 2022/01/04 17:15:28
5 |  | WOOD1-deep-3 | -10.572 | -80.286 | -11.8357 | -88.3532 | -0.835 | 77.785 | 2022/01/04 17:24:24
5 |  | WOOD1-deep-3 | -10.501 | -80.357 | -11.7621 | -88.4308 | -0.956 | 77.332 | 2022/01/04 17:33:21
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | WOOD1-deep-3 | -11.7629 | -87.8841 |  |  |  | 0.051409 | 0.69674
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | WOOD1-2030-2 | -9.11 | -74.301 | -10.3208 | -81.8087 | -0.895 | 77.881 | 2022/01/04 17:43:11
6 |  | WOOD1-2030-2 | -9.006 | -73.6 | -10.213 | -81.0422 | -1.108 | 77.617 | 2022/01/04 17:52:08
6 |  | WOOD1-2030-2 | -8.984 | -73.302 | -10.1902 | -80.7163 | -0.806 | 77.086 | 2022/01/04 18:01:03
6 |  | WOOD1-2030-2 | -8.971 | -73.412 | -10.1768 | -80.8366 | -0.788 | 76.847 | 2022/01/04 18:10:00
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | WOOD1-2030-2 | -10.2252 | -81.1009 |  |  |  | 0.063174 | 0.4487
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | SCBI1-C-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 18:19:50
7 |  | SCBI1-C-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 18:29:50
7 |  | SCBI1-C-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 18:39:49
7 |  | SCBI1-C-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 18:49:48
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 | * | SCBI1-C-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | UNDE1-3040-3 | -5.42 | -44.257 | -6.4972 | -48.9563 | -1.028 | 77.045 | 2022/01/04 19:00:38
8 |  | UNDE1-3040-3 | -5.099 | -41.07 | -6.1646 | -45.4714 | -1.149 | 76.271 | 2022/01/04 19:09:34
8 |  | UNDE1-3040-3 | -5.067 | -40.261 | -6.1314 | -44.5868 | -0.801 | 76.287 | 2022/01/04 19:18:31
8 |  | UNDE1-3040-3 | -5.086 | -39.955 | -6.1511 | -44.2522 | -0.981 | 76.785 | 2022/01/04 19:27:27
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | UNDE1-3040-3 | -6.2361 | -45.8167 |  |  |  | 0.16851 | 1.9711
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | SCBI1-D-4 | 1.6 | -26.44 | 0.77692 | -29.4739 | -1.067 | 77.005 | 2022/01/04 19:37:15
9 |  | SCBI1-D-4 | 1.982 | -24.785 | 1.1727 | -27.6642 | -1.165 | 76.687 | 2022/01/04 19:46:11
9 |  | SCBI1-D-4 | 2.045 | -24.426 | 1.238 | -27.2716 | -1.182 | 77.524 | 2022/01/04 19:55:08
9 |  | SCBI1-D-4 | 2.108 | -24.008 | 1.3033 | -26.8145 | -1.13 | 77.104 | 2022/01/04 20:04:04
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | SCBI1-D-4 | 1.1227 | -27.806 |  |  |  | 0.22837 | 1.0653
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | WOOD1-3040-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 20:13:53
10 |  | WOOD1-3040-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 20:23:52
10 |  | WOOD1-3040-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 20:33:51
10 |  | WOOD1-3040-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 20:43:49
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 | * | WOOD1-3040-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | UNDE1-deep-3 | -4.877 | -40.106 | -5.9346 | -44.4173 | -0.81 | 76.215 | 2022/01/04 20:54:41
11 |  | UNDE1-deep-3 | -4.81 | -39.147 | -5.8651 | -43.3687 | -0.828 | 76.337 | 2022/01/04 21:03:38
11 |  | UNDE1-deep-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 21:12:35
11 |  | UNDE1-deep-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 21:22:33
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 | * | UNDE1-deep-3 | -3.3904 | -22.2277 |  |  |  | 2.7965 | 22.8817
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | WOOD1-2030-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 21:33:24
12 |  | WOOD1-2030-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 21:43:23
12 |  | WOOD1-2030-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 21:53:22
12 |  | WOOD1-2030-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 22:03:20
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 | * | WOOD1-2030-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -9.002 | -64.996 | -10.2089 | -71.6339 | -0.883 | 76.302 | 2022/01/04 22:14:12
13 |  | EV | -8.969 | -65.963 | -10.1747 | -72.6913 | -0.983 | 76.411 | 2022/01/04 22:23:08
13 |  | EV | -8.923 | -66.236 | -10.127 | -72.9898 | -1.082 | 76.67 | 2022/01/04 22:32:05
13 |  | EV | -9.009 | -66.59 | -10.2161 | -73.3769 | -0.962 | 76.489 | 2022/01/04 22:41:01
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.1817 | -72.673 |  |  |  | 0.039255 | 0.68353
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD1-deep-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 22:50:51
14 |  | WOOD1-deep-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 23:00:49
14 |  | WOOD1-deep-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 23:10:48
14 |  | WOOD1-deep-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 23:20:47
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 | * | WOOD1-deep-2 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD1-3040-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 23:31:36
15 |  | WOOD1-3040-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 23:41:33
15 |  | WOOD1-3040-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/04 23:51:32
15 |  | WOOD1-3040-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 00:01:30
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 | * | WOOD1-3040-2 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | SCBI1-D-5 | -2.691 | -45.047 | -3.6694 | -49.8202 | -2.437 | 617.568 | 2022/01/05 00:12:19
16 |  | SCBI1-D-5 | -2.346 | -42.331 | -3.3119 | -46.8503 | -2.59 | 617.293 | 2022/01/05 00:21:15
16 |  | SCBI1-D-5 | -2.317 | -41.862 | -3.2819 | -46.3374 | -2.607 | 641.108 | 2022/01/05 00:30:12
16 |  | SCBI1-D-5 | -2.264 | -41.571 | -3.227 | -46.0192 | -2.6 | 640.465 | 2022/01/05 00:39:08
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | SCBI1-D-5 | -3.3725 | -47.2568 | -2.5585 | 551.4805 | 7.6659 | 0.19399 | 1.5939 |  | -11.6646
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | WOOD1-010-2 | -5.693 | -61.776 | -6.7801 | -68.1129 | -0.813 | 84.039 | 2022/01/05 00:48:56
17 |  | WOOD1-010-2 | -5.757 | -63.768 | -6.8464 | -70.2911 | -1.057 | 79.563 | 2022/01/05 00:57:53
17 |  | WOOD1-010-2 | -5.799 | -64.255 | -6.8899 | -70.8236 | -1.025 | 78.669 | 2022/01/05 01:06:49
17 |  | WOOD1-010-2 | -5.805 | -64.327 | -6.8961 | -70.9024 | -0.914 | 77.668 | 2022/01/05 01:15:46
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | WOOD1-010-2 | -6.8531 | -70.0325 |  | 2.3568 |  | 0.051624 | 1.1964
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | WOOD1-5060-3 | -11.088 | -84.849 | -12.3704 | -93.3427 | -0.892 | 76.876 | 2022/01/05 01:25:34
18 |  | WOOD1-5060-3 | -11.325 | -87.082 | -12.616 | -95.7844 | -0.991 | 76.917 | 2022/01/05 01:34:31
18 |  | WOOD1-5060-3 | -11.419 | -87.747 | -12.7134 | -96.5116 | -0.939 | 76.741 | 2022/01/05 01:43:27
18 |  | WOOD1-5060-3 | -11.469 | -88.026 | -12.7652 | -96.8167 | -0.928 | 76.234 | 2022/01/05 01:52:24
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | WOOD1-5060-3 | -12.6162 | -95.6138 |  |  |  | 0.16906 | 1.4402
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | WOOD1-3040-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 02:02:12
19 |  | WOOD1-3040-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 02:12:10
19 |  | WOOD1-3040-1 | -9.039 | -73.249 | -10.2472 | -80.6584 | -0.701 | 69.458 | 2022/01/05 02:22:08
19 |  | WOOD1-3040-1 | -9.039 | -73.249 | -10.2472 | -80.6584 | -0.701 | 69.458 | 2022/01/05 02:32:08
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 | * | WOOD1-3040-1 | -5.5641 | -40.6104 |  |  |  | 5.2187 | 42.2903
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | SCBI1-C-2 | -2.56 | -40.622 | -3.5337 | -44.9815 | -1.476 | 86.855 | 2022/01/05 02:41:57
20 |  | SCBI1-C-2 | -2.213 | -37.095 | -3.1741 | -41.1248 | -0.985 | 76.776 | 2022/01/05 02:50:52
20 |  | SCBI1-C-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 02:59:49
20 |  | SCBI1-C-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 03:09:48
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 | * | SCBI1-C-2 | -2.1174 | -21.8078 |  |  |  | 1.3851 | 22.4811
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | WOOD1-1020-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 03:20:39
21 |  | WOOD1-1020-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 03:30:37
21 |  | WOOD1-1020-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 03:40:36
21 |  | WOOD1-1020-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 03:50:34
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 | * | WOOD1-1020-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.987 | -65.151 | -10.1933 | -71.8034 | -0.901 | 76.904 | 2022/01/05 04:01:21
22 |  | EV | -9.01 | -66.385 | -10.2172 | -73.1527 | -0.854 | 76.322 | 2022/01/05 04:10:18
22 |  | EV | -8.968 | -66.802 | -10.1736 | -73.6087 | -0.956 | 76.598 | 2022/01/05 04:19:13
22 |  | EV | -9.019 | -66.97 | -10.2265 | -73.7924 | -0.895 | 76.448 | 2022/01/05 04:28:09
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.2027 | -73.0893 |  |  |  | 0.023022 | 0.82167
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | WOOD1-deep-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 04:37:55
23 |  | WOOD1-deep-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 04:47:54
23 |  | WOOD1-deep-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 04:57:52
23 |  | WOOD1-deep-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 05:07:51
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 | * | WOOD1-deep-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | WOOD1-010-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 05:18:40
24 |  | WOOD1-010-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 05:28:39
24 |  | WOOD1-010-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 05:38:37
24 |  | WOOD1-010-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 05:48:35
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 | * | WOOD1-010-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | WOOD1-010-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 05:59:24
25 |  | WOOD1-010-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 06:09:22
25 |  | WOOD1-010-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 06:19:21
25 |  | WOOD1-010-3 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 06:29:19
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 | * | WOOD1-010-3 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | WOOD1-6070-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 06:40:09
26 |  | WOOD1-6070-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 06:50:07
26 |  | WOOD1-6070-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 07:00:06
26 |  | WOOD1-6070-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 07:10:04
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 | * | WOOD1-6070-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | WOOD1-2030-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 07:20:54
27 |  | WOOD1-2030-1 | -7.594 | -67.408 | -8.7499 | -74.2714 | -1.096 | 75.829 | 2022/01/05 07:30:52
27 |  | WOOD1-2030-1 | -7.544 | -67.151 | -8.6981 | -73.9903 | -0.895 | 75.778 | 2022/01/05 07:39:48
27 |  | WOOD1-2030-1 | -7.514 | -67.241 | -8.667 | -74.0888 | -1.025 | 76.048 | 2022/01/05 07:48:44
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 | * | WOOD1-2030-1 | -6.749 | -55.7282 |  |  |  | 3.7755 | 33.6335
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | WOOD1-6070-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 07:58:32
28 |  | WOOD1-6070-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 08:07:29
28 |  | WOOD1-6070-2 | -9.554 | -68.784 | -10.7809 | -75.776 | -0.813 | 76.068 | 2022/01/05 08:16:25
28 |  | WOOD1-6070-2 | -9.572 | -68.75 | -10.7995 | -75.7388 | -1.142 | 76.372 | 2022/01/05 08:25:21
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 | * | WOOD1-6070-2 | -5.8356 | -38.1599 |  |  |  | 5.5212 | 39.7026
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | WOOD1-1020-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 08:35:06
29 |  | WOOD1-1020-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 08:45:05
29 |  | WOOD1-1020-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 08:55:03
29 |  | WOOD1-1020-2 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 09:05:01
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 | * | WOOD1-1020-2 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | WOOD1-1020-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 09:15:47
30 |  | WOOD1-1020-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 09:25:46
30 |  | WOOD1-1020-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 09:35:44
30 |  | WOOD1-1020-1 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 09:45:42
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 | * | WOOD1-1020-1 | -0.881 | -0.56235 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -9.19 | -67.719 | -10.4037 | -74.6114 | -1.185 | 76.404 | 2022/01/05 09:56:30
31 |  | EV | -9.041 | -67.142 | -10.2493 | -73.9805 | -0.911 | 75.704 | 2022/01/05 10:05:26
31 |  | EV | -9.014 | -67.11 | -10.2213 | -73.9455 | -1.183 | 76.223 | 2022/01/05 10:14:22
31 |  | EV | -9.038 | -66.944 | -10.2462 | -73.764 | -0.965 | 76.095 | 2022/01/05 10:23:18
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.2801 | -74.0754 |  |  |  | 0.080413 | 0.33816
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | SCBI1-B-1 | -2.809 | -38.097 | -3.7917 | -42.2205 | -1.177 | 77.445 | 2022/01/05 10:33:04
32 |  | SCBI1-B-1 | -2.44 | -34.865 | -3.4093 | -38.6864 | -0.921 | 76.466 | 2022/01/05 10:42:00
32 |  | SCBI1-B-1 | -2.351 | -34.269 | -3.3171 | -38.0347 | -0.973 | 76.889 | 2022/01/05 10:50:55
32 |  | SCBI1-B-1 | -2.38 | -33.818 | -3.3472 | -37.5415 | -1.101 | 77.085 | 2022/01/05 10:59:50
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | SCBI1-B-1 | -3.4663 | -39.1208 |  |  |  | 0.21259 | 1.9379
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | SCBI1-A-4 | 0.579 | -27.937 | -0.28104 | -31.1108 | -3.457 | 414.776 | 2022/01/05 11:09:37
33 |  | SCBI1-A-4 | 0.568 | -27.661 | -0.29244 | -30.809 | -3.478 | 412.647 | 2022/01/05 11:18:32
33 |  | SCBI1-A-4 |  |  | -0.881 | -0.56235 |  |  | 2022/01/05 11:27:38
33 |  | SCBI1-A-4 | 0.544 | -27.673 | -0.31731 | -30.8221 | -3.798 | 433.723 | 2022/01/05 11:37:37
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 | * | SCBI1-A-4 | -0.44295 | -23.3261 | -2.6833 | 237.6585 |  | 0.28221 | 13.8791 |  | -19.2966
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | SCBI1-B-3 | -2.347 | -37.099 | -3.313 | -41.1292 | -1.06 | 81.518 | 2022/01/05 11:47:23
34 |  | SCBI1-B-3 | -2.308 | -37.593 | -3.2726 | -41.6694 | -1.211 | 78.647 | 2022/01/05 11:56:19
34 |  | SCBI1-B-3 | -2.264 | -37.739 | -3.227 | -41.829 | -1.113 | 77.565 | 2022/01/05 12:05:16
34 |  | SCBI1-B-3 | -2.239 | -37.665 | -3.2011 | -41.7481 | -1.11 | 74.706 | 2022/01/05 12:14:11
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | SCBI1-B-3 | -3.2534 | -41.5939 |  |  |  | 0.047781 | 0.28954
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | EV | -8.538 | -62.699 | -9.7281 | -69.1222 | -0.867 | 77.116 | 2022/01/05 12:23:59
35 |  | EV | -8.842 | -65.234 | -10.0431 | -71.8942 | -0.775 | 76.793 | 2022/01/05 12:32:55
35 |  | EV | -8.947 | -65.97 | -10.1519 | -72.699 | -0.812 | 76.886 | 2022/01/05 12:41:51
35 |  | EV | -8.927 | -66.07 | -10.1312 | -72.8083 | -0.925 | 76.88 | 2022/01/05 12:50:48
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | EV | -10.0136 | -71.6309 |  |  |  | 0.18922 | 1.5743

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-18661 | 2022/01/04 11:05:10 | 1-01 | 1 | -10.796 | -77.868 | 13576 | -1 | 1 | EV | standard | H2O | 1641319815 | 0.638 | 0.732 | 67.986 | 0.013 | -0.005 | 1.659 | 98.482 | 30.58 | 2.284 | 70.816 | 0.856 | -1.297 | 0 | 0 | 1 | 44.438 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-18661 | 2022/01/04 11:14:06 | 1-01 | 2 | -9.608 | -72.591 | 13807 | -1 | 1 | EV | standard | H2O | 1641320351 | 0.194 | 0.763 | 61.481 | 0 | -0.001 | 1.497 | 104.579 | 32.166 | 2.332 | 68.727 | 0.856 | -1.123 | 0 | 0 | 1 | 43.812 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-18661 | 2022/01/04 11:23:02 | 1-01 | 3 | -9.333 | -70.124 | 12912 | -1 | 1 | EV | standard | H2O | 1641320888 | 0.188 | 0.736 | 66.949 | 0 | -0.002 | 1.647 | 87.424 | 29.969 | 2.198 | 65.081 | 0.856 | -1.089 | 0 | 0 | 1 | 43.562 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-18661 | 2022/01/04 11:31:59 | 1-01 | 4 | -9.258 | -69.023 | 12735 | 0 | 1 | EV | standard | H2O | 1641321425 | 0.19 | 0.713 | 60.533 | 0 | -0.001 | 1.486 | 82.082 | 29.502 | 2.161 | 64.987 | 0.856 | -0.863 | 0 | 0 | 1 | 43.375 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-18661 | 2022/01/04 11:40:56 | 1-01 | 5 | -9.192 | -68.633 | 12743 | 0 | 1 | EV | standard | H2O | 1641321961 | 0.214 | 0.805 | 60.765 | 0 | -0.005 | 1.492 | 79.415 | 29.486 | 2.164 | 64.977 | 0.856 | -0.915 | 0 | 0 | 1 | 43.312 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-18661 | 2022/01/04 11:49:53 | 1-01 | 6 | -9.166 | -68.277 | 12725 | 0 | 1 | EV | standard | H2O | 1641322499 | 0.212 | 0.717 | 61.948 | 0 | 0 | 1.519 | 78.827 | 29.873 | 2.149 | 63.539 | 0.856 | -1.199 | 0 | 0 | 1 | 43.812 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-18661 | 2022/01/04 11:58:51 | 1-01 | 7 | -9.163 | -68.241 | 12785 | 0 | 1 | EV | standard | H2O | 1641323036 | 0.195 | 0.776 | 59.391 | 0 | 0.003 | 1.46 | 78.623 | 29.652 | 2.173 | 65.298 | 0.856 | -1.014 | 0 | 0 | 1 | 43.688 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-18661 | 2022/01/04 12:07:46 | 1-01 | 8 | -9.086 | -67.947 | 12796 | 0 | 1 | EV | standard | H2O | 1641323571 | 0.214 | 0.778 | 56.463 | 0 | -0.001 | 1.381 | 78.439 | 29.567 | 2.184 | 65.541 | 0.855 | -0.862 | 0 | 0 | 1 | 43.375 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-18661 | 2022/01/04 12:16:43 | 1-01 | 9 | -9.146 | -68.123 | 12687 | 0 | 1 | EV | standard | H2O | 1641324108 | 0.199 | 0.882 | 61.105 | 0 | 0 | 1.502 | 77.923 | 29.431 | 2.163 | 65.871 | 0.855 | -0.736 | 0 | 0 | 1 | 43.312 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-18661 | 2022/01/04 12:25:40 | 1-01 | 10 | -9.121 | -67.74 | 12719 | 0 | 1 | EV | standard | H2O | 1641324645 | 0.195 | 0.809 | 60.415 | 0 | 0.003 | 1.482 | 78.085 | 29.644 | 2.147 | 64.428 | 0.856 | -0.985 | 0 | 0 | 1 | 43.5 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-18662 | 2022/01/04 12:35:29 | 1-02 | 1 | 1.94 | 6.748 | 12766 | -1 | 1 | PZ | standard | H2O | 1641325235 | 0.199 | 0.83 | 58.525 | 0 | 0.003 | 1.437 | 77.971 | 29.818 | 2.278 | 67.022 | 0.856 | -1.117 | 0 | 0 | 1 | 43.375 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-18662 | 2022/01/04 12:44:26 | 1-02 | 2 | 2.547 | 14.7 | 12670 | -1 | 1 | PZ | standard | H2O | 1641325771 | 0.222 | 0.894 | 57.221 | -0.001 | -0.008 | 1.412 | 77.638 | 29.551 | 2.287 | 67.276 | 0.856 | -0.94 | 0 | 0 | 1 | 43.375 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-18662 | 2022/01/04 12:53:22 | 1-02 | 3 | 2.715 | 16.79 | 12669 | -1 | 1 | PZ | standard | H2O | 1641326308 | 0.228 | 0.842 | 56.849 | 0 | -0.005 | 1.39 | 77.75 | 29.668 | 2.283 | 66.505 | 0.856 | -1.087 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-18662 | 2022/01/04 13:02:19 | 1-02 | 4 | 2.775 | 17.75 | 12672 | 0 | 1 | PZ | standard | H2O | 1641326845 | 0.191 | 0.826 | 61.272 | 0 | -0.002 | 1.506 | 77.716 | 29.878 | 2.285 | 66.223 | 0.856 | -1.161 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-18662 | 2022/01/04 13:11:16 | 1-02 | 5 | 2.777 | 18.305 | 12678 | 0 | 1 | PZ | standard | H2O | 1641327382 | 0.206 | 0.876 | 58.981 | 0 | -0.008 | 1.449 | 77.539 | 29.498 | 2.287 | 67.865 | 0.856 | -0.924 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-18662 | 2022/01/04 13:20:14 | 1-02 | 6 | 2.834 | 18.66 | 12682 | 0 | 1 | PZ | standard | H2O | 1641327919 | 0.206 | 0.855 | 56.405 | -0.001 | -0.004 | 1.39 | 77.559 | 29.73 | 2.289 | 66.904 | 0.856 | -0.953 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-18662 | 2022/01/04 13:29:10 | 1-02 | 7 | 2.855 | 19.201 | 12720 | 0 | 1 | PZ | standard | H2O | 1641328455 | 0.205 | 0.677 | 57.98 | 0.001 | -0.004 | 1.429 | 77.605 | 29.961 | 2.296 | 67.35 | 0.856 | -0.989 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-18662 | 2022/01/04 13:38:07 | 1-02 | 8 | 2.891 | 19.412 | 12706 | 0 | 1 | PZ | standard | H2O | 1641328992 | 0.195 | 0.85 | 60.005 | 0 | -0.004 | 1.475 | 77.605 | 29.796 | 2.298 | 67.109 | 0.856 | -0.998 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-18662 | 2022/01/04 13:47:04 | 1-02 | 9 | 2.897 | 19.44 | 12770 | 0 | 1 | PZ | standard | H2O | 1641329529 | 0.188 | 0.759 | 60.031 | 0 | -0.001 | 1.471 | 77.829 | 29.978 | 2.303 | 67.052 | 0.856 | -1.191 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-18662 | 2022/01/04 13:56:00 | 1-02 | 10 | 2.897 | 19.664 | 12738 | 0 | 1 | PZ | standard | H2O | 1641330066 | 0.2 | 0.846 | 56.418 | 0 | -0.001 | 1.384 | 77.774 | 30.093 | 2.292 | 66.239 | 0.856 | -1.192 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-18663 | 2022/01/04 14:05:50 | 1-03 | 1 | -13.381 | -93.373 | 12915 | -1 | 1 | UT2 | standard | H2O | 1641330656 | 0.205 | 0.835 | 57.916 | 0.001 | -0.004 | 1.422 | 78.125 | 30.344 | 2.145 | 63.95 | 0.856 | -1.031 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-18663 | 2022/01/04 14:14:47 | 1-03 | 2 | -14.276 | -105.128 | 12679 | -1 | 1 | UT2 | standard | H2O | 1641331192 | 0.184 | 0.798 | 56.85 | 0.001 | 0.005 | 1.395 | 77.647 | 29.167 | 2.101 | 64.602 | 0.856 | -0.906 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-18663 | 2022/01/04 14:23:43 | 1-03 | 3 | -14.433 | -108.108 | 12680 | -1 | 1 | UT2 | standard | H2O | 1641331728 | 0.178 | 0.773 | 57.052 | 0.001 | 0.006 | 1.4 | 77.512 | 29.61 | 2.092 | 63.769 | 0.856 | -0.908 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-18663 | 2022/01/04 14:32:41 | 1-03 | 4 | -14.491 | -109.403 | 12671 | 0 | 1 | UT2 | standard | H2O | 1641332266 | 0.197 | 0.707 | 57.431 | 0.001 | 0.006 | 1.407 | 77.516 | 29.499 | 2.094 | 63.946 | 0.856 | -0.806 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-18663 | 2022/01/04 14:41:38 | 1-03 | 5 | -14.542 | -110.162 | 12720 | 0 | 1 | UT2 | standard | H2O | 1641332803 | 0.187 | 0.85 | 58.974 | 0 | 0.008 | 1.448 | 77.457 | 29.481 | 2.112 | 65.138 | 0.856 | -0.564 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-18663 | 2022/01/04 14:50:35 | 1-03 | 6 | -14.614 | -110.794 | 12675 | 0 | 1 | UT2 | standard | H2O | 1641333340 | 0.192 | 0.78 | 58.196 | 0 | 0.003 | 1.439 | 77.451 | 29.357 | 2.093 | 64.154 | 0.856 | -0.753 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-18663 | 2022/01/04 14:59:31 | 1-03 | 7 | -14.64 | -111.243 | 12680 | 0 | 1 | UT2 | standard | H2O | 1641333876 | 0.182 | 0.896 | 57.886 | 0.001 | 0.006 | 1.422 | 77.54 | 29.531 | 2.083 | 63.262 | 0.856 | -0.85 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-18663 | 2022/01/04 15:08:28 | 1-03 | 8 | -14.623 | -111.467 | 12674 | 0 | 1 | UT2 | standard | H2O | 1641334413 | 0.165 | 0.729 | 51.932 | 0.001 | 0.004 | 1.272 | 77.395 | 29.412 | 2.098 | 64.341 | 0.856 | -0.688 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-18663 | 2022/01/04 15:17:25 | 1-03 | 9 | -14.683 | -111.603 | 12682 | 0 | 1 | UT2 | standard | H2O | 1641334950 | 0.193 | 0.709 | 60.095 | 0 | 0.002 | 1.466 | 77.493 | 29.426 | 2.105 | 63.87 | 0.856 | -0.834 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-18663 | 2022/01/04 15:26:21 | 1-03 | 10 | -14.681 | -111.849 | 12690 | 0 | 1 | UT2 | standard | H2O | 1641335486 | 0.194 | 0.825 | 57.387 | 0 | 0.003 | 1.41 | 77.437 | 29.665 | 2.088 | 64.241 | 0.856 | -0.823 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-18664 | 2022/01/04 15:36:10 | 1-04 | 1 | -9.453 | -73.348 | 12910 | -1 | 1 | EV | standard | H2O | 1641336076 | 0.211 | 0.767 | 58.401 | 0.001 | 0.001 | 1.432 | 78.029 | 30.097 | 2.175 | 64.969 | 0.857 | -0.965 | 0 | 0 | 1 | 43.25 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-18664 | 2022/01/04 15:45:07 | 1-04 | 2 | -9.19 | -69.45 | 12722 | -1 | 1 | EV | standard | H2O | 1641336613 | 0.211 | 0.826 | 57.14 | 0 | 0.004 | 1.402 | 77.586 | 29.554 | 2.162 | 64.602 | 0.856 | -0.938 | 0 | 0 | 1 | 43.188 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-18664 | 2022/01/04 15:54:03 | 1-04 | 3 | -9.125 | -68.299 | 12756 | -1 | 1 | EV | standard | H2O | 1641337148 | 0.202 | 0.724 | 59.399 | 0.001 | -0.002 | 1.455 | 77.638 | 29.535 | 2.18 | 66.414 | 0.856 | -0.877 | 0 | 0 | 1 | 43.125 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-18664 | 2022/01/04 16:02:59 | 1-04 | 4 | -9.133 | -68.096 | 12699 | 0 | 1 | EV | standard | H2O | 1641337685 | 0.194 | 0.74 | 55.819 | 0.001 | 0.001 | 1.37 | 77.439 | 29.708 | 2.151 | 64.779 | 0.856 | -1.003 | 0 | 0 | 1 | 43.062 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-18664 | 2022/01/04 16:11:56 | 1-04 | 5 | -9.082 | -67.871 | 12683 | 0 | 1 | EV | standard | H2O | 1641338222 | 0.216 | 0.775 | 60.246 | 0 | 0.001 | 1.49 | 77.412 | 29.563 | 2.166 | 65.003 | 0.856 | -0.871 | 0 | 0 | 1 | 42.938 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-18664 | 2022/01/04 16:20:53 | 1-04 | 6 | -9.076 | -67.651 | 12653 | 0 | 1 | EV | standard | H2O | 1641338759 | 0.216 | 0.786 | 58.235 | 0 | -0.001 | 1.429 | 77.305 | 29.272 | 2.157 | 65.324 | 0.856 | -0.795 | 0 | 0 | 1 | 42.812 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-18664 | 2022/01/04 16:29:50 | 1-04 | 7 | -9.088 | -67.466 | 12581 | 0 | 1 | EV | standard | H2O | 1641339296 | 0.199 | 0.743 | 59.818 | 0 | 0 | 1.47 | 77.248 | 29.264 | 2.145 | 64.654 | 0.856 | -0.913 | 0 | 0 | 1 | 42.812 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-18664 | 2022/01/04 16:38:47 | 1-04 | 8 | -9.092 | -67.566 | 12567 | 0 | 1 | EV | standard | H2O | 1641339833 | 0.199 | 0.712 | 60.56 | 0 | 0.003 | 1.488 | 77.181 | 29.226 | 2.151 | 64.588 | 0.856 | -0.9 | 0 | 0 | 1 | 42.812 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-18664 | 2022/01/04 16:47:44 | 1-04 | 9 | -9.068 | -67.445 | 12504 | 0 | 1 | EV | standard | H2O | 1641340369 | 0.202 | 0.785 | 60.022 | 0 | 0 | 1.485 | 77.039 | 29.097 | 2.127 | 64.343 | 0.856 | -0.924 | 0 | 0 | 1 | 42.812 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-18664 | 2022/01/04 16:56:40 | 1-04 | 10 | -9.039 | -67.319 | 12510 | 0 | 1 | EV | standard | H2O | 1641340906 | 0.194 | 0.709 | 55.259 | 0 | -0.001 | 1.363 | 77.096 | 29.175 | 2.116 | 64.101 | 0.856 | -1.047 | 0 | 0 | 1 | 42.875 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-18665 | 2022/01/04 17:06:31 | 1-05 | 1 | -10.45 | -78.849 | 12664 | -1 | 1 | WOOD1-deep-3 | 20-196 | H2O | 1641341496 | 0.209 | 0.79 | 56.315 | 0 | -0.001 | 1.383 | 77.497 | 29.505 | 2.129 | 64.18 | 0.856 | -1.072 | 0 | 0 | 1 | 42.875 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-18665 | 2022/01/04 17:15:28 | 1-05 | 2 | -10.484 | -79.936 | 12764 | -1 | 1 | WOOD1-deep-3 | 20-196 | H2O | 1641342033 | 0.205 | 0.798 | 59.496 | 0.001 | 0.002 | 1.461 | 77.63 | 29.591 | 2.154 | 65.588 | 0.856 | -0.829 | 0 | 0 | 1 | 42.812 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-18665 | 2022/01/04 17:24:24 | 1-05 | 3 | -10.572 | -80.286 | 12828 | -1 | 1 | WOOD1-deep-3 | 20-196 | H2O | 1641342570 | 0.2 | 0.736 | 56.097 | 0 | -0.002 | 1.379 | 77.785 | 29.502 | 2.166 | 65.697 | 0.856 | -0.835 | 0 | 0 | 1 | 42.875 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-18665 | 2022/01/04 17:33:21 | 1-05 | 4 | -10.501 | -80.357 | 12641 | 0 | 1 | WOOD1-deep-3 | 20-196 | H2O | 1641343107 | 0.199 | 0.679 | 62.862 | 0.001 | 0.002 | 1.554 | 77.332 | 29.665 | 2.127 | 64.117 | 0.856 | -0.956 | 0 | 0 | 1 | 42.875 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-18666 | 2022/01/04 17:43:11 | 1-06 | 1 | -9.11 | -74.301 | 12911 | -1 | 1 | WOOD1-2030-2 | 20-196 | H2O | 1641343696 | 0.183 | 0.792 | 58.032 | 0 | -0.001 | 1.426 | 77.881 | 30.287 | 2.184 | 64.885 | 0.856 | -0.895 | 0 | 0 | 1 | 43 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-18666 | 2022/01/04 17:52:08 | 1-06 | 2 | -9.006 | -73.6 | 12715 | -1 | 1 | WOOD1-2030-2 | 20-196 | H2O | 1641344233 | 0.194 | 0.802 | 55.52 | 0 | -0.001 | 1.372 | 77.617 | 29.755 | 2.139 | 63.68 | 0.855 | -1.108 | 0 | 0 | 1 | 43.062 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-18666 | 2022/01/04 18:01:03 | 1-06 | 3 | -8.984 | -73.302 | 12499 | -1 | 1 | WOOD1-2030-2 | 20-196 | H2O | 1641344769 | 0.2 | 0.719 | 60.811 | 0.001 | 0.001 | 1.494 | 77.086 | 28.942 | 2.122 | 64.752 | 0.856 | -0.806 | 0 | 0 | 1 | 43.062 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-18666 | 2022/01/04 18:10:00 | 1-06 | 4 | -8.971 | -73.412 | 12392 | 0 | 1 | WOOD1-2030-2 | 20-196 | H2O | 1641345306 | 0.218 | 0.842 | 63.583 | 0.001 | 0.003 | 1.566 | 76.847 | 28.751 | 2.102 | 63.276 | 0.856 | -0.788 | 0 | 0 | 1 | 43.062 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-18667 | 2022/01/04 18:19:50 | 1-07 | 1 | [5 empty] | SCBI1-C-3 | 20-196 | H2O | [17 empty] | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-18667 | 2022/01/04 18:29:50 | 1-07 | 2 | [5 empty] | SCBI1-C-3 | 20-196 | H2O | [17 empty] | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-18667 | 2022/01/04 18:39:49 | 1-07 | 3 | [5 empty] | SCBI1-C-3 | 20-196 | H2O | [17 empty] | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-18667 | 2022/01/04 18:49:48 | 1-07 | 4 | [5 empty] | SCBI1-C-3 | 20-196 | H2O | [17 empty] | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-18668 | 2022/01/04 19:00:38 | 1-08 | 1 | -5.42 | -44.257 | 12471 | -1 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641348343 | 0.194 | 0.789 | 52.746 | 0 | 0.004 | 1.305 | 77.045 | 29.35 | 2.151 | 63.831 | 0.856 | -1.028 | 0 | 0 | 1 | 43.438 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-18668 | 2022/01/04 19:09:34 | 1-08 | 2 | -5.099 | -41.07 | 12050 | -1 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641348880 | 0.212 | 0.84 | 50.552 | 0 | -0.004 | 1.242 | 76.271 | 28.114 | 2.088 | 61.998 | 0.856 | -1.149 | 0 | 0 | 1 | 43.188 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-18668 | 2022/01/04 19:18:31 | 1-08 | 3 | -5.067 | -40.261 | 12077 | -1 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641349417 | 0.198 | 0.778 | 56.571 | 0 | 0.001 | 1.399 | 76.287 | 27.651 | 2.102 | 64.487 | 0.855 | -0.801 | 0 | 0 | 1 | 43 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-18668 | 2022/01/04 19:27:27 | 1-08 | 4 | -5.086 | -39.955 | 12369 | 0 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641349953 | 0.201 | 0.735 | 55.269 | 0 | -0.003 | 1.359 | 76.785 | 29.03 | 2.158 | 63.91 | 0.855 | -0.981 | 0 | 0 | 1 | 43 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-18669 | 2022/01/04 19:37:15 | 1-09 | 1 | 1.6 | -26.44 | 12376 | -1 | 1 | SCBI1-D-4 | 20-196 | H2O | 1641350540 | 0.187 | 0.746 | 58.173 | 0 | 0 | 1.438 | 77.005 | 28.983 | 2.177 | 64.293 | 0.856 | -1.067 | 0 | 0 | 1 | 43 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-18669 | 2022/01/04 19:46:11 | 1-09 | 2 | 1.982 | -24.785 | 12171 | -1 | 1 | SCBI1-D-4 | 20-196 | H2O | 1641351077 | 0.187 | 0.856 | 59.613 | 0 | 0 | 1.474 | 76.687 | 28.256 | 2.158 | 63.099 | 0.856 | -1.165 | 0 | 0 | 1 | 43 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-18669 | 2022/01/04 19:55:08 | 1-09 | 3 | 2.045 | -24.426 | 12616 | -1 | 1 | SCBI1-D-4 | 20-196 | H2O | 1641351614 | 0.2 | 0.897 | 58.299 | 0.001 | -0.004 | 1.44 | 77.524 | 29.553 | 2.216 | 64.886 | 0.856 | -1.182 | 0 | 0 | 1 | 43.062 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-18669 | 2022/01/04 20:04:04 | 1-09 | 4 | 2.108 | -24.008 | 12343 | 0 | 1 | SCBI1-D-4 | 20-196 | H2O | 1641352150 | 0.232 | 0.808 | 58.062 | -0.001 | -0.003 | 1.434 | 77.104 | 28.611 | 2.185 | 64.36 | 0.856 | -1.13 | 0 | 0 | 1 | 43 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-18670 | 2022/01/04 20:13:53 | 1-10 | 1 | [5 empty] | WOOD1-3040-3 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-18670 | 2022/01/04 20:23:52 | 1-10 | 2 | [5 empty] | WOOD1-3040-3 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-18670 | 2022/01/04 20:33:51 | 1-10 | 3 | [5 empty] | WOOD1-3040-3 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-18670 | 2022/01/04 20:43:49 | 1-10 | 4 | [5 empty] | WOOD1-3040-3 | 20-196 | H2O | [17 empty] | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-18671 | 2022/01/04 20:54:41 | 1-11 | 1 | -4.877 | -40.106 | 12067 | -1 | 1 | UNDE1-deep-3 | 20-196 | H2O | 1641355186 | 0.193 | 0.792 | 49.326 | 0.001 | -0.002 | 1.218 | 76.215 | 27.632 | 2.105 | 63.479 | 0.855 | -0.81 | 0 | 0 | 1 | 43.438 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-18671 | 2022/01/04 21:03:38 | 1-11 | 2 | -4.81 | -39.147 | 12145 | -1 | 1 | UNDE1-deep-3 | 20-196 | H2O | 1641355724 | 0.203 | 0.733 | 56.948 | 0 | 0.002 | 1.409 | 76.337 | 27.845 | 2.122 | 64.464 | 0.855 | -0.828 | 0 | 0 | 1 | 43.312 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-18671 | 2022/01/04 21:12:35 | 1-11 | 3 | [5 empty] | UNDE1-deep-3 | 20-196 | H2O | [17 empty] | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-18671 | 2022/01/04 21:22:33 | 1-11 | 4 | [5 empty] | UNDE1-deep-3 | 20-196 | H2O | [17 empty] | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-18672 | 2022/01/04 21:33:24 | 1-12 | 1 | [5 empty] | WOOD1-2030-3 | 20-196 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-18672 | 2022/01/04 21:43:23 | 1-12 | 2 | [5 empty] | WOOD1-2030-3 | 20-196 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-18672 | 2022/01/04 21:53:22 | 1-12 | 3 | [5 empty] | WOOD1-2030-3 | 20-196 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-18672 | 2022/01/04 22:03:20 | 1-12 | 4 | [5 empty] | WOOD1-2030-3 | 20-196 | H2O | [17 empty] | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-18673 | 2022/01/04 22:14:12 | 1-13 | 1 | -9.002 | -64.996 | 12049 | -1 | 1 | EV | standard | H2O | 1641359958 | 0.181 | 0.783 | 47.828 | 0.001 | -0.001 | 1.179 | 76.302 | 27.623 | 2.043 | 62.724 | 0.855 | -0.883 | 0 | 0 | 1 | 43.625 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-18673 | 2022/01/04 22:23:08 | 1-13 | 2 | -8.969 | -65.963 | 12121 | -1 | 1 | EV | standard | H2O | 1641360494 | 0.196 | 0.76 | 53.972 | 0.001 | 0.002 | 1.336 | 76.411 | 27.914 | 2.059 | 62.573 | 0.856 | -0.983 | 0 | 0 | 1 | 43.375 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-18673 | 2022/01/04 22:32:05 | 1-13 | 3 | -8.923 | -66.236 | 12203 | -1 | 1 | EV | standard | H2O | 1641361030 | 0.176 | 0.758 | 50.356 | 0.001 | -0.003 | 1.243 | 76.67 | 28.268 | 2.072 | 62.324 | 0.856 | -1.082 | 0 | 0 | 1 | 43.188 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-18673 | 2022/01/04 22:41:01 | 1-13 | 4 | -9.009 | -66.59 | 12147 | 0 | 1 | EV | standard | H2O | 1641361567 | 0.2 | 0.728 | 54.452 | 0.001 | 0.001 | 1.344 | 76.489 | 27.97 | 2.06 | 62.911 | 0.856 | -0.962 | 0 | 0 | 1 | 43.188 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-18674 | 2022/01/04 22:50:51 | 1-14 | 1 | [5 empty] | WOOD1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-18674 | 2022/01/04 23:00:49 | 1-14 | 2 | [5 empty] | WOOD1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-18674 | 2022/01/04 23:10:48 | 1-14 | 3 | [5 empty] | WOOD1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-18674 | 2022/01/04 23:20:47 | 1-14 | 4 | [5 empty] | WOOD1-deep-2 | 20-196 | H2O | [17 empty] | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-18675 | 2022/01/04 23:31:36 | 1-15 | 1 | [5 empty] | WOOD1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-18675 | 2022/01/04 23:41:33 | 1-15 | 2 | [5 empty] | WOOD1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-18675 | 2022/01/04 23:51:32 | 1-15 | 3 | [5 empty] | WOOD1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-18675 | 2022/01/05 00:01:30 | 1-15 | 4 | [5 empty] | WOOD1-3040-2 | 20-196 | H2O | [17 empty] | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-18676 | 2022/01/05 00:12:19 | 1-16 | 1 | -2.691 | -45.047 | 11335 | -1 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641367046 | 0.267 | 1.081 | 49.902 | 0 | 0 | 1.251 | 617.568 | 5.396 | 2.314 | 52.686 | 0.857 | -2.437 | 0 | 0 | 1 | 43.438 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-18676 | 2022/01/05 00:21:15 | 1-16 | 2 | -2.346 | -42.331 | 11386 | -1 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641367581 | 0.255 | 1.054 | 48.344 | 0 | -0.001 | 1.205 | 617.293 | 5.594 | 2.318 | 52.649 | 0.857 | -2.59 | 0 | 0 | 1 | 43.188 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-18676 | 2022/01/05 00:30:12 | 1-16 | 3 | -2.317 | -41.862 | 11852 | -1 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641368117 | 0.244 | 0.988 | 51.494 | 0 | -0.002 | 1.279 | 641.108 | 6.146 | 2.413 | 54.827 | 0.858 | -2.607 | 0 | 0 | 1 | 43.062 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-18676 | 2022/01/05 00:39:08 | 1-16 | 4 | -2.264 | -41.571 | 11845 | 0 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641368654 | 0.271 | 1.011 | 48.444 | -0.001 | -0.001 | 1.197 | 640.465 | 5.644 | 2.41 | 55.343 | 0.858 | -2.6 | 0 | 0 | 1 | 43 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-18677 | 2022/01/05 00:48:56 | 1-17 | 1 | -5.693 | -61.776 | 12379 | -1 | 1 | WOOD1-010-2 | 20-196 | H2O | 1641369242 | 0.201 | 0.734 | 53.822 | 0.001 | -0.003 | 1.329 | 84.039 | 28.242 | 2.129 | 64.652 | 0.856 | -0.813 | 0 | 0 | 1 | 42.938 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-18677 | 2022/01/05 00:57:53 | 1-17 | 2 | -5.757 | -63.768 | 12183 | -1 | 1 | WOOD1-010-2 | 20-196 | H2O | 1641369778 | 0.197 | 0.794 | 54.969 | 0 | 0.001 | 1.359 | 79.563 | 28.016 | 2.081 | 62.393 | 0.856 | -1.057 | 0 | 0 | 1 | 43 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-18677 | 2022/01/05 01:06:49 | 1-17 | 3 | -5.799 | -64.255 | 12296 | -1 | 1 | WOOD1-010-2 | 20-196 | H2O | 1641370315 | 0.198 | 0.738 | 55.334 | 0.001 | 0.002 | 1.368 | 78.669 | 28.316 | 2.1 | 63.529 | 0.855 | -1.025 | 0 | 0 | 1 | 42.938 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-18677 | 2022/01/05 01:15:46 | 1-17 | 4 | -5.805 | -64.327 | 12203 | 0 | 1 | WOOD1-010-2 | 20-196 | H2O | 1641370851 | 0.184 | 0.773 | 59.275 | 0.001 | 0.004 | 1.458 | 77.668 | 28.414 | 2.085 | 62.314 | 0.856 | -0.914 | 0 | 0 | 1 | 43.062 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-18678 | 2022/01/05 01:25:34 | 1-18 | 1 | -11.088 | -84.849 | 12138 | -1 | 1 | WOOD1-5060-3 | 20-196 | H2O | 1641371439 | 0.199 | 0.732 | 54.781 | 0.001 | -0.001 | 1.354 | 76.876 | 28.06 | 2.033 | 62.499 | 0.856 | -0.892 | 0 | 0 | 1 | 43.125 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-18678 | 2022/01/05 01:34:31 | 1-18 | 2 | -11.325 | -87.082 | 11949 | -1 | 1 | WOOD1-5060-3 | 20-196 | H2O | 1641371976 | 0.182 | 0.796 | 51.883 | 0 | 0.003 | 1.292 | 76.917 | 27.429 | 2 | 60.936 | 0.855 | -0.991 | 0 | 0 | 1 | 43.125 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-18678 | 2022/01/05 01:43:27 | 1-18 | 3 | -11.419 | -87.747 | 11954 | -1 | 1 | WOOD1-5060-3 | 20-196 | H2O | 1641372513 | 0.216 | 0.795 | 56.411 | 0.001 | 0.004 | 1.402 | 76.741 | 27.426 | 2.002 | 61.821 | 0.856 | -0.939 | 0 | 0 | 1 | 43.125 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-18678 | 2022/01/05 01:52:24 | 1-18 | 4 | -11.469 | -88.026 | 11758 | 0 | 1 | WOOD1-5060-3 | 20-196 | H2O | 1641373050 | 0.192 | 0.768 | 53.641 | 0.001 | 0.004 | 1.334 | 76.234 | 26.889 | 1.975 | 60.036 | 0.856 | -0.928 | 0 | 0 | 1 | 43.188 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-18679 | 2022/01/05 02:02:12 | 1-19 | 1 | [5 empty] | WOOD1-3040-1 | 20-196 | H2O | [17 empty] | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-18679 | 2022/01/05 02:12:10 | 1-19 | 2 | [5 empty] | WOOD1-3040-1 | 20-196 | H2O | [17 empty] | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-18679 | 2022/01/05 02:22:08 | 1-19 | 3 | -9.039 | -73.249 | 8267 | 0 | 1 | WOOD1-3040-1 | 20-196 | H2O | 1641375438 | 0.218 | 0.912 | 42.079 | 0.002 | 0 | 1.149 | 69.458 | 17.006 | 1.429 | 46.79 | 0.854 | -0.701 | 0 | 0 | 1 | 43.625 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-18679 | 2022/01/05 02:32:08 | 1-19 | 4 | -9.039 | -73.249 | 8267 | 0 | 1 | WOOD1-3040-1 | 20-196 | H2O | 1641375438 | 0.218 | 0.912 | 42.079 | 0.002 | 0 | 1.149 | 69.458 | 17.006 | 1.429 | 46.79 | 0.854 | -0.701 | 0 | 0 | 1 | 43.625 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-18680 | 2022/01/05 02:41:57 | 1-20 | 1 | -2.56 | -40.622 | 12321 | -1 | 1 | SCBI1-C-2 | 20-196 | H2O | 1641376023 | 0.198 | 0.801 | 54.186 | 0 | 0.003 | 1.332 | 86.855 | 27.519 | 2.226 | 63.365 | 0.856 | -1.476 | 0 | 0 | 1 | 43.5 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-18680 | 2022/01/05 02:50:52 | 1-20 | 2 | -2.213 | -37.095 | 8601 | -1 | 1 | SCBI1-C-2 | 20-196 | H2O | 1641376561 | 0.227 | 0.974 | 45.425 | 0 | 0 | 1.216 | 76.776 | 16.86 | 1.592 | 49.53 | 0.854 | -0.985 | 0 | 0 | 1 | 43.438 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-18680 | 2022/01/05 02:59:49 | 1-20 | 3 | [5 empty] | SCBI1-C-2 | 20-196 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-18680 | 2022/01/05 03:09:48 | 1-20 | 4 | [5 empty] | SCBI1-C-2 | 20-196 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-18681 | 2022/01/05 03:20:39 | 1-21 | 1 | [5 empty] | WOOD1-1020-3 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-18681 | 2022/01/05 03:30:37 | 1-21 | 2 | [5 empty] | WOOD1-1020-3 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-18681 | 2022/01/05 03:40:36 | 1-21 | 3 | [5 empty] | WOOD1-1020-3 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-18681 | 2022/01/05 03:50:34 | 1-21 | 4 | [5 empty] | WOOD1-1020-3 | 20-196 | H2O | [17 empty] | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-18682 | 2022/01/05 04:01:21 | 1-22 | 1 | -8.987 | -65.151 | 12265 | -1 | 1 | EV | standard | H2O | 1641380788 | 0.207 | 0.807 | 52.272 | 0.001 | 0 | 1.281 | 76.904 | 28.292 | 2.098 | 62.824 | 0.856 | -0.901 | 0 | 0 | 1 | 43.625 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-18682 | 2022/01/05 04:10:18 | 1-22 | 2 | -9.01 | -66.385 | 12082 | -1 | 1 | EV | standard | H2O | 1641381323 | 0.218 | 0.859 | 53.028 | 0.001 | 0.005 | 1.316 | 76.322 | 27.994 | 2.067 | 62.641 | 0.856 | -0.854 | 0 | 0 | 1 | 43.375 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-18682 | 2022/01/05 04:19:13 | 1-22 | 3 | -8.968 | -66.802 | 12211 | -1 | 1 | EV | standard | H2O | 1641381859 | 0.191 | 0.768 | 57.029 | 0.001 | 0.001 | 1.411 | 76.598 | 28.288 | 2.084 | 62.932 | 0.856 | -0.956 | 0 | 0 | 1 | 43.25 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-18682 | 2022/01/05 04:28:09 | 1-22 | 4 | -9.019 | -66.97 | 12153 | 0 | 1 | EV | standard | H2O | 1641382395 | 0.206 | 0.799 | 57.436 | 0.001 | -0.002 | 1.42 | 76.448 | 28.187 | 2.069 | 62.659 | 0.856 | -0.895 | 0 | 0 | 1 | 43.188 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-18683 | 2022/01/05 04:37:55 | 1-23 | 1 | [5 empty] | WOOD1-deep-1 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-18683 | 2022/01/05 04:47:54 | 1-23 | 2 | [5 empty] | WOOD1-deep-1 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-18683 | 2022/01/05 04:57:52 | 1-23 | 3 | [5 empty] | WOOD1-deep-1 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-18683 | 2022/01/05 05:07:51 | 1-23 | 4 | [5 empty] | WOOD1-deep-1 | 20-196 | H2O | [17 empty] | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-18684 | 2022/01/05 05:18:40 | 1-24 | 1 | [5 empty] | WOOD1-010-1 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-18684 | 2022/01/05 05:28:39 | 1-24 | 2 | [5 empty] | WOOD1-010-1 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-18684 | 2022/01/05 05:38:37 | 1-24 | 3 | [5 empty] | WOOD1-010-1 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-18684 | 2022/01/05 05:48:35 | 1-24 | 4 | [5 empty] | WOOD1-010-1 | 20-196 | H2O | [17 empty] | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-18685 | 2022/01/05 05:59:24 | 1-25 | 1 | [5 empty] | WOOD1-010-3 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-18685 | 2022/01/05 06:09:22 | 1-25 | 2 | [5 empty] | WOOD1-010-3 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-18685 | 2022/01/05 06:19:21 | 1-25 | 3 | [5 empty] | WOOD1-010-3 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-18685 | 2022/01/05 06:29:19 | 1-25 | 4 | [5 empty] | WOOD1-010-3 | 20-196 | H2O | [17 empty] | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-18686 | 2022/01/05 06:40:09 | 1-26 | 1 | [5 empty] | WOOD1-6070-1 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-18686 | 2022/01/05 06:50:07 | 1-26 | 2 | [5 empty] | WOOD1-6070-1 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-18686 | 2022/01/05 07:00:06 | 1-26 | 3 | [5 empty] | WOOD1-6070-1 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-18686 | 2022/01/05 07:10:04 | 1-26 | 4 | [5 empty] | WOOD1-6070-1 | 20-196 | H2O | [17 empty] | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-18687 | 2022/01/05 07:20:54 | 1-27 | 1 | [5 empty] | WOOD1-2030-1 | 20-196 | H2O | [17 empty] | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-18687 | 2022/01/05 07:30:52 | 1-27 | 2 | -7.594 | -67.408 | 11751 | -1 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641393357 | 0.202 | 0.776 | 50.208 | 0.001 | 0.003 | 1.247 | 75.829 | 27.21 | 1.998 | 60.258 | 0.855 | -1.096 | 0 | 0 | 1 | 43.812 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-18687 | 2022/01/05 07:39:48 | 1-27 | 3 | -7.544 | -67.151 | 11813 | -1 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641393894 | 0.18 | 0.746 | 49.97 | 0 | 0.004 | 1.231 | 75.778 | 27.358 | 2.022 | 61.131 | 0.855 | -0.895 | 0 | 0 | 1 | 43.625 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-18687 | 2022/01/05 07:48:44 | 1-27 | 4 | -7.514 | -67.241 | 11901 | 0 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641394430 | 0.208 | 0.837 | 51.933 | 0.001 | 0.001 | 1.28 | 76.048 | 27.522 | 2.026 | 60.715 | 0.856 | -1.025 | 0 | 0 | 1 | 43.688 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-18688 | 2022/01/05 07:58:32 | 1-28 | 1 | [5 empty] | WOOD1-6070-2 | 20-196 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-18688 | 2022/01/05 08:07:29 | 1-28 | 2 | [5 empty] | WOOD1-6070-2 | 20-196 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-18688 | 2022/01/05 08:16:25 | 1-28 | 3 | -9.554 | -68.784 | 11952 | -1 | 1 | WOOD1-6070-2 | 20-196 | H2O | 1641396091 | 0.207 | 0.746 | 59.099 | 0.001 | 0.002 | 1.457 | 76.068 | 27.436 | 2.034 | 62.101 | 0.856 | -0.813 | 0 | 0 | 1 | 43.875 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-18688 | 2022/01/05 08:25:21 | 1-28 | 4 | -9.572 | -68.75 | 12029 | 0 | 1 | WOOD1-6070-2 | 20-196 | H2O | 1641396627 | 0.197 | 0.676 | 57.806 | 0.002 | 0.002 | 1.427 | 76.372 | 27.928 | 2.051 | 60.68 | 0.856 | -1.142 | 0 | 0 | 1 | 43.938 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-18689 | 2022/01/05 08:35:06 | 1-29 | 1 | [5 empty] | WOOD1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-18689 | 2022/01/05 08:45:05 | 1-29 | 2 | [5 empty] | WOOD1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-18689 | 2022/01/05 08:55:03 | 1-29 | 3 | [5 empty] | WOOD1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-18689 | 2022/01/05 09:05:01 | 1-29 | 4 | [5 empty] | WOOD1-1020-2 | 20-196 | H2O | [17 empty] | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-18690 | 2022/01/05 09:15:47 | 1-30 | 1 | [5 empty] | WOOD1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-18690 | 2022/01/05 09:25:46 | 1-30 | 2 | [5 empty] | WOOD1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-18690 | 2022/01/05 09:35:44 | 1-30 | 3 | [5 empty] | WOOD1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-18690 | 2022/01/05 09:45:42 | 1-30 | 4 | [5 empty] | WOOD1-1020-1 | 20-196 | H2O | [17 empty] | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-18691 | 2022/01/05 09:56:30 | 1-31 | 1 | -9.19 | -67.719 | 12025 | -1 | 1 | EV | standard | H2O | 1641402095 | 0.206 | 0.812 | 50.73 | 0.001 | -0.001 | 1.258 | 76.404 | 27.618 | 2.046 | 61.545 | 0.856 | -1.185 | 0 | 0 | 1 | 44.125 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-18691 | 2022/01/05 10:05:26 | 1-31 | 2 | -9.041 | -67.142 | 11769 | -1 | 1 | EV | standard | H2O | 1641402632 | 0.197 | 0.817 | 50.014 | 0 | 0.005 | 1.243 | 75.704 | 26.982 | 2.007 | 60.858 | 0.856 | -0.911 | 0 | 0 | 1 | 43.625 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-18691 | 2022/01/05 10:14:22 | 1-31 | 3 | -9.014 | -67.11 | 11937 | -1 | 1 | EV | standard | H2O | 1641403168 | 0.196 | 0.744 | 49.069 | 0.001 | 0 | 1.214 | 76.223 | 27.558 | 2.026 | 60.384 | 0.855 | -1.183 | 0 | 0 | 1 | 43.438 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-18691 | 2022/01/05 10:23:18 | 1-31 | 4 | -9.038 | -66.944 | 11968 | 0 | 1 | EV | standard | H2O | 1641403704 | 0.211 | 0.79 | 56.077 | 0 | -0.004 | 1.388 | 76.095 | 27.628 | 2.034 | 61.268 | 0.855 | -0.965 | 0 | 0 | 1 | 43.312 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-18692 | 2022/01/05 10:33:04 | 1-32 | 1 | -2.809 | -38.097 | 12555 | -1 | 1 | SCBI1-B-1 | 20-196 | H2O | 1641404289 | 0.21 | 0.827 | 56.699 | 0 | 0.002 | 1.393 | 77.445 | 29.431 | 2.17 | 64.85 | 0.856 | -1.177 | 0 | 0 | 1 | 43.438 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-18692 | 2022/01/05 10:42:00 | 1-32 | 2 | -2.44 | -34.865 | 12105 | -1 | 1 | SCBI1-B-1 | 20-196 | H2O | 1641404826 | 0.198 | 0.759 | 55.93 | 0 | 0.001 | 1.383 | 76.466 | 28.086 | 2.117 | 63.205 | 0.856 | -0.921 | 0 | 0 | 1 | 43.562 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-18692 | 2022/01/05 10:50:55 | 1-32 | 3 | -2.351 | -34.269 | 12330 | -1 | 1 | SCBI1-B-1 | 20-196 | H2O | 1641405360 | 0.198 | 0.763 | 57.649 | 0.001 | -0.001 | 1.418 | 76.889 | 28.577 | 2.155 | 64.466 | 0.856 | -0.973 | 0 | 0 | 1 | 43.688 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-18692 | 2022/01/05 10:59:50 | 1-32 | 4 | -2.38 | -33.818 | 12363 | 0 | 1 | SCBI1-B-1 | 20-196 | H2O | 1641405896 | 0.174 | 0.815 | 58.849 | 0 | 0.002 | 1.446 | 77.085 | 28.618 | 2.166 | 63.752 | 0.856 | -1.101 | 0 | 0 | 1 | 43.75 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-18693 | 2022/01/05 11:09:37 | 1-33 | 1 | 0.579 | -27.937 | 12250 | -1 | 1 | SCBI1-A-4 | 20-196 | H2O | 1641406483 | 0.23 | 1.082 | 55.186 | -0.001 | 0 | 1.359 | 414.776 | 12.468 | 2.735 | 61.417 | 0.857 | -3.457 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-18693 | 2022/01/05 11:18:32 | 1-33 | 2 | 0.568 | -27.661 | 12095 | -1 | 1 | SCBI1-A-4 | 20-196 | H2O | 1641407018 | 0.22 | 1.024 | 53.981 | -0.001 | 0.001 | 1.331 | 412.647 | 11.965 | 2.715 | 60.643 | 0.857 | -3.478 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-18693 | 2022/01/05 11:27:38 | 1-33 | 3 | [5 empty] | SCBI1-A-4 | 20-196 | H2O | [17 empty] | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-18693 | 2022/01/05 11:37:37 | 1-33 | 4 | 0.544 | -27.673 | 12667 | 0 | 1 | SCBI1-A-4 | 20-196 | H2O | 1641408162 | 0.246 | 1.104 | 54.873 | 0 | 0 | 1.35 | 433.723 | 12.92 | 2.846 | 62.917 | 0.857 | -3.798 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-18694 | 2022/01/05 11:47:23 | 1-34 | 1 | -2.347 | -37.099 | 12268 | -1 | 1 | SCBI1-B-3 | 20-196 | H2O | 1641408749 | 0.2 | 0.778 | 54.385 | 0.001 | -0.001 | 1.336 | 81.518 | 28.206 | 2.16 | 63.757 | 0.856 | -1.06 | 0 | 0 | 1 | 43.688 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-18694 | 2022/01/05 11:56:19 | 1-34 | 2 | -2.308 | -37.593 | 12136 | -1 | 1 | SCBI1-B-3 | 20-196 | H2O | 1641409284 | 0.23 | 0.731 | 55.326 | 0 | 0.002 | 1.364 | 78.647 | 27.649 | 2.125 | 63.301 | 0.856 | -1.211 | 0 | 0 | 1 | 43.562 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-18694 | 2022/01/05 12:05:16 | 1-34 | 3 | -2.264 | -37.739 | 12009 | -1 | 1 | SCBI1-B-3 | 20-196 | H2O | 1641409821 | 0.215 | 0.869 | 56.447 | 0 | 0.002 | 1.394 | 77.565 | 27.629 | 2.104 | 62.853 | 0.857 | -1.113 | 0 | 0 | 1 | 43.5 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-18694 | 2022/01/05 12:14:11 | 1-34 | 4 | -2.239 | -37.665 | 10652 | 0 | 1 | SCBI1-B-3 | 20-196 | H2O | 1641410358 | 0.198 | 0.852 | 52.398 | 0 | -0.002 | 1.323 | 74.706 | 23.662 | 1.882 | 56.84 | 0.855 | -1.11 | 0 | 0 | 1 | 43.5 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-18695 | 2022/01/05 12:23:59 | 1-35 | 1 | -8.538 | -62.699 | 12386 | -1 | 1 | EV | standard | H2O | 1641410944 | 0.195 | 0.804 | 58.394 | 0.001 | 0.001 | 1.442 | 77.116 | 28.852 | 2.122 | 63.534 | 0.856 | -0.867 | 0 | 0 | 1 | 43.5 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-18695 | 2022/01/05 12:32:55 | 1-35 | 2 | -8.842 | -65.234 | 12144 | -1 | 1 | EV | standard | H2O | 1641411481 | 0.202 | 0.691 | 52.962 | 0.001 | -0.003 | 1.311 | 76.793 | 27.94 | 2.085 | 63.375 | 0.856 | -0.775 | 0 | 0 | 1 | 43.438 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-18695 | 2022/01/05 12:41:51 | 1-35 | 3 | -8.947 | -65.97 | 12190 | -1 | 1 | EV | standard | H2O | 1641412017 | 0.2 | 0.781 | 60.215 | 0.001 | 0.003 | 1.488 | 76.886 | 27.799 | 2.084 | 63.841 | 0.856 | -0.812 | 0 | 0 | 1 | 43.438 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-18695 | 2022/01/05 12:50:48 | 1-35 | 4 | -8.927 | -66.07 | 12201 | 0 | 1 | EV | standard | H2O | 1641412553 | 0.187 | 0.772 | 56.536 | 0 | 0.004 | 1.394 | 76.88 | 28.123 | 2.084 | 62.558 | 0.856 | -0.925 | 0 | 0 | 1 | 43.438 | 1 | 35 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 77.62795
varray: [77.971, 77.638, 77.75, 77.716, 77.539, 77.559, 77.605, 77.605, 77.829, 77.774, 78.125, 77.647, 77.512, 77.516, 77.457, 77.451, 77.54, 77.395, 77.493, 77.437]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_shift
baseline_shift_STD | 0.187295565744
varray: [77.971, 77.638, 77.75, 77.716, 77.539, 77.559, 77.605, 77.605, 77.829, 77.774, 78.125, 77.647, 77.512, 77.516, 77.457, 77.451, 77.54, 77.395, 77.493, 77.437]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.93575
varray: [-1.117, -0.94, -1.087, -1.161, -0.924, -0.953, -0.989, -0.998, -1.191, -1.192, -1.031, -0.906, -0.908, -0.806, -0.564, -0.753, -0.85, -0.688, -0.834, -0.823]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.16786644911
varray: [-1.117, -0.94, -1.087, -1.161, -0.924, -0.953, -0.989, -0.998, -1.191, -1.192, -1.031, -0.906, -0.908, -0.806, -0.564, -0.753, -0.85, -0.688, -0.834, -0.823]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: ch4_ppm
residuals_MEAN | 2.19545
varray: [2.278, 2.287, 2.283, 2.285, 2.287, 2.289, 2.296, 2.298, 2.303, 2.292, 2.145, 2.101, 2.092, 2.094, 2.112, 2.093, 2.083, 2.098, 2.105, 2.088]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: residuals
residuals_STD | 0.0976882662781
varray: [2.278, 2.287, 2.283, 2.285, 2.287, 2.289, 2.296, 2.298, 2.303, 2.292, 2.145, 2.101, 2.092, 2.094, 2.112, 2.093, 2.083, 2.098, 2.105, 2.088]
inst: std_standards
operation: std_standards(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: residuals
slope_shift_MEAN | 29.67315
varray: [29.818, 29.551, 29.668, 29.878, 29.498, 29.73, 29.961, 29.796, 29.978, 30.093, 30.344, 29.167, 29.61, 29.499, 29.481, 29.357, 29.531, 29.412, 29.426, 29.665]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: slope_shift
slope_shift_STD | 0.281530642438
varray: [29.818, 29.551, 29.668, 29.878, 29.498, 29.73, 29.961, 29.796, 29.978, 30.093, 30.344, 29.167, 29.61, 29.499, 29.481, 29.357, 29.531, 29.412, 29.426, 29.665]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: slope_shift
baseline_curvature_MEAN | 65.5409
varray: [67.022, 67.276, 66.505, 66.223, 67.865, 66.904, 67.35, 67.109, 67.052, 66.239, 63.95, 64.602, 63.769, 63.946, 65.138, 64.154, 63.262, 64.341, 63.87, 64.241]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.53309888787
varray: [67.022, 67.276, 66.505, 66.223, 67.865, 66.904, 67.35, 67.109, 67.052, 66.239, 63.95, 64.602, 63.769, 63.946, 65.138, 64.154, 63.262, 64.341, 63.87, 64.241]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.19545
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 77.62795
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.93575
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 29.67315
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.146532399417
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.34198239942
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.195376532556
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.39082653256
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.293064798834
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.48851479883
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.219545
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.50359934733
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.43934934733
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.093647782872
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 77.7215977829
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 77.8152455657
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 0.561886697232
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 78.1898366972
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 3.37132018339
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 80.9992701834
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 1.40765321219
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 31.0808032122
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 4.59929666362
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 70.1401966636
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -5.34567281955
actual_array: [ 1.93 -15.84]
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121, 1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897, -13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681, -9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039, -10.45, -10.484, -10.572, -10.501, -9.11, -9.006, -8.984, -8.971, 0.0, 0.0, 0.0, 0.0, -5.42, -5.099, -5.067, -5.086, 1.6, 1.982, 2.045, 2.108, 0.0, 0.0, 0.0, 0.0, -4.877, -4.81, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -9.002, -8.969, -8.923, -9.009, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -2.691, -2.346, -2.317, -2.264, -5.693, -5.757, -5.799, -5.805, -11.088, -11.325, -11.419, -11.469, 0.0, 0.0, -9.039, -9.039, -2.56, -2.213, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -8.987, -9.01, -8.968, -9.019, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -7.594, -7.544, -7.514, 0.0, 0.0, -9.554, -9.572, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -9.19, -9.041, -9.014, -9.038, -2.809, -2.44, -2.351, -2.38, 0.579, 0.568, 0.0, 0.544, -2.347, -2.308, -2.264, -2.239, -8.538, -8.842, -8.947, -8.927]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -4.30869512195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0, 5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0, 16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0, 21682.0, 22219.0, 22755.0, 23292.0, 23882.0, 24419.0, 24954.0, 25491.0, 26081.0, 26681.0, 27280.0, 27879.0, 28529.0, 29065.0, 29602.0, 30138.0, 30726.0, 31262.0, 31799.0, 32335.0, 32924.0, 33523.0, 34122.0, 34720.0, 35372.0, 35909.0, 36446.0, 37044.0, 37695.0, 38294.0, 38893.0, 39491.0, 40143.0, 40679.0, 41216.0, 41752.0, 42342.0, 42940.0, 43539.0, 44138.0, 44787.0, 45384.0, 45983.0, 46581.0, 47230.0, 47766.0, 48303.0, 48839.0, 49427.0, 49964.0, 50500.0, 51037.0, 51625.0, 52162.0, 52698.0, 53235.0, 53823.0, 54421.0, 55019.0, 55619.0, 56208.0, 56743.0, 57280.0, 57879.0, 58530.0, 59128.0, 59727.0, 60325.0, 60972.0, 61509.0, 62044.0, 62580.0, 63166.0, 63765.0, 64363.0, 64962.0, 65611.0, 66210.0, 66808.0, 67406.0, 68055.0, 68653.0, 69252.0, 69850.0, 70500.0, 71098.0, 71697.0, 72295.0, 72945.0, 73543.0, 74079.0, 74615.0, 75203.0, 75740.0, 76276.0, 76812.0, 77397.0, 77996.0, 78594.0, 79192.0, 79838.0, 80437.0, 81035.0, 81633.0, 82281.0, 82817.0, 83353.0, 83889.0, 84475.0, 85011.0, 85546.0, 86081.0, 86668.0, 87203.0, 87749.0, 88348.0, 88934.0, 89470.0, 90007.0, 90542.0, 91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -40.7184051265
actual_array: [ 18.1 -119. ]
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74, 6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664, -93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849, -73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319, -78.849, -79.936, -80.286, -80.357, -74.301, -73.6, -73.302, -73.412, 0.0, 0.0, 0.0, 0.0, -44.257, -41.07, -40.261, -39.955, -26.44, -24.785, -24.426, -24.008, 0.0, 0.0, 0.0, 0.0, -40.106, -39.147, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -64.996, -65.963, -66.236, -66.59, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -45.047, -42.331, -41.862, -41.571, -61.776, -63.768, -64.255, -64.327, -84.849, -87.082, -87.747, -88.026, 0.0, 0.0, -73.249, -73.249, -40.622, -37.095, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -65.151, -66.385, -66.802, -66.97, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -67.408, -67.151, -67.241, 0.0, 0.0, -68.784, -68.75, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -67.719, -67.142, -67.11, -66.944, -38.097, -34.865, -34.269, -33.818, -27.937, -27.661, 0.0, -27.673, -37.099, -37.593, -37.739, -37.665, -62.699, -65.234, -65.97, -66.07]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -36.7233109756
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0, 5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0, 10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0, 16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0, 21682.0, 22219.0, 22755.0, 23292.0, 23882.0, 24419.0, 24954.0, 25491.0, 26081.0, 26681.0, 27280.0, 27879.0, 28529.0, 29065.0, 29602.0, 30138.0, 30726.0, 31262.0, 31799.0, 32335.0, 32924.0, 33523.0, 34122.0, 34720.0, 35372.0, 35909.0, 36446.0, 37044.0, 37695.0, 38294.0, 38893.0, 39491.0, 40143.0, 40679.0, 41216.0, 41752.0, 42342.0, 42940.0, 43539.0, 44138.0, 44787.0, 45384.0, 45983.0, 46581.0, 47230.0, 47766.0, 48303.0, 48839.0, 49427.0, 49964.0, 50500.0, 51037.0, 51625.0, 52162.0, 52698.0, 53235.0, 53823.0, 54421.0, 55019.0, 55619.0, 56208.0, 56743.0, 57280.0, 57879.0, 58530.0, 59128.0, 59727.0, 60325.0, 60972.0, 61509.0, 62044.0, 62580.0, 63166.0, 63765.0, 64363.0, 64962.0, 65611.0, 66210.0, 66808.0, 67406.0, 68055.0, 68653.0, 69252.0, 69850.0, 70500.0, 71098.0, 71697.0, 72295.0, 72945.0, 73543.0, 74079.0, 74615.0, 75203.0, 75740.0, 76276.0, 76812.0, 77397.0, 77996.0, 78594.0, 79192.0, 79838.0, 80437.0, 81035.0, 81633.0, 82281.0, 82817.0, 83353.0, 83889.0, 84475.0, 85011.0, 85546.0, 86081.0, 86668.0, 87203.0, 87749.0, 88348.0, 88934.0, 89470.0, 90007.0, 90542.0, 91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1955
varray: [2.284, 2.332, 2.198, 2.161, 2.164, 2.149, 2.173, 2.184, 2.163, 2.147]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: residuals
SAMP_residuals_STD | 0.0622383411661
varray: [2.284, 2.332, 2.198, 2.161, 2.164, 2.149, 2.173, 2.184, 2.163, 2.147]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0083
varray: [-1.297, -1.123, -1.089, -0.863, -0.915, -1.199, -1.014, -0.862, -0.736, -0.985]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.171699381996
varray: [-1.297, -1.123, -1.089, -0.863, -0.915, -1.199, -1.014, -0.862, -0.736, -0.985]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 84.3879
varray: [98.482, 104.579, 87.424, 82.082, 79.415, 78.827, 78.623, 78.439, 77.923, 78.085]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 9.58388438358
varray: [98.482, 104.579, 87.424, 82.082, 79.415, 78.827, 78.623, 78.439, 77.923, 78.085]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.987
varray: [30.58, 32.166, 29.969, 29.502, 29.486, 29.873, 29.652, 29.567, 29.431, 29.644]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.837631979651
varray: [30.58, 32.166, 29.969, 29.502, 29.486, 29.873, 29.652, 29.567, 29.431, 29.644]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: slope_shift
SAMP_d18O_STD | 0.517489549009
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.16509255789
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.9265
varray: [70.816, 68.727, 65.081, 64.987, 64.977, 63.539, 65.298, 65.541, 65.871, 64.428]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.17923157966
varray: [70.816, 68.727, 65.081, 64.987, 64.977, 63.539, 65.298, 65.541, 65.871, 64.428]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1955
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0083
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 84.3879
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.987
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 5.00000000003e-05
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 5.00000000003e-05
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.00227743742742
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.3856
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.3869
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
dD_mean | -69.8567
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
d18O_stdev | 0.517489549009
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
dD_stdev | 3.16509255789
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.3869
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.8567
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.517489549009
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.16509255789
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.6077084062
actual_array: [ 1.93 -15.84]
varray: [-10.796, -9.608, -9.333, -9.258, -9.192, -9.166, -9.163, -9.086, -9.146, -9.121]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.3869
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.9489653055
actual_array: [ 18.1 -119. ]
varray: [-77.868, -72.591, -70.124, -69.023, -68.633, -68.277, -68.241, -67.947, -68.123, -67.74]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -69.8567
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [1.0, 537.0, 1073.0, 1610.0, 2147.0, 2684.0, 3222.0, 3757.0, 4294.0, 4831.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.43245
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.75995
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2898
varray: [2.278, 2.287, 2.283, 2.285, 2.287, 2.289, 2.296, 2.298, 2.303, 2.292]
inst: mean_all
operation: mean_all(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: residuals
SAMP_residuals_STD | 0.00752477389841
varray: [2.278, 2.287, 2.283, 2.285, 2.287, 2.289, 2.296, 2.298, 2.303, 2.292]
inst: std_all
operation: std_all(residuals)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0552
varray: [-1.117, -0.94, -1.087, -1.161, -0.924, -0.953, -0.989, -0.998, -1.191, -1.192]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.106334691736
varray: [-1.117, -0.94, -1.087, -1.161, -0.924, -0.953, -0.989, -0.998, -1.191, -1.192]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.6986
varray: [77.971, 77.638, 77.75, 77.716, 77.539, 77.559, 77.605, 77.605, 77.829, 77.774]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.135806398147
varray: [77.971, 77.638, 77.75, 77.716, 77.539, 77.559, 77.605, 77.605, 77.829, 77.774]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.7971
varray: [29.818, 29.551, 29.668, 29.878, 29.498, 29.73, 29.961, 29.796, 29.978, 30.093]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.190123614759
varray: [29.818, 29.551, 29.668, 29.878, 29.498, 29.73, 29.961, 29.796, 29.978, 30.093]
inst: std_all
operation: std_all(slope_shift)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: slope_shift
SAMP_d18O_STD | 0.29197214479
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.93172961657
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.9545
varray: [67.022, 67.276, 66.505, 66.223, 67.865, 66.904, 67.35, 67.109, 67.052, 66.239]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.514064252361
varray: [67.022, 67.276, 66.505, 66.223, 67.865, 66.904, 67.35, 67.109, 67.052, 66.239]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2898
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0552
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.6986
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.7971
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.09435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.29752442552
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.4136
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.7128
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.067
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7128
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
dD_mean | 17.067
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
d18O_stdev | 0.29197214479
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
dD_stdev | 3.93172961657
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7128
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.067
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.29197214479
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.93172961657
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [1.94, 2.547, 2.715, 2.775, 2.777, 2.834, 2.855, 2.891, 2.897, 2.897]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.7128
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [6.748, 14.7, 16.79, 17.75, 18.305, 18.66, 19.201, 19.412, 19.44, 19.664]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 17.067
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [5420.0, 5957.0, 6493.0, 7030.0, 7567.0, 8105.0, 8641.0, 9178.0, 9715.0, 10251.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.33815
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.07065
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1011
varray: [2.145, 2.101, 2.092, 2.094, 2.112, 2.093, 2.083, 2.098, 2.105, 2.088]
inst: mean_all
operation: mean_all(residuals)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: residuals
SAMP_residuals_STD | 0.0175401127578
varray: [2.145, 2.101, 2.092, 2.094, 2.112, 2.093, 2.083, 2.098, 2.105, 2.088]
inst: std_all
operation: std_all(residuals)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8163
varray: [-1.031, -0.906, -0.908, -0.806, -0.564, -0.753, -0.85, -0.688, -0.834, -0.823]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.128354586985
varray: [-1.031, -0.906, -0.908, -0.806, -0.564, -0.753, -0.85, -0.688, -0.834, -0.823]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.5573
varray: [78.125, 77.647, 77.512, 77.516, 77.457, 77.451, 77.54, 77.395, 77.493, 77.437]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.211001342808
varray: [78.125, 77.647, 77.512, 77.516, 77.457, 77.451, 77.54, 77.395, 77.493, 77.437]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.5492
varray: [30.344, 29.167, 29.61, 29.499, 29.481, 29.357, 29.531, 29.412, 29.426, 29.665]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.311508087999
varray: [30.344, 29.167, 29.61, 29.499, 29.481, 29.357, 29.531, 29.412, 29.426, 29.665]
inst: std_all
operation: std_all(slope_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: slope_shift
SAMP_d18O_STD | 0.391889496896
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.63729765835
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.1273
varray: [63.95, 64.602, 63.769, 63.946, 65.138, 64.154, 63.262, 64.341, 63.87, 64.241]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.507038690174
varray: [63.95, 64.602, 63.769, 63.946, 65.138, 64.154, 63.262, 64.341, 63.87, 64.241]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1011
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8163
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.5573
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.5492
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.29752442552
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.4136
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.4364
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
dD_mean | -108.313
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
d18O_stdev | 0.391889496896
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
dD_stdev | 5.63729765835
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.4364
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.313
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.391889496896
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.63729765835
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.381, -14.276, -14.433, -14.491, -14.542, -14.614, -14.64, -14.623, -14.683, -14.681]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.4364
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-93.373, -105.128, -108.108, -109.403, -110.162, -110.794, -111.243, -111.467, -111.603, -111.849]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -108.313
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12989.0, 13526.0, 14062.0, 14599.0, 15136.0, 15672.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.52685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.07065
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.153
varray: [2.175, 2.162, 2.18, 2.151, 2.166, 2.157, 2.145, 2.151, 2.127, 2.116]
inst: mean_all
operation: mean_all(residuals)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: residuals
SAMP_residuals_STD | 0.0199888858008
varray: [2.175, 2.162, 2.18, 2.151, 2.166, 2.157, 2.145, 2.151, 2.127, 2.116]
inst: std_all
operation: std_all(residuals)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9233
varray: [-0.965, -0.938, -0.877, -1.003, -0.871, -0.795, -0.913, -0.9, -0.924, -1.047]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0711790387372
varray: [-0.965, -0.938, -0.877, -1.003, -0.871, -0.795, -0.913, -0.9, -0.924, -1.047]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.3973
varray: [78.029, 77.586, 77.638, 77.439, 77.412, 77.305, 77.248, 77.181, 77.039, 77.096]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.296516834823
varray: [78.029, 77.586, 77.638, 77.439, 77.412, 77.305, 77.248, 77.181, 77.039, 77.096]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.4491
varray: [30.097, 29.554, 29.535, 29.708, 29.563, 29.272, 29.264, 29.226, 29.097, 29.175]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.304168246571
varray: [30.097, 29.554, 29.535, 29.708, 29.563, 29.272, 29.264, 29.226, 29.097, 29.175]
inst: std_all
operation: std_all(slope_shift)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: slope_shift
SAMP_d18O_STD | 0.119398678571
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.83033235901
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.8777
varray: [64.969, 64.602, 66.414, 64.779, 65.003, 65.324, 64.654, 64.588, 64.343, 64.101]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.640514038358
varray: [64.969, 64.602, 66.414, 64.779, 65.003, 65.324, 64.654, 64.588, 64.343, 64.101]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.153
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9233
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.3973
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.4491
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04245
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04245
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.93354437587
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.6632
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1346
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
dD_mean | -68.4511
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
d18O_stdev | 0.119398678571
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
dD_stdev | 1.83033235901
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1346
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.4511
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.119398678571
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.83033235901
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3462751615
actual_array: [ 1.93 -15.84]
varray: [-9.453, -9.19, -9.125, -9.133, -9.082, -9.076, -9.088, -9.092, -9.068, -9.039]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.1346
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.411975674
actual_array: [ 18.1 -119. ]
varray: [-73.348, -69.45, -68.299, -68.096, -67.871, -67.651, -67.466, -67.566, -67.445, -67.319]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.4511
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [16261.0, 16798.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21091.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.47495
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.23065
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.144
varray: [2.129, 2.154, 2.166, 2.127]
inst: mean_all
operation: mean_all(residuals)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: residuals
SAMP_residuals_STD | 0.0191311264697
varray: [2.129, 2.154, 2.166, 2.127]
inst: std_all
operation: std_all(residuals)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.923
varray: [-1.072, -0.829, -0.835, -0.956]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115282262296
varray: [-1.072, -0.829, -0.835, -0.956]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.561
varray: [77.497, 77.63, 77.785, 77.332]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.192764104542
varray: [77.497, 77.63, 77.785, 77.332]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.56575
varray: [29.505, 29.591, 29.502, 29.665]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0779802325379
varray: [29.505, 29.591, 29.502, 29.665]
inst: std_all
operation: std_all(slope_shift)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0514093052537
varray: [-10.45, -10.484, -10.572, -10.501]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.696741463289
varray: [-78.849, -79.936, -80.286, -80.357]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.8955
varray: [64.18, 65.588, 65.697, 64.117]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.864091237467
varray: [64.18, 65.588, 65.697, 64.117]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.144
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.923
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.561
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.56575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05145
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05145
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.3434831128
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.6454
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-10.45, -10.484, -10.572, -10.501]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-78.849, -79.936, -80.286, -80.357]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -10.50175
varray: [-10.45, -10.484, -10.572, -10.501]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
dD_mean | -79.857
varray: [-78.849, -79.936, -80.286, -80.357]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0514093052537
varray: [-10.45, -10.484, -10.572, -10.501]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
dD_stdev | 0.696741463289
varray: [-78.849, -79.936, -80.286, -80.357]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
measured_d18O_mean | -10.50175
varray: [-10.45, -10.484, -10.572, -10.501]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
measured_dD_mean | -79.857
varray: [-78.849, -79.936, -80.286, -80.357]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0514093052537
varray: [-10.45, -10.484, -10.572, -10.501]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.696741463289
varray: [-78.849, -79.936, -80.286, -80.357]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.7629159086
actual_array: [ 1.93 -15.84]
varray: [-10.45, -10.484, -10.572, -10.501]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -10.50175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.8840516829
actual_array: [ 18.1 -119. ]
varray: [-78.849, -79.936, -80.286, -80.357]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -79.857
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [21682.0, 22219.0, 22755.0, 23292.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.48395
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.06695
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13675
varray: [2.184, 2.139, 2.122, 2.102]
inst: mean_all
operation: mean_all(residuals)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: residuals
SAMP_residuals_STD | 0.0349416179744
varray: [2.184, 2.139, 2.122, 2.102]
inst: std_all
operation: std_all(residuals)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.89925
varray: [-0.895, -1.108, -0.806, -0.788]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.146818198236
varray: [-0.895, -1.108, -0.806, -0.788]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.35775
varray: [77.881, 77.617, 77.086, 76.847]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.47459272013
varray: [77.881, 77.617, 77.086, 76.847]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.43375
varray: [30.287, 29.755, 28.942, 28.751]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.716287360864
varray: [30.287, 29.755, 28.942, 28.751]
inst: std_all
operation: std_all(slope_shift)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0631737023346
varray: [-9.11, -9.006, -8.984, -8.971]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.44869913825
varray: [-74.301, -73.6, -73.302, -73.412]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.14825
varray: [64.885, 63.68, 64.752, 63.276]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.79317773334
varray: [64.885, 63.68, 64.752, 63.276]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.89925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.35775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.43375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0587
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0587
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.67371153978
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.39265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-9.11, -9.006, -8.984, -8.971]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-74.301, -73.6, -73.302, -73.412]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.01775
varray: [-9.11, -9.006, -8.984, -8.971]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
dD_mean | -73.65375
varray: [-74.301, -73.6, -73.302, -73.412]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0631737023346
varray: [-9.11, -9.006, -8.984, -8.971]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
dD_stdev | 0.44869913825
varray: [-74.301, -73.6, -73.302, -73.412]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.01775
varray: [-9.11, -9.006, -8.984, -8.971]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.65375
varray: [-74.301, -73.6, -73.302, -73.412]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0631737023346
varray: [-9.11, -9.006, -8.984, -8.971]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.44869913825
varray: [-74.301, -73.6, -73.302, -73.412]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2251951986
actual_array: [ 1.93 -15.84]
varray: [-9.11, -9.006, -8.984, -8.971]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.01775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.1009477189
actual_array: [ 18.1 -119. ]
varray: [-74.301, -73.6, -73.302, -73.412]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.65375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [23882.0, 24419.0, 24954.0, 25491.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.4912
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.2702
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [26081.0, 26681.0, 27280.0, 27879.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12475
varray: [2.151, 2.088, 2.102, 2.158]
inst: mean_all
operation: mean_all(residuals)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: residuals
SAMP_residuals_STD | 0.0349416179744
varray: [2.151, 2.088, 2.102, 2.158]
inst: std_all
operation: std_all(residuals)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98975
varray: [-1.028, -1.149, -0.801, -0.981]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.144368452232
varray: [-1.028, -1.149, -0.801, -0.981]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.597
varray: [77.045, 76.271, 76.287, 76.785]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.382284361874
varray: [77.045, 76.271, 76.287, 76.785]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.53625
varray: [29.35, 28.114, 27.651, 29.03]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.789079791065
varray: [29.35, 28.114, 27.651, 29.03]
inst: std_all
operation: std_all(slope_shift)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: slope_shift
SAMP_d18O_STD | 0.168513105326
varray: [-5.42, -5.099, -5.067, -5.086]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.97111497635
varray: [-44.257, -41.07, -40.261, -39.955]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.5565
varray: [63.831, 61.998, 64.487, 63.91]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.07936169409
varray: [63.831, 61.998, 64.487, 63.91]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.597
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.53625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0707
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0707
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.22029652235
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.9844
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-5.42, -5.099, -5.067, -5.086]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-44.257, -41.07, -40.261, -39.955]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.168
varray: [-5.42, -5.099, -5.067, -5.086]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
dD_mean | -41.38575
varray: [-44.257, -41.07, -40.261, -39.955]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
d18O_stdev | 0.168513105326
varray: [-5.42, -5.099, -5.067, -5.086]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
dD_stdev | 1.97111497635
varray: [-44.257, -41.07, -40.261, -39.955]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.168
varray: [-5.42, -5.099, -5.067, -5.086]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.38575
varray: [-44.257, -41.07, -40.261, -39.955]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.168513105326
varray: [-5.42, -5.099, -5.067, -5.086]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.97111497635
varray: [-44.257, -41.07, -40.261, -39.955]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.23608448207
actual_array: [ 1.93 -15.84]
varray: [-5.42, -5.099, -5.067, -5.086]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.168
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -45.8166695246
actual_array: [ 18.1 -119. ]
varray: [-44.257, -41.07, -40.261, -39.955]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -41.38575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [28529.0, 29065.0, 29602.0, 30138.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.5032
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.03095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.184
varray: [2.177, 2.158, 2.216, 2.185]
inst: mean_all
operation: mean_all(residuals)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: residuals
SAMP_residuals_STD | 0.024152294577
varray: [2.177, 2.158, 2.216, 2.185]
inst: std_all
operation: std_all(residuals)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.136
varray: [-1.067, -1.165, -1.182, -1.13]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.050839617098
varray: [-1.067, -1.165, -1.182, -1.13]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.08
varray: [77.005, 76.687, 77.524, 77.104]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.345343307449
varray: [77.005, 76.687, 77.524, 77.104]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.85075
varray: [28.983, 28.256, 29.553, 28.611]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.554332255962
varray: [28.983, 28.256, 29.553, 28.611]
inst: std_all
operation: std_all(slope_shift)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: slope_shift
SAMP_d18O_STD | 0.228368671231
varray: [1.6, 1.982, 2.045, 2.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.06525345184
varray: [-26.44, -24.785, -24.426, -24.008]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.1595
varray: [64.293, 63.099, 64.886, 64.36]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.755090502302
varray: [64.293, 63.099, 64.886, 64.36]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.184
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.136
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.08
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.85075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01145
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01145
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.521533170876
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3814
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [1.6, 1.982, 2.045, 2.108]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-26.44, -24.785, -24.426, -24.008]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.93375
varray: [1.6, 1.982, 2.045, 2.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
dD_mean | -24.91475
varray: [-26.44, -24.785, -24.426, -24.008]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
d18O_stdev | 0.228368671231
varray: [1.6, 1.982, 2.045, 2.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
dD_stdev | 1.06525345184
varray: [-26.44, -24.785, -24.426, -24.008]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.93375
varray: [1.6, 1.982, 2.045, 2.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.91475
varray: [-26.44, -24.785, -24.426, -24.008]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.228368671231
varray: [1.6, 1.982, 2.045, 2.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.06525345184
varray: [-26.44, -24.785, -24.426, -24.008]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.12274843724
actual_array: [ 1.93 -15.84]
varray: [1.6, 1.982, 2.045, 2.108]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 1.93375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.8060290716
actual_array: [ 18.1 -119. ]
varray: [-26.44, -24.785, -24.426, -24.008]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -24.91475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [30726.0, 31262.0, 31799.0, 32335.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.44395
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.54795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [32924.0, 33523.0, 34122.0, 34720.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.05675
varray: [2.105, 2.122, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: residuals
SAMP_residuals_STD | 1.22024953049
varray: [2.105, 2.122, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.4095
varray: [-0.81, -0.828, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.472906967595
varray: [-0.81, -0.828, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 38.138
varray: [76.215, 76.337, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 44.0379972978
varray: [76.215, 76.337, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.86925
varray: [27.632, 27.845, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: slope_shift
SAMP_slope_shift_STD | 16.0150665182
varray: [27.632, 27.845, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: slope_shift
SAMP_d18O_STD | 2.79652979899
varray: [-4.877, -4.81, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 22.8817200894
varray: [-40.106, -39.147, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 31.98575
varray: [63.479, 64.464, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 36.9361517791
varray: [63.479, 64.464, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.05675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4095
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 38.138
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.86925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1387
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1387
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.8663599718
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.55515
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-4.877, -4.81, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-40.106, -39.147, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.42175
varray: [-4.877, -4.81, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
dD_mean | -19.81325
varray: [-40.106, -39.147, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
d18O_stdev | 2.79652979899
varray: [-4.877, -4.81, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
dD_stdev | 22.8817200894
varray: [-40.106, -39.147, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.42175
varray: [-4.877, -4.81, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.81325
varray: [-40.106, -39.147, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.79652979899
varray: [-4.877, -4.81, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
measured_dD_stdev | 22.8817200894
varray: [-40.106, -39.147, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.39042039862
actual_array: [ 1.93 -15.84]
varray: [-4.877, -4.81, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.42175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -22.2276621072
actual_array: [ 18.1 -119. ]
varray: [-40.106, -39.147, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -19.81325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [35372.0, 35909.0, 36446.0, 37044.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 76.5712
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -39.48995
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [37695.0, 38294.0, 38893.0, 39491.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0585
varray: [2.043, 2.059, 2.072, 2.06]
inst: mean_all
operation: mean_all(residuals)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: residuals
SAMP_residuals_STD | 0.0119023807142
varray: [2.043, 2.059, 2.072, 2.06]
inst: std_all
operation: std_all(residuals)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9775
varray: [-0.883, -0.983, -1.082, -0.962]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0818962758616
varray: [-0.883, -0.983, -1.082, -0.962]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.468
varray: [76.302, 76.411, 76.67, 76.489]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.154973115948
varray: [76.302, 76.411, 76.67, 76.489]
inst: std_all
operation: std_all(baseline_shift)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.94375
varray: [27.623, 27.914, 28.268, 27.97]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.264318463222
varray: [27.623, 27.914, 28.268, 27.97]
inst: std_all
operation: std_all(slope_shift)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: slope_shift
SAMP_d18O_STD | 0.039254511418
varray: [-9.002, -8.969, -8.923, -9.009]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.683526334533
varray: [-64.996, -65.963, -66.236, -66.59]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.633
varray: [62.724, 62.573, 62.324, 62.911]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.248090037419
varray: [62.724, 62.573, 62.324, 62.911]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0585
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.468
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.94375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13695
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13695
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.23790111367
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.9079
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-9.002, -8.969, -8.923, -9.009]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-64.996, -65.963, -66.236, -66.59]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.97575
varray: [-9.002, -8.969, -8.923, -9.009]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
dD_mean | -65.94625
varray: [-64.996, -65.963, -66.236, -66.59]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
d18O_stdev | 0.039254511418
varray: [-9.002, -8.969, -8.923, -9.009]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
dD_stdev | 0.683526334533
varray: [-64.996, -65.963, -66.236, -66.59]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.97575
varray: [-9.002, -8.969, -8.923, -9.009]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.94625
varray: [-64.996, -65.963, -66.236, -66.59]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.039254511418
varray: [-9.002, -8.969, -8.923, -9.009]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.683526334533
varray: [-64.996, -65.963, -66.236, -66.59]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1816748012
actual_array: [ 1.93 -15.84]
varray: [-9.002, -8.969, -8.923, -9.009]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.97575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.6729827325
actual_array: [ 18.1 -119. ]
varray: [-64.996, -65.963, -66.236, -66.59]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -65.94625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [40143.0, 40679.0, 41216.0, 41752.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.56945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.15995
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [42342.0, 42940.0, 43539.0, 44138.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [44787.0, 45384.0, 45983.0, 46581.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.36375
varray: [2.314, 2.318, 2.413, 2.41]
inst: mean_all
operation: mean_all(residuals)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: residuals
SAMP_residuals_STD | 0.0551747224732
varray: [2.314, 2.318, 2.413, 2.41]
inst: std_all
operation: std_all(residuals)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.5585
varray: [-2.437, -2.59, -2.607, -2.6]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0812998564984
varray: [-2.437, -2.59, -2.607, -2.6]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 629.1085
varray: [617.568, 617.293, 641.108, 640.465]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 13.4876149733
varray: [617.568, 617.293, 641.108, 640.465]
inst: std_all
operation: std_all(baseline_shift)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 5.695
varray: [5.396, 5.594, 6.146, 5.644]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.319167667535
varray: [5.396, 5.594, 6.146, 5.644]
inst: std_all
operation: std_all(slope_shift)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: slope_shift
SAMP_d18O_STD | 0.193993986161
varray: [-2.691, -2.346, -2.317, -2.264]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.5938867327
varray: [-45.047, -42.331, -41.862, -41.571]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 53.87625
varray: [52.686, 52.649, 54.827, 55.343]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.41163247696
varray: [52.686, 52.649, 54.827, 55.343]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.36375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.5585
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 629.1085
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 5.695
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1683
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1683
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.66585438065
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -11.66465
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-2.691, -2.346, -2.317, -2.264]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-45.047, -42.331, -41.862, -41.571]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.4045
varray: [-2.691, -2.346, -2.317, -2.264]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
dD_mean | -42.70275
varray: [-45.047, -42.331, -41.862, -41.571]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
d18O_stdev | 0.193993986161
varray: [-2.691, -2.346, -2.317, -2.264]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
dD_stdev | 1.5938867327
varray: [-45.047, -42.331, -41.862, -41.571]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.4045
varray: [-2.691, -2.346, -2.317, -2.264]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.70275
varray: [-45.047, -42.331, -41.862, -41.571]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.193993986161
varray: [-2.691, -2.346, -2.317, -2.264]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.5938867327
varray: [-45.047, -42.331, -41.862, -41.571]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.37254594967
actual_array: [ 1.93 -15.84]
varray: [-2.691, -2.346, -2.317, -2.264]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.4045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -47.2567771973
actual_array: [ 18.1 -119. ]
varray: [-45.047, -42.331, -41.862, -41.571]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -42.70275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [47230.0, 47766.0, 48303.0, 48839.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2642
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 551.48055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09875
varray: [2.129, 2.081, 2.1, 2.085]
inst: mean_all
operation: mean_all(residuals)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: residuals
SAMP_residuals_STD | 0.0217619698863
varray: [2.129, 2.081, 2.1, 2.085]
inst: std_all
operation: std_all(residuals)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95225
varray: [-0.813, -1.057, -1.025, -0.914]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.111233613025
varray: [-0.813, -1.057, -1.025, -0.914]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 79.98475
varray: [84.039, 79.563, 78.669, 77.668]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.81148506132
varray: [84.039, 79.563, 78.669, 77.668]
inst: std_all
operation: std_all(baseline_shift)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.247
varray: [28.242, 28.016, 28.316, 28.414]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.169347768414
varray: [28.242, 28.016, 28.316, 28.414]
inst: std_all
operation: std_all(slope_shift)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0516236379966
varray: [-5.693, -5.757, -5.799, -5.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.19638134946
varray: [-61.776, -63.768, -64.255, -64.327]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.222
varray: [64.652, 62.393, 63.529, 62.314]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.10315517192
varray: [64.652, 62.393, 63.529, 62.314]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 79.98475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.247
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0967
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0967
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.4045639846
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.3189
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-5.693, -5.757, -5.799, -5.805]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-61.776, -63.768, -64.255, -64.327]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.7635
varray: [-5.693, -5.757, -5.799, -5.805]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
dD_mean | -63.5315
varray: [-61.776, -63.768, -64.255, -64.327]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0516236379966
varray: [-5.693, -5.757, -5.799, -5.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
dD_stdev | 1.19638134946
varray: [-61.776, -63.768, -64.255, -64.327]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.7635
varray: [-5.693, -5.757, -5.799, -5.805]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.5315
varray: [-61.776, -63.768, -64.255, -64.327]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0516236379966
varray: [-5.693, -5.757, -5.799, -5.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.19638134946
varray: [-61.776, -63.768, -64.255, -64.327]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.85314154596
actual_array: [ 1.93 -15.84]
varray: [-5.693, -5.757, -5.799, -5.805]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.7635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.0325119636
actual_array: [ 18.1 -119. ]
varray: [-61.776, -63.768, -64.255, -64.327]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -63.5315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [49427.0, 49964.0, 50500.0, 51037.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.5292
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.3568
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0025
varray: [2.033, 2.0, 2.002, 1.975]
inst: mean_all
operation: mean_all(residuals)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: residuals
SAMP_residuals_STD | 0.0237557010701
varray: [2.033, 2.0, 2.002, 1.975]
inst: std_all
operation: std_all(residuals)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9375
varray: [-0.892, -0.991, -0.939, -0.928]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0409267638594
varray: [-0.892, -0.991, -0.939, -0.928]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.692
varray: [76.876, 76.917, 76.741, 76.234]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.314455084233
varray: [76.876, 76.917, 76.741, 76.234]
inst: std_all
operation: std_all(baseline_shift)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.451
varray: [28.06, 27.429, 27.426, 26.889]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.478829823633
varray: [28.06, 27.429, 27.426, 26.889]
inst: std_all
operation: std_all(slope_shift)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: slope_shift
SAMP_d18O_STD | 0.169056943069
varray: [-11.088, -11.325, -11.419, -11.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.44017429501
varray: [-84.849, -87.082, -87.747, -88.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.323
varray: [62.499, 60.936, 61.821, 60.036]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.07037656925
varray: [62.499, 60.936, 61.821, 60.036]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.692
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.451
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.19295
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.19295
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.78863103236
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.2179
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-11.088, -11.325, -11.419, -11.469]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-84.849, -87.082, -87.747, -88.026]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.32525
varray: [-11.088, -11.325, -11.419, -11.469]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
dD_mean | -86.926
varray: [-84.849, -87.082, -87.747, -88.026]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
d18O_stdev | 0.169056943069
varray: [-11.088, -11.325, -11.419, -11.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
dD_stdev | 1.44017429501
varray: [-84.849, -87.082, -87.747, -88.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.32525
varray: [-11.088, -11.325, -11.419, -11.469]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
measured_dD_mean | -86.926
varray: [-84.849, -87.082, -87.747, -88.026]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.169056943069
varray: [-11.088, -11.325, -11.419, -11.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.44017429501
varray: [-84.849, -87.082, -87.747, -88.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.6162265587
actual_array: [ 1.93 -15.84]
varray: [-11.088, -11.325, -11.419, -11.469]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -11.32525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -95.6138323497
actual_array: [ 18.1 -119. ]
varray: [-84.849, -87.082, -87.747, -88.026]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -86.926
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [51625.0, 52162.0, 52698.0, 53235.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.62545
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.93595
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.7145
varray: [0.0, 0.0, 1.429, 1.429]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: residuals
SAMP_residuals_STD | 0.825033534672
varray: [0.0, 0.0, 1.429, 1.429]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.3505
varray: [0.0, 0.0, -0.701, -0.701]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.404722538702
varray: [0.0, 0.0, -0.701, -0.701]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 34.729
varray: [0.0, 0.0, 69.458, 69.458]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 40.1015949974
varray: [0.0, 0.0, 69.458, 69.458]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 8.503
varray: [0.0, 0.0, 17.006, 17.006]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: slope_shift
SAMP_slope_shift_STD | 9.81841867784
varray: [0.0, 0.0, 17.006, 17.006]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: slope_shift
SAMP_d18O_STD | 5.21866908321
varray: [0.0, 0.0, -9.039, -9.039]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 42.2903298679
varray: [0.0, 0.0, -73.249, -73.249]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 23.395
varray: [0.0, 0.0, 46.79, 46.79]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 27.0142190954
varray: [0.0, 0.0, 46.79, 46.79]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.7145
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3505
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 34.729
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 8.503
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.48095
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.48095
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -67.4554191624
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -42.1459
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, -9.039, -9.039]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, -73.249, -73.249]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.5195
varray: [0.0, 0.0, -9.039, -9.039]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
dD_mean | -36.6245
varray: [0.0, 0.0, -73.249, -73.249]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
d18O_stdev | 5.21866908321
varray: [0.0, 0.0, -9.039, -9.039]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
dD_stdev | 42.2903298679
varray: [0.0, 0.0, -73.249, -73.249]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.5195
varray: [0.0, 0.0, -9.039, -9.039]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.6245
varray: [0.0, 0.0, -73.249, -73.249]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.21866908321
varray: [0.0, 0.0, -9.039, -9.039]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
measured_dD_stdev | 42.2903298679
varray: [0.0, 0.0, -73.249, -73.249]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.5641088214
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -9.039, -9.039]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -4.5195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.6103577126
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -73.249, -73.249]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -36.6245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [53823.0, 54421.0, 55019.0, 55619.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 76.91345
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -42.89895
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.9545
varray: [2.226, 1.592, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: residuals
SAMP_residuals_STD | 1.13214530869
varray: [2.226, 1.592, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.61525
varray: [-1.476, -0.985, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.73816681719
varray: [-1.476, -0.985, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.90775
varray: [86.855, 76.776, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 47.4150790387
varray: [86.855, 76.776, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 11.09475
varray: [27.519, 16.86, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: slope_shift
SAMP_slope_shift_STD | 13.529979684
varray: [27.519, 16.86, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: slope_shift
SAMP_d18O_STD | 1.38510971286
varray: [-2.56, -2.213, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 22.4811245771
varray: [-40.622, -37.095, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 28.22375
varray: [63.365, 49.53, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 33.0757910903
varray: [63.365, 49.53, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.9545
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.61525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.90775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 11.09475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.24095
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.24095
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -56.5237195108
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -37.31715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-2.56, -2.213, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-40.622, -37.095, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.19325
varray: [-2.56, -2.213, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
dD_mean | -19.42925
varray: [-40.622, -37.095, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
d18O_stdev | 1.38510971286
varray: [-2.56, -2.213, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
dD_stdev | 22.4811245771
varray: [-40.622, -37.095, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.19325
varray: [-2.56, -2.213, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.42925
varray: [-40.622, -37.095, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.38510971286
varray: [-2.56, -2.213, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
measured_dD_stdev | 22.4811245771
varray: [-40.622, -37.095, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.11744877312
actual_array: [ 1.93 -15.84]
varray: [-2.56, -2.213, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -1.19325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -21.8077673871
actual_array: [ 18.1 -119. ]
varray: [-40.622, -37.095, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -19.42925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [56208.0, 56743.0, 57280.0, 57879.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 76.67345
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -36.7202
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [58530.0, 59128.0, 59727.0, 60325.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0795
varray: [2.098, 2.067, 2.084, 2.069]
inst: mean_all
operation: mean_all(residuals)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: residuals
SAMP_residuals_STD | 0.014479871086
varray: [2.098, 2.067, 2.084, 2.069]
inst: std_all
operation: std_all(residuals)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9015
varray: [-0.901, -0.854, -0.956, -0.895]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0419086307738
varray: [-0.901, -0.854, -0.956, -0.895]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.568
varray: [76.904, 76.322, 76.598, 76.448]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.250806698475
varray: [76.904, 76.322, 76.598, 76.448]
inst: std_all
operation: std_all(baseline_shift)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.19025
varray: [28.292, 27.994, 28.288, 28.187]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.139562112815
varray: [28.292, 27.994, 28.288, 28.187]
inst: std_all
operation: std_all(slope_shift)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0230217288664
varray: [-8.987, -9.01, -8.968, -9.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.821667816091
varray: [-65.151, -66.385, -66.802, -66.97]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.764
varray: [62.824, 62.641, 62.932, 62.659]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.139017984448
varray: [62.824, 62.641, 62.932, 62.659]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0795
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9015
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.568
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.19025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.28137739416
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.7769
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-8.987, -9.01, -8.968, -9.019]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-65.151, -66.385, -66.802, -66.97]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.996
varray: [-8.987, -9.01, -8.968, -9.019]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
dD_mean | -66.327
varray: [-65.151, -66.385, -66.802, -66.97]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0230217288664
varray: [-8.987, -9.01, -8.968, -9.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
dD_stdev | 0.821667816091
varray: [-65.151, -66.385, -66.802, -66.97]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.996
varray: [-8.987, -9.01, -8.968, -9.019]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.327
varray: [-65.151, -66.385, -66.802, -66.97]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0230217288664
varray: [-8.987, -9.01, -8.968, -9.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.821667816091
varray: [-65.151, -66.385, -66.802, -66.97]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2026578499
actual_array: [ 1.93 -15.84]
varray: [-8.987, -9.01, -8.968, -9.019]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.0893236561
actual_array: [ 18.1 -119. ]
varray: [-65.151, -66.385, -66.802, -66.97]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.327
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [60972.0, 61509.0, 62044.0, 62580.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.54845
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.05995
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [63166.0, 63765.0, 64363.0, 64962.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [65611.0, 66210.0, 66808.0, 67406.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [68055.0, 68653.0, 69252.0, 69850.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [70500.0, 71098.0, 71697.0, 72295.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.5115
varray: [0.0, 1.998, 2.022, 2.026]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: residuals
SAMP_residuals_STD | 1.00774252664
varray: [0.0, 1.998, 2.022, 2.026]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.754
varray: [0.0, -1.096, -0.895, -1.025]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.509510222338
varray: [0.0, -1.096, -0.895, -1.025]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 56.91375
varray: [0.0, 75.829, 75.778, 76.048]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 37.9426807731
varray: [0.0, 75.829, 75.778, 76.048]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.5225
varray: [0.0, 27.21, 27.358, 27.522]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: slope_shift
SAMP_slope_shift_STD | 13.6822600838
varray: [0.0, 27.21, 27.358, 27.522]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: slope_shift
SAMP_d18O_STD | 3.77547754154
varray: [0.0, -7.594, -7.544, -7.514]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 33.6335018793
varray: [0.0, -67.408, -67.151, -67.241]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 45.526
varray: [0.0, 60.258, 61.131, 60.715]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 30.3527606982
varray: [0.0, 60.258, 61.131, 60.715]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.5115
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.754
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 56.91375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.5225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.68395
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.68395
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.1530665695
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -20.0149
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, -7.594, -7.544, -7.514]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, -67.408, -67.151, -67.241]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.663
varray: [0.0, -7.594, -7.544, -7.514]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
dD_mean | -50.45
varray: [0.0, -67.408, -67.151, -67.241]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
d18O_stdev | 3.77547754154
varray: [0.0, -7.594, -7.544, -7.514]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
dD_stdev | 33.6335018793
varray: [0.0, -67.408, -67.151, -67.241]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.663
varray: [0.0, -7.594, -7.544, -7.514]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.45
varray: [0.0, -67.408, -67.151, -67.241]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.77547754154
varray: [0.0, -7.594, -7.544, -7.514]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
measured_dD_stdev | 33.6335018793
varray: [0.0, -67.408, -67.151, -67.241]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.74900345206
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.594, -7.544, -7.514]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.663
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.7282078481
actual_array: [ 18.1 -119. ]
varray: [0.0, -67.408, -67.151, -67.241]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -50.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [72945.0, 73543.0, 74079.0, 74615.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 76.11645
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -20.7142
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.02125
varray: [0.0, 0.0, 2.034, 2.051]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: residuals
SAMP_residuals_STD | 1.17925834744
varray: [0.0, 0.0, 2.034, 2.051]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.48875
varray: [0.0, 0.0, -0.813, -1.142]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.580122616349
varray: [0.0, 0.0, -0.813, -1.142]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 38.11
varray: [0.0, 0.0, 76.068, 76.372]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 44.0058125252
varray: [0.0, 0.0, 76.068, 76.372]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 13.841
varray: [0.0, 0.0, 27.436, 27.928]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: slope_shift
SAMP_slope_shift_STD | 15.9834722552
varray: [0.0, 0.0, 27.436, 27.928]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: slope_shift
SAMP_d18O_STD | 5.5212055145
varray: [0.0, 0.0, -9.554, -9.572]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 39.7026483877
varray: [0.0, 0.0, -68.784, -68.75]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 30.69525
varray: [0.0, 0.0, 62.101, 60.68]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 35.448568898
varray: [0.0, 0.0, 62.101, 60.68]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.02125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.48875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 38.11
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 13.841
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1742
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1742
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -53.4833405452
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -34.84565
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, -9.554, -9.572]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, -68.784, -68.75]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.7815
varray: [0.0, 0.0, -9.554, -9.572]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
dD_mean | -34.3835
varray: [0.0, 0.0, -68.784, -68.75]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
d18O_stdev | 5.5212055145
varray: [0.0, 0.0, -9.554, -9.572]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
dD_stdev | 39.7026483877
varray: [0.0, 0.0, -68.784, -68.75]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.7815
varray: [0.0, 0.0, -9.554, -9.572]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
measured_dD_mean | -34.3835
varray: [0.0, 0.0, -68.784, -68.75]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.5212055145
varray: [0.0, 0.0, -9.554, -9.572]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
measured_dD_stdev | 39.7026483877
varray: [0.0, 0.0, -68.784, -68.75]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.83559320551
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -9.554, -9.572]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -4.7815
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -38.1598783698
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -68.784, -68.75]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -34.3835
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [75203.0, 75740.0, 76276.0, 76812.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 76.6067
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -39.51795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [77397.0, 77996.0, 78594.0, 79192.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.19545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.19545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.5409
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [79838.0, 80437.0, 81035.0, 81633.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 77.62795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -77.62795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.02825
varray: [2.046, 2.007, 2.026, 2.034]
inst: mean_all
operation: mean_all(residuals)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: residuals
SAMP_residuals_STD | 0.0163783393542
varray: [2.046, 2.007, 2.026, 2.034]
inst: std_all
operation: std_all(residuals)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.061
varray: [-1.185, -0.911, -1.183, -0.965]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.143731230659
varray: [-1.185, -0.911, -1.183, -0.965]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.1065
varray: [76.404, 75.704, 76.223, 76.095]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.296769832249
varray: [76.404, 75.704, 76.223, 76.095]
inst: std_all
operation: std_all(baseline_shift)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.4465
varray: [27.618, 26.982, 27.558, 27.628]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.311205719742
varray: [27.618, 26.982, 27.558, 27.628]
inst: std_all
operation: std_all(slope_shift)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: slope_shift
SAMP_d18O_STD | 0.080412996462
varray: [-9.19, -9.041, -9.014, -9.038]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.338159109493
varray: [-67.719, -67.142, -67.11, -66.944]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.01375
varray: [61.545, 60.858, 60.384, 61.268]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.505869795501
varray: [61.545, 60.858, 60.384, 61.268]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.02825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.061
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.1065
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.4465
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1672
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1672
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.61575075725
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.52715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-9.19, -9.041, -9.014, -9.038]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-67.719, -67.142, -67.11, -66.944]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.07075
varray: [-9.19, -9.041, -9.014, -9.038]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
dD_mean | -67.22875
varray: [-67.719, -67.142, -67.11, -66.944]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
d18O_stdev | 0.080412996462
varray: [-9.19, -9.041, -9.014, -9.038]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
dD_stdev | 0.338159109493
varray: [-67.719, -67.142, -67.11, -66.944]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.07075
varray: [-9.19, -9.041, -9.014, -9.038]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.22875
varray: [-67.719, -67.142, -67.11, -66.944]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.080412996462
varray: [-9.19, -9.041, -9.014, -9.038]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.338159109493
varray: [-67.719, -67.142, -67.11, -66.944]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2801137954
actual_array: [ 1.93 -15.84]
varray: [-9.19, -9.041, -9.014, -9.038]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.07075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.0753654889
actual_array: [ 18.1 -119. ]
varray: [-67.719, -67.142, -67.11, -66.944]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.22875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [82281.0, 82817.0, 83353.0, 83889.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.5997
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.52145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.152
varray: [2.17, 2.117, 2.155, 2.166]
inst: mean_all
operation: mean_all(residuals)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: residuals
SAMP_residuals_STD | 0.0241798814444
varray: [2.17, 2.117, 2.155, 2.166]
inst: std_all
operation: std_all(residuals)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.043
varray: [-1.177, -0.921, -0.973, -1.101]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117052694686
varray: [-1.177, -0.921, -0.973, -1.101]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.97125
varray: [77.445, 76.466, 76.889, 77.085]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.408011744112
varray: [77.445, 76.466, 76.889, 77.085]
inst: std_all
operation: std_all(baseline_shift)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.678
varray: [29.431, 28.086, 28.577, 28.618]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.557157667691
varray: [29.431, 28.086, 28.577, 28.618]
inst: std_all
operation: std_all(slope_shift)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: slope_shift
SAMP_d18O_STD | 0.212588804973
varray: [-2.809, -2.44, -2.351, -2.38]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.93786985717
varray: [-38.097, -34.865, -34.269, -33.818]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.06825
varray: [64.85, 63.205, 64.466, 63.752]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.733610421136
varray: [64.85, 63.205, 64.466, 63.752]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.152
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.043
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.97125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.678
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04345
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04345
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.97909312442
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.47265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-2.809, -2.44, -2.351, -2.38]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-38.097, -34.865, -34.269, -33.818]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.495
varray: [-2.809, -2.44, -2.351, -2.38]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
dD_mean | -35.26225
varray: [-38.097, -34.865, -34.269, -33.818]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
d18O_stdev | 0.212588804973
varray: [-2.809, -2.44, -2.351, -2.38]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
dD_stdev | 1.93786985717
varray: [-38.097, -34.865, -34.269, -33.818]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.495
varray: [-2.809, -2.44, -2.351, -2.38]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.26225
varray: [-38.097, -34.865, -34.269, -33.818]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.212588804973
varray: [-2.809, -2.44, -2.351, -2.38]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.93786985717
varray: [-38.097, -34.865, -34.269, -33.818]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.46632204418
actual_array: [ 1.93 -15.84]
varray: [-2.809, -2.44, -2.351, -2.38]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.495
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.1207702584
actual_array: [ 18.1 -119. ]
varray: [-38.097, -34.865, -34.269, -33.818]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -35.26225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [84475.0, 85011.0, 85546.0, 86081.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.47595
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.6567
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.074
varray: [2.735, 2.715, 0.0, 2.846]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: residuals
SAMP_residuals_STD | 1.38386680477
varray: [2.735, 2.715, 0.0, 2.846]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.68325
varray: [-3.457, -3.478, 0.0, -3.798]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.79562568018
varray: [-3.457, -3.478, 0.0, -3.798]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 315.2865
varray: [414.776, 412.647, 0.0, 433.723]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 210.404379854
varray: [414.776, 412.647, 0.0, 433.723]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 9.33825
varray: [12.468, 11.965, 0.0, 12.92]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: slope_shift
SAMP_slope_shift_STD | 6.23770782553
varray: [12.468, 11.965, 0.0, 12.92]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: slope_shift
SAMP_d18O_STD | 0.282211947538
varray: [0.579, 0.568, 0.0, 0.544]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 13.8790844889
varray: [-27.937, -27.661, 0.0, -27.673]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 46.24425
varray: [61.417, 60.643, 0.0, 62.917]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 30.8439491351
varray: [61.417, 60.643, 0.0, 62.917]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.074
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.68325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 315.2865
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.33825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12145
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12145
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.53189551117
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -19.29665
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [0.579, 0.568, 0.0, 0.544]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-27.937, -27.661, 0.0, -27.673]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.42275
varray: [0.579, 0.568, 0.0, 0.544]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
dD_mean | -20.81775
varray: [-27.937, -27.661, 0.0, -27.673]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
d18O_stdev | 0.282211947538
varray: [0.579, 0.568, 0.0, 0.544]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
dD_stdev | 13.8790844889
varray: [-27.937, -27.661, 0.0, -27.673]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.42275
varray: [0.579, 0.568, 0.0, 0.544]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.81775
varray: [-27.937, -27.661, 0.0, -27.673]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.282211947538
varray: [0.579, 0.568, 0.0, 0.544]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
measured_dD_stdev | 13.8790844889
varray: [-27.937, -27.661, 0.0, -27.673]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.442949671122
actual_array: [ 1.93 -15.84]
varray: [0.579, 0.568, 0.0, 0.544]
fit_rsq: 1.0
fit_intercept: 0.850225098481
except: True
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.42275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.3260585819
actual_array: [ 18.1 -119. ]
varray: [-27.937, -27.661, 0.0, -27.673]
fit_rsq: 1.0
fit_intercept: 0.514279358133
except: True
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -20.81775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [86668.0, 87203.0, 87749.0, 88348.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.55395
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 237.65855
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06775
varray: [2.16, 2.125, 2.104, 1.882]
inst: mean_all
operation: mean_all(residuals)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: residuals
SAMP_residuals_STD | 0.125969242278
varray: [2.16, 2.125, 2.104, 1.882]
inst: std_all
operation: std_all(residuals)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1235
varray: [-1.06, -1.211, -1.113, -1.11]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0631954639722
varray: [-1.06, -1.211, -1.113, -1.11]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 78.109
varray: [81.518, 78.647, 77.565, 74.706]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.81585215047
varray: [81.518, 78.647, 77.565, 74.706]
inst: std_all
operation: std_all(baseline_shift)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.7865
varray: [28.206, 27.649, 27.629, 23.662]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.10009468041
varray: [28.206, 27.649, 27.629, 23.662]
inst: std_all
operation: std_all(slope_shift)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0477807492616
varray: [-2.347, -2.308, -2.264, -2.239]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.289535259799
varray: [-37.099, -37.593, -37.739, -37.665]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.68775
varray: [63.757, 63.301, 62.853, 56.84]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.25283767143
varray: [63.757, 63.301, 62.853, 56.84]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1235
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 78.109
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.7865
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1277
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1277
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.8165751896
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.85315
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-2.347, -2.308, -2.264, -2.239]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-37.099, -37.593, -37.739, -37.665]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.2895
varray: [-2.347, -2.308, -2.264, -2.239]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
dD_mean | -37.524
varray: [-37.099, -37.593, -37.739, -37.665]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0477807492616
varray: [-2.347, -2.308, -2.264, -2.239]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
dD_stdev | 0.289535259799
varray: [-37.099, -37.593, -37.739, -37.665]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.2895
varray: [-2.347, -2.308, -2.264, -2.239]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.524
varray: [-37.099, -37.593, -37.739, -37.665]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0477807492616
varray: [-2.347, -2.308, -2.264, -2.239]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.289535259799
varray: [-37.099, -37.593, -37.739, -37.665]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.25338295664
actual_array: [ 1.93 -15.84]
varray: [-2.347, -2.308, -2.264, -2.239]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.2895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.5939392248
actual_array: [ 18.1 -119. ]
varray: [-37.099, -37.593, -37.739, -37.665]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -37.524
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [88934.0, 89470.0, 90007.0, 90542.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.5602
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.48105
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09375
varray: [2.122, 2.085, 2.084, 2.084]
inst: mean_all
operation: mean_all(residuals)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: residuals
SAMP_residuals_STD | 0.0188392321146
varray: [2.122, 2.085, 2.084, 2.084]
inst: std_all
operation: std_all(residuals)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.84475
varray: [-0.867, -0.775, -0.812, -0.925]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0655050888608
varray: [-0.867, -0.775, -0.812, -0.925]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.91875
varray: [77.116, 76.793, 76.886, 76.88]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.138196418188
varray: [77.116, 76.793, 76.886, 76.88]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.1785
varray: [28.852, 27.94, 27.799, 28.123]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.468182656663
varray: [28.852, 27.94, 27.799, 28.123]
inst: std_all
operation: std_all(slope_shift)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: slope_shift
SAMP_d18O_STD | 0.189223853324
varray: [-8.538, -8.842, -8.947, -8.927]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.57426964548
varray: [-62.699, -65.234, -65.97, -66.07]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.327
varray: [63.534, 63.375, 63.841, 62.558]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.547938561033
varray: [63.534, 63.375, 63.841, 62.558]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.84475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.91875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.1785
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1017
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1017
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.63230772735
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.2139
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4364
varray: [-8.538, -8.842, -8.947, -8.927]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.313
varray: [-62.699, -65.234, -65.97, -66.07]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8135
varray: [-8.538, -8.842, -8.947, -8.927]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
dD_mean | -64.99325
varray: [-62.699, -65.234, -65.97, -66.07]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
d18O_stdev | 0.189223853324
varray: [-8.538, -8.842, -8.947, -8.927]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
dD_stdev | 1.57426964548
varray: [-62.699, -65.234, -65.97, -66.07]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8135
varray: [-8.538, -8.842, -8.947, -8.927]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.99325
varray: [-62.699, -65.234, -65.97, -66.07]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.189223853324
varray: [-8.538, -8.842, -8.947, -8.927]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.57426964548
varray: [-62.699, -65.234, -65.97, -66.07]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.013551361
actual_array: [ 1.93 -15.84]
varray: [-8.538, -8.842, -8.947, -8.927]
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.8135
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.6309002632
actual_array: [ 18.1 -119. ]
varray: [-62.699, -65.234, -65.97, -66.07]
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -64.99325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: [91130.0, 91666.0, 92202.0, 92739.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.5342
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.7092
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.796
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.868
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.0678177408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -10.796
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.7091282501
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -77.868
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.608
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.591
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8368122128
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.608
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9388562769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -72.591
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.333
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.124
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5518572295
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.2412513958
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -70.124
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.023
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.474142234
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0373345031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -69.023
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.192
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.633
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.405753038
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.192
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6108789281
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.633
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.277
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3788118396
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2216015313
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.277
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.163
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3757032398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.163
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1822364013
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.086
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.947
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2959158445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.086
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8607545063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.947
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.146
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.123
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3580878408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.146
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.053206253
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.123
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.121
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.74
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3321828423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.121
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6344050088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.74
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.94
operation: col(d(18_16)Mean)
inst: col
dD_actual | 6.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.12922468687
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 1.94
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 6.81642287446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 6.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.547
operation: col(d(18_16)Mean)
inst: col
dD_actual | 14.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.75819805005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.547
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.5117427022
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 14.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.715
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.93227963987
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.715
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.7971071941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 16.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.775
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.75
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.99445163623
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.8468439943
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 17.75
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.777
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.99652403611
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.777
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.4537230818
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 18.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.834
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.66
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05558743265
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.834
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.8419070027
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 18.66
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.855
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.201
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07734763138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.855
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4334774286
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 19.201
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.891
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.11465082919
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.891
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6642008295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 19.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.12086802883
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6948181528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 19.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.897
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.664
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7128
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.664
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.12086802883
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.9397567395
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 19.664
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.381
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.373
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.746394584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -13.381
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.663471048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -93.373
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.276
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.128
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6737935297
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.276
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.51727947
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -105.128
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.433
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.108
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8364769202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.433
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.775837454
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -108.108
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.491
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.403
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8965765167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.491
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.191888658
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -109.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.542
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.162
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9494227136
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.542
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.021836816
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -110.162
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.614
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.794
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0240291092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.614
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.712913543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -110.794
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.64
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.243
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0509703077
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.64
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.203884192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -111.243
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.623
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.467
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0333549087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.623
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.448822779
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -111.467
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.683
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.603
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.095526905
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.683
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.597535492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -111.603
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.681
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.849
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0934545052
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -14.681
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.866530547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -111.849
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.453
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.348
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6762012222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.453
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.7666174829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.348
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.19
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4036806382
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.19
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.504248684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -69.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.299
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3363276421
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2456579997
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.299
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.133
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.096
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3446172416
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.133
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0236824055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.096
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.082
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.871
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2917710447
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.082
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.777650343
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.871
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.076
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.651
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2855538451
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.076
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5370856596
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.651
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.088
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.466
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2979882443
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.088
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3347926304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.466
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.092
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3021330441
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.092
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4441402138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.068
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2772642456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.068
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3118296379
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.319
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2472144473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1740516829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.319
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.45
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.849
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7092925618
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -10.45
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.7818280428
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -78.849
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.484
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.936
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7445233597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -10.484
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.9704362737
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -79.936
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.8357089544
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -10.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.3531528154
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -80.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.501
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.357
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7621387587
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -10.501
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.4307895996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -80.357
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.11
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.301
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.320784643
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.11
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8086999521
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -74.301
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.006
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.6
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2130198493
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.006
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.0421733929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.6
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.984
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.302
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1902234507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.984
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.7163175945
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.302
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.971
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1767528514
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.971
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.8365999362
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.257
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.49720686679
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.9563120115
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -44.257
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.099
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.07
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.16458668626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.099
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.4714045302
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -41.07
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.067
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.261
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.1314282882
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.067
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.586782581
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -40.261
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.086
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.955
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.15111608705
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.086
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.2521789759
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -39.955
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.6
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.776916707485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 1.6
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.4738530866
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -26.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.982
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.785
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.17274508432
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 1.982
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6641505822
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -24.785
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.426
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.2380256805
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.271592758
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -24.426
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.30330627668
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 2.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.8145198596
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -24.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.106
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.93455029972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -4.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.4172938268
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -40.106
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.81
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.86512490379
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -4.81
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.3686505025
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -39.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.996
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2088750496
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.6339073217
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -64.996
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.969
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.963
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1746804516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.969
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6912984527
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -65.963
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.236
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1270152544
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9898173552
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.236
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.009
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.59
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2161284491
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.009
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3769078003
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.59
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.047
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.66941723229
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.8201579199
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -45.047
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.346
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.331
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.31192825321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.346
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.8502775562
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -42.331
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.317
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.862
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.28187845497
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.317
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.3374373903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -41.862
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.264
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.571
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.22695985819
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.264
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.0192359228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -41.571
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.693
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.776
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.78008945024
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.693
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.112915138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -61.776
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.757
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.84640624636
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.757
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2911189982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -63.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.799
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.255
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.88992664381
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.799
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8236417291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -64.255
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.805
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.327
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.89614384344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -5.805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9023719892
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -64.327
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.088
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.849
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3703881231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -11.088
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.3426830435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -84.849
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.325
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.6159675087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -11.325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.7844145797
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -87.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.419
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.747
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.713370303
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -11.419
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.5115760089
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -87.747
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.026
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.7651803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -11.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.8166557665
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -88.026
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.249
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2472144473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6583633753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.249
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2472144473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6583633753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -73.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.53367504024
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.9815273568
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -40.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.213
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.095
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.17411366128
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.213
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.1248380922
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -37.095
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.987
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1933320505
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.987
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8033960759
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -65.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.01
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.385
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2171646491
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.01
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.1527452544
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.968
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.802
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1736442516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.968
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.608724677
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.802
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2264904485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.792428617
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.594
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.408
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.74990553495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -7.594
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2713710321
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.408
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.69809553798
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -7.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9903477429
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.514
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.6670095398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -7.514
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0887605679
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.554
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.784
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7808574161
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.554
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7759937789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.784
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.75
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.799509015
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7388156006
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -68.75
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.19
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.719
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4036806382
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.19
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6114420163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.719
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.142
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2492868472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9805064604
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.142
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.014
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.11
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2213094488
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.014
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9455152337
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -67.11
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.038
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.944
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2461782474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -9.038
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7639982453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.944
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.809
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.097
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.79168882513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.809
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.2205008773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -38.097
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.40933104751
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.6863869836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -34.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.351
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.269
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.31710925291
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.351
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.0346753868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -34.269
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.38
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.818
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.34715905115
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.38
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.5415177859
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -33.818
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.579
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.937
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.281043430597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.579
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.1107864093
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -27.937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.568
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.29244162993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.568
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.8089870793
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -27.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.881003195484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.562352049769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.317310428475
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: 0.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.8221087893
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -27.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.347
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.099
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.31296445315
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.347
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.1292119955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -37.099
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.308
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.593
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.27255265552
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.308
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.6693890573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -37.593
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.264
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.22695985819
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.264
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.829036529
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -37.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.2010548597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -2.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.7481193173
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -37.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 10:57:07
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.538
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7280782777
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.538
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1221933323
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -62.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.842
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.234
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0430830593
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.842
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8941545701
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -65.234
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.947
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1518840529
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.947
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6989527835
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -65.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7128
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.681
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4364
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.313
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.927
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.07
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1311600541
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.850225098481
measure_array: [ 2.7128 -14.4364]
inst: fit_type1
actual_mean: -8.927
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.965064715813
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8083003669
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.514279358133
measure_array: [ 17.067 -108.313]
inst: fit_type1
actual_mean: -66.07
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.914514952589
tarray: []
source_col: d(D_H)Mean