## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
220107_20-196_21-252_Precip_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.3724 | -75.8915 |  | 16.2351 |  | 0.30194 | 2.5125 |  | 1.2544
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27268 | 3.9052
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.38918 | 5.717
4 |  | EV | -10.3035 | -75.2037 |  |  |  | 0.12307 | 1.8219
5 |  | SCBI1-C-3 | 2.8995 | -34.1483 |  |  |  | 0.42654 | 2.4433
6 |  | UNDE1-3040-3 | -5.9365 | -43.3804 |  | 2.5019 | 10.3501 | 0.20671 | 0.24599 | 33.1762 | 5.967
7 |  | WOOD1-3040-3 | -13.1297 | -103.2417 |  |  |  | 0.21602 | 3.321
8 |  | UNDE1-deep-3 | -6.2212 | -47.1936 |  |  |  | 0.20122 | 3.3091
9 |  | WOOD1-2030-3 | -12.1801 | -98.4838 |  |  |  | 0.18014 | 2.9683
10 |  | WOOD1-deep-2 | -10.7187 | -78.8156 |  |  |  | 0.059579 | 1.1424
11 |  | WOOD1-3040-2 | -10.9706 | -83.5329 |  |  |  | 0.020467 | 0.22541
12 |  | SCBI1-D-5 | -3.4218 | -48.1655 | -2.8985 | 555.0391 | 9.6605 | 0.19043 | 1.9547 |  | -11.9817
13 |  | EV | -10.1137 | -72.6811 |  | 2.7859 |  | 0.11779 | 1.3189
14 |  | WOOD1-1020-3 | -10.3582 | -88.4136 |  |  |  | 0.023424 | 0.92281
15 |  | WOOD1-deep-1 | -10.1882 | -71.8285 |  |  |  | 0.024088 | 1.0034
16 |  | WOOD1-010-1 | -3.0432 | -63.4406 |  |  |  | 0.21437 | 0.55124
17 |  | WOOD1-010-3 | -7.484 | -77.0916 |  |  |  | 0.13609 | 0.87225
18 |  | WOOD1-6070-1 | -9.7568 | -75.2522 |  |  |  | 0.055088 | 0.13196
19 |  | WOOD1-2030-1 | -8.7255 | -74.254 |  |  |  | 0.035846 | 0.20987
20 |  | WOOD1-1020-2 | -9.4437 | -78.0863 |  |  |  | 0.024839 | 0.21611
21 |  | WOOD1-1020-1 | -8.0331 | -73.9851 |  |  |  | 0.063723 | 0.1989
22 |  | EV | -10.1882 | -73.9363 |  |  |  | 0.055476 | 0.063924
23 |  | SLCP-22-001 | -16.9541 | -125.6098 |  |  |  | 0.20246 | 2.8554
24 |  | 57 | -6.3816 | -43.5108 |  |  |  | 0.41481 | 4.8514
25 |  | 58 | -6.3589 | -39.403 |  |  |  | 0.019681 | 0.29761
26 |  | 59 | -6.1954 | -37.8363 |  |  |  | 0.040054 | 0.31176
27 |  | 60 | -6.4473 | -39.8263 |  |  |  | 0.024932 | 0.18942
28 |  | 61 | -6.5211 | -40.0906 |  |  |  | 0.014997 | 0.068946
29 |  | 62 | -6.5874 | -40.5994 |  |  |  | 0.012396 | 0.099965
30 |  | 63 | -6.3622 | -39.3782 |  |  |  | 0.022926 | 0.10226
31 |  | EV | -10.1178 | -71.5675 |  |  |  | 0.15022 | 1.7869
32 |  | 64 | -6.6309 | -42.1445 |  |  |  | 0.10941 | 1.7696
33 |  | 65 | -6.5874 | -40.6843 |  |  |  | 0.021764 | 0.079375
34 |  | 66 | -6.5987 | -40.5357 |  |  |  | 0.032604 | 0.14428
35 |  | 67 | -6.5631 | -40.4469 |  |  |  | 0.028024 | 0.1255
36 |  | 68 | -6.6343 | -40.4976 |  |  |  | 0.024739 | 0.11102
37 |  | 69 | -6.4607 | -39.8695 |  |  |  | 0.0232 | 0.044124
38 |  | 70 | -6.9765 | -43.0839 |  |  |  | 0.013451 | 0.19487
39 |  | 71 | -6.7534 | -41.7917 |  |  |  | 0.01698 | 0.032949
40 |  | EV | -10.1408 | -71.553 |  |  |  | 0.1139 | 1.5492
41 |  | 72 | -6.6998 | -42.849 |  |  |  | 0.093326 | 1.6074
42 |  | 73 | -6.7006 | -41.3142 |  |  |  | 0.01607 | 0.090903
43 |  | 74 | -6.7166 | -41.7193 |  |  |  | 0.020123 | 0.12324
44 |  | 75 | -6.6952 | -41.605 |  |  |  | 0.020116 | 0.097115
45 |  | 76 | -6.845 | -42.3721 |  |  |  | 0.033768 | 0.17605
46 |  | 77 | -6.673 | -41.152 |  |  |  | 0.030469 | 0.096426
47 |  | 78 | -6.9066 | -42.8174 |  |  |  | 0.041505 | 0.2662
48 |  | 79 | -6.7099 | -41.321 |  |  |  | 0.020998 | 0.078134
49 |  | EV | -10.123 | -71.5024 |  |  |  | 0.10942 | 1.6943

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.85 | -75.296 | -11.0968 | -82.9509 | -1.155 | 110.706 | 2022/01/07 10:56:50
1 |  | EV | -9.306 | -70.206 | -10.5356 | -77.4085 | -1.12 | 105.569 | 2022/01/07 11:05:45
1 |  | EV | -9.29 | -69.612 | -10.5191 | -76.7617 | -1.197 | 129.721 | 2022/01/07 11:14:42
1 |  | EV | -9.123 | -68.392 | -10.3468 | -75.4333 | -0.956 | 97.634 | 2022/01/07 11:23:38
1 |  | EV | -9.117 | -68.04 | -10.3406 | -75.05 | -1.095 | 89.742 | 2022/01/07 11:32:35
1 |  | EV | -8.678 | -66.923 | -9.8878 | -73.8337 | -1.052 | 83.997 | 2022/01/07 11:41:31
1 |  | EV | -8.989 | -67.386 | -10.2086 | -74.3379 | -0.925 | 81.283 | 2022/01/07 11:50:27
1 |  | EV | -9.065 | -67.561 | -10.287 | -74.5284 | -0.94 | 79.614 | 2022/01/07 11:59:24
1 |  | EV | -9.023 | -67.427 | -10.2437 | -74.3825 | -0.953 | 78.908 | 2022/01/07 12:08:21
1 |  | EV | -9.037 | -67.285 | -10.2581 | -74.2279 | -0.937 | 78.518 | 2022/01/07 12:17:17
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.3724 | -75.8915 |  | 16.2351 |  | 0.30194 | 2.5125 |  | 1.2544
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.041 | 7.133 | 1.1693 | 6.8046 | -1.018 | 77.777 | 2022/01/07 12:27:04
2 |  | PZ | 2.657 | 15.348 | 1.8048 | 15.7498 | -1.107 | 77.705 | 2022/01/07 12:36:00
2 |  | PZ | 2.776 | 17.392 | 1.9275 | 17.9754 | -1.118 | 77.489 | 2022/01/07 12:44:55
2 |  | PZ | 2.862 | 18.28 | 2.0162 | 18.9424 | -1.124 | 77.641 | 2022/01/07 12:53:52
2 |  | PZ | 2.89 | 18.829 | 2.0451 | 19.5402 | -1.072 | 77.453 | 2022/01/07 13:02:48
2 |  | PZ | 2.882 | 19.176 | 2.0369 | 19.918 | -1.141 | 77.368 | 2022/01/07 13:11:44
2 |  | PZ | 2.889 | 19.458 | 2.0441 | 20.2251 | -1.181 | 77.368 | 2022/01/07 13:20:41
2 |  | PZ | 2.936 | 19.673 | 2.0926 | 20.4592 | -1.118 | 77.375 | 2022/01/07 13:29:37
2 |  | PZ | 2.917 | 19.889 | 2.073 | 20.6944 | -0.977 | 77.165 | 2022/01/07 13:38:33
2 |  | PZ | 2.934 | 19.886 | 2.0905 | 20.6911 | -0.962 | 77.244 | 2022/01/07 13:47:30
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27268 | 3.9052
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.398 | -93.171 | -14.7567 | -102.4147 | -1.027 | 77.79 | 2022/01/07 13:57:18
3 |  | UT2 | -14.285 | -105.352 | -15.6717 | -115.6784 | -1.001 | 77.221 | 2022/01/07 14:06:15
3 |  | UT2 | -14.423 | -108.224 | -15.814 | -118.8056 | -1.028 | 77.756 | 2022/01/07 14:15:12
3 |  | UT2 | -14.547 | -109.564 | -15.9419 | -120.2647 | -0.749 | 77.148 | 2022/01/07 14:24:09
3 |  | UT2 | -14.574 | -110.399 | -15.9698 | -121.174 | -0.799 | 77.176 | 2022/01/07 14:33:05
3 |  | UT2 | -14.63 | -110.874 | -16.0275 | -121.6912 | -0.949 | 76.974 | 2022/01/07 14:42:02
3 |  | UT2 | -14.631 | -111.275 | -16.0286 | -122.1278 | -0.859 | 77.109 | 2022/01/07 14:50:58
3 |  | UT2 | -14.676 | -111.561 | -16.075 | -122.4392 | -0.846 | 76.86 | 2022/01/07 14:59:54
3 |  | UT2 | -14.651 | -111.689 | -16.0492 | -122.5786 | -0.845 | 77.134 | 2022/01/07 15:08:51
3 |  | UT2 | -14.667 | -111.916 | -16.0657 | -122.8258 | -1.044 | 76.93 | 2022/01/07 15:17:48
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.38918 | 5.717
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.407 | -73.079 | -10.6398 | -80.5369 | -1.141 | 77.671 | 2022/01/07 15:27:36
4 |  | EV | -9.138 | -69.115 | -10.3623 | -76.2205 | -0.826 | 76.661 | 2022/01/07 15:36:33
4 |  | EV | -9.099 | -68.106 | -10.3221 | -75.1218 | -0.981 | 77.141 | 2022/01/07 15:45:29
4 |  | EV | -9.071 | -67.684 | -10.2932 | -74.6623 | -0.947 | 77.045 | 2022/01/07 15:54:26
4 |  | EV | -9.028 | -67.542 | -10.2488 | -74.5077 | -1.018 | 76.701 | 2022/01/07 16:03:22
4 |  | EV | -9.02 | -67.438 | -10.2406 | -74.3945 | -0.98 | 76.672 | 2022/01/07 16:12:19
4 |  | EV | -9.015 | -67.381 | -10.2354 | -74.3324 | -1.112 | 76.683 | 2022/01/07 16:21:15
4 |  | EV | -9.02 | -67.106 | -10.2406 | -74.033 | -0.695 | 76.567 | 2022/01/07 16:30:12
4 |  | EV | -8.99 | -67.191 | -10.2096 | -74.1255 | -0.94 | 76.622 | 2022/01/07 16:39:09
4 |  | EV | -9.022 | -67.17 | -10.2426 | -74.1027 | -1.006 | 76.585 | 2022/01/07 16:48:05
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3035 | -75.2037 |  |  |  | 0.12307 | 1.8219
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | SCBI1-C-3 | 3.081 | -34.064 | 2.2421 | -38.0541 | -0.944 | 76.743 | 2022/01/07 16:57:55
5 |  | SCBI1-C-3 | 3.882 | -29.99 | 3.0684 | -33.618 | -0.929 | 76.805 | 2022/01/07 17:06:51
5 |  | SCBI1-C-3 | 3.975 | -28.941 | 3.1643 | -32.4758 | -1.116 | 76.684 | 2022/01/07 17:15:48
5 |  | SCBI1-C-3 | 3.935 | -28.913 | 3.1231 | -32.4453 | -1.027 | 76.95 | 2022/01/07 17:24:44
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | SCBI1-C-3 | 2.8995 | -34.1483 |  |  |  | 0.42654 | 2.4433
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | UNDE1-3040-3 | -4.54 | -38.607 | -5.6193 | -43.0009 | -0.987 | 84.908 | 2022/01/07 17:34:35
6 |  | UNDE1-3040-3 | -4.917 | -38.958 | -6.0081 | -43.3831 | -1.192 | 80.612 | 2022/01/07 17:43:31
6 |  | UNDE1-3040-3 | -4.982 | -39.118 | -6.0752 | -43.5573 | -0.971 | 77.219 | 2022/01/07 17:52:27
6 |  | UNDE1-3040-3 | -4.951 | -39.139 | -6.0432 | -43.5802 | -1.024 | 76.605 | 2022/01/07 18:01:23
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | UNDE1-3040-3 | -5.9365 | -43.3804 |  | 2.5019 | 10.3501 | 0.20671 | 0.24599 | 33.1762 | 5.967
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | WOOD1-3040-3 | -11.504 | -89.068 | -12.8029 | -97.947 | -1.075 | 76.677 | 2022/01/07 18:11:13
7 |  | WOOD1-3040-3 | -11.863 | -94.623 | -13.1733 | -103.9957 | -0.836 | 76.637 | 2022/01/07 18:20:10
7 |  | WOOD1-3040-3 | -11.948 | -95.648 | -13.2609 | -105.1118 | -0.87 | 76.489 | 2022/01/07 18:29:07
7 |  | WOOD1-3040-3 | -11.968 | -96.383 | -13.2816 | -105.9122 | -0.933 | 76.623 | 2022/01/07 18:38:03
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | WOOD1-3040-3 | -13.1297 | -103.2417 |  |  |  | 0.21602 | 3.321
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | UNDE1-deep-3 | -5.414 | -47.239 | -6.5208 | -52.4001 | -1.067 | 76.852 | 2022/01/07 18:47:50
8 |  | UNDE1-deep-3 | -5.099 | -42.083 | -6.1959 | -46.7859 | -1.126 | 76.865 | 2022/01/07 18:56:47
8 |  | UNDE1-deep-3 | -5.014 | -40.515 | -6.1082 | -45.0785 | -1.185 | 76.319 | 2022/01/07 19:05:44
8 |  | UNDE1-deep-3 | -4.967 | -39.993 | -6.0597 | -44.5101 | -1.131 | 76.476 | 2022/01/07 19:14:40
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | UNDE1-deep-3 | -6.2212 | -47.1936 |  |  |  | 0.20122 | 3.3091
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | WOOD1-2030-3 | -10.634 | -85.241 | -11.9055 | -93.7798 | -1.136 | 76.856 | 2022/01/07 19:24:27
9 |  | WOOD1-2030-3 | -10.948 | -90.009 | -12.2294 | -98.9716 | -1.067 | 76.571 | 2022/01/07 19:33:24
9 |  | WOOD1-2030-3 | -10.997 | -91.305 | -12.2799 | -100.3828 | -1.12 | 76.762 | 2022/01/07 19:42:20
9 |  | WOOD1-2030-3 | -11.022 | -91.689 | -12.3057 | -100.801 | -1.132 | 76.511 | 2022/01/07 19:51:17
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | WOOD1-2030-3 | -12.1801 | -98.4838 |  |  |  | 0.18014 | 2.9683
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | WOOD1-deep-2 | -9.572 | -73.175 | -10.81 | -80.6414 | -1.071 | 77.092 | 2022/01/07 20:01:06
10 |  | WOOD1-deep-2 | -9.463 | -71.211 | -10.6975 | -78.5028 | -0.969 | 76.163 | 2022/01/07 20:10:02
10 |  | WOOD1-deep-2 | -9.443 | -70.971 | -10.6769 | -78.2415 | -0.967 | 76.328 | 2022/01/07 20:18:59
10 |  | WOOD1-deep-2 | -9.456 | -70.636 | -10.6903 | -77.8767 | -0.968 | 76.065 | 2022/01/07 20:27:55
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | WOOD1-deep-2 | -10.7187 | -78.8156 |  |  |  | 0.059579 | 1.1424
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | WOOD1-3040-2 | -9.724 | -75.512 | -10.9668 | -83.1861 | -0.985 | 77.305 | 2022/01/07 20:37:44
11 |  | WOOD1-3040-2 | -9.724 | -75.831 | -10.9668 | -83.5335 | -0.625 | 75.987 | 2022/01/07 20:46:40
11 |  | WOOD1-3040-2 | -9.707 | -76.004 | -10.9492 | -83.7218 | -0.978 | 76.392 | 2022/01/07 20:55:36
11 |  | WOOD1-3040-2 | -9.756 | -75.975 | -10.9998 | -83.6903 | -0.995 | 76.147 | 2022/01/07 21:04:32
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | WOOD1-3040-2 | -10.9706 | -83.5329 |  |  |  | 0.020467 | 0.22541
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | SCBI1-D-5 | -2.688 | -46.215 | -3.7088 | -51.2851 | -2.812 | 638.339 | 2022/01/07 21:14:20
12 |  | SCBI1-D-5 | -2.377 | -42.97 | -3.388 | -47.7517 | -2.928 | 628.378 | 2022/01/07 21:23:16
12 |  | SCBI1-D-5 | -2.296 | -42.212 | -3.3045 | -46.9263 | -2.954 | 634.216 | 2022/01/07 21:32:12
12 |  | SCBI1-D-5 | -2.278 | -42.003 | -3.2859 | -46.6987 | -2.9 | 628.56 | 2022/01/07 21:41:09
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | SCBI1-D-5 | -3.4218 | -48.1655 | -2.8985 | 555.0391 | 9.6605 | 0.19043 | 1.9547 |  | -11.9817
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.724 | -63.968 | -9.9352 | -70.616 | -1.105 | 84.493 | 2022/01/07 21:50:58
13 |  | EV | -8.933 | -66.011 | -10.1508 | -72.8406 | -1.102 | 79.606 | 2022/01/07 21:59:53
13 |  | EV | -8.943 | -66.554 | -10.1611 | -73.4319 | -1.197 | 78.483 | 2022/01/07 22:08:49
13 |  | EV | -8.988 | -66.925 | -10.2076 | -73.8359 | -1.171 | 77.898 | 2022/01/07 22:17:46
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.1137 | -72.6811 |  | 2.7859 |  | 0.11779 | 1.3189
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD1-1020-3 | -9.134 | -78.967 | -10.3582 | -86.9482 | -0.942 | 77.283 | 2022/01/07 22:27:36
14 |  | WOOD1-1020-3 | -9.148 | -80.478 | -10.3726 | -88.5935 | -0.809 | 77.103 | 2022/01/07 22:36:33
14 |  | WOOD1-1020-3 | -9.101 | -80.804 | -10.3241 | -88.9485 | -0.921 | 76.997 | 2022/01/07 22:45:29
14 |  | WOOD1-1020-3 | -9.153 | -81.002 | -10.3778 | -89.1641 | -1.012 | 76.497 | 2022/01/07 22:54:26
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | WOOD1-1020-3 | -10.3582 | -88.4136 |  |  |  | 0.023424 | 0.92281
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD1-deep-1 | -8.957 | -66.556 | -10.1756 | -73.4341 | -0.832 | 76.844 | 2022/01/07 23:04:13
15 |  | WOOD1-deep-1 | -8.948 | -64.871 | -10.1663 | -71.5993 | -0.87 | 76.679 | 2022/01/07 23:13:08
15 |  | WOOD1-deep-1 | -8.969 | -64.49 | -10.188 | -71.1844 | -1.005 | 76.611 | 2022/01/07 23:22:05
15 |  | WOOD1-deep-1 | -9.003 | -64.409 | -10.223 | -71.0962 | -0.873 | 76.407 | 2022/01/07 23:31:01
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | WOOD1-deep-1 | -10.1882 | -71.8285 |  |  |  | 0.024088 | 1.0034
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | WOOD1-010-1 | -2.358 | -58.189 | -3.3684 | -64.3234 | -0.953 | 77.017 | 2022/01/07 23:40:48
16 |  | WOOD1-010-1 | -1.98 | -57.237 | -2.9785 | -63.2868 | -1.018 | 76.299 | 2022/01/07 23:49:45
16 |  | WOOD1-010-1 | -1.953 | -57.115 | -2.9506 | -63.1539 | -0.947 | 76.607 | 2022/01/07 23:58:40
16 |  | WOOD1-010-1 | -1.88 | -56.972 | -2.8753 | -62.9982 | -0.978 | 76.384 | 2022/01/08 00:07:36
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | WOOD1-010-1 | -3.0432 | -63.4406 |  |  |  | 0.21437 | 0.55124
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | WOOD1-010-3 | -6.145 | -68.652 | -7.2749 | -75.7164 | -0.921 | 77.014 | 2022/01/08 00:17:23
17 |  | WOOD1-010-3 | -6.393 | -70.015 | -7.5307 | -77.2005 | -0.888 | 76.586 | 2022/01/08 00:26:19
17 |  | WOOD1-010-3 | -6.425 | -70.474 | -7.5637 | -77.7003 | -0.967 | 76.497 | 2022/01/08 00:35:16
17 |  | WOOD1-010-3 | -6.428 | -70.519 | -7.5668 | -77.7493 | -0.855 | 76.379 | 2022/01/08 00:44:12
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | WOOD1-010-3 | -7.484 | -77.0916 |  |  |  | 0.13609 | 0.87225
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | WOOD1-6070-1 | -8.469 | -68.364 | -9.6722 | -75.4028 | -1.02 | 76.919 | 2022/01/08 00:53:59
18 |  | WOOD1-6070-1 | -8.569 | -68.213 | -9.7753 | -75.2384 | -1.045 | 76.452 | 2022/01/08 01:02:55
18 |  | WOOD1-6070-1 | -8.581 | -68.051 | -9.7877 | -75.062 | -1.004 | 76.386 | 2022/01/08 01:11:51
18 |  | WOOD1-6070-1 | -8.585 | -68.275 | -9.7918 | -75.3059 | -1.007 | 76.352 | 2022/01/08 01:20:47
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | WOOD1-6070-1 | -9.7568 | -75.2522 |  |  |  | 0.055088 | 0.13196
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | WOOD1-2030-1 | -7.602 | -67.599 | -8.7778 | -74.5698 | -0.926 | 76.301 | 2022/01/08 01:30:34
19 |  | WOOD1-2030-1 | -7.549 | -67.297 | -8.7232 | -74.2409 | -0.906 | 76.306 | 2022/01/08 01:39:30
19 |  | WOOD1-2030-1 | -7.52 | -67.238 | -8.6933 | -74.1767 | -1.004 | 76.075 | 2022/01/08 01:48:27
19 |  | WOOD1-2030-1 | -7.534 | -67.102 | -8.7077 | -74.0286 | -1.086 | 76.047 | 2022/01/08 01:57:24
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | WOOD1-2030-1 | -8.7255 | -74.254 |  |  |  | 0.035846 | 0.20987
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | WOOD1-1020-2 | -8.234 | -70.514 | -9.4298 | -77.7439 | -0.967 | 76.68 | 2022/01/08 02:07:12
20 |  | WOOD1-1020-2 | -8.22 | -70.867 | -9.4153 | -78.1283 | -1.222 | 76.37 | 2022/01/08 02:16:08
20 |  | WOOD1-1020-2 | -8.274 | -70.938 | -9.471 | -78.2056 | -1.059 | 76.187 | 2022/01/08 02:25:05
20 |  | WOOD1-1020-2 | -8.262 | -70.995 | -9.4587 | -78.2676 | -1.059 | 76.168 | 2022/01/08 02:34:02
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | WOOD1-1020-2 | -9.4437 | -78.0863 |  |  |  | 0.024839 | 0.21611
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | WOOD1-1020-1 | -6.973 | -67.348 | -8.129 | -74.2965 | -1.127 | 76.89 | 2022/01/08 02:43:50
21 |  | WOOD1-1020-1 | -6.854 | -66.985 | -8.0062 | -73.9012 | -0.951 | 76.254 | 2022/01/08 02:52:47
21 |  | WOOD1-1020-1 | -6.864 | -67.025 | -8.0166 | -73.9448 | -0.989 | 76.161 | 2022/01/08 03:01:43
21 |  | WOOD1-1020-1 | -6.829 | -66.89 | -7.9805 | -73.7978 | -1.157 | 76.235 | 2022/01/08 03:10:40
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | WOOD1-1020-1 | -8.0331 | -73.9851 |  |  |  | 0.063723 | 0.1989
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.897 | -67.062 | -10.1137 | -73.9851 | -0.937 | 76.522 | 2022/01/08 03:20:26
22 |  | EV | -8.966 | -66.928 | -10.1849 | -73.8391 | -1.288 | 76.845 | 2022/01/08 03:29:22
22 |  | EV | -9.031 | -67.065 | -10.2519 | -73.9883 | -1.018 | 76.464 | 2022/01/08 03:38:17
22 |  | EV | -8.983 | -67.014 | -10.2024 | -73.9328 | -0.993 | 76.63 | 2022/01/08 03:47:13
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.1882 | -73.9363 |  |  |  | 0.055476 | 0.063924
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | SLCP-22-001 | -15.233 | -110.324 | -16.6496 | -121.0923 | -0.936 | 76.927 | 2022/01/08 03:56:59
23 |  | SLCP-22-001 | -15.56 | -114.864 | -16.9869 | -126.0358 | -0.98 | 76.542 | 2022/01/08 04:05:55
23 |  | SLCP-22-001 | -15.654 | -116.219 | -17.0838 | -127.5113 | -0.911 | 76.86 | 2022/01/08 04:14:51
23 |  | SLCP-22-001 | -15.666 | -116.484 | -17.0962 | -127.7998 | -0.906 | 76.899 | 2022/01/08 04:23:47
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | SLCP-22-001 | -16.9541 | -125.6098 |  |  |  | 0.20246 | 2.8554
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | 57 | -5.898 | -46.216 | -7.0201 | -51.2862 | -0.844 | 76.324 | 2022/01/08 04:33:33
24 |  | 57 | -5.075 | -37.987 | -6.1711 | -42.3258 | -0.977 | 76.502 | 2022/01/08 04:42:30
24 |  | 57 | -5.02 | -36.292 | -6.1144 | -40.4801 | -1.145 | 77.21 | 2022/01/08 04:51:25
24 |  | 57 | -5.123 | -35.806 | -6.2206 | -39.9509 | -0.959 | 76.123 | 2022/01/08 05:00:21
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | 57 | -6.3816 | -43.5108 |  |  |  | 0.41481 | 4.8514
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | 58 | -5.228 | -35.654 | -6.329 | -39.7854 | -1.009 | 75.979 | 2022/01/08 05:10:07
25 |  | 58 | -5.262 | -35.444 | -6.364 | -39.5568 | -1.021 | 75.189 | 2022/01/08 05:19:03
25 |  | 58 | -5.267 | -35.08 | -6.3692 | -39.1604 | -0.947 | 75.749 | 2022/01/08 05:27:59
25 |  | 58 | -5.271 | -35.033 | -6.3733 | -39.1092 | -0.965 | 75.62 | 2022/01/08 05:36:54
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | 58 | -6.3589 | -39.403 |  |  |  | 0.019681 | 0.29761
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | 59 | -5.134 | -34.252 | -6.232 | -38.2588 | -0.984 | 76.126 | 2022/01/08 05:46:41
26 |  | 59 | -5.041 | -33.526 | -6.1361 | -37.4683 | -0.962 | 75.554 | 2022/01/08 05:55:37
26 |  | 59 | -5.111 | -33.726 | -6.2083 | -37.6861 | -1.095 | 75.788 | 2022/01/08 06:04:33
26 |  | 59 | -5.108 | -33.952 | -6.2052 | -37.9322 | -0.709 | 75.615 | 2022/01/08 06:13:29
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | 59 | -6.1954 | -37.8363 |  |  |  | 0.040054 | 0.31176
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | 60 | -5.349 | -35.493 | -6.4538 | -39.6101 | -0.966 | 76.297 | 2022/01/08 06:23:16
27 |  | 60 | -5.365 | -35.942 | -6.4703 | -40.099 | -1.14 | 75.901 | 2022/01/08 06:32:11
27 |  | 60 | -5.307 | -35.62 | -6.4104 | -39.7484 | -0.853 | 75.551 | 2022/01/08 06:41:07
27 |  | 60 | -5.35 | -35.711 | -6.4548 | -39.8475 | -1.048 | 75.754 | 2022/01/08 06:50:03
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | 60 | -6.4473 | -39.8263 |  |  |  | 0.024932 | 0.18942
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | 61 | -5.406 | -36.018 | -6.5126 | -40.1818 | -1.01 | 76.251 | 2022/01/08 06:59:51
28 |  | 61 | -5.399 | -35.852 | -6.5053 | -40.001 | -1.057 | 76.099 | 2022/01/08 07:08:47
28 |  | 61 | -5.433 | -35.918 | -6.5404 | -40.0729 | -1.048 | 76.197 | 2022/01/08 07:17:43
28 |  | 61 | -5.419 | -35.949 | -6.526 | -40.1067 | -1.03 | 76.335 | 2022/01/08 07:26:40
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | 61 | -6.5211 | -40.0906 |  |  |  | 0.014997 | 0.068946
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | 62 | -5.48 | -36.283 | -6.5889 | -40.4703 | -0.927 | 76.2 | 2022/01/08 07:36:24
29 |  | 62 | -5.466 | -36.389 | -6.5745 | -40.5858 | -1.058 | 76.282 | 2022/01/08 07:45:18
29 |  | 62 | -5.495 | -36.527 | -6.6044 | -40.736 | -0.799 | 75.287 | 2022/01/08 07:54:15
29 |  | 62 | -5.473 | -36.407 | -6.5817 | -40.6054 | -0.873 | 76.174 | 2022/01/08 08:03:11
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | 62 | -6.5874 | -40.5994 |  |  |  | 0.012396 | 0.099965
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | 63 | -5.269 | -35.428 | -6.3712 | -39.5393 | -1.019 | 76.14 | 2022/01/08 08:12:55
30 |  | 63 | -5.286 | -35.193 | -6.3888 | -39.2835 | -1.143 | 76.151 | 2022/01/08 08:21:51
30 |  | 63 | -5.254 | -35.245 | -6.3558 | -39.3401 | -0.936 | 75.996 | 2022/01/08 08:30:47
30 |  | 63 | -5.232 | -35.254 | -6.3331 | -39.3499 | -0.763 | 75.892 | 2022/01/08 08:39:43
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | 63 | -6.3622 | -39.3782 |  |  |  | 0.022926 | 0.10226
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.684 | -62.256 | -9.894 | -68.7519 | -1.051 | 76.97 | 2022/01/08 08:49:28
31 |  | EV | -8.92 | -65.068 | -10.1374 | -71.8138 | -1.175 | 76.943 | 2022/01/08 08:58:24
31 |  | EV | -8.982 | -65.851 | -10.2014 | -72.6664 | -0.932 | 76.373 | 2022/01/08 09:07:21
31 |  | EV | -9.018 | -66.192 | -10.2385 | -73.0377 | -1.013 | 76.455 | 2022/01/08 09:16:17
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.1178 | -71.5675 |  |  |  | 0.15022 | 1.7869
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | 64 | -5.679 | -40.386 | -6.7942 | -44.938 | -1.114 | 76.189 | 2022/01/08 09:26:02
32 |  | 64 | -5.504 | -37.577 | -6.6137 | -41.8794 | -1.037 | 76.092 | 2022/01/08 09:34:59
32 |  | 64 | -5.467 | -36.829 | -6.5755 | -41.0649 | -1.106 | 75.951 | 2022/01/08 09:43:55
32 |  | 64 | -5.433 | -36.49 | -6.5404 | -40.6957 | -1.064 | 76.062 | 2022/01/08 09:52:51
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | 64 | -6.6309 | -42.1445 |  |  |  | 0.10941 | 1.7696
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | 65 | -5.461 | -36.59 | -6.5693 | -40.8046 | -1.128 | 76.105 | 2022/01/08 10:02:37
33 |  | 65 | -5.51 | -36.432 | -6.6199 | -40.6326 | -0.934 | 75.297 | 2022/01/08 10:11:33
33 |  | 65 | -5.468 | -36.413 | -6.5765 | -40.6119 | -0.889 | 75.352 | 2022/01/08 10:20:29
33 |  | 65 | -5.475 | -36.483 | -6.5837 | -40.6881 | -0.875 | 75.193 | 2022/01/08 10:29:25
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | 65 | -6.5874 | -40.6843 |  |  |  | 0.021764 | 0.079375
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | 66 | -5.461 | -36.142 | -6.5693 | -40.3168 | -0.982 | 75.419 | 2022/01/08 10:39:12
34 |  | 66 | -5.535 | -36.476 | -6.6456 | -40.6805 | -0.863 | 76.143 | 2022/01/08 10:48:07
34 |  | 66 | -5.49 | -36.409 | -6.5992 | -40.6075 | -0.922 | 74.864 | 2022/01/08 10:57:02
34 |  | 66 | -5.472 | -36.345 | -6.5807 | -40.5378 | -0.957 | 75.79 | 2022/01/08 11:05:58
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | 66 | -6.5987 | -40.5357 |  |  |  | 0.032604 | 0.14428
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | 67 | -5.46 | -36.286 | -6.5683 | -40.4736 | -1.025 | 76.005 | 2022/01/08 11:15:45
35 |  | 67 | -5.486 | -36.39 | -6.5951 | -40.5868 | -0.956 | 74.759 | 2022/01/08 11:24:42
35 |  | 67 | -5.418 | -36.089 | -6.5249 | -40.2591 | -0.981 | 74.21 | 2022/01/08 11:33:38
35 |  | 67 | -5.456 | -36.281 | -6.5641 | -40.4682 | -1.094 | 74.55 | 2022/01/08 11:42:34
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | 67 | -6.5631 | -40.4469 |  |  |  | 0.028024 | 0.1255
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | 68 | -5.523 | -36.343 | -6.6333 | -40.5357 | -0.966 | 76.254 | 2022/01/08 11:52:16
36 |  | 68 | -5.511 | -36.192 | -6.6209 | -40.3712 | -1.044 | 76.297 | 2022/01/08 12:01:12
36 |  | 68 | -5.503 | -36.251 | -6.6126 | -40.4355 | -0.955 | 76.028 | 2022/01/08 12:10:08
36 |  | 68 | -5.559 | -36.446 | -6.6704 | -40.6478 | -1.08 | 76.001 | 2022/01/08 12:19:03
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | 68 | -6.6343 | -40.4976 |  |  |  | 0.024739 | 0.11102
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | 69 | -5.345 | -35.681 | -6.4496 | -39.8148 | -0.945 | 75.199 | 2022/01/08 12:28:46
37 |  | 69 | -5.339 | -35.752 | -6.4435 | -39.8921 | -1.034 | 75.738 | 2022/01/08 12:37:41
37 |  | 69 | -5.349 | -35.781 | -6.4538 | -39.9237 | -0.926 | 75.438 | 2022/01/08 12:46:37
37 |  | 69 | -5.39 | -35.711 | -6.4961 | -39.8475 | -0.769 | 75.814 | 2022/01/08 12:55:33
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | 69 | -6.4607 | -39.8695 |  |  |  | 0.0232 | 0.044124
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | 70 | -5.847 | -38.418 | -6.9675 | -42.7951 | -0.804 | 75.711 | 2022/01/08 13:05:18
38 |  | 70 | -5.846 | -38.763 | -6.9665 | -43.1708 | -0.99 | 75.609 | 2022/01/08 13:14:13
38 |  | 70 | -5.875 | -38.676 | -6.9964 | -43.076 | -1.07 | 75.655 | 2022/01/08 13:23:09
38 |  | 70 | -5.855 | -38.876 | -6.9757 | -43.2938 | -0.982 | 75.554 | 2022/01/08 13:32:05
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | 70 | -6.9765 | -43.0839 |  |  |  | 0.013451 | 0.19487
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | 71 | -5.643 | -37.485 | -6.757 | -41.7792 | -0.864 | 76.124 | 2022/01/08 13:41:49
39 |  | 71 | -5.636 | -37.544 | -6.7498 | -41.8434 | -0.83 | 76.348 | 2022/01/08 13:50:46
39 |  | 71 | -5.619 | -37.468 | -6.7323 | -41.7607 | -0.877 | 76.196 | 2022/01/08 13:59:42
39 |  | 71 | -5.66 | -37.489 | -6.7746 | -41.7835 | -0.989 | 75.986 | 2022/01/08 14:08:38
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | 71 | -6.7534 | -41.7917 |  |  |  | 0.01698 | 0.032949
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -8.754 | -62.574 | -9.9662 | -69.0981 | -0.967 | 76.897 | 2022/01/08 14:18:23
40 |  | EV | -8.958 | -65.065 | -10.1766 | -71.8106 | -0.751 | 76.061 | 2022/01/08 14:27:19
40 |  | EV | -8.987 | -65.726 | -10.2065 | -72.5303 | -0.892 | 76.226 | 2022/01/08 14:36:14
40 |  | EV | -8.994 | -65.949 | -10.2138 | -72.7731 | -0.865 | 76.151 | 2022/01/08 14:45:10
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.1408 | -71.553 |  |  |  | 0.1139 | 1.5492
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | 72 | -5.716 | -40.798 | -6.8323 | -45.3866 | -1.01 | 75.932 | 2022/01/08 14:54:55
41 |  | 72 | -5.597 | -38.259 | -6.7096 | -42.622 | -0.815 | 75.543 | 2022/01/08 15:03:51
41 |  | 72 | -5.517 | -37.518 | -6.6271 | -41.8151 | -0.93 | 75.182 | 2022/01/08 15:12:46
41 |  | 72 | -5.52 | -37.295 | -6.6302 | -41.5723 | -0.975 | 75.671 | 2022/01/08 15:21:41
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | 72 | -6.6998 | -42.849 |  |  |  | 0.093326 | 1.6074
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | 73 | -5.593 | -37.081 | -6.7055 | -41.3393 | -0.817 | 75.47 | 2022/01/08 15:31:27
42 |  | 73 | -5.602 | -37.164 | -6.7148 | -41.4296 | -0.951 | 75.214 | 2022/01/08 15:40:23
42 |  | 73 | -5.565 | -37.042 | -6.6766 | -41.2968 | -1.181 | 75.179 | 2022/01/08 15:49:18
42 |  | 73 | -5.593 | -36.945 | -6.7055 | -41.1912 | -0.812 | 75.125 | 2022/01/08 15:58:13
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | 73 | -6.7006 | -41.3142 |  |  |  | 0.01607 | 0.090903
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | 74 | -5.589 | -37.354 | -6.7013 | -41.6365 | -0.802 | 76.113 | 2022/01/08 16:07:57
43 |  | 74 | -5.609 | -37.486 | -6.722 | -41.7803 | -0.984 | 75.949 | 2022/01/08 16:16:52
43 |  | 74 | -5.587 | -37.305 | -6.6993 | -41.5832 | -0.95 | 75.624 | 2022/01/08 16:25:48
43 |  | 74 | -5.63 | -37.575 | -6.7436 | -41.8772 | -0.805 | 75.851 | 2022/01/08 16:34:43
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | 74 | -6.7166 | -41.7193 |  |  |  | 0.020123 | 0.12324
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | 75 | -5.61 | -37.437 | -6.723 | -41.7269 | -1.048 | 75.868 | 2022/01/08 16:44:26
44 |  | 75 | -5.574 | -37.236 | -6.6859 | -41.508 | -0.971 | 76.374 | 2022/01/08 16:53:21
44 |  | 75 | -5.563 | -37.375 | -6.6745 | -41.6594 | -0.792 | 75.774 | 2022/01/08 17:02:17
44 |  | 75 | -5.585 | -37.252 | -6.6972 | -41.5255 | -0.915 | 75.902 | 2022/01/08 17:11:12
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | 75 | -6.6952 | -41.605 |  |  |  | 0.020116 | 0.097115
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | 76 | -5.766 | -38.251 | -6.8839 | -42.6133 | -0.799 | 75.057 | 2022/01/08 17:20:58
45 |  | 76 | -5.718 | -37.98 | -6.8344 | -42.3182 | -0.793 | 75.198 | 2022/01/08 17:29:53
45 |  | 76 | -5.742 | -38.059 | -6.8592 | -42.4042 | -0.901 | 75.156 | 2022/01/08 17:38:49
45 |  | 76 | -5.687 | -37.828 | -6.8024 | -42.1527 | -0.775 | 75.224 | 2022/01/08 17:47:45
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | 76 | -6.845 | -42.3721 |  |  |  | 0.033768 | 0.17605
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | 77 | -5.591 | -36.925 | -6.7034 | -41.1694 | -1.081 | 75.977 | 2022/01/08 17:57:29
46 |  | 77 | -5.523 | -36.798 | -6.6333 | -41.0311 | -1.004 | 75.411 | 2022/01/08 18:06:25
46 |  | 77 | -5.552 | -37.03 | -6.6632 | -41.2837 | -1.092 | 76.282 | 2022/01/08 18:15:20
46 |  | 77 | -5.58 | -36.883 | -6.6921 | -41.1237 | -1.014 | 76.356 | 2022/01/08 18:24:15
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | 77 | -6.673 | -41.152 |  |  |  | 0.030469 | 0.096426
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | 78 | -5.73 | -38.08 | -6.8468 | -42.4271 | -0.78 | 71.88 | 2022/01/08 18:34:01
47 |  | 78 | -5.786 | -38.497 | -6.9046 | -42.8811 | -0.789 | 72.242 | 2022/01/08 18:42:57
47 |  | 78 | -5.818 | -38.722 | -6.9376 | -43.1261 | -1.017 | 74.521 | 2022/01/08 18:51:53
47 |  | 78 | -5.818 | -38.455 | -6.9376 | -42.8354 | -0.95 | 72.769 | 2022/01/08 19:00:49
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | 78 | -6.9066 | -42.8174 |  |  |  | 0.041505 | 0.2662
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | 79 | -5.586 | -37.167 | -6.6982 | -41.4329 | -1.123 | 74.79 | 2022/01/08 19:10:35
48 |  | 79 | -5.593 | -37.053 | -6.7055 | -41.3088 | -0.945 | 75.544 | 2022/01/08 19:19:31
48 |  | 79 | -5.582 | -37.06 | -6.6941 | -41.3164 | -1.009 | 75.496 | 2022/01/08 19:28:26
48 |  | 79 | -5.628 | -36.977 | -6.7416 | -41.226 | -1.09 | 75.709 | 2022/01/08 19:37:22
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | 79 | -6.7099 | -41.321 |  |  |  | 0.020998 | 0.078134
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -8.752 | -62.305 | -9.9641 | -68.8052 | -1.067 | 76.395 | 2022/01/08 19:47:08
49 |  | EV | -8.904 | -65.092 | -10.1209 | -71.84 | -0.903 | 75.975 | 2022/01/08 19:56:03
49 |  | EV | -8.98 | -65.738 | -10.1993 | -72.5434 | -1.06 | 76.016 | 2022/01/08 20:04:58
49 |  | EV | -8.988 | -65.993 | -10.2076 | -72.821 | -0.882 | 75.981 | 2022/01/08 20:13:53
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.123 | -71.5024 |  |  |  | 0.10942 | 1.6943

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-18696 | 2022/01/07 10:56:50 | 1-01 | 1 | -9.85 | -75.296 | 13455 | -1 | 1 | EV | standard | H2O | 1641578515 | 0.209 | 0.826 | 68.959 | 0.001 | 0.003 | 1.695 | 110.706 | 30.566 | 2.262 | 66.448 | 0.856 | -1.155 | 0 | 0 | 1 | 45.312 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-18696 | 2022/01/07 11:05:45 | 1-01 | 2 | -9.306 | -70.206 | 12998 | -1 | 1 | EV | standard | H2O | 1641579051.03 | 0.205 | 0.726 | 61.226 | 0 | -0.002 | 1.507 | 105.569 | 29.827 | 2.206 | 65.064 | 0.856 | -1.12 | 0 | 0 | 1 | 44.375 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-18696 | 2022/01/07 11:14:42 | 1-01 | 3 | -9.29 | -69.612 | 13865 | -1 | 1 | EV | standard | H2O | 1641579587.23 | 0.194 | 0.768 | 58.539 | 0 | -0.001 | 1.424 | 129.721 | 31.4 | 2.35 | 68.543 | 0.857 | -1.197 | 0 | 0 | 1 | 44 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-18696 | 2022/01/07 11:23:38 | 1-01 | 4 | -9.123 | -68.392 | 12890 | 0 | 1 | EV | standard | H2O | 1641580123.39 | 0.207 | 0.762 | 60.082 | 0 | 0 | 1.47 | 97.634 | 29.838 | 2.19 | 64.452 | 0.856 | -0.956 | 0 | 0 | 1 | 43.875 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-18696 | 2022/01/07 11:32:35 | 1-01 | 5 | -9.117 | -68.04 | 13046 | 0 | 1 | EV | standard | H2O | 1641580660.2 | 0.201 | 0.82 | 58.79 | 0 | -0.003 | 1.443 | 89.742 | 30.46 | 2.221 | 64.962 | 0.856 | -1.095 | 0 | 0 | 1 | 43.688 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-18696 | 2022/01/07 11:41:31 | 1-01 | 6 | -8.678 | -66.923 | 13089 | 0 | 1 | EV | standard | H2O | 1641581196.28 | 0.209 | 0.786 | 64.234 | 0.001 | -0.002 | 1.578 | 83.997 | 30.386 | 2.218 | 66.429 | 0.856 | -1.052 | 0 | 0 | 1 | 43.5 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-18696 | 2022/01/07 11:50:27 | 1-01 | 7 | -8.989 | -67.386 | 12955 | 0 | 1 | EV | standard | H2O | 1641581732.87 | 0.2 | 0.741 | 57.146 | 0 | -0.001 | 1.404 | 81.283 | 29.989 | 2.193 | 66.741 | 0.856 | -0.925 | 0 | 0 | 1 | 43.312 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-18696 | 2022/01/07 11:59:24 | 1-01 | 8 | -9.065 | -67.561 | 12906 | 0 | 1 | EV | standard | H2O | 1641582269.37 | 0.165 | 0.72 | 62.267 | 0.001 | -0.001 | 1.536 | 79.614 | 30.044 | 2.191 | 66.092 | 0.857 | -0.94 | 0 | 0 | 1 | 43.188 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-18696 | 2022/01/07 12:08:21 | 1-01 | 9 | -9.023 | -67.427 | 12889 | 0 | 1 | EV | standard | H2O | 1641582806.55 | 0.22 | 0.809 | 57.98 | 0 | 0 | 1.417 | 78.908 | 29.848 | 2.182 | 66.431 | 0.855 | -0.953 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-18696 | 2022/01/07 12:17:17 | 1-01 | 10 | -9.037 | -67.285 | 12870 | 0 | 1 | EV | standard | H2O | 1641583342.57 | 0.178 | 0.809 | 60.734 | 0 | 0.001 | 1.491 | 78.518 | 30.123 | 2.183 | 65.829 | 0.856 | -0.937 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-18697 | 2022/01/07 12:27:04 | 1-02 | 1 | 2.041 | 7.133 | 12786 | -1 | 1 | PZ | standard | H2O | 1641583929.77 | 0.189 | 0.835 | 59.791 | 0 | 0.002 | 1.466 | 77.777 | 29.812 | 2.291 | 67.854 | 0.856 | -1.018 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-18697 | 2022/01/07 12:36:00 | 1-02 | 2 | 2.657 | 15.348 | 12657 | -1 | 1 | PZ | standard | H2O | 1641584465.96 | 0.217 | 0.751 | 63.36 | 0 | -0.002 | 1.569 | 77.705 | 29.582 | 2.286 | 66.739 | 0.856 | -1.107 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-18697 | 2022/01/07 12:44:55 | 1-02 | 3 | 2.776 | 17.392 | 12559 | -1 | 1 | PZ | standard | H2O | 1641585000.69 | 0.187 | 0.861 | 57.211 | 0 | -0.003 | 1.414 | 77.489 | 29.462 | 2.261 | 66.297 | 0.856 | -1.118 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-18697 | 2022/01/07 12:53:52 | 1-02 | 4 | 2.862 | 18.28 | 12612 | 0 | 1 | PZ | standard | H2O | 1641585537.65 | 0.229 | 0.883 | 52.629 | 0 | -0.004 | 1.298 | 77.641 | 29.611 | 2.272 | 65.865 | 0.856 | -1.124 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-18697 | 2022/01/07 13:02:48 | 1-02 | 5 | 2.89 | 18.829 | 12564 | 0 | 1 | PZ | standard | H2O | 1641586074 | 0.193 | 0.87 | 57.898 | 0 | -0.006 | 1.42 | 77.453 | 29.267 | 2.273 | 66.19 | 0.856 | -1.072 | 0 | 0 | 1 | 43.375 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-18697 | 2022/01/07 13:11:44 | 1-02 | 6 | 2.882 | 19.176 | 12581 | 0 | 1 | PZ | standard | H2O | 1641586609.15 | 0.203 | 0.794 | 57.891 | 0.001 | -0.003 | 1.422 | 77.368 | 29.5 | 2.267 | 65.978 | 0.856 | -1.141 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-18697 | 2022/01/07 13:20:41 | 1-02 | 7 | 2.889 | 19.458 | 12593 | 0 | 1 | PZ | standard | H2O | 1641587146.62 | 0.197 | 0.776 | 58.825 | 0 | -0.003 | 1.451 | 77.368 | 29.487 | 2.276 | 66.51 | 0.856 | -1.181 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-18697 | 2022/01/07 13:29:37 | 1-02 | 8 | 2.936 | 19.673 | 12605 | 0 | 1 | PZ | standard | H2O | 1641587682.61 | 0.199 | 0.809 | 54.998 | 0.001 | -0.004 | 1.352 | 77.375 | 29.525 | 2.271 | 66.677 | 0.856 | -1.118 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-18697 | 2022/01/07 13:38:33 | 1-02 | 9 | 2.917 | 19.889 | 12583 | 0 | 1 | PZ | standard | H2O | 1641588218.51 | 0.197 | 0.781 | 56.454 | 0 | -0.003 | 1.387 | 77.165 | 29.437 | 2.276 | 67.465 | 0.856 | -0.977 | 0 | 0 | 1 | 43.188 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-18697 | 2022/01/07 13:47:30 | 1-02 | 10 | 2.934 | 19.886 | 12573 | 0 | 1 | PZ | standard | H2O | 1641588755.5 | 0.2 | 0.856 | 58.955 | 0 | 0 | 1.45 | 77.244 | 29.204 | 2.271 | 66.911 | 0.856 | -0.962 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-18698 | 2022/01/07 13:57:18 | 1-03 | 1 | -13.398 | -93.171 | 12793 | -1 | 1 | UT2 | standard | H2O | 1641589343.13 | 0.218 | 0.747 | 58.588 | 0.002 | -0.003 | 1.447 | 77.79 | 29.481 | 2.126 | 64.682 | 0.856 | -1.027 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-18698 | 2022/01/07 14:06:15 | 1-03 | 2 | -14.285 | -105.352 | 12529 | -1 | 1 | UT2 | standard | H2O | 1641589880.36 | 0.206 | 0.833 | 55.777 | 0.001 | 0.005 | 1.377 | 77.221 | 28.958 | 2.082 | 62.795 | 0.856 | -1.001 | 0 | 0 | 1 | 43.375 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-18698 | 2022/01/07 14:15:12 | 1-03 | 3 | -14.423 | -108.224 | 12779 | -1 | 1 | UT2 | standard | H2O | 1641590417.56 | 0.199 | 0.815 | 56.967 | 0 | 0.007 | 1.399 | 77.756 | 29.73 | 2.088 | 63.362 | 0.856 | -1.028 | 0 | 0 | 1 | 43.375 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-18698 | 2022/01/07 14:24:09 | 1-03 | 4 | -14.547 | -109.564 | 12590 | 0 | 1 | UT2 | standard | H2O | 1641590954.31 | 0.191 | 0.745 | 58.399 | 0.001 | 0.006 | 1.434 | 77.148 | 29.351 | 2.085 | 63.324 | 0.856 | -0.749 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-18698 | 2022/01/07 14:33:05 | 1-03 | 5 | -14.574 | -110.399 | 12582 | 0 | 1 | UT2 | standard | H2O | 1641591490.22 | 0.2 | 0.811 | 56.073 | 0.001 | 0.001 | 1.377 | 77.176 | 29.211 | 2.074 | 63.062 | 0.856 | -0.799 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-18698 | 2022/01/07 14:42:02 | 1-03 | 6 | -14.63 | -110.874 | 12484 | 0 | 1 | UT2 | standard | H2O | 1641592027.23 | 0.194 | 0.836 | 58.26 | 0.001 | 0.006 | 1.433 | 76.974 | 28.744 | 2.063 | 63.003 | 0.856 | -0.949 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-18698 | 2022/01/07 14:50:58 | 1-03 | 7 | -14.631 | -111.275 | 12582 | 0 | 1 | UT2 | standard | H2O | 1641592563.49 | 0.183 | 0.751 | 56.061 | 0 | 0.006 | 1.382 | 77.109 | 29.337 | 2.069 | 63.096 | 0.856 | -0.859 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-18698 | 2022/01/07 14:59:54 | 1-03 | 8 | -14.676 | -111.561 | 12462 | 0 | 1 | UT2 | standard | H2O | 1641593099.89 | 0.201 | 0.832 | 59.319 | 0.001 | 0.004 | 1.466 | 76.86 | 28.954 | 2.052 | 62.564 | 0.856 | -0.846 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-18698 | 2022/01/07 15:08:51 | 1-03 | 9 | -14.651 | -111.689 | 12594 | 0 | 1 | UT2 | standard | H2O | 1641593636.96 | 0.187 | 0.798 | 56.277 | 0 | 0.002 | 1.379 | 77.134 | 29.283 | 2.076 | 62.608 | 0.856 | -0.845 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-18698 | 2022/01/07 15:17:48 | 1-03 | 10 | -14.667 | -111.916 | 12459 | 0 | 1 | UT2 | standard | H2O | 1641594173.71 | 0.191 | 0.817 | 55.187 | 0.001 | 0.003 | 1.355 | 76.93 | 28.928 | 2.048 | 61.912 | 0.856 | -1.044 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-18699 | 2022/01/07 15:27:36 | 1-04 | 1 | -9.407 | -73.079 | 12778 | -1 | 1 | EV | standard | H2O | 1641594761.79 | 0.206 | 0.786 | 63.695 | 0 | 0.006 | 1.567 | 77.671 | 29.812 | 2.149 | 64.253 | 0.856 | -1.141 | 0 | 0 | 1 | 43.375 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-18699 | 2022/01/07 15:36:33 | 1-04 | 2 | -9.138 | -69.115 | 12402 | -1 | 1 | EV | standard | H2O | 1641595298.83 | 0.212 | 0.682 | 62.809 | 0 | 0.002 | 1.554 | 76.661 | 28.619 | 2.116 | 64.908 | 0.855 | -0.826 | 0 | 0 | 1 | 43.375 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-18699 | 2022/01/07 15:45:29 | 1-04 | 3 | -9.099 | -68.106 | 12552 | -1 | 1 | EV | standard | H2O | 1641595835.14 | 0.205 | 0.777 | 56.937 | 0 | 0.001 | 1.402 | 77.141 | 29.023 | 2.134 | 64.23 | 0.855 | -0.981 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-18699 | 2022/01/07 15:54:26 | 1-04 | 4 | -9.071 | -67.684 | 12520 | 0 | 1 | EV | standard | H2O | 1641596371.56 | 0.179 | 0.699 | 54.417 | 0 | 0.001 | 1.337 | 77.045 | 28.848 | 2.134 | 64.487 | 0.856 | -0.947 | 0 | 0 | 1 | 43.5 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-18699 | 2022/01/07 16:03:22 | 1-04 | 5 | -9.028 | -67.542 | 12419 | 0 | 1 | EV | standard | H2O | 1641596907.87 | 0.193 | 0.703 | 57.261 | 0.001 | 0.001 | 1.414 | 76.701 | 29.022 | 2.115 | 63.151 | 0.856 | -1.018 | 0 | 0 | 1 | 43.5 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-18699 | 2022/01/07 16:12:19 | 1-04 | 6 | -9.02 | -67.438 | 12358 | 0 | 1 | EV | standard | H2O | 1641597444.14 | 0.201 | 0.785 | 56.813 | 0 | -0.001 | 1.414 | 76.672 | 28.7 | 2.112 | 63.15 | 0.856 | -0.98 | 0 | 0 | 1 | 43.5 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-18699 | 2022/01/07 16:21:15 | 1-04 | 7 | -9.015 | -67.381 | 12357 | 0 | 1 | EV | standard | H2O | 1641597981.28 | 0.203 | 0.924 | 59.859 | 0.001 | 0.006 | 1.471 | 76.683 | 29.005 | 2.092 | 62.105 | 0.856 | -1.112 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-18699 | 2022/01/07 16:30:12 | 1-04 | 8 | -9.02 | -67.106 | 12342 | 0 | 1 | EV | standard | H2O | 1641598517.93 | 0.193 | 0.746 | 57.819 | 0.001 | 0.003 | 1.431 | 76.567 | 28.195 | 2.094 | 65.052 | 0.856 | -0.695 | 0 | 0 | 1 | 43.312 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-18699 | 2022/01/07 16:39:09 | 1-04 | 9 | -8.99 | -67.191 | 12323 | 0 | 1 | EV | standard | H2O | 1641599054.3 | 0.195 | 0.724 | 58.6 | 0.001 | -0.001 | 1.449 | 76.622 | 28.467 | 2.09 | 63.338 | 0.856 | -0.94 | 0 | 0 | 1 | 43.188 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-18699 | 2022/01/07 16:48:05 | 1-04 | 10 | -9.022 | -67.17 | 12326 | 0 | 1 | EV | standard | H2O | 1641599591.5 | 0.216 | 0.813 | 58.055 | 0.001 | 0.001 | 1.432 | 76.585 | 28.759 | 2.096 | 62.933 | 0.856 | -1.006 | 0 | 0 | 1 | 43.062 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-18700 | 2022/01/07 16:57:55 | 1-05 | 1 | 3.081 | -34.064 | 12417 | -1 | 1 | SCBI1-C-3 | 20-196 | H2O | 1641600180.44 | 0.218 | 0.785 | 53.617 | 0 | 0.003 | 1.325 | 76.743 | 29.038 | 2.168 | 64.672 | 0.856 | -0.944 | 0 | 0 | 1 | 42.812 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-18700 | 2022/01/07 17:06:51 | 1-05 | 2 | 3.882 | -29.99 | 12423 | -1 | 1 | SCBI1-C-3 | 20-196 | H2O | 1641600716.83 | 0.203 | 0.91 | 58.748 | -0.001 | -0.002 | 1.444 | 76.805 | 28.696 | 2.176 | 65.235 | 0.856 | -0.929 | 0 | 0 | 1 | 42.75 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-18700 | 2022/01/07 17:15:48 | 1-05 | 3 | 3.975 | -28.941 | 12305 | -1 | 1 | SCBI1-C-3 | 20-196 | H2O | 1641601253.42 | 0.197 | 0.908 | 52.636 | 0 | 0.001 | 1.311 | 76.684 | 28.397 | 2.157 | 63.792 | 0.856 | -1.116 | 0 | 0 | 1 | 42.688 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-18700 | 2022/01/07 17:24:44 | 1-05 | 4 | 3.935 | -28.913 | 12460 | 0 | 1 | SCBI1-C-3 | 20-196 | H2O | 1641601789.95 | 0.2 | 0.749 | 59.181 | 0.001 | -0.001 | 1.459 | 76.95 | 28.868 | 2.183 | 64.895 | 0.856 | -1.027 | 0 | 0 | 1 | 42.688 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-18701 | 2022/01/07 17:34:35 | 1-06 | 1 | -4.54 | -38.607 | 16412 | -1 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641602379.85 | 0.19 | 0.75 | 66.196 | 0 | 0 | 1.579 | 84.908 | 40.289 | 2.808 | 81.903 | 0.859 | -0.987 | 0 | 0 | 1 | 42.688 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-18701 | 2022/01/07 17:43:31 | 1-06 | 2 | -4.917 | -38.958 | 14287 | -1 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641602916.01 | 0.21 | 0.701 | 63.088 | 0 | -0.001 | 1.518 | 80.612 | 34.12 | 2.451 | 71.978 | 0.857 | -1.192 | 0 | 0 | 1 | 42.625 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-18701 | 2022/01/07 17:52:27 | 1-06 | 3 | -4.982 | -39.118 | 12624 | -1 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641603452.43 | 0.19 | 0.822 | 56.414 | 0.001 | -0.001 | 1.393 | 77.219 | 29.148 | 2.192 | 65.692 | 0.856 | -0.971 | 0 | 0 | 1 | 42.688 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-18701 | 2022/01/07 18:01:23 | 1-06 | 4 | -4.951 | -39.139 | 12385 | 0 | 1 | UNDE1-3040-3 | 20-196 | H2O | 1641603988.98 | 0.186 | 0.776 | 62.459 | 0 | 0.001 | 1.543 | 76.605 | 29.148 | 2.151 | 63.674 | 0.857 | -1.024 | 0 | 0 | 1 | 42.812 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-18702 | 2022/01/07 18:11:13 | 1-07 | 1 | -11.504 | -89.068 | 12353 | -1 | 1 | WOOD1-3040-3 | 20-196 | H2O | 1641604579.11 | 0.203 | 0.638 | 61.4 | 0.001 | 0.002 | 1.509 | 76.677 | 28.727 | 2.062 | 62.257 | 0.856 | -1.075 | 0 | 0 | 1 | 42.938 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-18702 | 2022/01/07 18:20:10 | 1-07 | 2 | -11.863 | -94.623 | 12334 | -1 | 1 | WOOD1-3040-3 | 20-196 | H2O | 1641605115.81 | 0.195 | 0.91 | 57.257 | 0 | 0.004 | 1.409 | 76.637 | 28.269 | 2.058 | 63.261 | 0.856 | -0.836 | 0 | 0 | 1 | 43.062 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-18702 | 2022/01/07 18:29:07 | 1-07 | 3 | -11.948 | -95.648 | 12278 | -1 | 1 | WOOD1-3040-3 | 20-196 | H2O | 1641605652.09 | 0.187 | 0.807 | 58.506 | 0.001 | -0.002 | 1.446 | 76.489 | 28.375 | 2.045 | 62.042 | 0.856 | -0.87 | 0 | 0 | 1 | 43.062 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-18702 | 2022/01/07 18:38:03 | 1-07 | 4 | -11.968 | -96.383 | 12354 | 0 | 1 | WOOD1-3040-3 | 20-196 | H2O | 1641606188.31 | 0.212 | 0.756 | 56.132 | 0.001 | 0.002 | 1.389 | 76.623 | 28.439 | 2.063 | 63.183 | 0.856 | -0.933 | 0 | 0 | 1 | 43.125 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-18703 | 2022/01/07 18:47:50 | 1-08 | 1 | -5.414 | -47.239 | 12434 | -1 | 1 | UNDE1-deep-3 | 20-196 | H2O | 1641606775.3 | 0.198 | 0.849 | 56.694 | 0.001 | 0.004 | 1.404 | 76.852 | 28.971 | 2.132 | 63.254 | 0.856 | -1.067 | 0 | 0 | 1 | 43.25 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-18703 | 2022/01/07 18:56:47 | 1-08 | 2 | -5.099 | -42.083 | 12393 | -1 | 1 | UNDE1-deep-3 | 20-196 | H2O | 1641607312.5 | 0.193 | 0.853 | 56.884 | 0 | 0.001 | 1.41 | 76.865 | 28.604 | 2.144 | 63.283 | 0.856 | -1.126 | 0 | 0 | 1 | 43.312 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-18703 | 2022/01/07 19:05:44 | 1-08 | 3 | -5.014 | -40.515 | 12143 | -1 | 1 | UNDE1-deep-3 | 20-196 | H2O | 1641607849.32 | 0.204 | 0.84 | 57.782 | 0 | -0.001 | 1.428 | 76.319 | 28.147 | 2.101 | 62.6 | 0.856 | -1.185 | 0 | 0 | 1 | 43.312 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-18703 | 2022/01/07 19:14:40 | 1-08 | 4 | -4.967 | -39.993 | 12210 | 0 | 1 | UNDE1-deep-3 | 20-196 | H2O | 1641608385.41 | 0.191 | 0.718 | 54.801 | 0 | 0 | 1.352 | 76.476 | 28.332 | 2.12 | 63.4 | 0.856 | -1.131 | 0 | 0 | 1 | 43.375 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-18704 | 2022/01/07 19:24:27 | 1-09 | 1 | -10.634 | -85.241 | 12392 | -1 | 1 | WOOD1-2030-3 | 20-196 | H2O | 1641608972.59 | 0.213 | 0.856 | 56.274 | 0.001 | -0.004 | 1.393 | 76.856 | 28.934 | 2.078 | 61.687 | 0.856 | -1.136 | 0 | 0 | 1 | 43.5 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-18704 | 2022/01/07 19:33:24 | 1-09 | 2 | -10.948 | -90.009 | 12257 | -1 | 1 | WOOD1-2030-3 | 20-196 | H2O | 1641609509.85 | 0.194 | 0.736 | 58.482 | 0.001 | 0.003 | 1.438 | 76.571 | 28.476 | 2.043 | 61.765 | 0.856 | -1.067 | 0 | 0 | 1 | 43.562 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-18704 | 2022/01/07 19:42:20 | 1-09 | 3 | -10.997 | -91.305 | 12352 | -1 | 1 | WOOD1-2030-3 | 20-196 | H2O | 1641610046.04 | 0.195 | 0.78 | 59.543 | 0 | 0.006 | 1.471 | 76.762 | 28.74 | 2.062 | 61.931 | 0.856 | -1.12 | 0 | 0 | 1 | 43.625 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-18704 | 2022/01/07 19:51:17 | 1-09 | 4 | -11.022 | -91.689 | 12176 | 0 | 1 | WOOD1-2030-3 | 20-196 | H2O | 1641610582.79 | 0.2 | 0.702 | 55.577 | 0 | 0.002 | 1.372 | 76.511 | 28.125 | 2.027 | 61.073 | 0.856 | -1.132 | 0 | 0 | 1 | 43.562 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-18705 | 2022/01/07 20:01:06 | 1-10 | 1 | -9.572 | -73.175 | 12555 | -1 | 1 | WOOD1-deep-2 | 20-196 | H2O | 1641611171.14 | 0.208 | 0.778 | 59.261 | 0.001 | 0.004 | 1.453 | 77.092 | 29.443 | 2.13 | 63.361 | 0.856 | -1.071 | 0 | 0 | 1 | 43.688 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-18705 | 2022/01/07 20:10:02 | 1-10 | 2 | -9.463 | -71.211 | 12113 | -1 | 1 | WOOD1-deep-2 | 20-196 | H2O | 1641611707.92 | 0.212 | 0.747 | 63.194 | 0 | 0.001 | 1.572 | 76.163 | 28.029 | 2.056 | 61.918 | 0.856 | -0.969 | 0 | 0 | 1 | 43.625 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-18705 | 2022/01/07 20:18:59 | 1-10 | 3 | -9.443 | -70.971 | 12211 | -1 | 1 | WOOD1-deep-2 | 20-196 | H2O | 1641612243.96 | 0.22 | 0.707 | 61.858 | 0.002 | 0.003 | 1.529 | 76.328 | 28.344 | 2.082 | 62.527 | 0.856 | -0.967 | 0 | 0 | 1 | 43.562 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-18705 | 2022/01/07 20:27:55 | 1-10 | 4 | -9.456 | -70.636 | 12088 | 0 | 1 | WOOD1-deep-2 | 20-196 | H2O | 1641612780.74 | 0.2 | 0.806 | 56.717 | 0 | -0.001 | 1.399 | 76.065 | 27.904 | 2.055 | 61.924 | 0.856 | -0.968 | 0 | 0 | 1 | 43.562 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-18706 | 2022/01/07 20:37:44 | 1-11 | 1 | -9.724 | -75.512 | 12641 | -1 | 1 | WOOD1-3040-2 | 20-196 | H2O | 1641613368.68 | 0.201 | 0.732 | 58.522 | 0.001 | -0.001 | 1.45 | 77.305 | 29.312 | 2.126 | 64.109 | 0.856 | -0.985 | 0 | 0 | 1 | 43.5 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-18706 | 2022/01/07 20:46:40 | 1-11 | 2 | -9.724 | -75.831 | 12093 | -1 | 1 | WOOD1-3040-2 | 20-196 | H2O | 1641613905.44 | 0.182 | 0.844 | 56.466 | 0 | 0.002 | 1.391 | 75.987 | 27.553 | 2.071 | 63.563 | 0.856 | -0.625 | 0 | 0 | 1 | 43.5 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-18706 | 2022/01/07 20:55:36 | 1-11 | 3 | -9.707 | -76.004 | 12197 | -1 | 1 | WOOD1-3040-2 | 20-196 | H2O | 1641614441.4 | 0.209 | 0.807 | 59.466 | 0.001 | -0.001 | 1.471 | 76.392 | 28.269 | 2.058 | 62.073 | 0.856 | -0.978 | 0 | 0 | 1 | 43.438 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-18706 | 2022/01/07 21:04:32 | 1-11 | 4 | -9.756 | -75.975 | 12055 | 0 | 1 | WOOD1-3040-2 | 20-196 | H2O | 1641614977.47 | 0.203 | 0.871 | 56.587 | 0.001 | 0.002 | 1.405 | 76.147 | 27.675 | 2.036 | 61.228 | 0.856 | -0.995 | 0 | 0 | 1 | 43.375 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-18707 | 2022/01/07 21:14:20 | 1-12 | 1 | -2.688 | -46.215 | 11981 | -1 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641615566.02 | 0.256 | 1.064 | 59.288 | 0 | 0.002 | 1.472 | 638.339 | 6.579 | 2.422 | 53.787 | 0.858 | -2.812 | 0 | 0 | 1 | 43.438 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-18707 | 2022/01/07 21:23:16 | 1-12 | 2 | -2.377 | -42.97 | 11697 | -1 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641616102 | 0.243 | 1.119 | 58.345 | 0 | 0.002 | 1.456 | 628.378 | 6.162 | 2.366 | 52.114 | 0.857 | -2.928 | 0 | 0 | 1 | 43.5 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-18707 | 2022/01/07 21:32:12 | 1-12 | 3 | -2.296 | -42.212 | 11792 | -1 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641616638.35 | 0.263 | 1.024 | 52.829 | 0 | 0.001 | 1.318 | 634.216 | 6.498 | 2.385 | 52.805 | 0.857 | -2.954 | 0 | 0 | 1 | 43.562 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-18707 | 2022/01/07 21:41:09 | 1-12 | 4 | -2.278 | -42.003 | 11667 | 0 | 1 | SCBI1-D-5 | 20-196 | H2O | 1641617174.28 | 0.262 | 1.112 | 52.91 | 0 | 0.001 | 1.308 | 628.56 | 6.037 | 2.369 | 52.746 | 0.858 | -2.9 | 0 | 0 | 1 | 43.625 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-18708 | 2022/01/07 21:50:58 | 1-13 | 1 | -8.724 | -63.968 | 12770 | -1 | 1 | EV | standard | H2O | 1641617762.8 | 0.196 | 0.675 | 59.299 | 0.001 | 0.003 | 1.456 | 84.493 | 29.613 | 2.174 | 64.825 | 0.856 | -1.105 | 0 | 0 | 1 | 43.688 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-18708 | 2022/01/07 21:59:53 | 1-13 | 2 | -8.933 | -66.011 | 12345 | -1 | 1 | EV | standard | H2O | 1641618298.27 | 0.197 | 0.796 | 56.504 | 0.001 | 0.002 | 1.395 | 79.606 | 28.915 | 2.086 | 62.198 | 0.856 | -1.102 | 0 | 0 | 1 | 43.812 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-18708 | 2022/01/07 22:08:49 | 1-13 | 3 | -8.943 | -66.554 | 12309 | -1 | 1 | EV | standard | H2O | 1641618834.94 | 0.21 | 0.762 | 61.104 | 0.001 | 0.002 | 1.512 | 78.483 | 28.853 | 2.091 | 61.402 | 0.856 | -1.197 | 0 | 0 | 1 | 43.875 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-18708 | 2022/01/07 22:17:46 | 1-13 | 4 | -8.988 | -66.925 | 12312 | 0 | 1 | EV | standard | H2O | 1641619371.65 | 0.215 | 0.838 | 57.359 | 0 | 0.004 | 1.421 | 77.898 | 28.958 | 2.092 | 61.237 | 0.855 | -1.171 | 0 | 0 | 1 | 43.812 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-18709 | 2022/01/07 22:27:36 | 1-14 | 1 | -9.134 | -78.967 | 12456 | -1 | 1 | WOOD1-1020-3 | 20-196 | H2O | 1641619961.7 | 0.18 | 0.722 | 55.832 | 0 | 0 | 1.371 | 77.283 | 28.981 | 2.109 | 63.647 | 0.856 | -0.942 | 0 | 0 | 1 | 43.938 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-18709 | 2022/01/07 22:36:33 | 1-14 | 2 | -9.148 | -80.478 | 12289 | -1 | 1 | WOOD1-1020-3 | 20-196 | H2O | 1641620498.55 | 0.17 | 0.821 | 58.741 | 0 | 0.003 | 1.449 | 77.103 | 28.348 | 2.092 | 63.739 | 0.856 | -0.809 | 0 | 0 | 1 | 43.938 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-18709 | 2022/01/07 22:45:29 | 1-14 | 3 | -9.101 | -80.804 | 12253 | -1 | 1 | WOOD1-1020-3 | 20-196 | H2O | 1641621034.5 | 0.218 | 0.857 | 56.437 | 0 | 0.002 | 1.384 | 76.997 | 28.288 | 2.086 | 63.063 | 0.856 | -0.921 | 0 | 0 | 1 | 44 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-18709 | 2022/01/07 22:54:26 | 1-14 | 4 | -9.153 | -81.002 | 12054 | 0 | 1 | WOOD1-1020-3 | 20-196 | H2O | 1641621572.02 | 0.202 | 0.683 | 55.636 | 0.001 | 0.001 | 1.374 | 76.497 | 27.684 | 2.034 | 61.975 | 0.855 | -1.012 | 0 | 0 | 1 | 44.062 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-18710 | 2022/01/07 23:04:13 | 1-15 | 1 | -8.957 | -66.556 | 12393 | -1 | 1 | WOOD1-deep-1 | 20-196 | H2O | 1641622158.09 | 0.215 | 0.744 | 58.583 | 0.001 | 0.001 | 1.451 | 76.844 | 28.726 | 2.113 | 64.046 | 0.855 | -0.832 | 0 | 0 | 1 | 44.062 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-18710 | 2022/01/07 23:13:08 | 1-15 | 2 | -8.948 | -64.871 | 12243 | -1 | 1 | WOOD1-deep-1 | 20-196 | H2O | 1641622693.94 | 0.201 | 0.83 | 57.4 | 0.001 | -0.001 | 1.419 | 76.679 | 28.368 | 2.091 | 62.947 | 0.855 | -0.87 | 0 | 0 | 1 | 44.062 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-18710 | 2022/01/07 23:22:05 | 1-15 | 3 | -8.969 | -64.49 | 12176 | -1 | 1 | WOOD1-deep-1 | 20-196 | H2O | 1641623230.63 | 0.189 | 0.758 | 55.636 | 0.001 | 0 | 1.374 | 76.611 | 28.121 | 2.085 | 62.44 | 0.856 | -1.005 | 0 | 0 | 1 | 44.062 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-18710 | 2022/01/07 23:31:01 | 1-15 | 4 | -9.003 | -64.409 | 12125 | 0 | 1 | WOOD1-deep-1 | 20-196 | H2O | 1641623766.84 | 0.214 | 0.845 | 54.598 | 0.001 | -0.001 | 1.344 | 76.407 | 27.858 | 2.075 | 63.22 | 0.856 | -0.873 | 0 | 0 | 1 | 44 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-18711 | 2022/01/07 23:40:48 | 1-16 | 1 | -2.358 | -58.189 | 12513 | -1 | 1 | WOOD1-010-1 | 20-196 | H2O | 1641624354 | 0.197 | 0.849 | 59.607 | 0.001 | -0.002 | 1.472 | 77.017 | 29.165 | 2.152 | 64.077 | 0.856 | -0.953 | 0 | 0 | 1 | 44 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-18711 | 2022/01/07 23:49:45 | 1-16 | 2 | -1.98 | -57.237 | 12132 | -1 | 1 | WOOD1-010-1 | 20-196 | H2O | 1641624890.21 | 0.192 | 0.681 | 57.13 | 0 | 0 | 1.411 | 76.299 | 28.319 | 2.082 | 61.727 | 0.855 | -1.018 | 0 | 0 | 1 | 44 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-18711 | 2022/01/07 23:58:40 | 1-16 | 3 | -1.953 | -57.115 | 12245 | -1 | 1 | WOOD1-010-1 | 20-196 | H2O | 1641625425.69 | 0.184 | 0.804 | 55.16 | 0.001 | -0.003 | 1.362 | 76.607 | 28.227 | 2.106 | 63.381 | 0.856 | -0.947 | 0 | 0 | 1 | 44.062 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-18711 | 2022/01/08 00:07:36 | 1-16 | 4 | -1.88 | -56.972 | 12126 | 0 | 1 | WOOD1-010-1 | 20-196 | H2O | 1641625962.32 | 0.194 | 0.83 | 57.195 | 0 | 0 | 1.416 | 76.384 | 27.907 | 2.084 | 63.022 | 0.855 | -0.978 | 0 | 0 | 1 | 44 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-18712 | 2022/01/08 00:17:23 | 1-17 | 1 | -6.145 | -68.652 | 12484 | -1 | 1 | WOOD1-010-3 | 20-196 | H2O | 1641626549.16 | 0.174 | 0.793 | 57.674 | 0.001 | 0.003 | 1.414 | 77.014 | 29.017 | 2.117 | 63.805 | 0.856 | -0.921 | 0 | 0 | 1 | 44.062 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-18712 | 2022/01/08 00:26:19 | 1-17 | 2 | -6.393 | -70.015 | 12288 | -1 | 1 | WOOD1-010-3 | 20-196 | H2O | 1641627085.1 | 0.209 | 0.769 | 55.233 | 0 | 0.002 | 1.362 | 76.586 | 28.492 | 2.089 | 62.781 | 0.855 | -0.888 | 0 | 0 | 1 | 44 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-18712 | 2022/01/08 00:35:16 | 1-17 | 3 | -6.425 | -70.474 | 12218 | -1 | 1 | WOOD1-010-3 | 20-196 | H2O | 1641627621.81 | 0.191 | 0.78 | 54.459 | 0 | 0 | 1.342 | 76.497 | 28.296 | 2.089 | 61.941 | 0.856 | -0.967 | 0 | 0 | 1 | 43.938 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-18712 | 2022/01/08 00:44:12 | 1-17 | 4 | -6.428 | -70.519 | 12205 | 0 | 1 | WOOD1-010-3 | 20-196 | H2O | 1641628158.12 | 0.202 | 0.781 | 57.439 | 0.001 | 0.002 | 1.419 | 76.379 | 28.155 | 2.077 | 63.054 | 0.856 | -0.855 | 0 | 0 | 1 | 44 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-18713 | 2022/01/08 00:53:59 | 1-18 | 1 | -8.469 | -68.364 | 12452 | -1 | 1 | WOOD1-6070-1 | 20-196 | H2O | 1641628744.25 | 0.209 | 0.744 | 56.11 | 0 | 0.001 | 1.383 | 76.919 | 28.995 | 2.11 | 62.986 | 0.856 | -1.02 | 0 | 0 | 1 | 44.125 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-18713 | 2022/01/08 01:02:55 | 1-18 | 2 | -8.569 | -68.213 | 12219 | -1 | 1 | WOOD1-6070-1 | 20-196 | H2O | 1641629280.92 | 0.214 | 0.769 | 56.348 | 0.001 | 0.004 | 1.403 | 76.452 | 28.212 | 2.075 | 62.385 | 0.856 | -1.045 | 0 | 0 | 1 | 44.125 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-18713 | 2022/01/08 01:11:51 | 1-18 | 3 | -8.581 | -68.051 | 12130 | -1 | 1 | WOOD1-6070-1 | 20-196 | H2O | 1641629816.44 | 0.199 | 0.785 | 58.248 | 0.001 | 0.006 | 1.44 | 76.386 | 27.595 | 2.06 | 62.502 | 0.857 | -1.004 | 0 | 0 | 1 | 44.125 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-18713 | 2022/01/08 01:20:47 | 1-18 | 4 | -8.585 | -68.275 | 12169 | 0 | 1 | WOOD1-6070-1 | 20-196 | H2O | 1641630352.66 | 0.204 | 0.809 | 57.62 | 0.001 | 0 | 1.424 | 76.352 | 28.07 | 2.077 | 63.039 | 0.856 | -1.007 | 0 | 0 | 1 | 44.062 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-18714 | 2022/01/08 01:30:34 | 1-19 | 1 | -7.602 | -67.599 | 12195 | -1 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641630940.31 | 0.205 | 0.794 | 59.864 | 0 | 0.002 | 1.476 | 76.301 | 28.185 | 2.088 | 63.044 | 0.856 | -0.926 | 0 | 0 | 1 | 44.125 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-18714 | 2022/01/08 01:39:30 | 1-19 | 2 | -7.549 | -67.297 | 12157 | -1 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641631476.13 | 0.201 | 0.746 | 56.778 | 0 | -0.002 | 1.402 | 76.306 | 28.063 | 2.068 | 62.581 | 0.855 | -0.906 | 0 | 0 | 1 | 44.062 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-18714 | 2022/01/08 01:48:27 | 1-19 | 3 | -7.52 | -67.238 | 12088 | -1 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641632012.51 | 0.205 | 0.814 | 55.924 | 0 | -0.002 | 1.387 | 76.075 | 28.358 | 2.062 | 61.232 | 0.856 | -1.004 | 0 | 0 | 1 | 44 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-18714 | 2022/01/08 01:57:24 | 1-19 | 4 | -7.534 | -67.102 | 12068 | 0 | 1 | WOOD1-2030-1 | 20-196 | H2O | 1641632549.39 | 0.197 | 0.759 | 52.896 | 0.001 | 0.001 | 1.317 | 76.047 | 28.253 | 2.047 | 61.386 | 0.856 | -1.086 | 0 | 0 | 1 | 43.938 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-18715 | 2022/01/08 02:07:12 | 1-20 | 1 | -8.234 | -70.514 | 12380 | -1 | 1 | WOOD1-1020-2 | 20-196 | H2O | 1641633137.25 | 0.209 | 0.759 | 59.48 | 0.001 | -0.001 | 1.463 | 76.68 | 28.898 | 2.094 | 63.272 | 0.856 | -0.967 | 0 | 0 | 1 | 43.875 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-18715 | 2022/01/08 02:16:08 | 1-20 | 2 | -8.22 | -70.867 | 12092 | -1 | 1 | WOOD1-1020-2 | 20-196 | H2O | 1641633674.14 | 0.203 | 0.768 | 54.74 | 0.001 | 0.001 | 1.353 | 76.37 | 27.977 | 2.043 | 60.997 | 0.856 | -1.222 | 0 | 0 | 1 | 43.812 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-18715 | 2022/01/08 02:25:05 | 1-20 | 3 | -8.274 | -70.938 | 12071 | -1 | 1 | WOOD1-1020-2 | 20-196 | H2O | 1641634210.66 | 0.203 | 0.791 | 56.445 | 0.001 | 0.003 | 1.396 | 76.187 | 27.938 | 2.045 | 61.428 | 0.856 | -1.059 | 0 | 0 | 1 | 43.688 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-18715 | 2022/01/08 02:34:02 | 1-20 | 4 | -8.262 | -70.995 | 12081 | 0 | 1 | WOOD1-1020-2 | 20-196 | H2O | 1641634746.89 | 0.174 | 0.895 | 58.095 | 0.001 | 0.001 | 1.444 | 76.168 | 28.041 | 2.035 | 61.458 | 0.856 | -1.059 | 0 | 0 | 1 | 43.688 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-18716 | 2022/01/08 02:43:50 | 1-21 | 1 | -6.973 | -67.348 | 12415 | -1 | 1 | WOOD1-1020-1 | 20-196 | H2O | 1641635335.89 | 0.189 | 0.751 | 56.193 | 0 | 0.002 | 1.377 | 76.89 | 28.934 | 2.107 | 62.877 | 0.856 | -1.127 | 0 | 0 | 1 | 43.625 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-18716 | 2022/01/08 02:52:47 | 1-21 | 2 | -6.854 | -66.985 | 12144 | -1 | 1 | WOOD1-1020-1 | 20-196 | H2O | 1641635872.7 | 0.204 | 0.846 | 53.389 | 0 | 0.001 | 1.325 | 76.254 | 27.859 | 2.073 | 62.775 | 0.856 | -0.951 | 0 | 0 | 1 | 43.562 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-18716 | 2022/01/08 03:01:43 | 1-21 | 3 | -6.864 | -67.025 | 12098 | -1 | 1 | WOOD1-1020-1 | 20-196 | H2O | 1641636408.54 | 0.179 | 0.684 | 58.734 | 0 | 0.001 | 1.462 | 76.161 | 27.845 | 2.064 | 62.619 | 0.856 | -0.989 | 0 | 0 | 1 | 43.562 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-18716 | 2022/01/08 03:10:40 | 1-21 | 4 | -6.829 | -66.89 | 12130 | 0 | 1 | WOOD1-1020-1 | 20-196 | H2O | 1641636945.88 | 0.2 | 0.828 | 51.368 | 0.001 | 0.001 | 1.264 | 76.235 | 28.551 | 2.049 | 60.761 | 0.856 | -1.157 | 0 | 0 | 1 | 43.562 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-18717 | 2022/01/08 03:20:26 | 1-22 | 1 | -8.897 | -67.062 | 12303 | -1 | 1 | EV | standard | H2O | 1641637530.96 | 0.2 | 0.772 | 52.615 | 0.001 | -0.003 | 1.299 | 76.522 | 28.338 | 2.093 | 63.037 | 0.856 | -0.937 | 0 | 0 | 1 | 43.562 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-18717 | 2022/01/08 03:29:22 | 1-22 | 2 | -8.966 | -66.928 | 12360 | -1 | 1 | EV | standard | H2O | 1641638067.31 | 0.197 | 0.851 | 56.688 | 0 | 0.004 | 1.4 | 76.845 | 28.729 | 2.095 | 62.385 | 0.856 | -1.288 | 0 | 0 | 1 | 43.5 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-18717 | 2022/01/08 03:38:17 | 1-22 | 3 | -9.031 | -67.065 | 12306 | -1 | 1 | EV | standard | H2O | 1641638602.35 | 0.208 | 0.759 | 56.857 | 0.001 | -0.001 | 1.407 | 76.464 | 28.844 | 2.092 | 62.722 | 0.856 | -1.018 | 0 | 0 | 1 | 43.5 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-18717 | 2022/01/08 03:47:13 | 1-22 | 4 | -8.983 | -67.014 | 12324 | 0 | 1 | EV | standard | H2O | 1641639138.32 | 0.189 | 0.777 | 50.036 | 0 | -0.002 | 1.233 | 76.63 | 28.548 | 2.096 | 62.944 | 0.855 | -0.993 | 0 | 0 | 1 | 43.562 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-18718 | 2022/01/08 03:56:59 | 1-23 | 1 | -15.233 | -110.324 | 12521 | -1 | 1 | SLCP-22-001 | Precipitation | H2O | 1641639724.33 | 0.222 | 0.804 | 57.727 | 0.001 | 0 | 1.427 | 76.927 | 29.182 | 2.063 | 62.704 | 0.856 | -0.936 | 0 | 0 | 1 | 43.562 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-18718 | 2022/01/08 04:05:55 | 1-23 | 2 | -15.56 | -114.864 | 12228 | -1 | 1 | SLCP-22-001 | Precipitation | H2O | 1641640260.35 | 0.209 | 0.812 | 55.639 | 0.001 | 0.002 | 1.371 | 76.542 | 28.122 | 2.001 | 61.251 | 0.856 | -0.98 | 0 | 0 | 1 | 43.438 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-18718 | 2022/01/08 04:14:51 | 1-23 | 3 | -15.654 | -116.219 | 12493 | -1 | 1 | SLCP-22-001 | Precipitation | H2O | 1641640796.3 | 0.217 | 0.766 | 55.422 | 0.001 | 0.003 | 1.372 | 76.86 | 29.317 | 2.039 | 61.851 | 0.856 | -0.911 | 0 | 0 | 1 | 43.438 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-18718 | 2022/01/08 04:23:47 | 1-23 | 4 | -15.666 | -116.484 | 12462 | 0 | 1 | SLCP-22-001 | Precipitation | H2O | 1641641332.79 | 0.203 | 0.841 | 57.163 | 0.001 | 0.003 | 1.404 | 76.899 | 28.95 | 2.047 | 62.19 | 0.856 | -0.906 | 0 | 0 | 1 | 43.5 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-18719 | 2022/01/08 04:33:33 | 1-24 | 1 | -5.898 | -46.216 | 12245 | -1 | 1 | 57 | 21-252 | H2O | 1641641918.84 | 0.177 | 0.837 | 60.347 | -0.001 | 0.006 | 1.502 | 76.324 | 28.481 | 2.109 | 63.728 | 0.856 | -0.844 | 0 | 0 | 1 | 43.562 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-18719 | 2022/01/08 04:42:30 | 1-24 | 2 | -5.075 | -37.987 | 12275 | -1 | 1 | 57 | 21-252 | H2O | 1641642455.38 | 0.18 | 0.825 | 62.059 | 0 | -0.002 | 1.533 | 76.502 | 28.306 | 2.122 | 64.042 | 0.856 | -0.977 | 0 | 0 | 1 | 43.562 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-18719 | 2022/01/08 04:51:25 | 1-24 | 3 | -5.02 | -36.292 | 12614 | -1 | 1 | 57 | 21-252 | H2O | 1641642990.84 | 0.206 | 0.749 | 58.168 | 0 | 0.001 | 1.427 | 77.21 | 29.682 | 2.188 | 64.592 | 0.856 | -1.145 | 0 | 0 | 1 | 43.562 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-18719 | 2022/01/08 05:00:21 | 1-24 | 4 | -5.123 | -35.806 | 12117 | 0 | 1 | 57 | 21-252 | H2O | 1641643526.87 | 0.23 | 0.737 | 57.572 | 0.001 | -0.002 | 1.424 | 76.123 | 28.003 | 2.115 | 63.507 | 0.856 | -0.959 | 0 | 0 | 1 | 43.562 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-18720 | 2022/01/08 05:10:07 | 1-25 | 1 | -5.228 | -35.654 | 12022 | -1 | 1 | 58 | 21-252 | H2O | 1641644112.84 | 0.208 | 0.801 | 57.923 | 0.001 | -0.001 | 1.434 | 75.979 | 27.95 | 2.088 | 62.116 | 0.855 | -1.009 | 0 | 0 | 1 | 43.625 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-18720 | 2022/01/08 05:19:03 | 1-25 | 2 | -5.262 | -35.444 | 11636 | -1 | 1 | 58 | 21-252 | H2O | 1641644648.62 | 0.199 | 0.708 | 53.425 | 0.001 | 0.001 | 1.329 | 75.189 | 26.927 | 2.028 | 60.654 | 0.855 | -1.021 | 0 | 0 | 1 | 43.625 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-18720 | 2022/01/08 05:27:59 | 1-25 | 3 | -5.267 | -35.08 | 11949 | -1 | 1 | 58 | 21-252 | H2O | 1641645184.43 | 0.204 | 0.789 | 59.408 | 0.001 | -0.003 | 1.477 | 75.749 | 27.572 | 2.084 | 62.321 | 0.856 | -0.947 | 0 | 0 | 1 | 43.688 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-18720 | 2022/01/08 05:36:54 | 1-25 | 4 | -5.271 | -35.033 | 11865 | 0 | 1 | 58 | 21-252 | H2O | 1641645719.86 | 0.195 | 0.769 | 54.382 | 0 | 0.001 | 1.342 | 75.62 | 27.32 | 2.07 | 62.126 | 0.856 | -0.965 | 0 | 0 | 1 | 43.688 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-18721 | 2022/01/08 05:46:41 | 1-26 | 1 | -5.134 | -34.252 | 12109 | -1 | 1 | 59 | 21-252 | H2O | 1641646306.79 | 0.218 | 0.749 | 61.134 | 0.001 | 0.002 | 1.511 | 76.126 | 27.945 | 2.106 | 62.97 | 0.856 | -0.984 | 0 | 0 | 1 | 43.75 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-18721 | 2022/01/08 05:55:37 | 1-26 | 2 | -5.041 | -33.526 | 11819 | -1 | 1 | 59 | 21-252 | H2O | 1641646843.06 | 0.212 | 0.819 | 58.459 | 0.001 | 0.004 | 1.45 | 75.554 | 27.074 | 2.071 | 62.1 | 0.855 | -0.962 | 0 | 0 | 1 | 43.688 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-18721 | 2022/01/08 06:04:33 | 1-26 | 3 | -5.111 | -33.726 | 11897 | -1 | 1 | 59 | 21-252 | H2O | 1641647378.43 | 0.219 | 0.761 | 54.823 | 0 | 0.001 | 1.37 | 75.788 | 27.55 | 2.076 | 61.514 | 0.856 | -1.095 | 0 | 0 | 1 | 43.75 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-18721 | 2022/01/08 06:13:29 | 1-26 | 4 | -5.108 | -33.952 | 11941 | 0 | 1 | 59 | 21-252 | H2O | 1641647915.45 | 0.215 | 0.857 | 59.324 | 0 | 0.001 | 1.467 | 75.615 | 27.462 | 2.09 | 63.248 | 0.855 | -0.709 | 0 | 0 | 1 | 43.812 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-18722 | 2022/01/08 06:23:16 | 1-27 | 1 | -5.349 | -35.493 | 12197 | -1 | 1 | 60 | 21-252 | H2O | 1641648501.45 | 0.182 | 0.784 | 56.179 | 0.001 | -0.001 | 1.388 | 76.297 | 28.378 | 2.128 | 63.066 | 0.856 | -0.966 | 0 | 0 | 1 | 43.938 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-18722 | 2022/01/08 06:32:11 | 1-27 | 2 | -5.365 | -35.942 | 11975 | -1 | 1 | 60 | 21-252 | H2O | 1641649037.24 | 0.208 | 0.871 | 59.865 | 0.001 | 0.004 | 1.477 | 75.901 | 28.018 | 2.074 | 61.213 | 0.856 | -1.14 | 0 | 0 | 1 | 44 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-18722 | 2022/01/08 06:41:07 | 1-27 | 3 | -5.307 | -35.62 | 11804 | -1 | 1 | 60 | 21-252 | H2O | 1641649573.04 | 0.222 | 0.853 | 52.563 | 0 | 0 | 1.297 | 75.551 | 26.728 | 2.072 | 62.365 | 0.855 | -0.853 | 0 | 0 | 1 | 43.938 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-18722 | 2022/01/08 06:50:03 | 1-27 | 4 | -5.35 | -35.711 | 11883 | 0 | 1 | 60 | 21-252 | H2O | 1641650109.8 | 0.212 | 0.721 | 55.781 | 0.001 | -0.002 | 1.386 | 75.754 | 27.372 | 2.076 | 61.869 | 0.856 | -1.048 | 0 | 0 | 1 | 43.875 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-18723 | 2022/01/08 06:59:51 | 1-28 | 1 | -5.406 | -36.018 | 12154 | -1 | 1 | 61 | 21-252 | H2O | 1641650697.12 | 0.203 | 0.8 | 58.024 | 0 | -0.001 | 1.435 | 76.251 | 28.248 | 2.114 | 62.558 | 0.856 | -1.01 | 0 | 0 | 1 | 43.812 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-18723 | 2022/01/08 07:08:47 | 1-28 | 2 | -5.399 | -35.852 | 12016 | -1 | 1 | 61 | 21-252 | H2O | 1641651232.97 | 0.208 | 0.819 | 54.758 | 0.001 | 0.001 | 1.35 | 76.099 | 27.546 | 2.081 | 62.137 | 0.855 | -1.057 | 0 | 0 | 1 | 43.75 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-18723 | 2022/01/08 07:17:43 | 1-28 | 3 | -5.433 | -35.918 | 12153 | -1 | 1 | 61 | 21-252 | H2O | 1641651768.66 | 0.206 | 0.844 | 58.654 | 0.001 | 0.003 | 1.451 | 76.197 | 28.213 | 2.111 | 62.804 | 0.856 | -1.048 | 0 | 0 | 1 | 43.75 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-18723 | 2022/01/08 07:26:40 | 1-28 | 4 | -5.419 | -35.949 | 12232 | 0 | 1 | 61 | 21-252 | H2O | 1641652305.43 | 0.192 | 0.837 | 54.658 | 0 | -0.001 | 1.351 | 76.335 | 28.49 | 2.123 | 62.983 | 0.855 | -1.03 | 0 | 0 | 1 | 43.75 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-18724 | 2022/01/08 07:36:24 | 1-29 | 1 | -5.48 | -36.283 | 12137 | -1 | 1 | 62 | 21-252 | H2O | 1641652889.11 | 0.207 | 0.739 | 58.114 | 0.001 | 0.001 | 1.434 | 76.2 | 28.108 | 2.112 | 62.759 | 0.856 | -0.927 | 0 | 0 | 1 | 43.75 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-18724 | 2022/01/08 07:45:18 | 1-29 | 2 | -5.466 | -36.389 | 12179 | -1 | 1 | 62 | 21-252 | H2O | 1641653424.22 | 0.203 | 0.793 | 53.598 | 0.002 | -0.002 | 1.325 | 76.282 | 28.398 | 2.11 | 62.725 | 0.856 | -1.058 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-18724 | 2022/01/08 07:54:15 | 1-29 | 3 | -5.495 | -36.527 | 11735 | -1 | 1 | 62 | 21-252 | H2O | 1641653960.5 | 0.218 | 0.843 | 59.353 | 0.001 | -0.001 | 1.469 | 75.287 | 26.988 | 2.058 | 61.711 | 0.855 | -0.799 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-18724 | 2022/01/08 08:03:11 | 1-29 | 4 | -5.473 | -36.407 | 12164 | 0 | 1 | 62 | 21-252 | H2O | 1641654495.75 | 0.205 | 0.807 | 54.764 | 0.001 | 0.002 | 1.352 | 76.174 | 28.353 | 2.115 | 63.03 | 0.856 | -0.873 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-18725 | 2022/01/08 08:12:55 | 1-30 | 1 | -5.269 | -35.428 | 12146 | -1 | 1 | 63 | 21-252 | H2O | 1641655080.83 | 0.205 | 0.792 | 54.571 | 0 | 0.003 | 1.345 | 76.14 | 28.312 | 2.12 | 63.057 | 0.855 | -1.019 | 0 | 0 | 1 | 43.688 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-18725 | 2022/01/08 08:21:51 | 1-30 | 2 | -5.286 | -35.193 | 12069 | -1 | 1 | 63 | 21-252 | H2O | 1641655616.65 | 0.202 | 0.874 | 60.668 | 0 | 0.004 | 1.491 | 76.151 | 28.106 | 2.097 | 61.63 | 0.856 | -1.143 | 0 | 0 | 1 | 43.688 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-18725 | 2022/01/08 08:30:47 | 1-30 | 3 | -5.254 | -35.245 | 12091 | -1 | 1 | 63 | 21-252 | H2O | 1641656152.64 | 0.221 | 0.809 | 54.014 | 0.001 | -0.001 | 1.337 | 75.996 | 28.316 | 2.101 | 62.777 | 0.856 | -0.936 | 0 | 0 | 1 | 43.688 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-18725 | 2022/01/08 08:39:43 | 1-30 | 4 | -5.232 | -35.254 | 12057 | 0 | 1 | 63 | 21-252 | H2O | 1641656688.44 | 0.228 | 0.742 | 58.336 | 0.001 | 0.004 | 1.441 | 75.892 | 27.856 | 2.109 | 63.712 | 0.856 | -0.763 | 0 | 0 | 1 | 43.625 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-18726 | 2022/01/08 08:49:28 | 1-31 | 1 | -8.684 | -62.256 | 12517 | -1 | 1 | EV | standard | H2O | 1641657273.52 | 0.2 | 0.742 | 60.087 | 0.001 | 0.001 | 1.475 | 76.97 | 29.283 | 2.134 | 63.765 | 0.856 | -1.051 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-18726 | 2022/01/08 08:58:24 | 1-31 | 2 | -8.92 | -65.068 | 12447 | -1 | 1 | EV | standard | H2O | 1641657809.79 | 0.186 | 0.669 | 54.873 | 0.001 | 0.001 | 1.354 | 76.943 | 29.097 | 2.108 | 62.844 | 0.856 | -1.175 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-18726 | 2022/01/08 09:07:21 | 1-31 | 3 | -8.982 | -65.851 | 12224 | -1 | 1 | EV | standard | H2O | 1641658346.06 | 0.2 | 0.689 | 54.242 | 0.001 | 0 | 1.346 | 76.373 | 28.148 | 2.089 | 63.02 | 0.855 | -0.932 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-18726 | 2022/01/08 09:16:17 | 1-31 | 4 | -9.018 | -66.192 | 12256 | 0 | 1 | EV | standard | H2O | 1641658882.26 | 0.202 | 0.794 | 59.245 | 0.001 | 0.003 | 1.455 | 76.455 | 28.344 | 2.085 | 62.553 | 0.855 | -1.013 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-18727 | 2022/01/08 09:26:02 | 1-32 | 1 | -5.679 | -40.386 | 12162 | -1 | 1 | 64 | 21-252 | H2O | 1641659467.64 | 0.192 | 0.71 | 57.91 | 0 | 0.002 | 1.427 | 76.189 | 28.324 | 2.111 | 62.575 | 0.856 | -1.114 | 0 | 0 | 1 | 43.688 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-18727 | 2022/01/08 09:34:59 | 1-32 | 2 | -5.504 | -37.577 | 12037 | -1 | 1 | 64 | 21-252 | H2O | 1641660004.35 | 0.199 | 0.855 | 54.759 | 0.001 | 0.002 | 1.354 | 76.092 | 27.578 | 2.096 | 62.634 | 0.856 | -1.037 | 0 | 0 | 1 | 43.688 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-18727 | 2022/01/08 09:43:55 | 1-32 | 3 | -5.467 | -36.829 | 11992 | -1 | 1 | 64 | 21-252 | H2O | 1641660540.38 | 0.197 | 0.856 | 55.298 | 0.001 | -0.001 | 1.372 | 75.951 | 27.916 | 2.078 | 61.526 | 0.856 | -1.106 | 0 | 0 | 1 | 43.625 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-18727 | 2022/01/08 09:52:51 | 1-32 | 4 | -5.433 | -36.49 | 12026 | 0 | 1 | 64 | 21-252 | H2O | 1641661076.33 | 0.198 | 0.742 | 55.815 | 0 | 0.001 | 1.379 | 76.062 | 27.662 | 2.098 | 63.073 | 0.856 | -1.064 | 0 | 0 | 1 | 43.625 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-18728 | 2022/01/08 10:02:37 | 1-33 | 1 | -5.461 | -36.59 | 12015 | -1 | 1 | 65 | 21-252 | H2O | 1641661662.98 | 0.196 | 0.77 | 54.496 | 0 | 0 | 1.347 | 76.105 | 27.682 | 2.096 | 62.125 | 0.855 | -1.128 | 0 | 0 | 1 | 43.562 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-18728 | 2022/01/08 10:11:33 | 1-33 | 2 | -5.51 | -36.432 | 11728 | -1 | 1 | 65 | 21-252 | H2O | 1641662198.85 | 0.199 | 0.782 | 55.544 | 0 | -0.001 | 1.383 | 75.297 | 26.876 | 2.051 | 61.876 | 0.855 | -0.934 | 0 | 0 | 1 | 43.5 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-18728 | 2022/01/08 10:20:29 | 1-33 | 3 | -5.468 | -36.413 | 11730 | -1 | 1 | 65 | 21-252 | H2O | 1641662735.11 | 0.177 | 0.863 | 54.053 | 0 | -0.002 | 1.336 | 75.352 | 26.847 | 2.037 | 61.778 | 0.856 | -0.889 | 0 | 0 | 1 | 43.438 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-18728 | 2022/01/08 10:29:25 | 1-33 | 4 | -5.475 | -36.483 | 11727 | 0 | 1 | 65 | 21-252 | H2O | 1641663271 | 0.201 | 0.844 | 53.255 | 0.001 | 0.003 | 1.325 | 75.193 | 27.087 | 2.038 | 61.561 | 0.855 | -0.875 | 0 | 0 | 1 | 43.438 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-18729 | 2022/01/08 10:39:12 | 1-34 | 1 | -5.461 | -36.142 | 11736 | -1 | 1 | 66 | 21-252 | H2O | 1641663857.54 | 0.193 | 0.786 | 59.376 | 0 | 0.004 | 1.487 | 75.419 | 26.887 | 2.042 | 60.62 | 0.855 | -0.982 | 0 | 0 | 1 | 43.312 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-18729 | 2022/01/08 10:48:07 | 1-34 | 2 | -5.535 | -36.476 | 12152 | -1 | 1 | 66 | 21-252 | H2O | 1641664392.44 | 0.178 | 0.812 | 52.545 | -0.001 | 0.001 | 1.297 | 76.143 | 27.949 | 2.116 | 63.731 | 0.856 | -0.863 | 0 | 0 | 1 | 43.312 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-18729 | 2022/01/08 10:57:02 | 1-34 | 3 | -5.49 | -36.409 | 11540 | -1 | 1 | 66 | 21-252 | H2O | 1641664927.96 | 0.21 | 0.768 | 57.241 | 0.001 | 0 | 1.427 | 74.864 | 26.451 | 2.01 | 61.103 | 0.856 | -0.922 | 0 | 0 | 1 | 43.312 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-18729 | 2022/01/08 11:05:58 | 1-34 | 4 | -5.472 | -36.345 | 11963 | 0 | 1 | 66 | 21-252 | H2O | 1641665463.62 | 0.182 | 0.811 | 56.041 | 0.001 | 0 | 1.388 | 75.79 | 27.55 | 2.083 | 62.453 | 0.855 | -0.957 | 0 | 0 | 1 | 43.312 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-18730 | 2022/01/08 11:15:45 | 1-35 | 1 | -5.46 | -36.286 | 12011 | -1 | 1 | 67 | 21-252 | H2O | 1641666050.89 | 0.199 | 0.814 | 55.324 | 0 | 0 | 1.376 | 76.005 | 27.529 | 2.092 | 62.613 | 0.856 | -1.025 | 0 | 0 | 1 | 43.375 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-18730 | 2022/01/08 11:24:42 | 1-35 | 2 | -5.486 | -36.39 | 11498 | -1 | 1 | 67 | 21-252 | H2O | 1641666587.96 | 0.202 | 0.8 | 53.389 | 0.001 | 0.001 | 1.318 | 74.759 | 26.431 | 2.014 | 60.499 | 0.856 | -0.956 | 0 | 0 | 1 | 43.438 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-18730 | 2022/01/08 11:33:38 | 1-35 | 3 | -5.418 | -36.089 | 11123 | -1 | 1 | 67 | 21-252 | H2O | 1641667123.83 | 0.206 | 0.864 | 52.283 | 0 | 0.002 | 1.315 | 74.21 | 25.154 | 1.953 | 58.632 | 0.855 | -0.981 | 0 | 0 | 1 | 43.438 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-18730 | 2022/01/08 11:42:34 | 1-35 | 4 | -5.456 | -36.281 | 11260 | 0 | 1 | 67 | 21-252 | H2O | 1641667659.61 | 0.21 | 0.989 | 57.294 | 0.001 | 0.002 | 1.436 | 74.55 | 25.603 | 1.963 | 58.722 | 0.855 | -1.094 | 0 | 0 | 1 | 43.375 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-18731 | 2022/01/08 11:52:16 | 1-36 | 1 | -5.523 | -36.343 | 12230 | -1 | 1 | 68 | 21-252 | H2O | 1641668242 | 0.213 | 0.811 | 59.42 | 0.001 | 0.005 | 1.468 | 76.254 | 28.509 | 2.129 | 63.208 | 0.856 | -0.966 | 0 | 0 | 1 | 43.375 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-18731 | 2022/01/08 12:01:12 | 1-36 | 2 | -5.511 | -36.192 | 12217 | -1 | 1 | 68 | 21-252 | H2O | 1641668777.47 | 0.204 | 0.725 | 52.551 | 0.001 | -0.001 | 1.291 | 76.297 | 28.295 | 2.136 | 63.595 | 0.856 | -1.044 | 0 | 0 | 1 | 43.375 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-18731 | 2022/01/08 12:10:08 | 1-36 | 3 | -5.503 | -36.251 | 12062 | -1 | 1 | 68 | 21-252 | H2O | 1641669313.55 | 0.211 | 0.83 | 55.141 | 0.001 | 0.003 | 1.363 | 76.028 | 27.697 | 2.107 | 62.932 | 0.855 | -0.955 | 0 | 0 | 1 | 43.312 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-18731 | 2022/01/08 12:19:03 | 1-36 | 4 | -5.559 | -36.446 | 12039 | 0 | 1 | 68 | 21-252 | H2O | 1641669848.11 | 0.2 | 0.823 | 57.892 | 0.001 | 0 | 1.439 | 76.001 | 27.725 | 2.101 | 62.232 | 0.856 | -1.08 | 0 | 0 | 1 | 43.375 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-18732 | 2022/01/08 12:28:46 | 1-37 | 1 | -5.345 | -35.681 | 11648 | -1 | 1 | 69 | 21-252 | H2O | 1641670431.8 | 0.194 | 0.721 | 53.533 | 0.001 | 0.001 | 1.331 | 75.199 | 26.534 | 2.043 | 61.806 | 0.856 | -0.945 | 0 | 0 | 1 | 43.438 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-18732 | 2022/01/08 12:37:41 | 1-37 | 2 | -5.339 | -35.752 | 11923 | -1 | 1 | 69 | 21-252 | H2O | 1641670966.61 | 0.219 | 0.886 | 56.443 | 0.001 | 0.004 | 1.405 | 75.738 | 27.654 | 2.081 | 61.446 | 0.855 | -1.034 | 0 | 0 | 1 | 43.5 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-18732 | 2022/01/08 12:46:37 | 1-37 | 3 | -5.349 | -35.781 | 11806 | -1 | 1 | 69 | 21-252 | H2O | 1641671502.49 | 0.191 | 0.817 | 53.03 | 0 | 0 | 1.319 | 75.438 | 27.361 | 2.06 | 60.851 | 0.856 | -0.926 | 0 | 0 | 1 | 43.5 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-18732 | 2022/01/08 12:55:33 | 1-37 | 4 | -5.39 | -35.711 | 11985 | 0 | 1 | 69 | 21-252 | H2O | 1641672038.72 | 0.209 | 0.789 | 55.016 | 0 | 0.002 | 1.357 | 75.814 | 27.455 | 2.091 | 63.153 | 0.855 | -0.769 | 0 | 0 | 1 | 43.562 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-18733 | 2022/01/08 13:05:18 | 1-38 | 1 | -5.847 | -38.418 | 11948 | -1 | 1 | 70 | 21-252 | H2O | 1641672623.62 | 0.189 | 0.884 | 58.268 | 0 | 0.005 | 1.44 | 75.711 | 27.613 | 2.089 | 62.567 | 0.856 | -0.804 | 0 | 0 | 1 | 43.688 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-18733 | 2022/01/08 13:14:13 | 1-38 | 2 | -5.846 | -38.763 | 11868 | -1 | 1 | 70 | 21-252 | H2O | 1641673158.41 | 0.206 | 0.725 | 60.133 | 0.001 | 0 | 1.488 | 75.609 | 27.3 | 2.062 | 61.991 | 0.855 | -0.99 | 0 | 0 | 1 | 43.75 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-18733 | 2022/01/08 13:23:09 | 1-38 | 3 | -5.875 | -38.676 | 11858 | -1 | 1 | 70 | 21-252 | H2O | 1641673694.54 | 0.178 | 0.766 | 57.036 | 0 | 0 | 1.418 | 75.655 | 27.415 | 2.059 | 61.166 | 0.856 | -1.07 | 0 | 0 | 1 | 43.812 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-18733 | 2022/01/08 13:32:05 | 1-38 | 4 | -5.855 | -38.876 | 11863 | 0 | 1 | 70 | 21-252 | H2O | 1641674230.39 | 0.22 | 0.832 | 53.204 | 0 | 0.003 | 1.322 | 75.554 | 27.477 | 2.058 | 61.288 | 0.856 | -0.982 | 0 | 0 | 1 | 43.938 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-18734 | 2022/01/08 13:41:49 | 1-39 | 1 | -5.643 | -37.485 | 12101 | -1 | 1 | 71 | 21-252 | H2O | 1641674815.1 | 0.208 | 0.837 | 55.786 | 0 | -0.001 | 1.371 | 76.124 | 27.601 | 2.117 | 63.755 | 0.856 | -0.864 | 0 | 0 | 1 | 43.938 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-18734 | 2022/01/08 13:50:46 | 1-39 | 2 | -5.636 | -37.544 | 12280 | -1 | 1 | 71 | 21-252 | H2O | 1641675351.05 | 0.207 | 0.713 | 57.91 | 0.001 | -0.001 | 1.432 | 76.348 | 28.395 | 2.142 | 64.711 | 0.855 | -0.83 | 0 | 0 | 1 | 43.938 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-18734 | 2022/01/08 13:59:42 | 1-39 | 3 | -5.619 | -37.468 | 12187 | -1 | 1 | 71 | 21-252 | H2O | 1641675887.24 | 0.214 | 0.888 | 58.555 | 0.001 | 0.004 | 1.447 | 76.196 | 28.143 | 2.123 | 63.518 | 0.856 | -0.877 | 0 | 0 | 1 | 43.875 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-18734 | 2022/01/08 14:08:38 | 1-39 | 4 | -5.66 | -37.489 | 12098 | 0 | 1 | 71 | 21-252 | H2O | 1641676424.15 | 0.201 | 0.739 | 54.892 | 0.001 | 0 | 1.357 | 75.986 | 28.153 | 2.105 | 62.396 | 0.856 | -0.989 | 0 | 0 | 1 | 43.938 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-18735 | 2022/01/08 14:18:23 | 1-40 | 1 | -8.754 | -62.574 | 12529 | -1 | 1 | EV | standard | H2O | 1641677008.66 | 0.199 | 0.73 | 56.659 | 0.001 | -0.001 | 1.398 | 76.897 | 29.487 | 2.13 | 63.202 | 0.856 | -0.967 | 0 | 0 | 1 | 43.875 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-18735 | 2022/01/08 14:27:19 | 1-40 | 2 | -8.958 | -65.065 | 12151 | -1 | 1 | EV | standard | H2O | 1641677544.9 | 0.182 | 0.786 | 56.346 | 0 | 0.003 | 1.383 | 76.061 | 27.948 | 2.076 | 63.401 | 0.856 | -0.751 | 0 | 0 | 1 | 43.938 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-18735 | 2022/01/08 14:36:14 | 1-40 | 3 | -8.987 | -65.726 | 12145 | -1 | 1 | EV | standard | H2O | 1641678079.99 | 0.188 | 0.756 | 53.292 | 0.001 | 0.002 | 1.314 | 76.226 | 27.881 | 2.071 | 62.674 | 0.856 | -0.892 | 0 | 0 | 1 | 43.875 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-18735 | 2022/01/08 14:45:10 | 1-40 | 4 | -8.994 | -65.949 | 12151 | 0 | 1 | EV | standard | H2O | 1641678615.46 | 0.185 | 0.722 | 58.575 | 0 | 0.004 | 1.447 | 76.151 | 27.983 | 2.066 | 63.291 | 0.855 | -0.865 | 0 | 0 | 1 | 43.812 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-18736 | 2022/01/08 14:54:55 | 1-41 | 1 | -5.716 | -40.798 | 11980 | -1 | 1 | 72 | 21-252 | H2O | 1641679200.34 | 0.208 | 0.838 | 57.665 | 0 | 0.004 | 1.436 | 75.932 | 27.648 | 2.064 | 61.866 | 0.855 | -1.01 | 0 | 0 | 1 | 43.812 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-18736 | 2022/01/08 15:03:51 | 1-41 | 2 | -5.597 | -38.259 | 11840 | -1 | 1 | 72 | 21-252 | H2O | 1641679736.43 | 0.207 | 0.851 | 52.43 | 0.001 | 0 | 1.288 | 75.543 | 27.041 | 2.077 | 62.001 | 0.855 | -0.815 | 0 | 0 | 1 | 43.75 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-18736 | 2022/01/08 15:12:46 | 1-41 | 3 | -5.517 | -37.518 | 11589 | -1 | 1 | 72 | 21-252 | H2O | 1641680271.69 | 0.173 | 0.921 | 55.985 | 0 | 0.005 | 1.393 | 75.182 | 26.308 | 2.024 | 61.305 | 0.855 | -0.93 | 0 | 0 | 1 | 43.75 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-18736 | 2022/01/08 15:21:41 | 1-41 | 4 | -5.52 | -37.295 | 11846 | 0 | 1 | 72 | 21-252 | H2O | 1641680807.37 | 0.22 | 0.943 | 54.717 | 0 | -0.004 | 1.356 | 75.671 | 27.217 | 2.064 | 61.512 | 0.856 | -0.975 | 0 | 0 | 1 | 43.812 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-18737 | 2022/01/08 15:31:27 | 1-42 | 1 | -5.593 | -37.081 | 11857 | -1 | 1 | 73 | 21-252 | H2O | 1641681391.88 | 0.222 | 0.799 | 55.083 | 0.001 | 0.005 | 1.371 | 75.47 | 27.21 | 2.082 | 63.027 | 0.856 | -0.817 | 0 | 0 | 1 | 43.812 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-18737 | 2022/01/08 15:40:23 | 1-42 | 2 | -5.602 | -37.164 | 11692 | -1 | 1 | 73 | 21-252 | H2O | 1641681928.23 | 0.2 | 0.784 | 54.3 | 0 | -0.001 | 1.343 | 75.214 | 26.836 | 2.04 | 61.617 | 0.855 | -0.951 | 0 | 0 | 1 | 43.75 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-18737 | 2022/01/08 15:49:18 | 1-42 | 3 | -5.565 | -37.042 | 11568 | -1 | 1 | 73 | 21-252 | H2O | 1641682464.32 | 0.191 | 0.825 | 55.217 | 0 | 0 | 1.373 | 75.179 | 26.526 | 1.999 | 59.76 | 0.855 | -1.181 | 0 | 0 | 1 | 43.812 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-18737 | 2022/01/08 15:58:13 | 1-42 | 4 | -5.593 | -36.945 | 11673 | 0 | 1 | 73 | 21-252 | H2O | 1641682998.5 | 0.181 | 0.764 | 54.118 | 0 | 0.002 | 1.349 | 75.125 | 26.89 | 2.04 | 61.812 | 0.856 | -0.812 | 0 | 0 | 1 | 43.812 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-18738 | 2022/01/08 16:07:57 | 1-43 | 1 | -5.589 | -37.354 | 12167 | -1 | 1 | 74 | 21-252 | H2O | 1641683582.14 | 0.221 | 0.9 | 60.968 | 0 | 0 | 1.517 | 76.113 | 28.285 | 2.108 | 63.211 | 0.855 | -0.802 | 0 | 0 | 1 | 43.812 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-18738 | 2022/01/08 16:16:52 | 1-43 | 2 | -5.609 | -37.486 | 12013 | -1 | 1 | 74 | 21-252 | H2O | 1641684117.77 | 0.2 | 0.732 | 54.452 | 0.001 | -0.002 | 1.347 | 75.949 | 27.618 | 2.082 | 62.482 | 0.855 | -0.984 | 0 | 0 | 1 | 43.812 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-18738 | 2022/01/08 16:25:48 | 1-43 | 3 | -5.587 | -37.305 | 11873 | -1 | 1 | 74 | 21-252 | H2O | 1641684653.5 | 0.202 | 0.68 | 58.141 | 0 | 0.002 | 1.439 | 75.624 | 27.431 | 2.063 | 61.598 | 0.855 | -0.95 | 0 | 0 | 1 | 43.75 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-18738 | 2022/01/08 16:34:43 | 1-43 | 4 | -5.63 | -37.575 | 12024 | 0 | 1 | 74 | 21-252 | H2O | 1641685189.12 | 0.206 | 0.849 | 55.645 | 0 | 0.003 | 1.376 | 75.851 | 27.869 | 2.101 | 62.124 | 0.856 | -0.805 | 0 | 0 | 1 | 43.75 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-18739 | 2022/01/08 16:44:26 | 1-44 | 1 | -5.61 | -37.437 | 11954 | -1 | 1 | 75 | 21-252 | H2O | 1641685771.49 | 0.208 | 0.769 | 55.285 | 0 | 0.003 | 1.376 | 75.868 | 27.638 | 2.078 | 61.256 | 0.856 | -1.048 | 0 | 0 | 1 | 43.75 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-18739 | 2022/01/08 16:53:21 | 1-44 | 2 | -5.574 | -37.236 | 12236 | -1 | 1 | 75 | 21-252 | H2O | 1641686306.68 | 0.21 | 0.765 | 55.314 | 0 | 0.001 | 1.376 | 76.374 | 28.443 | 2.127 | 63.418 | 0.856 | -0.971 | 0 | 0 | 1 | 43.75 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-18739 | 2022/01/08 17:02:17 | 1-44 | 3 | -5.563 | -37.375 | 12013 | -1 | 1 | 75 | 21-252 | H2O | 1641686842.72 | 0.185 | 0.851 | 53.224 | 0.001 | -0.001 | 1.321 | 75.774 | 27.762 | 2.098 | 62.765 | 0.855 | -0.792 | 0 | 0 | 1 | 43.688 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-18739 | 2022/01/08 17:11:12 | 1-44 | 4 | -5.585 | -37.252 | 12047 | 0 | 1 | 75 | 21-252 | H2O | 1641687377.93 | 0.189 | 0.791 | 50.427 | 0 | -0.002 | 1.243 | 75.902 | 27.854 | 2.109 | 63.183 | 0.855 | -0.915 | 0 | 0 | 1 | 43.688 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-18740 | 2022/01/08 17:20:58 | 1-45 | 1 | -5.766 | -38.251 | 11611 | -1 | 1 | 76 | 21-252 | H2O | 1641687963.11 | 0.195 | 0.901 | 59.953 | 0 | 0.003 | 1.491 | 75.057 | 26.447 | 2.018 | 61.172 | 0.855 | -0.799 | 0 | 0 | 1 | 43.75 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-18740 | 2022/01/08 17:29:53 | 1-45 | 2 | -5.718 | -37.98 | 11712 | -1 | 1 | 76 | 21-252 | H2O | 1641688498.83 | 0.196 | 0.785 | 51.805 | 0 | -0.001 | 1.284 | 75.198 | 26.75 | 2.042 | 62.42 | 0.855 | -0.793 | 0 | 0 | 1 | 43.688 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-18740 | 2022/01/08 17:38:49 | 1-45 | 3 | -5.742 | -38.059 | 11679 | -1 | 1 | 76 | 21-252 | H2O | 1641689034.63 | 0.203 | 0.736 | 56.434 | 0.001 | 0.001 | 1.406 | 75.156 | 26.839 | 2.034 | 61.668 | 0.856 | -0.901 | 0 | 0 | 1 | 43.688 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-18740 | 2022/01/08 17:47:45 | 1-45 | 4 | -5.687 | -37.828 | 11687 | 0 | 1 | 76 | 21-252 | H2O | 1641689570.71 | 0.2 | 0.774 | 55.86 | 0 | -0.001 | 1.379 | 75.224 | 26.79 | 2.045 | 61.37 | 0.856 | -0.775 | 0 | 0 | 1 | 43.688 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-18741 | 2022/01/08 17:57:29 | 1-46 | 1 | -5.591 | -36.925 | 12023 | -1 | 1 | 77 | 21-252 | H2O | 1641690154.97 | 0.191 | 0.789 | 58.226 | 0 | 0.002 | 1.44 | 75.977 | 27.734 | 2.097 | 62.539 | 0.855 | -1.081 | 0 | 0 | 1 | 43.75 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-18741 | 2022/01/08 18:06:25 | 1-46 | 2 | -5.523 | -36.798 | 11730 | -1 | 1 | 77 | 21-252 | H2O | 1641690690.79 | 0.207 | 0.807 | 58.825 | 0.001 | 0.001 | 1.456 | 75.411 | 26.933 | 2.055 | 60.597 | 0.856 | -1.004 | 0 | 0 | 1 | 43.688 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-18741 | 2022/01/08 18:15:20 | 1-46 | 3 | -5.552 | -37.03 | 12175 | -1 | 1 | 77 | 21-252 | H2O | 1641691225.81 | 0.202 | 0.759 | 56.927 | 0.001 | 0.001 | 1.408 | 76.282 | 28.03 | 2.118 | 63.232 | 0.856 | -1.092 | 0 | 0 | 1 | 43.688 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-18741 | 2022/01/08 18:24:15 | 1-46 | 4 | -5.58 | -36.883 | 12274 | 0 | 1 | 77 | 21-252 | H2O | 1641691760.57 | 0.205 | 0.917 | 53.536 | 0 | 0.004 | 1.323 | 76.356 | 28.485 | 2.127 | 63.371 | 0.856 | -1.014 | 0 | 0 | 1 | 43.75 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-18742 | 2022/01/08 18:34:01 | 1-47 | 1 | -5.73 | -38.08 | 9945 | -1 | 1 | 78 | 21-252 | H2O | 1641692348.28 | 0.192 | 0.759 | 51.281 | 0.001 | 0.001 | 1.317 | 71.88 | 21.546 | 1.754 | 55.217 | 0.855 | -0.78 | 0 | 0 | 1 | 43.688 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-18742 | 2022/01/08 18:42:57 | 1-47 | 2 | -5.786 | -38.497 | 10183 | -1 | 1 | 78 | 21-252 | H2O | 1641692883.96 | 0.215 | 0.891 | 54.465 | 0 | 0.004 | 1.381 | 72.242 | 22.511 | 1.786 | 55.285 | 0.856 | -0.789 | 0 | 0 | 1 | 43.625 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-18742 | 2022/01/08 18:51:53 | 1-47 | 3 | -5.818 | -38.722 | 11320 | -1 | 1 | 78 | 21-252 | H2O | 1641693419.35 | 0.211 | 0.833 | 54.435 | 0 | 0.003 | 1.354 | 74.521 | 25.741 | 1.967 | 59.712 | 0.855 | -1.017 | 0 | 0 | 1 | 43.625 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-18742 | 2022/01/08 19:00:49 | 1-47 | 4 | -5.818 | -38.455 | 10439 | 0 | 1 | 78 | 21-252 | H2O | 1641693955.79 | 0.198 | 0.822 | 51.589 | 0.001 | 0 | 1.318 | 72.769 | 23.464 | 1.83 | 55.486 | 0.855 | -0.95 | 0 | 0 | 1 | 43.5 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-18743 | 2022/01/08 19:10:35 | 1-48 | 1 | -5.586 | -37.167 | 11454 | -1 | 1 | 79 | 21-252 | H2O | 1641694540.66 | 0.194 | 0.779 | 55.54 | 0.001 | 0.001 | 1.391 | 74.79 | 26.453 | 1.998 | 59.317 | 0.856 | -1.123 | 0 | 0 | 1 | 43.5 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-18743 | 2022/01/08 19:19:31 | 1-48 | 2 | -5.593 | -37.053 | 11878 | -1 | 1 | 79 | 21-252 | H2O | 1641695076.89 | 0.19 | 0.796 | 53.779 | 0 | 0 | 1.333 | 75.544 | 27.337 | 2.077 | 62.176 | 0.856 | -0.945 | 0 | 0 | 1 | 43.375 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-18743 | 2022/01/08 19:28:26 | 1-48 | 3 | -5.582 | -37.06 | 11831 | -1 | 1 | 79 | 21-252 | H2O | 1641695611.77 | 0.2 | 0.826 | 57.813 | 0 | 0 | 1.428 | 75.496 | 27.301 | 2.073 | 61.355 | 0.855 | -1.009 | 0 | 0 | 1 | 43.312 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-18743 | 2022/01/08 19:37:22 | 1-48 | 4 | -5.628 | -36.977 | 11924 | 0 | 1 | 79 | 21-252 | H2O | 1641696147.53 | 0.197 | 0.717 | 52.685 | 0 | -0.001 | 1.304 | 75.709 | 27.662 | 2.08 | 61.186 | 0.856 | -1.09 | 0 | 0 | 1 | 43.25 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-18744 | 2022/01/08 19:47:08 | 1-49 | 1 | -8.752 | -62.305 | 12287 | -1 | 1 | EV | standard | H2O | 1641696733.29 | 0.191 | 0.797 | 59.835 | 0.001 | -0.001 | 1.473 | 76.395 | 28.783 | 2.088 | 61.945 | 0.855 | -1.067 | 0 | 0 | 1 | 43.312 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-18744 | 2022/01/08 19:56:03 | 1-49 | 2 | -8.904 | -65.092 | 12150 | -1 | 1 | EV | standard | H2O | 1641697268.69 | 0.196 | 0.841 | 53.213 | 0.001 | 0.002 | 1.306 | 75.975 | 28.593 | 2.07 | 61.342 | 0.856 | -0.903 | 0 | 0 | 1 | 43.25 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-18744 | 2022/01/08 20:04:58 | 1-49 | 3 | -8.98 | -65.738 | 12106 | -1 | 1 | EV | standard | H2O | 1641697803.03 | 0.195 | 0.805 | 55.338 | 0 | 0.003 | 1.361 | 76.016 | 28.066 | 2.061 | 61.677 | 0.856 | -1.06 | 0 | 0 | 1 | 43.188 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-18744 | 2022/01/08 20:13:53 | 1-49 | 4 | -8.988 | -65.993 | 12111 | 0 | 1 | EV | standard | H2O | 1641698338.32 | 0.192 | 0.741 | 57.968 | 0.001 | 0.004 | 1.424 | 75.981 | 28.059 | 2.069 | 62.355 | 0.855 | -0.882 | 0 | 0 | 1 | 43.188 | 1 | 49 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 77.33415
varray: [77.777, 77.705, 77.489, 77.641, 77.453, 77.368, 77.368, 77.375, 77.165, 77.244, 77.79, 77.221, 77.756, 77.148, 77.176, 76.974, 77.109, 76.86, 77.134, 76.93]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_shift
baseline_shift_STD | 0.28777225553
varray: [77.777, 77.705, 77.489, 77.641, 77.453, 77.368, 77.368, 77.375, 77.165, 77.244, 77.79, 77.221, 77.756, 77.148, 77.176, 76.974, 77.109, 76.86, 77.134, 76.93]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.99825
varray: [-1.018, -1.107, -1.118, -1.124, -1.072, -1.141, -1.181, -1.118, -0.977, -0.962, -1.027, -1.001, -1.028, -0.749, -0.799, -0.949, -0.859, -0.846, -0.845, -1.044]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.12392904379
varray: [-1.018, -1.107, -1.118, -1.124, -1.072, -1.141, -1.181, -1.118, -0.977, -0.962, -1.027, -1.001, -1.028, -0.749, -0.799, -0.949, -0.859, -0.846, -0.845, -1.044]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: ch4_ppm
residuals_MEAN | 2.17535
varray: [2.291, 2.286, 2.261, 2.272, 2.273, 2.267, 2.276, 2.271, 2.276, 2.271, 2.126, 2.082, 2.088, 2.085, 2.074, 2.063, 2.069, 2.052, 2.076, 2.048]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: residuals
residuals_STD | 0.102915537674
varray: [2.291, 2.286, 2.261, 2.272, 2.273, 2.267, 2.276, 2.271, 2.276, 2.271, 2.126, 2.082, 2.088, 2.085, 2.074, 2.063, 2.069, 2.052, 2.076, 2.048]
inst: std_standards
operation: std_standards(residuals)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: residuals
slope_shift_MEAN | 29.3432
varray: [29.812, 29.582, 29.462, 29.611, 29.267, 29.5, 29.487, 29.525, 29.437, 29.204, 29.481, 28.958, 29.73, 29.351, 29.211, 28.744, 29.337, 28.954, 29.283, 28.928]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: slope_shift
slope_shift_STD | 0.280602472135
varray: [29.812, 29.582, 29.462, 29.611, 29.267, 29.5, 29.487, 29.525, 29.437, 29.204, 29.481, 28.958, 29.73, 29.351, 29.211, 28.744, 29.337, 28.954, 29.283, 28.928]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: slope_shift
baseline_curvature_MEAN | 64.8447
varray: [67.854, 66.739, 66.297, 65.865, 66.19, 65.978, 66.51, 66.677, 67.465, 66.911, 64.682, 62.795, 63.362, 63.324, 63.062, 63.003, 63.096, 62.564, 62.608, 61.912]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.964506235
varray: [67.854, 66.739, 66.297, 65.865, 66.19, 65.978, 66.51, 66.677, 67.465, 66.911, 64.682, 62.795, 63.362, 63.324, 63.062, 63.003, 63.096, 62.564, 62.608, 61.912]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.17535
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 77.33415
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.99825
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 29.3432
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.154373306511
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.32972330651
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.205831075348
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.38118107535
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.308746613022
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.48409661302
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.217535
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.37178713137
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.37003713137
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.143886127765
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 77.4780361278
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 77.6219222555
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 0.863316766589
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 78.1974667666
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 5.17990059953
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 82.5140505995
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 1.40301236068
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 30.7462123607
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 5.89351870501
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 70.738218705
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -7.86944521316
actual_array: [ 1.93 -15.84]
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037, 2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934, -13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667, -9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022, 3.081, 3.882, 3.975, 3.935, -4.54, -4.917, -4.982, -4.951, -11.504, -11.863, -11.948, -11.968, -5.414, -5.099, -5.014, -4.967, -10.634, -10.948, -10.997, -11.022, -9.572, -9.463, -9.443, -9.456, -9.724, -9.724, -9.707, -9.756, -2.688, -2.377, -2.296, -2.278, -8.724, -8.933, -8.943, -8.988, -9.134, -9.148, -9.101, -9.153, -8.957, -8.948, -8.969, -9.003, -2.358, -1.98, -1.953, -1.88, -6.145, -6.393, -6.425, -6.428, -8.469, -8.569, -8.581, -8.585, -7.602, -7.549, -7.52, -7.534, -8.234, -8.22, -8.274, -8.262, -6.973, -6.854, -6.864, -6.829, -8.897, -8.966, -9.031, -8.983, -15.233, -15.56, -15.654, -15.666, -5.898, -5.075, -5.02, -5.123, -5.228, -5.262, -5.267, -5.271, -5.134, -5.041, -5.111, -5.108, -5.349, -5.365, -5.307, -5.35, -5.406, -5.399, -5.433, -5.419, -5.48, -5.466, -5.495, -5.473, -5.269, -5.286, -5.254, -5.232, -8.684, -8.92, -8.982, -9.018, -5.679, -5.504, -5.467, -5.433, -5.461, -5.51, -5.468, -5.475, -5.461, -5.535, -5.49, -5.472, -5.46, -5.486, -5.418, -5.456, -5.523, -5.511, -5.503, -5.559, -5.345, -5.339, -5.349, -5.39, -5.847, -5.846, -5.875, -5.855, -5.643, -5.636, -5.619, -5.66, -8.754, -8.958, -8.987, -8.994, -5.716, -5.597, -5.517, -5.52, -5.593, -5.602, -5.565, -5.593, -5.589, -5.609, -5.587, -5.63, -5.61, -5.574, -5.563, -5.585, -5.766, -5.718, -5.742, -5.687, -5.591, -5.523, -5.552, -5.58, -5.73, -5.786, -5.818, -5.818, -5.586, -5.593, -5.582, -5.628, -8.752, -8.904, -8.98, -8.988]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.72138181818
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0, 5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0, 16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0, 21666.0, 22202.0, 22739.0, 23275.0, 23866.0, 24402.0, 24938.0, 25474.0, 26064.0, 26601.0, 27138.0, 27674.0, 28261.0, 28798.0, 29335.0, 29871.0, 30458.0, 30995.0, 31531.0, 32068.0, 32657.0, 33193.0, 33730.0, 34266.0, 34855.0, 35391.0, 35927.0, 36463.0, 37051.0, 37587.0, 38123.0, 38660.0, 39249.0, 39784.0, 40320.0, 40857.0, 41447.0, 41984.0, 42520.0, 43057.0, 43644.0, 44179.0, 44716.0, 45252.0, 45839.0, 46376.0, 46911.0, 47447.0, 48034.0, 48570.0, 49107.0, 49643.0, 50230.0, 50766.0, 51302.0, 51838.0, 52425.0, 52961.0, 53498.0, 54035.0, 54623.0, 55159.0, 55696.0, 56233.0, 56821.0, 57358.0, 57894.0, 58431.0, 59017.0, 59553.0, 60088.0, 60624.0, 61210.0, 61746.0, 62282.0, 62818.0, 63404.0, 63941.0, 64476.0, 65012.0, 65598.0, 66134.0, 66670.0, 67205.0, 67792.0, 68328.0, 68864.0, 69400.0, 69987.0, 70522.0, 71058.0, 71594.0, 72182.0, 72718.0, 73254.0, 73791.0, 74375.0, 74909.0, 75446.0, 75982.0, 76566.0, 77102.0, 77638.0, 78174.0, 78759.0, 79295.0, 79832.0, 80368.0, 80953.0, 81490.0, 82026.0, 82562.0, 83148.0, 83684.0, 84220.0, 84756.0, 85343.0, 85878.0, 86413.0, 86949.0, 87536.0, 88073.0, 88609.0, 89145.0, 89727.0, 90263.0, 90799.0, 91334.0, 91917.0, 92452.0, 92988.0, 93524.0, 94109.0, 94644.0, 95180.0, 95716.0, 96300.0, 96837.0, 97373.0, 97909.0, 98494.0, 99030.0, 99565.0, 100101.0, 100686.0, 101222.0, 101757.0, 102292.0, 102878.0, 103414.0, 103949.0, 104484.0, 105068.0, 105603.0, 106139.0, 106674.0, 107257.0, 107792.0, 108328.0, 108863.0, 109449.0, 109984.0, 110520.0, 111056.0, 111640.0, 112176.0, 112711.0, 113246.0, 113832.0, 114368.0, 114904.0, 115440.0, 116026.0, 116562.0, 117097.0, 117633.0, 118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -58.2000074802
actual_array: [ 18.1 -119. ]
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285, 7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886, -93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916, -73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17, -34.064, -29.99, -28.941, -28.913, -38.607, -38.958, -39.118, -39.139, -89.068, -94.623, -95.648, -96.383, -47.239, -42.083, -40.515, -39.993, -85.241, -90.009, -91.305, -91.689, -73.175, -71.211, -70.971, -70.636, -75.512, -75.831, -76.004, -75.975, -46.215, -42.97, -42.212, -42.003, -63.968, -66.011, -66.554, -66.925, -78.967, -80.478, -80.804, -81.002, -66.556, -64.871, -64.49, -64.409, -58.189, -57.237, -57.115, -56.972, -68.652, -70.015, -70.474, -70.519, -68.364, -68.213, -68.051, -68.275, -67.599, -67.297, -67.238, -67.102, -70.514, -70.867, -70.938, -70.995, -67.348, -66.985, -67.025, -66.89, -67.062, -66.928, -67.065, -67.014, -110.324, -114.864, -116.219, -116.484, -46.216, -37.987, -36.292, -35.806, -35.654, -35.444, -35.08, -35.033, -34.252, -33.526, -33.726, -33.952, -35.493, -35.942, -35.62, -35.711, -36.018, -35.852, -35.918, -35.949, -36.283, -36.389, -36.527, -36.407, -35.428, -35.193, -35.245, -35.254, -62.256, -65.068, -65.851, -66.192, -40.386, -37.577, -36.829, -36.49, -36.59, -36.432, -36.413, -36.483, -36.142, -36.476, -36.409, -36.345, -36.286, -36.39, -36.089, -36.281, -36.343, -36.192, -36.251, -36.446, -35.681, -35.752, -35.781, -35.711, -38.418, -38.763, -38.676, -38.876, -37.485, -37.544, -37.468, -37.489, -62.574, -65.065, -65.726, -65.949, -40.798, -38.259, -37.518, -37.295, -37.081, -37.164, -37.042, -36.945, -37.354, -37.486, -37.305, -37.575, -37.437, -37.236, -37.375, -37.252, -38.251, -37.98, -38.059, -37.828, -36.925, -36.798, -37.03, -36.883, -38.08, -38.497, -38.722, -38.455, -37.167, -37.053, -37.06, -36.977, -62.305, -65.092, -65.738, -65.993]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -52.5654454545
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0, 5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0, 10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0, 16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0, 21666.0, 22202.0, 22739.0, 23275.0, 23866.0, 24402.0, 24938.0, 25474.0, 26064.0, 26601.0, 27138.0, 27674.0, 28261.0, 28798.0, 29335.0, 29871.0, 30458.0, 30995.0, 31531.0, 32068.0, 32657.0, 33193.0, 33730.0, 34266.0, 34855.0, 35391.0, 35927.0, 36463.0, 37051.0, 37587.0, 38123.0, 38660.0, 39249.0, 39784.0, 40320.0, 40857.0, 41447.0, 41984.0, 42520.0, 43057.0, 43644.0, 44179.0, 44716.0, 45252.0, 45839.0, 46376.0, 46911.0, 47447.0, 48034.0, 48570.0, 49107.0, 49643.0, 50230.0, 50766.0, 51302.0, 51838.0, 52425.0, 52961.0, 53498.0, 54035.0, 54623.0, 55159.0, 55696.0, 56233.0, 56821.0, 57358.0, 57894.0, 58431.0, 59017.0, 59553.0, 60088.0, 60624.0, 61210.0, 61746.0, 62282.0, 62818.0, 63404.0, 63941.0, 64476.0, 65012.0, 65598.0, 66134.0, 66670.0, 67205.0, 67792.0, 68328.0, 68864.0, 69400.0, 69987.0, 70522.0, 71058.0, 71594.0, 72182.0, 72718.0, 73254.0, 73791.0, 74375.0, 74909.0, 75446.0, 75982.0, 76566.0, 77102.0, 77638.0, 78174.0, 78759.0, 79295.0, 79832.0, 80368.0, 80953.0, 81490.0, 82026.0, 82562.0, 83148.0, 83684.0, 84220.0, 84756.0, 85343.0, 85878.0, 86413.0, 86949.0, 87536.0, 88073.0, 88609.0, 89145.0, 89727.0, 90263.0, 90799.0, 91334.0, 91917.0, 92452.0, 92988.0, 93524.0, 94109.0, 94644.0, 95180.0, 95716.0, 96300.0, 96837.0, 97373.0, 97909.0, 98494.0, 99030.0, 99565.0, 100101.0, 100686.0, 101222.0, 101757.0, 102292.0, 102878.0, 103414.0, 103949.0, 104484.0, 105068.0, 105603.0, 106139.0, 106674.0, 107257.0, 107792.0, 108328.0, 108863.0, 109449.0, 109984.0, 110520.0, 111056.0, 111640.0, 112176.0, 112711.0, 113246.0, 113832.0, 114368.0, 114904.0, 115440.0, 116026.0, 116562.0, 117097.0, 117633.0, 118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2196
varray: [2.262, 2.206, 2.35, 2.19, 2.221, 2.218, 2.193, 2.191, 2.182, 2.183]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: residuals
SAMP_residuals_STD | 0.0517970398125
varray: [2.262, 2.206, 2.35, 2.19, 2.221, 2.218, 2.193, 2.191, 2.182, 2.183]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.033
varray: [-1.155, -1.12, -1.197, -0.956, -1.095, -1.052, -0.925, -0.94, -0.953, -0.937]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.102962776445
varray: [-1.155, -1.12, -1.197, -0.956, -1.095, -1.052, -0.925, -0.94, -0.953, -0.937]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.5692
varray: [110.706, 105.569, 129.721, 97.634, 89.742, 83.997, 81.283, 79.614, 78.908, 78.518]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 17.1742784212
varray: [110.706, 105.569, 129.721, 97.634, 89.742, 83.997, 81.283, 79.614, 78.908, 78.518]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.2481
varray: [30.566, 29.827, 31.4, 29.838, 30.46, 30.386, 29.989, 30.044, 29.848, 30.123]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.485555683956
varray: [30.566, 29.827, 31.4, 29.838, 30.46, 30.386, 29.989, 30.044, 29.848, 30.123]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: slope_shift
SAMP_d18O_STD | 0.301944365736
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.51251581753
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.0991
varray: [66.448, 65.064, 68.543, 64.452, 64.962, 66.429, 66.741, 66.092, 66.431, 65.829]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14870830356
varray: [66.448, 65.064, 68.543, 64.452, 64.962, 66.429, 66.741, 66.092, 66.431, 65.829]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2196
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.033
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.5692
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.2481
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.03415542327
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.2544
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1478
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
dD_mean | -68.8128
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
d18O_stdev | 0.301944365736
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
dD_stdev | 2.51251581753
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1478
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.8128
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.301944365736
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.51251581753
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3724029118
actual_array: [ 1.93 -15.84]
varray: [-9.85, -9.306, -9.29, -9.123, -9.117, -8.678, -8.989, -9.065, -9.023, -9.037]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.1478
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.8914677993
actual_array: [ 18.1 -119. ]
varray: [-75.296, -70.206, -69.612, -68.392, -68.04, -66.923, -67.386, -67.561, -67.427, -67.285]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.8128
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [1.0, 536.0, 1073.0, 1609.0, 2146.0, 2682.0, 3218.0, 3755.0, 4292.0, 4828.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.11455
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 16.23505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2744
varray: [2.291, 2.286, 2.261, 2.272, 2.273, 2.267, 2.276, 2.271, 2.276, 2.271]
inst: mean_all
operation: mean_all(residuals)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: residuals
SAMP_residuals_STD | 0.00869482604771
varray: [2.291, 2.286, 2.261, 2.272, 2.273, 2.267, 2.276, 2.271, 2.276, 2.271]
inst: std_all
operation: std_all(residuals)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0818
varray: [-1.018, -1.107, -1.118, -1.124, -1.072, -1.141, -1.181, -1.118, -0.977, -0.962]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0729410721062
varray: [-1.018, -1.107, -1.118, -1.124, -1.072, -1.141, -1.181, -1.118, -0.977, -0.962]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.4585
varray: [77.777, 77.705, 77.489, 77.641, 77.453, 77.368, 77.368, 77.375, 77.165, 77.244]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.197758691339
varray: [77.777, 77.705, 77.489, 77.641, 77.453, 77.368, 77.368, 77.375, 77.165, 77.244]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.4887
varray: [29.812, 29.582, 29.462, 29.611, 29.267, 29.5, 29.487, 29.525, 29.437, 29.204]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.171017900298
varray: [29.812, 29.582, 29.462, 29.611, 29.267, 29.5, 29.487, 29.525, 29.437, 29.204]
inst: std_all
operation: std_all(slope_shift)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: slope_shift
SAMP_d18O_STD | 0.272676120456
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.90521978895
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.6486
varray: [67.854, 66.739, 66.297, 65.865, 66.19, 65.978, 66.51, 66.677, 67.465, 66.911]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.63370255904
varray: [67.854, 66.739, 66.297, 65.865, 66.19, 65.978, 66.51, 66.677, 67.465, 66.911]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2744
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0818
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.4585
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.4887
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.09905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.55329027513
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8039
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.7784
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.5064
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7784
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
dD_mean | 17.5064
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
d18O_stdev | 0.272676120456
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
dD_stdev | 3.90521978895
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7784
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.5064
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.272676120456
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.90521978895
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.041, 2.657, 2.776, 2.862, 2.89, 2.882, 2.889, 2.936, 2.917, 2.934]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.7784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [7.133, 15.348, 17.392, 18.28, 18.829, 19.176, 19.458, 19.673, 19.889, 19.886]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 17.5064
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [5415.0, 5951.0, 6486.0, 7023.0, 7559.0, 8095.0, 8632.0, 9168.0, 9704.0, 10241.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.05975
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.12435
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0763
varray: [2.126, 2.082, 2.088, 2.085, 2.074, 2.063, 2.069, 2.052, 2.076, 2.048]
inst: mean_all
operation: mean_all(residuals)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: residuals
SAMP_residuals_STD | 0.0219648709433
varray: [2.126, 2.082, 2.088, 2.085, 2.074, 2.063, 2.069, 2.052, 2.076, 2.048]
inst: std_all
operation: std_all(residuals)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9147
varray: [-1.027, -1.001, -1.028, -0.749, -0.799, -0.949, -0.859, -0.846, -0.845, -1.044]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107658978054
varray: [-1.027, -1.001, -1.028, -0.749, -0.799, -0.949, -0.859, -0.846, -0.845, -1.044]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 77.2098
varray: [77.79, 77.221, 77.756, 77.148, 77.176, 76.974, 77.109, 76.86, 77.134, 76.93]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.318365548108
varray: [77.79, 77.221, 77.756, 77.148, 77.176, 76.974, 77.109, 76.86, 77.134, 76.93]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.1977
varray: [29.481, 28.958, 29.73, 29.351, 29.211, 28.744, 29.337, 28.954, 29.283, 28.928]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.299886663777
varray: [29.481, 28.958, 29.73, 29.351, 29.211, 28.744, 29.337, 28.954, 29.283, 28.928]
inst: std_all
operation: std_all(slope_shift)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: slope_shift
SAMP_d18O_STD | 0.389178279627
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.71695873404
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.0408
varray: [64.682, 62.795, 63.362, 63.324, 63.062, 63.003, 63.096, 62.564, 62.608, 61.912]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.717332682775
varray: [64.682, 62.795, 63.362, 63.324, 63.062, 63.003, 63.096, 62.564, 62.608, 61.912]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0763
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9147
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 77.2098
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.1977
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.55329027513
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8039
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.4482
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
dD_mean | -108.4025
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
d18O_stdev | 0.389178279627
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
dD_stdev | 5.71695873404
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.4482
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.4025
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.389178279627
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.71695873404
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.398, -14.285, -14.423, -14.547, -14.574, -14.63, -14.631, -14.676, -14.651, -14.667]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.4482
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-93.171, -105.352, -108.224, -109.564, -110.399, -110.874, -111.275, -111.561, -111.689, -111.916]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -108.4025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [10829.0, 11366.0, 11903.0, 12440.0, 12976.0, 13513.0, 14049.0, 14585.0, 15122.0, 15659.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.25785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.12435
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1132
varray: [2.149, 2.116, 2.134, 2.134, 2.115, 2.112, 2.092, 2.094, 2.09, 2.096]
inst: mean_all
operation: mean_all(residuals)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: residuals
SAMP_residuals_STD | 0.0205793855863
varray: [2.149, 2.116, 2.134, 2.134, 2.115, 2.112, 2.092, 2.094, 2.09, 2.096]
inst: std_all
operation: std_all(residuals)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9646
varray: [-1.141, -0.826, -0.981, -0.947, -1.018, -0.98, -1.112, -0.695, -0.94, -1.006]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129410801541
varray: [-1.141, -0.826, -0.981, -0.947, -1.018, -0.98, -1.112, -0.695, -0.94, -1.006]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.8348
varray: [77.671, 76.661, 77.141, 77.045, 76.701, 76.672, 76.683, 76.567, 76.622, 76.585]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.351856915363
varray: [77.671, 76.661, 77.141, 77.045, 76.701, 76.672, 76.683, 76.567, 76.622, 76.585]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.845
varray: [29.812, 28.619, 29.023, 28.848, 29.022, 28.7, 29.005, 28.195, 28.467, 28.759]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.430579970634
varray: [29.812, 28.619, 29.023, 28.848, 29.022, 28.7, 29.005, 28.195, 28.467, 28.759]
inst: std_all
operation: std_all(slope_shift)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: slope_shift
SAMP_d18O_STD | 0.123070891946
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.82189186165
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.7607
varray: [64.253, 64.908, 64.23, 64.487, 63.151, 63.15, 62.105, 65.052, 63.338, 62.933]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.961691802098
varray: [64.253, 64.908, 64.23, 64.487, 63.151, 63.15, 62.105, 65.052, 63.338, 62.933]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1132
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9646
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.8348
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.845
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06215
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06215
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.85701151539
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.084
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.081
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
dD_mean | -68.1812
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
d18O_stdev | 0.123070891946
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
dD_stdev | 1.82189186165
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.081
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.1812
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.123070891946
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.82189186165
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3034957566
actual_array: [ 1.93 -15.84]
varray: [-9.407, -9.138, -9.099, -9.071, -9.028, -9.02, -9.015, -9.02, -8.99, -9.022]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.2037296013
actual_array: [ 18.1 -119. ]
varray: [-73.079, -69.115, -68.106, -67.684, -67.542, -67.438, -67.381, -67.106, -67.191, -67.17]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.1812
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [16247.0, 16784.0, 17320.0, 17857.0, 18393.0, 18930.0, 19466.0, 20003.0, 20540.0, 21076.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.22095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.49935
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.171
varray: [2.168, 2.176, 2.157, 2.183]
inst: mean_all
operation: mean_all(residuals)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: residuals
SAMP_residuals_STD | 0.0111654228163
varray: [2.168, 2.176, 2.157, 2.183]
inst: std_all
operation: std_all(residuals)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.004
varray: [-0.944, -0.929, -1.116, -1.027]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0862129147325
varray: [-0.944, -0.929, -1.116, -1.027]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.7955
varray: [76.743, 76.805, 76.684, 76.95]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.114235137618
varray: [76.743, 76.805, 76.684, 76.95]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.74975
varray: [29.038, 28.696, 28.397, 28.868]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.273491468484
varray: [29.038, 28.696, 28.397, 28.868]
inst: std_all
operation: std_all(slope_shift)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: slope_shift
SAMP_d18O_STD | 0.426537513004
varray: [3.081, 3.882, 3.975, 3.935]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.4432989993
varray: [-34.064, -29.99, -28.941, -28.913]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.6485
varray: [64.672, 65.235, 63.792, 64.895]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.616140947078
varray: [64.672, 65.235, 63.792, 64.895]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.171
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.004
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.7955
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.74975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.19996782127
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.1962
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [3.081, 3.882, 3.975, 3.935]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-34.064, -29.99, -28.941, -28.913]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.71825
varray: [3.081, 3.882, 3.975, 3.935]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
dD_mean | -30.477
varray: [-34.064, -29.99, -28.941, -28.913]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
d18O_stdev | 0.426537513004
varray: [3.081, 3.882, 3.975, 3.935]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
dD_stdev | 2.4432989993
varray: [-34.064, -29.99, -28.941, -28.913]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.71825
varray: [3.081, 3.882, 3.975, 3.935]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
measured_dD_mean | -30.477
varray: [-34.064, -29.99, -28.941, -28.913]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.426537513004
varray: [3.081, 3.882, 3.975, 3.935]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.4432989993
varray: [-34.064, -29.99, -28.941, -28.913]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.8994968537
actual_array: [ 1.93 -15.84]
varray: [3.081, 3.882, 3.975, 3.935]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 3.71825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -34.1482853873
actual_array: [ 18.1 -119. ]
varray: [-34.064, -29.99, -28.941, -28.913]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -30.477
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [21666.0, 22202.0, 22739.0, 23275.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.16315
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.53865
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.4005
varray: [2.808, 2.451, 2.192, 2.151]
inst: mean_all
operation: mean_all(residuals)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: residuals
SAMP_residuals_STD | 0.302395436474
varray: [2.808, 2.451, 2.192, 2.151]
inst: std_all
operation: std_all(residuals)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0435
varray: [-0.987, -1.192, -0.971, -1.024]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101457708759
varray: [-0.987, -1.192, -0.971, -1.024]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 79.836
varray: [84.908, 80.612, 77.219, 76.605]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.8129345654
varray: [84.908, 80.612, 77.219, 76.605]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 33.17625
varray: [40.289, 34.12, 29.148, 29.148]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: slope_shift
SAMP_slope_shift_STD | 5.28946981433
varray: [40.289, 34.12, 29.148, 29.148]
inst: std_all
operation: std_all(slope_shift)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: slope_shift
SAMP_d18O_STD | 0.206711554265
varray: [-4.54, -4.917, -4.982, -4.951]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.245992547313
varray: [-38.607, -38.958, -39.118, -39.139]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 70.81175
varray: [81.903, 71.978, 65.692, 63.674]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 8.1962446838
varray: [81.903, 71.978, 65.692, 63.674]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.4005
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0435
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 79.836
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 33.17625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.22515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.22515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 10.3500586113
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.96705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-4.54, -4.917, -4.982, -4.951]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-38.607, -38.958, -39.118, -39.139]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.8475
varray: [-4.54, -4.917, -4.982, -4.951]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
dD_mean | -38.9555
varray: [-38.607, -38.958, -39.118, -39.139]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
d18O_stdev | 0.206711554265
varray: [-4.54, -4.917, -4.982, -4.951]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
dD_stdev | 0.245992547313
varray: [-38.607, -38.958, -39.118, -39.139]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.8475
varray: [-4.54, -4.917, -4.982, -4.951]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.9555
varray: [-38.607, -38.958, -39.118, -39.139]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.206711554265
varray: [-4.54, -4.917, -4.982, -4.951]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.245992547313
varray: [-38.607, -38.958, -39.118, -39.139]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.9364532177
actual_array: [ 1.93 -15.84]
varray: [-4.54, -4.917, -4.982, -4.951]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -4.8475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.3803758114
actual_array: [ 18.1 -119. ]
varray: [-38.607, -38.958, -39.118, -39.139]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.9555
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [23866.0, 24402.0, 24938.0, 25474.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 74.93365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.50185
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.057
varray: [2.062, 2.058, 2.045, 2.063]
inst: mean_all
operation: mean_all(residuals)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: residuals
SAMP_residuals_STD | 0.0082865352631
varray: [2.062, 2.058, 2.045, 2.063]
inst: std_all
operation: std_all(residuals)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9285
varray: [-1.075, -0.836, -0.87, -0.933]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105610921152
varray: [-1.075, -0.836, -0.87, -0.933]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.6065
varray: [76.677, 76.637, 76.489, 76.623]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.081606780764
varray: [76.677, 76.637, 76.489, 76.623]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4525
varray: [28.727, 28.269, 28.375, 28.439]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.195968534889
varray: [28.727, 28.269, 28.375, 28.439]
inst: std_all
operation: std_all(slope_shift)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: slope_shift
SAMP_d18O_STD | 0.216017553299
varray: [-11.504, -11.863, -11.948, -11.968]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.32104526718
varray: [-89.068, -94.623, -95.648, -96.383]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.68575
varray: [62.257, 63.261, 62.042, 63.183]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.626208365216
varray: [62.257, 63.261, 62.042, 63.183]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.057
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9285
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.6065
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11835
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11835
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.44050382697
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.15895
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-11.504, -11.863, -11.948, -11.968]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-89.068, -94.623, -95.648, -96.383]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.82075
varray: [-11.504, -11.863, -11.948, -11.968]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
dD_mean | -93.9305
varray: [-89.068, -94.623, -95.648, -96.383]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
d18O_stdev | 0.216017553299
varray: [-11.504, -11.863, -11.948, -11.968]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
dD_stdev | 3.32104526718
varray: [-89.068, -94.623, -95.648, -96.383]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.82075
varray: [-11.504, -11.863, -11.948, -11.968]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
measured_dD_mean | -93.9305
varray: [-89.068, -94.623, -95.648, -96.383]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.216017553299
varray: [-11.504, -11.863, -11.948, -11.968]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.32104526718
varray: [-89.068, -94.623, -95.648, -96.383]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -13.1296690873
actual_array: [ 1.93 -15.84]
varray: [-11.504, -11.863, -11.948, -11.968]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -11.82075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -103.241692208
actual_array: [ 18.1 -119. ]
varray: [-89.068, -94.623, -95.648, -96.383]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -93.9305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [26064.0, 26601.0, 27138.0, 27674.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.27715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.72765
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12425
varray: [2.132, 2.144, 2.101, 2.12]
inst: mean_all
operation: mean_all(residuals)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: residuals
SAMP_residuals_STD | 0.0183371208209
varray: [2.132, 2.144, 2.101, 2.12]
inst: std_all
operation: std_all(residuals)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.12725
varray: [-1.067, -1.126, -1.185, -1.131]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0482381246181
varray: [-1.067, -1.126, -1.185, -1.131]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.628
varray: [76.852, 76.865, 76.319, 76.476]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.273818674795
varray: [76.852, 76.865, 76.319, 76.476]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.5135
varray: [28.971, 28.604, 28.147, 28.332]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.358125210878
varray: [28.971, 28.604, 28.147, 28.332]
inst: std_all
operation: std_all(slope_shift)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: slope_shift
SAMP_d18O_STD | 0.201223756053
varray: [-5.414, -5.099, -5.014, -4.967]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.30908139318
varray: [-47.239, -42.083, -40.515, -39.993]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.13425
varray: [63.254, 63.283, 62.6, 63.4]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.361714781746
varray: [63.254, 63.283, 62.6, 63.4]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.12725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.628
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.5135
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0511
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0511
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.34904727975
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.71045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.414, -5.099, -5.014, -4.967]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-47.239, -42.083, -40.515, -39.993]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.1235
varray: [-5.414, -5.099, -5.014, -4.967]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
dD_mean | -42.4575
varray: [-47.239, -42.083, -40.515, -39.993]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
d18O_stdev | 0.201223756053
varray: [-5.414, -5.099, -5.014, -4.967]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
dD_stdev | 3.30908139318
varray: [-47.239, -42.083, -40.515, -39.993]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.1235
varray: [-5.414, -5.099, -5.014, -4.967]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.4575
varray: [-47.239, -42.083, -40.515, -39.993]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.201223756053
varray: [-5.414, -5.099, -5.014, -4.967]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.30908139318
varray: [-47.239, -42.083, -40.515, -39.993]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.22115942786
actual_array: [ 1.93 -15.84]
varray: [-5.414, -5.099, -5.014, -4.967]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.1235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -47.1936423875
actual_array: [ 18.1 -119. ]
varray: [-47.239, -42.083, -40.515, -39.993]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -42.4575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [28261.0, 28798.0, 29335.0, 29871.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2099
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.70615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0525
varray: [2.078, 2.043, 2.062, 2.027]
inst: mean_all
operation: mean_all(residuals)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: residuals
SAMP_residuals_STD | 0.0222186108177
varray: [2.078, 2.043, 2.062, 2.027]
inst: std_all
operation: std_all(residuals)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.11375
varray: [-1.136, -1.067, -1.12, -1.132]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0318995820244
varray: [-1.136, -1.067, -1.12, -1.132]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.675
varray: [76.856, 76.571, 76.762, 76.511]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.161288561281
varray: [76.856, 76.571, 76.762, 76.511]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.56875
varray: [28.934, 28.476, 28.74, 28.125]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.350357317606
varray: [28.934, 28.476, 28.74, 28.125]
inst: std_all
operation: std_all(slope_shift)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: slope_shift
SAMP_d18O_STD | 0.180141379662
varray: [-10.634, -10.948, -10.997, -11.022]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.96833421299
varray: [-85.241, -90.009, -91.305, -91.689]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.614
varray: [61.687, 61.765, 61.931, 61.073]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.374744357307
varray: [61.687, 61.765, 61.931, 61.073]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.11375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.56875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12285
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12285
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.64736709035
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.2307
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-10.634, -10.948, -10.997, -11.022]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-85.241, -90.009, -91.305, -91.689]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -10.90025
varray: [-10.634, -10.948, -10.997, -11.022]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
dD_mean | -89.561
varray: [-85.241, -90.009, -91.305, -91.689]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
d18O_stdev | 0.180141379662
varray: [-10.634, -10.948, -10.997, -11.022]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
dD_stdev | 2.96833421299
varray: [-85.241, -90.009, -91.305, -91.689]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
measured_d18O_mean | -10.90025
varray: [-10.634, -10.948, -10.997, -11.022]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
measured_dD_mean | -89.561
varray: [-85.241, -90.009, -91.305, -91.689]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.180141379662
varray: [-10.634, -10.948, -10.997, -11.022]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.96833421299
varray: [-85.241, -90.009, -91.305, -91.689]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.1801326147
actual_array: [ 1.93 -15.84]
varray: [-10.634, -10.948, -10.997, -11.022]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -10.90025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -98.4838200477
actual_array: [ 18.1 -119. ]
varray: [-85.241, -90.009, -91.305, -91.689]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -89.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [30458.0, 30995.0, 31531.0, 32068.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.28165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.65915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08075
varray: [2.13, 2.056, 2.082, 2.055]
inst: mean_all
operation: mean_all(residuals)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: residuals
SAMP_residuals_STD | 0.0351318943412
varray: [2.13, 2.056, 2.082, 2.055]
inst: std_all
operation: std_all(residuals)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.99375
varray: [-1.071, -0.969, -0.967, -0.968]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0515064720852
varray: [-1.071, -0.969, -0.967, -0.968]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.412
varray: [77.092, 76.163, 76.328, 76.065]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.466142324475
varray: [77.092, 76.163, 76.328, 76.065]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.43
varray: [29.443, 28.029, 28.344, 27.904]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.700248051288
varray: [29.443, 28.029, 28.344, 27.904]
inst: std_all
operation: std_all(slope_shift)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0595790791022
varray: [-9.572, -9.463, -9.443, -9.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.14243464437
varray: [-73.175, -71.211, -70.971, -70.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.4325
varray: [63.361, 61.918, 62.527, 61.924]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.681744086883
varray: [63.361, 61.918, 62.527, 61.924]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.99375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.412
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.43
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0946
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0946
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.34872549245
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.4122
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-9.572, -9.463, -9.443, -9.456]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-73.175, -71.211, -70.971, -70.636]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.4835
varray: [-9.572, -9.463, -9.443, -9.456]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
dD_mean | -71.49825
varray: [-73.175, -71.211, -70.971, -70.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0595790791022
varray: [-9.572, -9.463, -9.443, -9.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
dD_stdev | 1.14243464437
varray: [-73.175, -71.211, -70.971, -70.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.4835
varray: [-9.572, -9.463, -9.443, -9.456]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.49825
varray: [-73.175, -71.211, -70.971, -70.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0595790791022
varray: [-9.572, -9.463, -9.443, -9.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.14243464437
varray: [-73.175, -71.211, -70.971, -70.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.7186923131
actual_array: [ 1.93 -15.84]
varray: [-9.572, -9.463, -9.443, -9.456]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.4835
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.8156073558
actual_array: [ 18.1 -119. ]
varray: [-73.175, -71.211, -70.971, -70.636]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -71.49825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [32657.0, 33193.0, 33730.0, 34266.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2534
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.92215
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.07275
varray: [2.126, 2.071, 2.058, 2.036]
inst: mean_all
operation: mean_all(residuals)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: residuals
SAMP_residuals_STD | 0.0383264486571
varray: [2.126, 2.071, 2.058, 2.036]
inst: std_all
operation: std_all(residuals)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.89575
varray: [-0.985, -0.625, -0.978, -0.995]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.180634760405
varray: [-0.985, -0.625, -0.978, -0.995]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.45775
varray: [77.305, 75.987, 76.392, 76.147]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.588876543372
varray: [77.305, 75.987, 76.392, 76.147]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.20225
varray: [29.312, 27.553, 28.269, 27.675]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.803226566211
varray: [29.312, 27.553, 28.269, 27.675]
inst: std_all
operation: std_all(slope_shift)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0204674538394
varray: [-9.724, -9.724, -9.707, -9.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.225407039228
varray: [-75.512, -75.831, -76.004, -75.975]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.74325
varray: [64.109, 63.563, 62.073, 61.228]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.32696153549
varray: [64.109, 63.563, 62.073, 61.228]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.07275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.89575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.45775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.20225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1026
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1026
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.71648240513
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.10145
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-9.724, -9.724, -9.707, -9.756]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-75.512, -75.831, -76.004, -75.975]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.72775
varray: [-9.724, -9.724, -9.707, -9.756]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
dD_mean | -75.8305
varray: [-75.512, -75.831, -76.004, -75.975]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0204674538394
varray: [-9.724, -9.724, -9.707, -9.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
dD_stdev | 0.225407039228
varray: [-75.512, -75.831, -76.004, -75.975]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.72775
varray: [-9.724, -9.724, -9.707, -9.756]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
measured_dD_mean | -75.8305
varray: [-75.512, -75.831, -76.004, -75.975]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0204674538394
varray: [-9.724, -9.724, -9.707, -9.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.225407039228
varray: [-75.512, -75.831, -76.004, -75.975]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.9706469936
actual_array: [ 1.93 -15.84]
varray: [-9.724, -9.724, -9.707, -9.756]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.72775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.532918642
actual_array: [ 18.1 -119. ]
varray: [-75.512, -75.831, -76.004, -75.975]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -75.8305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [34855.0, 35391.0, 35927.0, 36463.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2614
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8764
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3855
varray: [2.422, 2.366, 2.385, 2.369]
inst: mean_all
operation: mean_all(residuals)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: residuals
SAMP_residuals_STD | 0.025722882161
varray: [2.422, 2.366, 2.385, 2.369]
inst: std_all
operation: std_all(residuals)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.8985
varray: [-2.812, -2.928, -2.954, -2.9]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0617386966713
varray: [-2.812, -2.928, -2.954, -2.9]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 632.37325
varray: [638.339, 628.378, 634.216, 628.56]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.8127884762
varray: [638.339, 628.378, 634.216, 628.56]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 6.319
varray: [6.579, 6.162, 6.498, 6.037]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.260649189525
varray: [6.579, 6.162, 6.498, 6.037]
inst: std_all
operation: std_all(slope_shift)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: slope_shift
SAMP_d18O_STD | 0.190431746303
varray: [-2.688, -2.377, -2.296, -2.278]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.95465922009
varray: [-46.215, -42.97, -42.212, -42.003]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 52.863
varray: [53.787, 52.114, 52.805, 52.746]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.690852130826
varray: [53.787, 52.114, 52.805, 52.746]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3855
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.8985
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 632.37325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 6.319
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.21015
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.21015
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 9.66051439998
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -11.9817
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-2.688, -2.377, -2.296, -2.278]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-46.215, -42.97, -42.212, -42.003]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.40975
varray: [-2.688, -2.377, -2.296, -2.278]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
dD_mean | -43.35
varray: [-46.215, -42.97, -42.212, -42.003]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
d18O_stdev | 0.190431746303
varray: [-2.688, -2.377, -2.296, -2.278]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
dD_stdev | 1.95465922009
varray: [-46.215, -42.97, -42.212, -42.003]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.40975
varray: [-2.688, -2.377, -2.296, -2.278]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.35
varray: [-46.215, -42.97, -42.212, -42.003]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.190431746303
varray: [-2.688, -2.377, -2.296, -2.278]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.95465922009
varray: [-46.215, -42.97, -42.212, -42.003]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.42180624732
actual_array: [ 1.93 -15.84]
varray: [-2.688, -2.377, -2.296, -2.278]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.40975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.1654700343
actual_array: [ 18.1 -119. ]
varray: [-46.215, -42.97, -42.212, -42.003]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -43.35
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [37051.0, 37587.0, 38123.0, 38660.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 74.94865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 555.0391
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11075
varray: [2.174, 2.086, 2.091, 2.092]
inst: mean_all
operation: mean_all(residuals)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: residuals
SAMP_residuals_STD | 0.0422482741265
varray: [2.174, 2.086, 2.091, 2.092]
inst: std_all
operation: std_all(residuals)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.14375
varray: [-1.105, -1.102, -1.197, -1.171]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0476890972865
varray: [-1.105, -1.102, -1.197, -1.171]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 80.12
varray: [84.493, 79.606, 78.483, 77.898]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.00024321236
varray: [84.493, 79.606, 78.483, 77.898]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.08475
varray: [29.613, 28.915, 28.853, 28.958]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.354794189167
varray: [29.613, 28.915, 28.853, 28.958]
inst: std_all
operation: std_all(slope_shift)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: slope_shift
SAMP_d18O_STD | 0.117787945054
varray: [-8.724, -8.933, -8.943, -8.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.31886883351
varray: [-63.968, -66.011, -66.554, -66.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.4155
varray: [64.825, 62.198, 61.402, 61.237]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.66022498475
varray: [64.825, 62.198, 61.402, 61.237]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.14375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 80.12
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.08475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0646
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0646
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.9696370699
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.4292
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.724, -8.933, -8.943, -8.988]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-63.968, -66.011, -66.554, -66.925]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.897
varray: [-8.724, -8.933, -8.943, -8.988]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
dD_mean | -65.8645
varray: [-63.968, -66.011, -66.554, -66.925]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
d18O_stdev | 0.117787945054
varray: [-8.724, -8.933, -8.943, -8.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
dD_stdev | 1.31886883351
varray: [-63.968, -66.011, -66.554, -66.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.897
varray: [-8.724, -8.933, -8.943, -8.988]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.8645
varray: [-63.968, -66.011, -66.554, -66.925]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.117787945054
varray: [-8.724, -8.933, -8.943, -8.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.31886883351
varray: [-63.968, -66.011, -66.554, -66.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1136916165
actual_array: [ 1.93 -15.84]
varray: [-8.724, -8.933, -8.943, -8.988]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.6811154732
actual_array: [ 18.1 -119. ]
varray: [-63.968, -66.011, -66.554, -66.925]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.8645
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [39249.0, 39784.0, 40320.0, 40857.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2234
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.78585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08025
varray: [2.109, 2.092, 2.086, 2.034]
inst: mean_all
operation: mean_all(residuals)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: residuals
SAMP_residuals_STD | 0.0323354810283
varray: [2.109, 2.092, 2.086, 2.034]
inst: std_all
operation: std_all(residuals)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.921
varray: [-0.942, -0.809, -0.921, -1.012]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0841942199125
varray: [-0.942, -0.809, -0.921, -1.012]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.97
varray: [77.283, 77.103, 76.997, 76.497]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.336707588272
varray: [77.283, 77.103, 76.997, 76.497]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.32525
varray: [28.981, 28.348, 28.288, 27.684]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.530130408485
varray: [28.981, 28.348, 28.288, 27.684]
inst: std_all
operation: std_all(slope_shift)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0234236347877
varray: [-9.134, -9.148, -9.101, -9.153]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.922811419883
varray: [-78.967, -80.478, -80.804, -81.002]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.106
varray: [63.647, 63.739, 63.063, 61.975]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.811250064201
varray: [63.647, 63.739, 63.063, 61.975]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.921
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.97
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.32525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0951
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0951
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.37171029949
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7387
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-9.134, -9.148, -9.101, -9.153]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-78.967, -80.478, -80.804, -81.002]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.134
varray: [-9.134, -9.148, -9.101, -9.153]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
dD_mean | -80.31275
varray: [-78.967, -80.478, -80.804, -81.002]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0234236347877
varray: [-9.134, -9.148, -9.101, -9.153]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
dD_stdev | 0.922811419883
varray: [-78.967, -80.478, -80.804, -81.002]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.134
varray: [-9.134, -9.148, -9.101, -9.153]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
measured_dD_mean | -80.31275
varray: [-78.967, -80.478, -80.804, -81.002]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0234236347877
varray: [-9.134, -9.148, -9.101, -9.153]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.922811419883
varray: [-78.967, -80.478, -80.804, -81.002]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3581676013
actual_array: [ 1.93 -15.84]
varray: [-9.134, -9.148, -9.101, -9.153]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -88.4135623058
actual_array: [ 18.1 -119. ]
varray: [-78.967, -80.478, -80.804, -81.002]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -80.31275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [41447.0, 41984.0, 42520.0, 43057.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2539
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.36415
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.091
varray: [2.113, 2.091, 2.085, 2.075]
inst: mean_all
operation: mean_all(residuals)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: residuals
SAMP_residuals_STD | 0.0160831174424
varray: [2.113, 2.091, 2.085, 2.075]
inst: std_all
operation: std_all(residuals)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.895
varray: [-0.832, -0.87, -1.005, -0.873]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0756703376496
varray: [-0.832, -0.87, -1.005, -0.873]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.63525
varray: [76.844, 76.679, 76.611, 76.407]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.180902137448
varray: [76.844, 76.679, 76.611, 76.407]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.26825
varray: [28.726, 28.368, 28.121, 27.858]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.369446771087
varray: [28.726, 28.368, 28.121, 27.858]
inst: std_all
operation: std_all(slope_shift)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0240883789409
varray: [-8.957, -8.948, -8.969, -9.003]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.00342563252
varray: [-66.556, -64.871, -64.49, -64.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.16325
varray: [64.046, 62.947, 62.44, 63.22]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.671397237111
varray: [64.046, 62.947, 62.44, 63.22]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.091
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.895
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.63525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.26825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.87753694808
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.68145
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.957, -8.948, -8.969, -9.003]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-66.556, -64.871, -64.49, -64.409]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.96925
varray: [-8.957, -8.948, -8.969, -9.003]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
dD_mean | -65.0815
varray: [-66.556, -64.871, -64.49, -64.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0240883789409
varray: [-8.957, -8.948, -8.969, -9.003]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
dD_stdev | 1.00342563252
varray: [-66.556, -64.871, -64.49, -64.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.96925
varray: [-8.957, -8.948, -8.969, -9.003]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.0815
varray: [-66.556, -64.871, -64.49, -64.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0240883789409
varray: [-8.957, -8.948, -8.969, -9.003]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.00342563252
varray: [-66.556, -64.871, -64.49, -64.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1882206878
actual_array: [ 1.93 -15.84]
varray: [-8.957, -8.948, -8.969, -9.003]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.96925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.828520462
actual_array: [ 18.1 -119. ]
varray: [-66.556, -64.871, -64.49, -64.409]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.0815
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [43644.0, 44179.0, 44716.0, 45252.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.24315
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.6989
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.106
varray: [2.152, 2.082, 2.106, 2.084]
inst: mean_all
operation: mean_all(residuals)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: residuals
SAMP_residuals_STD | 0.0325371582451
varray: [2.152, 2.082, 2.106, 2.084]
inst: std_all
operation: std_all(residuals)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.974
varray: [-0.953, -1.018, -0.947, -0.978]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0322593655652
varray: [-0.953, -1.018, -0.947, -0.978]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.57675
varray: [77.017, 76.299, 76.607, 76.384]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.320953137805
varray: [77.017, 76.299, 76.607, 76.384]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4045
varray: [29.165, 28.319, 28.227, 27.907]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.536868388589
varray: [29.165, 28.319, 28.227, 27.907]
inst: std_all
operation: std_all(slope_shift)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: slope_shift
SAMP_d18O_STD | 0.214369424126
varray: [-2.358, -1.98, -1.953, -1.88]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.55124306496
varray: [-58.189, -57.237, -57.115, -56.972]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.05175
varray: [64.077, 61.727, 63.381, 63.022]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.985797604312
varray: [64.077, 61.727, 63.381, 63.022]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.106
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.974
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.57675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4045
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06935
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06935
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.1879927368
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.79295
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-2.358, -1.98, -1.953, -1.88]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-58.189, -57.237, -57.115, -56.972]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.04275
varray: [-2.358, -1.98, -1.953, -1.88]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
dD_mean | -57.37825
varray: [-58.189, -57.237, -57.115, -56.972]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
d18O_stdev | 0.214369424126
varray: [-2.358, -1.98, -1.953, -1.88]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
dD_stdev | 0.55124306496
varray: [-58.189, -57.237, -57.115, -56.972]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.04275
varray: [-2.358, -1.98, -1.953, -1.88]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.37825
varray: [-58.189, -57.237, -57.115, -56.972]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.214369424126
varray: [-2.358, -1.98, -1.953, -1.88]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.55124306496
varray: [-58.189, -57.237, -57.115, -56.972]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.04322951134
actual_array: [ 1.93 -15.84]
varray: [-2.358, -1.98, -1.953, -1.88]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.04275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.4405862096
actual_array: [ 18.1 -119. ]
varray: [-58.189, -57.237, -57.115, -56.972]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -57.37825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [45839.0, 46376.0, 46911.0, 47447.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.22815
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.7574
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.093
varray: [2.117, 2.089, 2.089, 2.077]
inst: mean_all
operation: mean_all(residuals)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: residuals
SAMP_residuals_STD | 0.0169705627485
varray: [2.117, 2.089, 2.089, 2.077]
inst: std_all
operation: std_all(residuals)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.90775
varray: [-0.921, -0.888, -0.967, -0.855]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0478147466792
varray: [-0.921, -0.888, -0.967, -0.855]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.619
varray: [77.014, 76.586, 76.497, 76.379]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.276645380707
varray: [77.014, 76.586, 76.497, 76.379]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.49
varray: [29.017, 28.492, 28.296, 28.155]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.377533221496
varray: [29.017, 28.492, 28.296, 28.155]
inst: std_all
operation: std_all(slope_shift)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: slope_shift
SAMP_d18O_STD | 0.136091574562
varray: [-6.145, -6.393, -6.425, -6.428]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.872251110633
varray: [-68.652, -70.015, -70.474, -70.519]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.89525
varray: [63.805, 62.781, 61.941, 63.054]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.769522091951
varray: [63.805, 62.781, 61.941, 63.054]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.093
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.90775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.619
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.49
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08235
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08235
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.78559771991
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.94945
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-6.145, -6.393, -6.425, -6.428]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-68.652, -70.015, -70.474, -70.519]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.34775
varray: [-6.145, -6.393, -6.425, -6.428]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
dD_mean | -69.915
varray: [-68.652, -70.015, -70.474, -70.519]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
d18O_stdev | 0.136091574562
varray: [-6.145, -6.393, -6.425, -6.428]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
dD_stdev | 0.872251110633
varray: [-68.652, -70.015, -70.474, -70.519]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.34775
varray: [-6.145, -6.393, -6.425, -6.428]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.915
varray: [-68.652, -70.015, -70.474, -70.519]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.136091574562
varray: [-6.145, -6.393, -6.425, -6.428]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.872251110633
varray: [-68.652, -70.015, -70.474, -70.519]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.48402746334
actual_array: [ 1.93 -15.84]
varray: [-6.145, -6.393, -6.425, -6.428]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.34775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.0916341101
actual_array: [ 18.1 -119. ]
varray: [-68.652, -70.015, -70.474, -70.519]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -69.915
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [48034.0, 48570.0, 49107.0, 49643.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.24115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.71515
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0805
varray: [2.11, 2.075, 2.06, 2.077]
inst: mean_all
operation: mean_all(residuals)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: residuals
SAMP_residuals_STD | 0.0210792156717
varray: [2.11, 2.075, 2.06, 2.077]
inst: std_all
operation: std_all(residuals)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.019
varray: [-1.02, -1.045, -1.004, -1.007]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0186726180989
varray: [-1.02, -1.045, -1.004, -1.007]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.52725
varray: [76.919, 76.452, 76.386, 76.352]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.264445804151
varray: [76.919, 76.452, 76.386, 76.352]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.218
varray: [28.995, 28.212, 27.595, 28.07]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.581319762839
varray: [28.995, 28.212, 27.595, 28.07]
inst: std_all
operation: std_all(slope_shift)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0550878086936
varray: [-8.469, -8.569, -8.581, -8.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.131958010998
varray: [-68.364, -68.213, -68.051, -68.275]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.728
varray: [62.986, 62.385, 62.502, 63.039]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.332671008656
varray: [62.986, 62.385, 62.502, 63.039]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0805
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.019
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.52725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.218
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.36021789597
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.1167
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.469, -8.569, -8.581, -8.585]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-68.364, -68.213, -68.051, -68.275]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.551
varray: [-8.469, -8.569, -8.581, -8.585]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
dD_mean | -68.22575
varray: [-68.364, -68.213, -68.051, -68.275]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0550878086936
varray: [-8.469, -8.569, -8.581, -8.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
dD_stdev | 0.131958010998
varray: [-68.364, -68.213, -68.051, -68.275]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.551
varray: [-8.469, -8.569, -8.581, -8.585]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.22575
varray: [-68.364, -68.213, -68.051, -68.275]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0550878086936
varray: [-8.469, -8.569, -8.581, -8.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.131958010998
varray: [-68.364, -68.213, -68.051, -68.275]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.75677730951
actual_array: [ 1.93 -15.84]
varray: [-8.469, -8.569, -8.581, -8.585]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.551
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.2522393175
actual_array: [ 18.1 -119. ]
varray: [-68.364, -68.213, -68.051, -68.275]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.22575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [50230.0, 50766.0, 51302.0, 51838.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.25365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8069
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06625
varray: [2.088, 2.068, 2.062, 2.047]
inst: mean_all
operation: mean_all(residuals)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: residuals
SAMP_residuals_STD | 0.0169779268463
varray: [2.088, 2.068, 2.062, 2.047]
inst: std_all
operation: std_all(residuals)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9805
varray: [-0.926, -0.906, -1.004, -1.086]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0820629839412
varray: [-0.926, -0.906, -1.004, -1.086]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.18225
varray: [76.301, 76.306, 76.075, 76.047]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.140488137103
varray: [76.301, 76.306, 76.075, 76.047]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.21475
varray: [28.185, 28.063, 28.358, 28.253]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.123688789576
varray: [28.185, 28.063, 28.358, 28.253]
inst: std_all
operation: std_all(slope_shift)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0358457342883
varray: [-7.602, -7.549, -7.52, -7.534]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.209868212616
varray: [-67.599, -67.297, -67.238, -67.102]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.06075
varray: [63.044, 62.581, 61.232, 61.386]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.890609295183
varray: [63.044, 62.581, 61.232, 61.386]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9805
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.18225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.21475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1091
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1091
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.01528489668
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.78395
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-7.602, -7.549, -7.52, -7.534]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-67.599, -67.297, -67.238, -67.102]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.55125
varray: [-7.602, -7.549, -7.52, -7.534]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
dD_mean | -67.309
varray: [-67.599, -67.297, -67.238, -67.102]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0358457342883
varray: [-7.602, -7.549, -7.52, -7.534]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
dD_stdev | 0.209868212616
varray: [-67.599, -67.297, -67.238, -67.102]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.55125
varray: [-7.602, -7.549, -7.52, -7.534]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.309
varray: [-67.599, -67.297, -67.238, -67.102]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0358457342883
varray: [-7.602, -7.549, -7.52, -7.534]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.209868212616
varray: [-67.599, -67.297, -67.238, -67.102]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.72549095585
actual_array: [ 1.93 -15.84]
varray: [-7.602, -7.549, -7.52, -7.534]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -7.55125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.2540062696
actual_array: [ 18.1 -119. ]
varray: [-67.599, -67.297, -67.238, -67.102]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.309
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [52425.0, 52961.0, 53498.0, 54035.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2679
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.1519
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.05425
varray: [2.094, 2.043, 2.045, 2.035]
inst: mean_all
operation: mean_all(residuals)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: residuals
SAMP_residuals_STD | 0.0268498913716
varray: [2.094, 2.043, 2.045, 2.035]
inst: std_all
operation: std_all(residuals)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.07675
varray: [-0.967, -1.222, -1.059, -1.059]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.106101759332
varray: [-0.967, -1.222, -1.059, -1.059]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.35125
varray: [76.68, 76.37, 76.187, 76.168]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.237337137142
varray: [76.68, 76.37, 76.187, 76.168]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.2135
varray: [28.898, 27.977, 27.938, 28.041]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.458304483941
varray: [28.898, 27.977, 27.938, 28.041]
inst: std_all
operation: std_all(slope_shift)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0248394846967
varray: [-8.234, -8.22, -8.274, -8.262]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.216105684022
varray: [-70.514, -70.867, -70.938, -70.995]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.78875
varray: [63.272, 60.997, 61.428, 61.458]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.01101182815
varray: [63.272, 60.997, 61.428, 61.458]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.05425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.07675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.35125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.2135
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1211
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1211
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.5669202657
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.05595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.234, -8.22, -8.274, -8.262]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-70.514, -70.867, -70.938, -70.995]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.2475
varray: [-8.234, -8.22, -8.274, -8.262]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
dD_mean | -70.8285
varray: [-70.514, -70.867, -70.938, -70.995]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0248394846967
varray: [-8.234, -8.22, -8.274, -8.262]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
dD_stdev | 0.216105684022
varray: [-70.514, -70.867, -70.938, -70.995]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.2475
varray: [-8.234, -8.22, -8.274, -8.262]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.8285
varray: [-70.514, -70.867, -70.938, -70.995]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0248394846967
varray: [-8.234, -8.22, -8.274, -8.262]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.216105684022
varray: [-70.514, -70.867, -70.938, -70.995]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.44370363275
actual_array: [ 1.93 -15.84]
varray: [-8.234, -8.22, -8.274, -8.262]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.2475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.0863282897
actual_array: [ 18.1 -119. ]
varray: [-70.514, -70.867, -70.938, -70.995]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.8285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [54623.0, 55159.0, 55696.0, 56233.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2799
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9829
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.07325
varray: [2.107, 2.073, 2.064, 2.049]
inst: mean_all
operation: mean_all(residuals)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: residuals
SAMP_residuals_STD | 0.0245814971066
varray: [2.107, 2.073, 2.064, 2.049]
inst: std_all
operation: std_all(residuals)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.056
varray: [-1.127, -0.951, -0.989, -1.157]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101252160471
varray: [-1.127, -0.951, -0.989, -1.157]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.385
varray: [76.89, 76.254, 76.161, 76.235]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.339048669073
varray: [76.89, 76.254, 76.161, 76.235]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.29725
varray: [28.934, 27.859, 27.845, 28.551]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.537411310512
varray: [28.934, 27.859, 27.845, 28.551]
inst: std_all
operation: std_all(slope_shift)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0637233604471
varray: [-6.973, -6.854, -6.864, -6.829]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.198895282331
varray: [-67.348, -66.985, -67.025, -66.89]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.258
varray: [62.877, 62.775, 62.619, 60.761]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.00362343536
varray: [62.877, 62.775, 62.619, 60.761]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.07325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.056
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.385
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.29725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1021
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1021
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.69349759809
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.5867
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-6.973, -6.854, -6.864, -6.829]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-67.348, -66.985, -67.025, -66.89]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.88
varray: [-6.973, -6.854, -6.864, -6.829]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
dD_mean | -67.062
varray: [-67.348, -66.985, -67.025, -66.89]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0637233604471
varray: [-6.973, -6.854, -6.864, -6.829]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
dD_stdev | 0.198895282331
varray: [-67.348, -66.985, -67.025, -66.89]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.88
varray: [-6.973, -6.854, -6.864, -6.829]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.062
varray: [-67.348, -66.985, -67.025, -66.89]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0637233604471
varray: [-6.973, -6.854, -6.864, -6.829]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.198895282331
varray: [-67.348, -66.985, -67.025, -66.89]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.03306688493
actual_array: [ 1.93 -15.84]
varray: [-6.973, -6.854, -6.864, -6.829]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.88
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.9850522878
actual_array: [ 18.1 -119. ]
varray: [-67.348, -66.985, -67.025, -66.89]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.062
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [56821.0, 57358.0, 57894.0, 58431.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2609
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.94915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.094
varray: [2.093, 2.095, 2.092, 2.096]
inst: mean_all
operation: mean_all(residuals)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: residuals
SAMP_residuals_STD | 0.00182574185835
varray: [2.093, 2.095, 2.092, 2.096]
inst: std_all
operation: std_all(residuals)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.059
varray: [-0.937, -1.288, -1.018, -0.993]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.156377747778
varray: [-0.937, -1.288, -1.018, -0.993]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.61525
varray: [76.522, 76.845, 76.464, 76.63]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.167903494107
varray: [76.522, 76.845, 76.464, 76.63]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.61475
varray: [28.338, 28.729, 28.844, 28.548]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.221099336649
varray: [28.338, 28.729, 28.844, 28.548]
inst: std_all
operation: std_all(slope_shift)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0554759707741
varray: [-8.897, -8.966, -9.031, -8.983]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0639237827416
varray: [-67.062, -66.928, -67.065, -67.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.772
varray: [63.037, 62.385, 62.722, 62.944]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.289872385715
varray: [63.037, 62.385, 62.722, 62.944]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.094
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.059
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.61525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.61475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08135
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08135
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.73962810582
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0727
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.897, -8.966, -9.031, -8.983]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-67.062, -66.928, -67.065, -67.014]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.96925
varray: [-8.897, -8.966, -9.031, -8.983]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
dD_mean | -67.01725
varray: [-67.062, -66.928, -67.065, -67.014]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0554759707741
varray: [-8.897, -8.966, -9.031, -8.983]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
dD_stdev | 0.0639237827416
varray: [-67.062, -66.928, -67.065, -67.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.96925
varray: [-8.897, -8.966, -9.031, -8.983]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.01725
varray: [-67.062, -66.928, -67.065, -67.014]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0554759707741
varray: [-8.897, -8.966, -9.031, -8.983]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0639237827416
varray: [-67.062, -66.928, -67.065, -67.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1882206878
actual_array: [ 1.93 -15.84]
varray: [-8.897, -8.966, -9.031, -8.983]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.96925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.9363247951
actual_array: [ 18.1 -119. ]
varray: [-67.062, -66.928, -67.065, -67.014]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.01725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [59017.0, 59553.0, 60088.0, 60624.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.24015
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.7189
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0375
varray: [2.063, 2.001, 2.039, 2.047]
inst: mean_all
operation: mean_all(residuals)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: residuals
SAMP_residuals_STD | 0.0262995563968
varray: [2.063, 2.001, 2.039, 2.047]
inst: std_all
operation: std_all(residuals)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.93325
varray: [-0.936, -0.98, -0.911, -0.906]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0338169089855
varray: [-0.936, -0.98, -0.911, -0.906]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.807
varray: [76.927, 76.542, 76.86, 76.899]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.178790380054
varray: [76.927, 76.542, 76.86, 76.899]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.89275
varray: [29.182, 28.122, 29.317, 28.95]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.535719687274
varray: [29.182, 28.122, 29.317, 28.95]
inst: std_all
operation: std_all(slope_shift)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: slope_shift
SAMP_d18O_STD | 0.202458843554
varray: [-15.233, -15.56, -15.654, -15.666]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.85538781663
varray: [-110.324, -114.864, -116.219, -116.484]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.999
varray: [62.704, 61.251, 61.851, 62.19]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.609621193857
varray: [62.704, 61.251, 61.851, 62.19]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.93325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.807
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.89275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.33691130163
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.8457
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-15.233, -15.56, -15.654, -15.666]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-110.324, -114.864, -116.219, -116.484]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -15.52825
varray: [-15.233, -15.56, -15.654, -15.666]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
dD_mean | -114.47275
varray: [-110.324, -114.864, -116.219, -116.484]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
d18O_stdev | 0.202458843554
varray: [-15.233, -15.56, -15.654, -15.666]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
dD_stdev | 2.85538781663
varray: [-110.324, -114.864, -116.219, -116.484]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
measured_d18O_mean | -15.52825
varray: [-15.233, -15.56, -15.654, -15.666]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
measured_dD_mean | -114.47275
varray: [-110.324, -114.864, -116.219, -116.484]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.202458843554
varray: [-15.233, -15.56, -15.654, -15.666]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.85538781663
varray: [-110.324, -114.864, -116.219, -116.484]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -16.9541193561
actual_array: [ 1.93 -15.84]
varray: [-15.233, -15.56, -15.654, -15.666]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -15.52825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -125.609789101
actual_array: [ 18.1 -119. ]
varray: [-110.324, -114.864, -116.219, -116.484]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -114.47275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [61210.0, 61746.0, 62282.0, 62818.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.29665
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.52715
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1335
varray: [2.109, 2.122, 2.188, 2.115]
inst: mean_all
operation: mean_all(residuals)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: residuals
SAMP_residuals_STD | 0.0367196586767
varray: [2.109, 2.122, 2.188, 2.115]
inst: std_all
operation: std_all(residuals)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98125
varray: [-0.844, -0.977, -1.145, -0.959]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.124049385327
varray: [-0.844, -0.977, -1.145, -0.959]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.53975
varray: [76.324, 76.502, 77.21, 76.123]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.472894896709
varray: [76.324, 76.502, 77.21, 76.123]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.618
varray: [28.481, 28.306, 29.682, 28.003]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.736304737637
varray: [28.481, 28.306, 29.682, 28.003]
inst: std_all
operation: std_all(slope_shift)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: slope_shift
SAMP_d18O_STD | 0.41480678233
varray: [-5.898, -5.075, -5.02, -5.123]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.8514284752
varray: [-46.216, -37.987, -36.292, -35.806]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.96725
varray: [63.728, 64.042, 64.592, 63.507]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.470804541893
varray: [63.728, 64.042, 64.592, 63.507]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1335
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.53975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.618
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.92382834946
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.87745
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.898, -5.075, -5.02, -5.123]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-46.216, -37.987, -36.292, -35.806]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.279
varray: [-5.898, -5.075, -5.02, -5.123]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
dD_mean | -39.07525
varray: [-46.216, -37.987, -36.292, -35.806]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
d18O_stdev | 0.41480678233
varray: [-5.898, -5.075, -5.02, -5.123]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
dD_stdev | 4.8514284752
varray: [-46.216, -37.987, -36.292, -35.806]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.279
varray: [-5.898, -5.075, -5.02, -5.123]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.07525
varray: [-46.216, -37.987, -36.292, -35.806]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.41480678233
varray: [-5.898, -5.075, -5.02, -5.123]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.8514284752
varray: [-46.216, -37.987, -36.292, -35.806]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.38156455714
actual_array: [ 1.93 -15.84]
varray: [-5.898, -5.075, -5.02, -5.123]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.279
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.5107694929
actual_array: [ 18.1 -119. ]
varray: [-46.216, -37.987, -36.292, -35.806]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -39.07525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [63404.0, 63941.0, 64476.0, 65012.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.20065
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.7944
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0675
varray: [2.088, 2.028, 2.084, 2.07]
inst: mean_all
operation: mean_all(residuals)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: residuals
SAMP_residuals_STD | 0.027440845468
varray: [2.088, 2.028, 2.084, 2.07]
inst: std_all
operation: std_all(residuals)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9855
varray: [-1.009, -1.021, -0.947, -0.965]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0351899606896
varray: [-1.009, -1.021, -0.947, -0.965]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.63425
varray: [75.979, 75.189, 75.749, 75.62]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.331898955507
varray: [75.979, 75.189, 75.749, 75.62]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.44225
varray: [27.95, 26.927, 27.572, 27.32]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.430144452481
varray: [27.95, 26.927, 27.572, 27.32]
inst: std_all
operation: std_all(slope_shift)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0196807858922
varray: [-5.228, -5.262, -5.267, -5.271]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.297607543587
varray: [-35.654, -35.444, -35.08, -35.033]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.80425
varray: [62.116, 60.654, 62.321, 62.126]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.772618221806
varray: [62.116, 60.654, 62.321, 62.126]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9855
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.63425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.44225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.95782287908
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.04045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.228, -5.262, -5.267, -5.271]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-35.654, -35.444, -35.08, -35.033]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.257
varray: [-5.228, -5.262, -5.267, -5.271]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
dD_mean | -35.30275
varray: [-35.654, -35.444, -35.08, -35.033]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0196807858922
varray: [-5.228, -5.262, -5.267, -5.271]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
dD_stdev | 0.297607543587
varray: [-35.654, -35.444, -35.08, -35.033]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.257
varray: [-5.228, -5.262, -5.267, -5.271]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.30275
varray: [-35.654, -35.444, -35.08, -35.033]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0196807858922
varray: [-5.228, -5.262, -5.267, -5.271]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.297607543587
varray: [-35.654, -35.444, -35.08, -35.033]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.35887058386
actual_array: [ 1.93 -15.84]
varray: [-5.228, -5.262, -5.267, -5.271]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.257
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.4029601958
actual_array: [ 18.1 -119. ]
varray: [-35.654, -35.444, -35.08, -35.033]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.30275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [65598.0, 66134.0, 66670.0, 67205.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.26665
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.6999
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08575
varray: [2.106, 2.071, 2.076, 2.09]
inst: mean_all
operation: mean_all(residuals)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: residuals
SAMP_residuals_STD | 0.0157135822353
varray: [2.106, 2.071, 2.076, 2.09]
inst: std_all
operation: std_all(residuals)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9375
varray: [-0.984, -0.962, -1.095, -0.709]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.163075647068
varray: [-0.984, -0.962, -1.095, -0.709]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.77075
varray: [76.126, 75.554, 75.788, 75.615]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.256735109922
varray: [76.126, 75.554, 75.788, 75.615]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.50775
varray: [27.945, 27.074, 27.55, 27.462]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.357400219175
varray: [27.945, 27.074, 27.55, 27.462]
inst: std_all
operation: std_all(slope_shift)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0400541300409
varray: [-5.134, -5.041, -5.111, -5.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.311756315092
varray: [-34.252, -33.526, -33.726, -33.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.458
varray: [62.97, 62.1, 61.514, 63.248]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.796982643057
varray: [62.97, 62.1, 61.514, 63.248]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.77075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.50775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0896
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0896
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.11887742202
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.3867
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.134, -5.041, -5.111, -5.108]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-34.252, -33.526, -33.726, -33.952]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.0985
varray: [-5.134, -5.041, -5.111, -5.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
dD_mean | -33.864
varray: [-34.252, -33.526, -33.726, -33.952]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0400541300409
varray: [-5.134, -5.041, -5.111, -5.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
dD_stdev | 0.311756315092
varray: [-34.252, -33.526, -33.726, -33.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.0985
varray: [-5.134, -5.041, -5.111, -5.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
measured_dD_mean | -33.864
varray: [-34.252, -33.526, -33.726, -33.952]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0400541300409
varray: [-5.134, -5.041, -5.111, -5.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.311756315092
varray: [-34.252, -33.526, -33.726, -33.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.19537082187
actual_array: [ 1.93 -15.84]
varray: [-5.134, -5.041, -5.111, -5.108]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.0985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.8363304739
actual_array: [ 18.1 -119. ]
varray: [-34.252, -33.526, -33.726, -33.952]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -33.864
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [67792.0, 68328.0, 68864.0, 69400.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2484
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.5634
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0875
varray: [2.128, 2.074, 2.072, 2.076]
inst: mean_all
operation: mean_all(residuals)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: residuals
SAMP_residuals_STD | 0.0270493376382
varray: [2.128, 2.074, 2.072, 2.076]
inst: std_all
operation: std_all(residuals)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00175
varray: [-0.966, -1.14, -0.853, -1.048]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.122006488898
varray: [-0.966, -1.14, -0.853, -1.048]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.87575
varray: [76.297, 75.901, 75.551, 75.754]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.315370020769
varray: [76.297, 75.901, 75.551, 75.754]
inst: std_all
operation: std_all(baseline_shift)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.624
varray: [28.378, 28.018, 26.728, 27.372]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.728027472009
varray: [28.378, 28.018, 26.728, 27.372]
inst: std_all
operation: std_all(slope_shift)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0249315730216
varray: [-5.349, -5.365, -5.307, -5.35]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.18942456722
varray: [-35.493, -35.942, -35.62, -35.711]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.12825
varray: [63.066, 61.213, 62.365, 61.869]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.783223839865
varray: [63.066, 61.213, 62.365, 61.869]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.87575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.624
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.03843059738
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.71645
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.349, -5.365, -5.307, -5.35]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-35.493, -35.942, -35.62, -35.711]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.34275
varray: [-5.349, -5.365, -5.307, -5.35]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
dD_mean | -35.6915
varray: [-35.493, -35.942, -35.62, -35.711]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0249315730216
varray: [-5.349, -5.365, -5.307, -5.35]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
dD_stdev | 0.18942456722
varray: [-35.493, -35.942, -35.62, -35.711]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.34275
varray: [-5.349, -5.365, -5.307, -5.35]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.6915
varray: [-35.493, -35.942, -35.62, -35.711]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0249315730216
varray: [-5.349, -5.365, -5.307, -5.35]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.18942456722
varray: [-35.493, -35.942, -35.62, -35.711]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.44732550242
actual_array: [ 1.93 -15.84]
varray: [-5.349, -5.365, -5.307, -5.35]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.34275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.8262632745
actual_array: [ 18.1 -119. ]
varray: [-35.493, -35.942, -35.62, -35.711]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.6915
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [69987.0, 70522.0, 71058.0, 71594.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.24665
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4584
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10725
varray: [2.114, 2.081, 2.111, 2.123]
inst: mean_all
operation: mean_all(residuals)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: residuals
SAMP_residuals_STD | 0.0182277261336
varray: [2.114, 2.081, 2.111, 2.123]
inst: std_all
operation: std_all(residuals)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03625
varray: [-1.01, -1.057, -1.048, -1.03]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0207906228863
varray: [-1.01, -1.057, -1.048, -1.03]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.2205
varray: [76.251, 76.099, 76.197, 76.335]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0989191589127
varray: [76.251, 76.099, 76.197, 76.335]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.12425
varray: [28.248, 27.546, 28.213, 28.49]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.404696079381
varray: [28.248, 27.546, 28.213, 28.49]
inst: std_all
operation: std_all(slope_shift)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: slope_shift
SAMP_d18O_STD | 0.014997221965
varray: [-5.406, -5.399, -5.433, -5.419]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0689462350918
varray: [-36.018, -35.852, -35.918, -35.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.6205
varray: [62.558, 62.137, 62.804, 62.983]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.366404603319
varray: [62.558, 62.137, 62.804, 62.983]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.2205
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.12425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0681
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0681
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.13053071919
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.2242
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.406, -5.399, -5.433, -5.419]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.018, -35.852, -35.918, -35.949]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.41425
varray: [-5.406, -5.399, -5.433, -5.419]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
dD_mean | -35.93425
varray: [-36.018, -35.852, -35.918, -35.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
d18O_stdev | 0.014997221965
varray: [-5.406, -5.399, -5.433, -5.419]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
dD_stdev | 0.0689462350918
varray: [-36.018, -35.852, -35.918, -35.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.41425
varray: [-5.406, -5.399, -5.433, -5.419]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.93425
varray: [-36.018, -35.852, -35.918, -35.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.014997221965
varray: [-5.406, -5.399, -5.433, -5.419]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0689462350918
varray: [-36.018, -35.852, -35.918, -35.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.52108091556
actual_array: [ 1.93 -15.84]
varray: [-5.406, -5.399, -5.433, -5.419]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.41425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.0905895056
actual_array: [ 18.1 -119. ]
varray: [-36.018, -35.852, -35.918, -35.949]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.93425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [72182.0, 72718.0, 73254.0, 73791.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2269
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.11365
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09875
varray: [2.112, 2.11, 2.058, 2.115]
inst: mean_all
operation: mean_all(residuals)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: residuals
SAMP_residuals_STD | 0.0272442654517
varray: [2.112, 2.11, 2.058, 2.115]
inst: std_all
operation: std_all(residuals)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.91425
varray: [-0.927, -1.058, -0.799, -0.873]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109256197383
varray: [-0.927, -1.058, -0.799, -0.873]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.98575
varray: [76.2, 76.282, 75.287, 76.174]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.468101395711
varray: [76.2, 76.282, 75.287, 76.174]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.96175
varray: [28.108, 28.398, 26.988, 28.353]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.66155593112
varray: [28.108, 28.398, 26.988, 28.353]
inst: std_all
operation: std_all(slope_shift)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0123962359879
varray: [-5.48, -5.466, -5.495, -5.473]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0999649938729
varray: [-36.283, -36.389, -36.527, -36.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.55625
varray: [62.759, 62.725, 61.711, 63.03]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.579790407532
varray: [62.759, 62.725, 61.711, 63.03]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.91425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.98575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.96175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0766
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0766
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.52127243892
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.28845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.48, -5.466, -5.495, -5.473]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.283, -36.389, -36.527, -36.407]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.4785
varray: [-5.48, -5.466, -5.495, -5.473]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
dD_mean | -36.4015
varray: [-36.283, -36.389, -36.527, -36.407]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0123962359879
varray: [-5.48, -5.466, -5.495, -5.473]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
dD_stdev | 0.0999649938729
varray: [-36.283, -36.389, -36.527, -36.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.4785
varray: [-5.48, -5.466, -5.495, -5.473]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.4015
varray: [-36.283, -36.389, -36.527, -36.407]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0123962359879
varray: [-5.48, -5.466, -5.495, -5.473]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0999649938729
varray: [-36.283, -36.389, -36.527, -36.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.58735763296
actual_array: [ 1.93 -15.84]
varray: [-5.48, -5.466, -5.495, -5.473]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.4785
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.5993698619
actual_array: [ 18.1 -119. ]
varray: [-36.283, -36.389, -36.527, -36.407]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.4015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [74375.0, 74909.0, 75446.0, 75982.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2354
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.3484
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10675
varray: [2.12, 2.097, 2.101, 2.109]
inst: mean_all
operation: mean_all(residuals)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: residuals
SAMP_residuals_STD | 0.0101447851957
varray: [2.12, 2.097, 2.101, 2.109]
inst: std_all
operation: std_all(residuals)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.96525
varray: [-1.019, -1.143, -0.936, -0.763]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.159420565382
varray: [-1.019, -1.143, -0.936, -0.763]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.04475
varray: [76.14, 76.151, 75.996, 75.892]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.123923027185
varray: [76.14, 76.151, 75.996, 75.892]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.1475
varray: [28.312, 28.106, 28.316, 27.856]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.217674833946
varray: [28.312, 28.106, 28.316, 27.856]
inst: std_all
operation: std_all(slope_shift)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0229256043177
varray: [-5.269, -5.286, -5.254, -5.232]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.102264363294
varray: [-35.428, -35.193, -35.245, -35.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.794
varray: [63.057, 61.63, 62.777, 63.712]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.869305853349
varray: [63.057, 61.63, 62.777, 63.712]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.96525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.04475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.1475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0686
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0686
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.15351552624
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0507
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.269, -5.286, -5.254, -5.232]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-35.428, -35.193, -35.245, -35.254]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.26025
varray: [-5.269, -5.286, -5.254, -5.232]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
dD_mean | -35.28
varray: [-35.428, -35.193, -35.245, -35.254]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0229256043177
varray: [-5.269, -5.286, -5.254, -5.232]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
dD_stdev | 0.102264363294
varray: [-35.428, -35.193, -35.245, -35.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.26025
varray: [-5.269, -5.286, -5.254, -5.232]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.28
varray: [-35.428, -35.193, -35.245, -35.254]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0229256043177
varray: [-5.269, -5.286, -5.254, -5.232]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.102264363294
varray: [-35.428, -35.193, -35.245, -35.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.36222310264
actual_array: [ 1.93 -15.84]
varray: [-5.269, -5.286, -5.254, -5.232]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.26025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.3781881186
actual_array: [ 18.1 -119. ]
varray: [-35.428, -35.193, -35.245, -35.254]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.28
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [76566.0, 77102.0, 77638.0, 78174.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2274
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.2894
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.104
varray: [2.134, 2.108, 2.089, 2.085]
inst: mean_all
operation: mean_all(residuals)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: residuals
SAMP_residuals_STD | 0.0223755819291
varray: [2.134, 2.108, 2.089, 2.085]
inst: std_all
operation: std_all(residuals)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04275
varray: [-1.051, -1.175, -0.932, -1.013]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101174354458
varray: [-1.051, -1.175, -0.932, -1.013]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.68525
varray: [76.97, 76.943, 76.373, 76.455]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.315189228877
varray: [76.97, 76.943, 76.373, 76.455]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.718
varray: [29.283, 29.097, 28.148, 28.344]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.556070139461
varray: [29.283, 29.097, 28.148, 28.344]
inst: std_all
operation: std_all(slope_shift)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: slope_shift
SAMP_d18O_STD | 0.150222057857
varray: [-8.684, -8.92, -8.982, -9.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.78690167142
varray: [-62.256, -65.068, -65.851, -66.192]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.0455
varray: [63.765, 62.844, 63.02, 62.553]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.516878128769
varray: [63.765, 62.844, 63.02, 62.553]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.104
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.68525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.718
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07135
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07135
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.27993196497
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7992
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.684, -8.92, -8.982, -9.018]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-62.256, -65.068, -65.851, -66.192]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.901
varray: [-8.684, -8.92, -8.982, -9.018]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
dD_mean | -64.84175
varray: [-62.256, -65.068, -65.851, -66.192]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
d18O_stdev | 0.150222057857
varray: [-8.684, -8.92, -8.982, -9.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
dD_stdev | 1.78690167142
varray: [-62.256, -65.068, -65.851, -66.192]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.901
varray: [-8.684, -8.92, -8.982, -9.018]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.84175
varray: [-62.256, -65.068, -65.851, -66.192]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.150222057857
varray: [-8.684, -8.92, -8.982, -9.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.78690167142
varray: [-62.256, -65.068, -65.851, -66.192]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1178177934
actual_array: [ 1.93 -15.84]
varray: [-8.684, -8.92, -8.982, -9.018]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.901
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.5674608785
actual_array: [ 18.1 -119. ]
varray: [-62.256, -65.068, -65.851, -66.192]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -64.84175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [78759.0, 79295.0, 79832.0, 80368.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.23015
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.6489
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09575
varray: [2.111, 2.096, 2.078, 2.098]
inst: mean_all
operation: mean_all(residuals)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: residuals
SAMP_residuals_STD | 0.0135738719605
varray: [2.111, 2.096, 2.078, 2.098]
inst: std_all
operation: std_all(residuals)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.08025
varray: [-1.114, -1.037, -1.106, -1.064]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0362249913734
varray: [-1.114, -1.037, -1.106, -1.064]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.0735
varray: [76.189, 76.092, 75.951, 76.062]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0980153049273
varray: [76.189, 76.092, 75.951, 76.062]
inst: std_all
operation: std_all(baseline_shift)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.87
varray: [28.324, 27.578, 27.916, 27.662]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.335042285888
varray: [28.324, 27.578, 27.916, 27.662]
inst: std_all
operation: std_all(slope_shift)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: slope_shift
SAMP_d18O_STD | 0.109411684324
varray: [-5.679, -5.504, -5.467, -5.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.7695936445
varray: [-40.386, -37.577, -36.829, -36.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.452
varray: [62.575, 62.634, 61.526, 63.073]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.656091965098
varray: [62.575, 62.634, 61.526, 63.073]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.08025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.0735
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.87
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0796
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0796
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.65918128117
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.3927
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.679, -5.504, -5.467, -5.433]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-40.386, -37.577, -36.829, -36.49]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.52075
varray: [-5.679, -5.504, -5.467, -5.433]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
dD_mean | -37.8205
varray: [-40.386, -37.577, -36.829, -36.49]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
d18O_stdev | 0.109411684324
varray: [-5.679, -5.504, -5.467, -5.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
dD_stdev | 1.7695936445
varray: [-40.386, -37.577, -36.829, -36.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.52075
varray: [-5.679, -5.504, -5.467, -5.433]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.8205
varray: [-40.386, -37.577, -36.829, -36.49]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.109411684324
varray: [-5.679, -5.504, -5.467, -5.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.7695936445
varray: [-40.386, -37.577, -36.829, -36.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.63094037709
actual_array: [ 1.93 -15.84]
varray: [-5.679, -5.504, -5.467, -5.433]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.52075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.1444941541
actual_array: [ 18.1 -119. ]
varray: [-40.386, -37.577, -36.829, -36.49]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.8205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [80953.0, 81490.0, 82026.0, 82562.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2384
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.26065
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0555
varray: [2.096, 2.051, 2.037, 2.038]
inst: mean_all
operation: mean_all(residuals)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: residuals
SAMP_residuals_STD | 0.0277428669511
varray: [2.096, 2.051, 2.037, 2.038]
inst: std_all
operation: std_all(residuals)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9565
varray: [-1.128, -0.934, -0.889, -0.875]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117071203405
varray: [-1.128, -0.934, -0.889, -0.875]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.48675
varray: [76.105, 75.297, 75.352, 75.193]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.417406576374
varray: [76.105, 75.297, 75.352, 75.193]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.123
varray: [27.682, 26.876, 26.847, 27.087]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.387712092495
varray: [27.682, 26.876, 26.847, 27.087]
inst: std_all
operation: std_all(slope_shift)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0217638844572
varray: [-5.461, -5.51, -5.468, -5.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0793746391068
varray: [-36.59, -36.432, -36.413, -36.483]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.835
varray: [62.125, 61.876, 61.778, 61.561]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.233884586923
varray: [62.125, 61.876, 61.778, 61.561]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0555
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.48675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.123
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.5094582481
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.0097
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.461, -5.51, -5.468, -5.475]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.59, -36.432, -36.413, -36.483]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.4785
varray: [-5.461, -5.51, -5.468, -5.475]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
dD_mean | -36.4795
varray: [-36.59, -36.432, -36.413, -36.483]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0217638844572
varray: [-5.461, -5.51, -5.468, -5.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
dD_stdev | 0.0793746391068
varray: [-36.59, -36.432, -36.413, -36.483]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.4785
varray: [-5.461, -5.51, -5.468, -5.475]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.4795
varray: [-36.59, -36.432, -36.413, -36.483]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0217638844572
varray: [-5.461, -5.51, -5.468, -5.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0793746391068
varray: [-36.59, -36.432, -36.413, -36.483]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.58735763296
actual_array: [ 1.93 -15.84]
varray: [-5.461, -5.51, -5.468, -5.475]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.4785
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.6843026982
actual_array: [ 18.1 -119. ]
varray: [-36.59, -36.432, -36.413, -36.483]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.4795
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [83148.0, 83684.0, 84220.0, 84756.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.27865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.8474
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06275
varray: [2.042, 2.116, 2.01, 2.083]
inst: mean_all
operation: mean_all(residuals)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: residuals
SAMP_residuals_STD | 0.0463995330436
varray: [2.042, 2.116, 2.01, 2.083]
inst: std_all
operation: std_all(residuals)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.931
varray: [-0.982, -0.863, -0.922, -0.957]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0515816504841
varray: [-0.982, -0.863, -0.922, -0.957]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.554
varray: [75.419, 76.143, 74.864, 75.79]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.546791245967
varray: [75.419, 76.143, 74.864, 75.79]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.20925
varray: [26.887, 27.949, 26.451, 27.55]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.668861906525
varray: [26.887, 27.949, 26.451, 27.55]
inst: std_all
operation: std_all(slope_shift)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0326036807738
varray: [-5.461, -5.535, -5.49, -5.472]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.144279820719
varray: [-36.142, -36.476, -36.409, -36.345]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.97675
varray: [60.62, 63.731, 61.103, 62.453]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.40337886902
varray: [60.62, 63.731, 61.103, 62.453]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.931
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.554
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.20925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1126
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1126
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.17617854598
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.86795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.461, -5.535, -5.49, -5.472]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.142, -36.476, -36.409, -36.345]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.4895
varray: [-5.461, -5.535, -5.49, -5.472]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
dD_mean | -36.343
varray: [-36.142, -36.476, -36.409, -36.345]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0326036807738
varray: [-5.461, -5.535, -5.49, -5.472]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
dD_stdev | 0.144279820719
varray: [-36.142, -36.476, -36.409, -36.345]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.4895
varray: [-5.461, -5.535, -5.49, -5.472]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.343
varray: [-36.142, -36.476, -36.409, -36.345]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0326036807738
varray: [-5.461, -5.535, -5.49, -5.472]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.144279820719
varray: [-36.142, -36.476, -36.409, -36.345]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.5987046196
actual_array: [ 1.93 -15.84]
varray: [-5.461, -5.535, -5.49, -5.472]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.4895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.5356702346
actual_array: [ 18.1 -119. ]
varray: [-36.142, -36.476, -36.409, -36.345]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.343
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [85343.0, 85878.0, 86413.0, 86949.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2714
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.78015
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0055
varray: [2.092, 2.014, 1.953, 1.963]
inst: mean_all
operation: mean_all(residuals)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: residuals
SAMP_residuals_STD | 0.0635531273817
varray: [2.092, 2.014, 1.953, 1.963]
inst: std_all
operation: std_all(residuals)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.014
varray: [-1.025, -0.956, -0.981, -1.094]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0604814021001
varray: [-1.025, -0.956, -0.981, -1.094]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 74.881
varray: [76.005, 74.759, 74.21, 74.55]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.782743465852
varray: [76.005, 74.759, 74.21, 74.55]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.17925
varray: [27.529, 26.431, 25.154, 25.603]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.04377595776
varray: [27.529, 26.431, 25.154, 25.603]
inst: std_all
operation: std_all(slope_shift)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0280237994093
varray: [-5.46, -5.486, -5.418, -5.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.125497675941
varray: [-36.286, -36.39, -36.089, -36.281]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.1165
varray: [62.613, 60.499, 58.632, 58.722]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.87324931603
varray: [62.613, 60.499, 58.632, 58.722]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0055
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.014
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 74.881
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.17925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.80793895235
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.7282
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.46, -5.486, -5.418, -5.456]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.286, -36.39, -36.089, -36.281]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.455
varray: [-5.46, -5.486, -5.418, -5.456]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
dD_mean | -36.2615
varray: [-36.286, -36.39, -36.089, -36.281]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0280237994093
varray: [-5.46, -5.486, -5.418, -5.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
dD_stdev | 0.125497675941
varray: [-36.286, -36.39, -36.089, -36.281]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.455
varray: [-5.46, -5.486, -5.418, -5.456]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.2615
varray: [-36.286, -36.39, -36.089, -36.281]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0280237994093
varray: [-5.46, -5.486, -5.418, -5.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.125497675941
varray: [-36.286, -36.39, -36.089, -36.281]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.56311634333
actual_array: [ 1.93 -15.84]
varray: [-5.46, -5.486, -5.418, -5.456]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.4469263094
actual_array: [ 18.1 -119. ]
varray: [-36.286, -36.39, -36.089, -36.281]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.2615
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [87536.0, 88073.0, 88609.0, 89145.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.32865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.45315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11825
varray: [2.129, 2.136, 2.107, 2.101]
inst: mean_all
operation: mean_all(residuals)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: residuals
SAMP_residuals_STD | 0.0168794747154
varray: [2.129, 2.136, 2.107, 2.101]
inst: std_all
operation: std_all(residuals)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.01125
varray: [-0.966, -1.044, -0.955, -1.08]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0605825882577
varray: [-0.966, -1.044, -0.955, -1.08]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.145
varray: [76.254, 76.297, 76.028, 76.001]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.152107418184
varray: [76.254, 76.297, 76.028, 76.001]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.0565
varray: [28.509, 28.295, 27.697, 27.725]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.40856292865
varray: [28.509, 28.295, 27.697, 27.725]
inst: std_all
operation: std_all(slope_shift)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0247386337537
varray: [-5.523, -5.511, -5.503, -5.559]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.111016515288
varray: [-36.343, -36.192, -36.251, -36.446]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.99175
varray: [63.208, 63.595, 62.932, 62.232]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.574881074658
varray: [63.208, 63.595, 62.932, 62.232]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.01125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.145
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.0565
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0571
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0571
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.62486496426
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.85295
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.523, -5.511, -5.503, -5.559]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.343, -36.192, -36.251, -36.446]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.524
varray: [-5.523, -5.511, -5.503, -5.559]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
dD_mean | -36.308
varray: [-36.343, -36.192, -36.251, -36.446]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0247386337537
varray: [-5.523, -5.511, -5.503, -5.559]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
dD_stdev | 0.111016515288
varray: [-36.343, -36.192, -36.251, -36.446]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.524
varray: [-5.523, -5.511, -5.503, -5.559]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.308
varray: [-36.343, -36.192, -36.251, -36.446]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0247386337537
varray: [-5.523, -5.511, -5.503, -5.559]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.111016515288
varray: [-36.343, -36.192, -36.251, -36.446]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.63429289587
actual_array: [ 1.93 -15.84]
varray: [-5.523, -5.511, -5.503, -5.559]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.524
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.4975593465
actual_array: [ 18.1 -119. ]
varray: [-36.343, -36.192, -36.251, -36.446]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [89727.0, 90263.0, 90799.0, 91334.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2159
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.18915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06875
varray: [2.043, 2.081, 2.06, 2.091]
inst: mean_all
operation: mean_all(residuals)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: residuals
SAMP_residuals_STD | 0.02148449053
varray: [2.043, 2.081, 2.06, 2.091]
inst: std_all
operation: std_all(residuals)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9185
varray: [-0.945, -1.034, -0.926, -0.769]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.110225526384
varray: [-0.945, -1.034, -0.926, -0.769]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.54725
varray: [75.199, 75.738, 75.438, 75.814]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.283287104072
varray: [75.199, 75.738, 75.438, 75.814]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.251
varray: [26.534, 27.654, 27.361, 27.455]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.493360584833
varray: [26.534, 27.654, 27.361, 27.455]
inst: std_all
operation: std_all(slope_shift)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0232002155162
varray: [-5.345, -5.339, -5.349, -5.39]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0441238786449
varray: [-35.681, -35.752, -35.781, -35.711]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.814
varray: [61.806, 61.446, 60.851, 63.153]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.975666951372
varray: [61.806, 61.446, 60.851, 63.153]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9185
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.54725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.251
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1066
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1066
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.90036086147
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.0307
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.345, -5.339, -5.349, -5.39]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-35.681, -35.752, -35.781, -35.711]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.35575
varray: [-5.345, -5.339, -5.349, -5.39]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
dD_mean | -35.73125
varray: [-35.681, -35.752, -35.781, -35.711]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0232002155162
varray: [-5.345, -5.339, -5.349, -5.39]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
dD_stdev | 0.0441238786449
varray: [-35.681, -35.752, -35.781, -35.711]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.35575
varray: [-5.345, -5.339, -5.349, -5.39]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.73125
varray: [-35.681, -35.752, -35.781, -35.711]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0232002155162
varray: [-5.345, -5.339, -5.349, -5.39]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0441238786449
varray: [-35.681, -35.752, -35.781, -35.711]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.46073557754
actual_array: [ 1.93 -15.84]
varray: [-5.345, -5.339, -5.349, -5.39]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.35575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.8695463545
actual_array: [ 18.1 -119. ]
varray: [-35.681, -35.752, -35.781, -35.711]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.73125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [91917.0, 92452.0, 92988.0, 93524.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2654
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.7869
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.067
varray: [2.089, 2.062, 2.059, 2.058]
inst: mean_all
operation: mean_all(residuals)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: residuals
SAMP_residuals_STD | 0.0147648230602
varray: [2.089, 2.062, 2.059, 2.058]
inst: std_all
operation: std_all(residuals)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9615
varray: [-0.804, -0.99, -1.07, -0.982]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.112266053047
varray: [-0.804, -0.99, -1.07, -0.982]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.63225
varray: [75.711, 75.609, 75.655, 75.554]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0667900940759
varray: [75.711, 75.609, 75.655, 75.554]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.45125
varray: [27.613, 27.3, 27.415, 27.477]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.130405457452
varray: [27.613, 27.3, 27.415, 27.477]
inst: std_all
operation: std_all(slope_shift)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0134505266316
varray: [-5.847, -5.846, -5.875, -5.855]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.194869828347
varray: [-38.418, -38.763, -38.676, -38.876]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.753
varray: [62.567, 61.991, 61.166, 61.288]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.653205939961
varray: [62.567, 61.991, 61.166, 61.288]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.067
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9615
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.63225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.45125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10835
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10835
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.98080768612
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.0917
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.847, -5.846, -5.875, -5.855]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-38.418, -38.763, -38.676, -38.876]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.85575
varray: [-5.847, -5.846, -5.875, -5.855]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
dD_mean | -38.68325
varray: [-38.418, -38.763, -38.676, -38.876]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0134505266316
varray: [-5.847, -5.846, -5.875, -5.855]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
dD_stdev | 0.194869828347
varray: [-38.418, -38.763, -38.676, -38.876]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.85575
varray: [-5.847, -5.846, -5.875, -5.855]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.68325
varray: [-38.418, -38.763, -38.676, -38.876]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0134505266316
varray: [-5.847, -5.846, -5.875, -5.855]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.194869828347
varray: [-38.418, -38.763, -38.676, -38.876]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.9765076974
actual_array: [ 1.93 -15.84]
varray: [-5.847, -5.846, -5.875, -5.855]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.85575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.083927546
actual_array: [ 18.1 -119. ]
varray: [-38.418, -38.763, -38.676, -38.876]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.68325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [94109.0, 94644.0, 95180.0, 95716.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.26715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.7019
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12175
varray: [2.117, 2.142, 2.123, 2.105]
inst: mean_all
operation: mean_all(residuals)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: residuals
SAMP_residuals_STD | 0.0154353490404
varray: [2.117, 2.142, 2.123, 2.105]
inst: std_all
operation: std_all(residuals)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.89
varray: [-0.864, -0.83, -0.877, -0.989]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0689105700649
varray: [-0.864, -0.83, -0.877, -0.989]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.1635
varray: [76.124, 76.348, 76.196, 75.986]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.15073486657
varray: [76.124, 76.348, 76.196, 75.986]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.073
varray: [27.601, 28.395, 28.143, 28.153]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.335543340469
varray: [27.601, 28.395, 28.143, 28.153]
inst: std_all
operation: std_all(slope_shift)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0169803808359
varray: [-5.643, -5.636, -5.619, -5.66]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0329494562423
varray: [-37.485, -37.544, -37.468, -37.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.595
varray: [63.755, 64.711, 63.518, 62.396]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.951242345567
varray: [63.755, 64.711, 63.518, 62.396]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.89
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.1635
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.073
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0536
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0536
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.46397131496
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2497
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.643, -5.636, -5.619, -5.66]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-37.485, -37.544, -37.468, -37.489]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.6395
varray: [-5.643, -5.636, -5.619, -5.66]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
dD_mean | -37.4965
varray: [-37.485, -37.544, -37.468, -37.489]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0169803808359
varray: [-5.643, -5.636, -5.619, -5.66]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
dD_stdev | 0.0329494562423
varray: [-37.485, -37.544, -37.468, -37.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.6395
varray: [-5.643, -5.636, -5.619, -5.66]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.4965
varray: [-37.485, -37.544, -37.468, -37.489]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0169803808359
varray: [-5.643, -5.636, -5.619, -5.66]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0329494562423
varray: [-37.485, -37.544, -37.468, -37.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.75343625556
actual_array: [ 1.93 -15.84]
varray: [-5.643, -5.636, -5.619, -5.66]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.6395
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.7916962185
actual_array: [ 18.1 -119. ]
varray: [-37.485, -37.544, -37.468, -37.489]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.4965
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [96300.0, 96837.0, 97373.0, 97909.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2124
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.17065
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08575
varray: [2.13, 2.076, 2.071, 2.066]
inst: mean_all
operation: mean_all(residuals)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: residuals
SAMP_residuals_STD | 0.0297811461611
varray: [2.13, 2.076, 2.071, 2.066]
inst: std_all
operation: std_all(residuals)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.86875
varray: [-0.967, -0.751, -0.892, -0.865]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0895781781462
varray: [-0.967, -0.751, -0.892, -0.865]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.33375
varray: [76.897, 76.061, 76.226, 76.151]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.381510484784
varray: [76.897, 76.061, 76.226, 76.151]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.32475
varray: [29.487, 27.948, 27.881, 27.983]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.7759881335
varray: [29.487, 27.948, 27.881, 27.983]
inst: std_all
operation: std_all(slope_shift)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: slope_shift
SAMP_d18O_STD | 0.113904565317
varray: [-8.754, -8.958, -8.987, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.54916418325
varray: [-62.574, -65.065, -65.726, -65.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.142
varray: [63.202, 63.401, 62.674, 63.291]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.322441726001
varray: [63.202, 63.401, 62.674, 63.291]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.86875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.33375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.32475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0896
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0896
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.11887742202
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7027
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.754, -8.958, -8.987, -8.994]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-62.574, -65.065, -65.726, -65.949]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.92325
varray: [-8.754, -8.958, -8.987, -8.994]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
dD_mean | -64.8285
varray: [-62.574, -65.065, -65.726, -65.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
d18O_stdev | 0.113904565317
varray: [-8.754, -8.958, -8.987, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
dD_stdev | 1.54916418325
varray: [-62.574, -65.065, -65.726, -65.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.92325
varray: [-8.754, -8.958, -8.987, -8.994]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.8285
varray: [-62.574, -65.065, -65.726, -65.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.113904565317
varray: [-8.754, -8.958, -8.987, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.54916418325
varray: [-62.574, -65.065, -65.726, -65.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1407696527
actual_array: [ 1.93 -15.84]
varray: [-8.754, -8.958, -8.987, -8.994]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.92325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.5530331851
actual_array: [ 18.1 -119. ]
varray: [-62.574, -65.065, -65.726, -65.949]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -64.8285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [98494.0, 99030.0, 99565.0, 100101.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2484
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0004
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.05725
varray: [2.064, 2.077, 2.024, 2.064]
inst: mean_all
operation: mean_all(residuals)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: residuals
SAMP_residuals_STD | 0.0229981883344
varray: [2.064, 2.077, 2.024, 2.064]
inst: std_all
operation: std_all(residuals)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9325
varray: [-1.01, -0.815, -0.93, -0.975]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0849019041797
varray: [-1.01, -0.815, -0.93, -0.975]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.582
varray: [75.932, 75.543, 75.182, 75.671]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.311951919372
varray: [75.932, 75.543, 75.182, 75.671]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.0535
varray: [27.648, 27.041, 26.308, 27.217]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.558596157046
varray: [27.648, 27.041, 26.308, 27.217]
inst: std_all
operation: std_all(slope_shift)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0933255949173
varray: [-5.716, -5.597, -5.517, -5.52]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.60737975185
varray: [-40.798, -38.259, -37.518, -37.295]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.671
varray: [61.866, 62.001, 61.305, 61.512]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.319458917547
varray: [61.866, 62.001, 61.305, 61.512]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.05725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.582
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.0535
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1181
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1181
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.42901142345
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.1737
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.716, -5.597, -5.517, -5.52]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-40.798, -38.259, -37.518, -37.295]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.5875
varray: [-5.716, -5.597, -5.517, -5.52]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
dD_mean | -38.4675
varray: [-40.798, -38.259, -37.518, -37.295]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0933255949173
varray: [-5.716, -5.597, -5.517, -5.52]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
dD_stdev | 1.60737975185
varray: [-40.798, -38.259, -37.518, -37.295]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.5875
varray: [-5.716, -5.597, -5.517, -5.52]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.4675
varray: [-40.798, -38.259, -37.518, -37.295]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0933255949173
varray: [-5.716, -5.597, -5.517, -5.52]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.60737975185
varray: [-40.798, -38.259, -37.518, -37.295]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.69979595509
actual_array: [ 1.93 -15.84]
varray: [-5.716, -5.597, -5.517, -5.52]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.5875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.8490011429
actual_array: [ 18.1 -119. ]
varray: [-40.798, -38.259, -37.518, -37.295]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.4675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [100686.0, 101222.0, 101757.0, 102292.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2769
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.75215
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.04025
varray: [2.082, 2.04, 1.999, 2.04]
inst: mean_all
operation: mean_all(residuals)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: residuals
SAMP_residuals_STD | 0.0338858377497
varray: [2.082, 2.04, 1.999, 2.04]
inst: std_all
operation: std_all(residuals)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.94025
varray: [-0.817, -0.951, -1.181, -0.812]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.172930380982
varray: [-0.817, -0.951, -1.181, -0.812]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.247
varray: [75.47, 75.214, 75.179, 75.125]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.153107805157
varray: [75.47, 75.214, 75.179, 75.125]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.8655
varray: [27.21, 26.836, 26.526, 26.89]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.280125567083
varray: [27.21, 26.836, 26.526, 26.89]
inst: std_all
operation: std_all(slope_shift)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: slope_shift
SAMP_d18O_STD | 0.01607015868
varray: [-5.593, -5.602, -5.565, -5.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0909028785756
varray: [-37.081, -37.164, -37.042, -36.945]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.554
varray: [63.027, 61.617, 59.76, 61.812]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.34891289563
varray: [63.027, 61.617, 59.76, 61.812]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.04025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.94025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.247
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.8655
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1351
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1351
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.2104948629
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.2907
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.593, -5.602, -5.565, -5.593]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-37.081, -37.164, -37.042, -36.945]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.58825
varray: [-5.593, -5.602, -5.565, -5.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
dD_mean | -37.058
varray: [-37.081, -37.164, -37.042, -36.945]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
d18O_stdev | 0.01607015868
varray: [-5.593, -5.602, -5.565, -5.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
dD_stdev | 0.0909028785756
varray: [-37.081, -37.164, -37.042, -36.945]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.58825
varray: [-5.593, -5.602, -5.565, -5.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.058
varray: [-37.081, -37.164, -37.042, -36.945]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.01607015868
varray: [-5.593, -5.602, -5.565, -5.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0909028785756
varray: [-37.081, -37.164, -37.042, -36.945]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.70056961327
actual_array: [ 1.93 -15.84]
varray: [-5.593, -5.602, -5.565, -5.593]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.58825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.3142212346
actual_array: [ 18.1 -119. ]
varray: [-37.081, -37.164, -37.042, -36.945]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.058
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [102878.0, 103414.0, 103949.0, 104484.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2939
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.08715
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0885
varray: [2.108, 2.082, 2.063, 2.101]
inst: mean_all
operation: mean_all(residuals)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: residuals
SAMP_residuals_STD | 0.0202402239777
varray: [2.108, 2.082, 2.063, 2.101]
inst: std_all
operation: std_all(residuals)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.88525
varray: [-0.802, -0.984, -0.95, -0.805]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0954196869973
varray: [-0.802, -0.984, -0.95, -0.805]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.88425
varray: [76.113, 75.949, 75.624, 75.851]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.204413592177
varray: [76.113, 75.949, 75.624, 75.851]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.80075
varray: [28.285, 27.618, 27.431, 27.869]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.369354730126
varray: [28.285, 27.618, 27.431, 27.869]
inst: std_all
operation: std_all(slope_shift)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0201225412577
varray: [-5.589, -5.609, -5.587, -5.63]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.123236899236
varray: [-37.354, -37.486, -37.305, -37.575]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.35375
varray: [63.211, 62.482, 61.598, 62.124]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.67706935883
varray: [63.211, 62.482, 61.598, 62.124]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0885
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.88525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.88425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.80075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.99246098329
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.49095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.589, -5.609, -5.587, -5.63]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-37.354, -37.486, -37.305, -37.575]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.60375
varray: [-5.589, -5.609, -5.587, -5.63]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
dD_mean | -37.43
varray: [-37.354, -37.486, -37.305, -37.575]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0201225412577
varray: [-5.589, -5.609, -5.587, -5.63]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
dD_stdev | 0.123236899236
varray: [-37.354, -37.486, -37.305, -37.575]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.60375
varray: [-5.589, -5.609, -5.587, -5.63]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.43
varray: [-37.354, -37.486, -37.305, -37.575]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0201225412577
varray: [-5.589, -5.609, -5.587, -5.63]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.123236899236
varray: [-37.354, -37.486, -37.305, -37.575]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.71655854899
actual_array: [ 1.93 -15.84]
varray: [-5.589, -5.609, -5.587, -5.63]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.60375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.719285531
actual_array: [ 18.1 -119. ]
varray: [-37.354, -37.486, -37.305, -37.575]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.43
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [105068.0, 105603.0, 106139.0, 106674.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.24565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4499
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.103
varray: [2.078, 2.127, 2.098, 2.109]
inst: mean_all
operation: mean_all(residuals)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: residuals
SAMP_residuals_STD | 0.0205101600839
varray: [2.078, 2.127, 2.098, 2.109]
inst: std_all
operation: std_all(residuals)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9315
varray: [-1.048, -0.971, -0.792, -0.915]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107803834193
varray: [-1.048, -0.971, -0.792, -0.915]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.9795
varray: [75.868, 76.374, 75.774, 75.902]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.268513810942
varray: [75.868, 76.374, 75.774, 75.902]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.92425
varray: [27.638, 28.443, 27.762, 27.854]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.356978407377
varray: [27.638, 28.443, 27.762, 27.854]
inst: std_all
operation: std_all(slope_shift)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0201163283595
varray: [-5.61, -5.574, -5.563, -5.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0971150520431
varray: [-37.437, -37.236, -37.375, -37.252]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.6555
varray: [61.256, 63.418, 62.765, 63.183]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.97129689934
varray: [61.256, 63.418, 62.765, 63.183]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.103
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9315
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.9795
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.92425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07235
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07235
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.32590157906
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.1892
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.61, -5.574, -5.563, -5.585]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-37.437, -37.236, -37.375, -37.252]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.583
varray: [-5.61, -5.574, -5.563, -5.585]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
dD_mean | -37.325
varray: [-37.437, -37.236, -37.375, -37.252]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0201163283595
varray: [-5.61, -5.574, -5.563, -5.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
dD_stdev | 0.0971150520431
varray: [-37.437, -37.236, -37.375, -37.252]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.583
varray: [-5.61, -5.574, -5.563, -5.585]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.325
varray: [-37.437, -37.236, -37.375, -37.252]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0201163283595
varray: [-5.61, -5.574, -5.563, -5.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0971150520431
varray: [-37.437, -37.236, -37.375, -37.252]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.69515400601
actual_array: [ 1.93 -15.84]
varray: [-5.61, -5.574, -5.563, -5.585]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.583
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.6049528667
actual_array: [ 18.1 -119. ]
varray: [-37.437, -37.236, -37.375, -37.252]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [107257.0, 107792.0, 108328.0, 108863.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.23115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.35465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.03475
varray: [2.018, 2.042, 2.034, 2.045]
inst: mean_all
operation: mean_all(residuals)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: residuals
SAMP_residuals_STD | 0.0120933866224
varray: [2.018, 2.042, 2.034, 2.045]
inst: std_all
operation: std_all(residuals)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.817
varray: [-0.799, -0.793, -0.901, -0.775]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.056920997883
varray: [-0.799, -0.793, -0.901, -0.775]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.15875
varray: [75.057, 75.198, 75.156, 75.224]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0733910757518
varray: [75.057, 75.198, 75.156, 75.224]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.7065
varray: [26.447, 26.75, 26.839, 26.79]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.176787065892
varray: [26.447, 26.75, 26.839, 26.79]
inst: std_all
operation: std_all(slope_shift)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0337675880098
varray: [-5.766, -5.718, -5.742, -5.687]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.176053968998
varray: [-38.251, -37.98, -38.059, -37.828]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.6575
varray: [61.172, 62.42, 61.668, 61.37]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.547686954382
varray: [61.172, 62.42, 61.668, 61.37]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.03475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.817
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.15875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.7065
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1406
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1406
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.46332774036
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.1872
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.766, -5.718, -5.742, -5.687]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-38.251, -37.98, -38.059, -37.828]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.72825
varray: [-5.766, -5.718, -5.742, -5.687]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
dD_mean | -38.0295
varray: [-38.251, -37.98, -38.059, -37.828]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0337675880098
varray: [-5.766, -5.718, -5.742, -5.687]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
dD_stdev | 0.176053968998
varray: [-38.251, -37.98, -38.059, -37.828]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.72825
varray: [-5.766, -5.718, -5.742, -5.687]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.0295
varray: [-38.251, -37.98, -38.059, -37.828]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0337675880098
varray: [-5.766, -5.718, -5.742, -5.687]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.176053968998
varray: [-38.251, -37.98, -38.059, -37.828]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.84498580683
actual_array: [ 1.93 -15.84]
varray: [-5.766, -5.718, -5.742, -5.687]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.72825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.3720706003
actual_array: [ 18.1 -119. ]
varray: [-38.251, -37.98, -38.059, -37.828]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.0295
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [109449.0, 109984.0, 110520.0, 111056.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2994
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.1754
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09925
varray: [2.097, 2.055, 2.118, 2.127]
inst: mean_all
operation: mean_all(residuals)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: residuals
SAMP_residuals_STD | 0.0320663374896
varray: [2.097, 2.055, 2.118, 2.127]
inst: std_all
operation: std_all(residuals)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04775
varray: [-1.081, -1.004, -1.092, -1.014]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0451543648684
varray: [-1.081, -1.004, -1.092, -1.014]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.0065
varray: [75.977, 75.411, 76.282, 76.356]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.429550540294
varray: [75.977, 75.411, 76.282, 76.356]
inst: std_all
operation: std_all(baseline_shift)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.7955
varray: [27.734, 26.933, 28.03, 28.485]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.652709481061
varray: [27.734, 26.933, 28.03, 28.485]
inst: std_all
operation: std_all(slope_shift)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0304685630336
varray: [-5.591, -5.523, -5.552, -5.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0964261375354
varray: [-36.925, -36.798, -37.03, -36.883]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.43475
varray: [62.539, 60.597, 63.232, 63.371]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.27806686706
varray: [62.539, 60.597, 63.232, 63.371]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.0065
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.7955
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0761
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0761
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.49828763188
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.40995
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.591, -5.523, -5.552, -5.58]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-36.925, -36.798, -37.03, -36.883]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.5615
varray: [-5.591, -5.523, -5.552, -5.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
dD_mean | -36.909
varray: [-36.925, -36.798, -37.03, -36.883]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0304685630336
varray: [-5.591, -5.523, -5.552, -5.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
dD_stdev | 0.0964261375354
varray: [-36.925, -36.798, -37.03, -36.883]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.5615
varray: [-5.591, -5.523, -5.552, -5.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.909
varray: [-36.925, -36.798, -37.03, -36.883]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0304685630336
varray: [-5.591, -5.523, -5.552, -5.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0964261375354
varray: [-36.925, -36.798, -37.03, -36.883]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.67297580486
actual_array: [ 1.93 -15.84]
varray: [-5.591, -5.523, -5.552, -5.58]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.5615
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.1519777395
actual_array: [ 18.1 -119. ]
varray: [-36.925, -36.798, -37.03, -36.883]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [111640.0, 112176.0, 112711.0, 113246.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.2349
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.32765
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.83425
varray: [1.754, 1.786, 1.967, 1.83]
inst: mean_all
operation: mean_all(residuals)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: residuals
SAMP_residuals_STD | 0.0938238597941
varray: [1.754, 1.786, 1.967, 1.83]
inst: std_all
operation: std_all(residuals)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.884
varray: [-0.78, -0.789, -1.017, -0.95]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.118160907241
varray: [-0.78, -0.789, -1.017, -0.95]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 72.853
varray: [71.88, 72.242, 74.521, 72.769]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.17037458391
varray: [71.88, 72.242, 74.521, 72.769]
inst: std_all
operation: std_all(baseline_shift)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.3155
varray: [21.546, 22.511, 25.741, 23.464]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.7966128316
varray: [21.546, 22.511, 25.741, 23.464]
inst: std_all
operation: std_all(slope_shift)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0415050197767
varray: [-5.73, -5.786, -5.818, -5.818]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.266201051839
varray: [-38.08, -38.497, -38.722, -38.455]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 56.425
varray: [55.217, 55.285, 59.712, 55.486]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.19430733186
varray: [55.217, 55.285, 59.712, 55.486]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.83425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.884
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 72.853
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.3155
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3411
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3411
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -15.6802353644
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.4197
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.73, -5.786, -5.818, -5.818]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-38.08, -38.497, -38.722, -38.455]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.788
varray: [-5.73, -5.786, -5.818, -5.818]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
dD_mean | -38.4385
varray: [-38.08, -38.497, -38.722, -38.455]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0415050197767
varray: [-5.73, -5.786, -5.818, -5.818]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
dD_stdev | 0.266201051839
varray: [-38.08, -38.497, -38.722, -38.455]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.788
varray: [-5.73, -5.786, -5.818, -5.818]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.4385
varray: [-38.08, -38.497, -38.722, -38.455]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0415050197767
varray: [-5.73, -5.786, -5.818, -5.818]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.266201051839
varray: [-38.08, -38.497, -38.722, -38.455]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.90662057516
actual_array: [ 1.93 -15.84]
varray: [-5.73, -5.786, -5.818, -5.818]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.788
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.8174235499
actual_array: [ 18.1 -119. ]
varray: [-38.08, -38.497, -38.722, -38.455]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.4385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [113832.0, 114368.0, 114904.0, 115440.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.4999
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.48115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.057
varray: [1.998, 2.077, 2.073, 2.08]
inst: mean_all
operation: mean_all(residuals)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: residuals
SAMP_residuals_STD | 0.0394377146059
varray: [1.998, 2.077, 2.073, 2.08]
inst: std_all
operation: std_all(residuals)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04175
varray: [-1.123, -0.945, -1.009, -1.09]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0803383470081
varray: [-1.123, -0.945, -1.009, -1.09]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 75.38475
varray: [74.79, 75.544, 75.496, 75.709]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.406859005062
varray: [74.79, 75.544, 75.496, 75.709]
inst: std_all
operation: std_all(baseline_shift)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.18825
varray: [26.453, 27.337, 27.301, 27.662]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.516356062551
varray: [26.453, 27.337, 27.301, 27.662]
inst: std_all
operation: std_all(slope_shift)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0209980157793
varray: [-5.586, -5.593, -5.582, -5.628]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0781339661522
varray: [-37.167, -37.053, -37.06, -36.977]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.0085
varray: [59.317, 62.176, 61.355, 61.186]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.20772472029
varray: [59.317, 62.176, 61.355, 61.186]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.057
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 75.38475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.18825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11835
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11835
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.44050382697
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.8362
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-5.586, -5.593, -5.582, -5.628]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-37.167, -37.053, -37.06, -36.977]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.59725
varray: [-5.586, -5.593, -5.582, -5.628]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
dD_mean | -37.06425
varray: [-37.167, -37.053, -37.06, -36.977]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0209980157793
varray: [-5.586, -5.593, -5.582, -5.628]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
dD_stdev | 0.0781339661522
varray: [-37.167, -37.053, -37.06, -36.977]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.59725
varray: [-5.586, -5.593, -5.582, -5.628]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.06425
varray: [-37.167, -37.053, -37.06, -36.977]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0209980157793
varray: [-5.586, -5.593, -5.582, -5.628]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0781339661522
varray: [-37.167, -37.053, -37.06, -36.977]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.70985351143
actual_array: [ 1.93 -15.84]
varray: [-5.586, -5.593, -5.582, -5.628]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.59725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.3210267503
actual_array: [ 18.1 -119. ]
varray: [-37.167, -37.053, -37.06, -36.977]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.06425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [116026.0, 116562.0, 117097.0, 117633.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.27715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.9494
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.072
varray: [2.088, 2.07, 2.061, 2.069]
inst: mean_all
operation: mean_all(residuals)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: residuals
SAMP_residuals_STD | 0.011401754251
varray: [2.088, 2.07, 2.061, 2.069]
inst: std_all
operation: std_all(residuals)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.978
varray: [-1.067, -0.903, -1.06, -0.882]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0991396321693
varray: [-1.067, -0.903, -1.06, -0.882]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 76.09175
varray: [76.395, 75.975, 76.016, 75.981]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.20297352044
varray: [76.395, 75.975, 76.016, 75.981]
inst: std_all
operation: std_all(baseline_shift)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.37525
varray: [28.783, 28.593, 28.066, 28.059]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.369379998556
varray: [28.783, 28.593, 28.066, 28.059]
inst: std_all
operation: std_all(slope_shift)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: slope_shift
SAMP_d18O_STD | 0.109422727682
varray: [-8.752, -8.904, -8.98, -8.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.69431264333
varray: [-62.305, -65.092, -65.738, -65.993]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.82975
varray: [61.945, 61.342, 61.677, 62.355]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.428331160825
varray: [61.945, 61.342, 61.677, 62.355]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.072
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.978
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 76.09175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.37525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10335
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10335
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.75095961569
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.01495
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4482
varray: [-8.752, -8.904, -8.98, -8.988]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.4025
varray: [-62.305, -65.092, -65.738, -65.993]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.906
varray: [-8.752, -8.904, -8.98, -8.988]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
dD_mean | -64.782
varray: [-62.305, -65.092, -65.738, -65.993]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
d18O_stdev | 0.109422727682
varray: [-8.752, -8.904, -8.98, -8.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
dD_stdev | 1.69431264333
varray: [-62.305, -65.092, -65.738, -65.993]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.906
varray: [-8.752, -8.904, -8.98, -8.988]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.782
varray: [-62.305, -65.092, -65.738, -65.993]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.109422727682
varray: [-8.752, -8.904, -8.98, -8.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.69431264333
varray: [-62.305, -65.092, -65.738, -65.993]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1229755146
actual_array: [ 1.93 -15.84]
varray: [-8.752, -8.904, -8.98, -8.988]
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.906
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.502400148
actual_array: [ 18.1 -119. ]
varray: [-62.305, -65.092, -65.738, -65.993]
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -64.782
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: [118219.0, 118754.0, 119289.0, 119824.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 75.26215
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.2424
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.85
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.296
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0967532769
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.85
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9509109364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -75.296
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.306
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.206
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5355932105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.306
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.4084989226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.206
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.29
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.612
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5190885027
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.29
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.7617027073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -69.612
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.123
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.392
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3468206146
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.123
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.433266036
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.392
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3406313492
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0499793899
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.923
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.88778342795
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8336976179
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.923
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.989
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.386
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2085936865
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.989
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3378502235
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.386
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.065
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2869910487
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.065
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.528404664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.023
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.427
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2436661907
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3824944067
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.427
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.037
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.285
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.25810781
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.037
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2278730892
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.133
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.16933927763
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 6.8045860936
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 7.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.657
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.348
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.8047705293
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.657
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.7497559744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 15.348
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.776
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.392
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.92752429382
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.776
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.97543184
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 17.392
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.862
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.28
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.01623709844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.9423595155
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 18.28
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04512033715
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5401560176
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 18.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.882
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.176
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03686798324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.882
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.9179982511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 19.176
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.458
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04408879291
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.225063121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 19.458
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.936
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.09257137218
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.936
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4591728623
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 19.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.917
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.889
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07297203163
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.917
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6943714861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 19.889
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.934
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.886
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7784
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.5064
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.934
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.886
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0905082837
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 2.934
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6911048385
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: 19.886
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.398
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.7566722394
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -13.398
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.414685936
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -93.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.285
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.352
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6716519801
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.678363881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -105.352
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.423
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.224
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8140050852
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.423
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.805634471
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -108.224
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.547
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.564
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9419165709
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.547
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.264737044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -109.564
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.574
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9697682654
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.574
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.173953946
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -110.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0275347428
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.691173142
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -110.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.631
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.028566287
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.631
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.127815031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -111.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.676
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0749857778
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.676
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.439235431
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -111.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.651
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.689
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0491971718
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.651
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.578612394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -111.689
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.667
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.916
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0657018797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -14.667
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.825788725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -111.916
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.407
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6397791787
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.407
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.5368583952
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -73.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.138
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.115
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3622937782
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.138
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2205280961
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -69.115
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.099
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.106
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3220635529
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.099
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.121845636
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.106
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.684
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2931803142
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6623372137
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.684
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.028
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.542
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2488239119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.028
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5077158962
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.542
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.438
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2405715579
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3944721144
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.438
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.381
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2354138367
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3324058109
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.381
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.106
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2405715579
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0329631186
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.106
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.99
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.191
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2096252307
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.99
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1255181326
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.191
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.022
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.17
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2426346464
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.022
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1026515997
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.17
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.24214528694
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 3.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.0541069773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -34.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.882
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.06841222296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 3.882
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.6179996013
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -29.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.975
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.941
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.16434583725
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 3.975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.4757618405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -28.941
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.935
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.913
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.12308406766
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: 3.935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.44527313
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -28.913
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.607
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.61925336398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -4.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.0009002541
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.607
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.917
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.958
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.00814554236
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -4.917
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.3830980177
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.958
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.982
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.118
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.07519591794
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -4.982
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.5573192205
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -39.118
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.951
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.139
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.04321804651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -4.951
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.5801857534
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -39.139
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.504
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.068
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.8029274494
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -11.504
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.9470009666
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -89.068
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.863
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.623
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.1732518315
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -11.863
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.995743351
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -94.623
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.948
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.648
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.2609330918
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -11.948
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -105.111847931
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -95.648
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.968
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.2815639766
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -11.968
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -105.912176582
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -96.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.414
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.239
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5208230295
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.414
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.4001341446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -47.239
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.099
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.083
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.19588659399
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.099
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.7858558847
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -42.083
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.014
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.515
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.10820533361
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.014
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.0784880973
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -40.515
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.967
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.05972275435
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -4.967
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.5100914232
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -39.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.634
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.9054839609
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -10.634
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.7798475723
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -85.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.948
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.009
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2293888521
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -10.948
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.9716394155
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -90.009
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.997
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2799345199
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -10.997
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.382831158
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -91.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.022
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.689
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3057231259
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -11.022
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.800962045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -91.689
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8099839783
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6413911169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -73.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.463
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.211
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6975456561
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.463
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.5028258527
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -71.211
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.443
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.971
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6769147713
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.443
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2414940485
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.971
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6903248465
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.8767184051
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.512
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9667787027
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.1861095602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -75.512
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9667787027
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5334630832
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -75.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.707
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.004
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9492424506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.707
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.7218397587
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -76.004
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.756
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.975
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9997881184
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.756
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.6902621657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -75.975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.688
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.215
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.70883343202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.688
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.2851184467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -46.215
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.377
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.38802317346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.377
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.7516946777
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -42.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.296
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.212
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.30446809005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.296
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.9263217294
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -42.212
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.278
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.003
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.28590029373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.278
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.6987452833
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -42.003
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.968
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.93523446298
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.6160497788
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -63.968
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.933
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.011
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1508272091
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.933
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.840636762
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.011
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1611426515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4318999689
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.988
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2075621423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.988
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8358753829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.134
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.967
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3581676013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.9481986579
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -78.967
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.148
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.478
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3726092206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.148
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.5935001418
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -80.478
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.101
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.804
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3241266414
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.101
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.9484758425
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -80.804
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.153
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.002
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3777669418
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.153
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.1640745809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -81.002
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.957
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.556
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1755842708
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.957
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.434077734
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.556
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.948
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.871
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1663003727
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.948
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5993106921
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -64.871
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.969
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.49
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1879628017
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.969
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.184446453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -64.49
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.003
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.409
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2230353059
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.003
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.096246969
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -64.409
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.358
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.189
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.36842383291
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -2.358
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3233977106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -58.189
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.98
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.97850011029
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -1.98
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.286781554
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -57.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.115
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.95064841582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -1.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.1539378868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -57.115
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.88
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.972
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.87534568632
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -1.88
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.9982276868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -56.972
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.145
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.652
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.27488186874
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.145
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7163754905
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.652
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.393
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.53070484019
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.393
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.2005223618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.425
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.474
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56371425586
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.7003194373
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.428
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.519
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56680888858
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.428
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.7493191506
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.519
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.67219068185
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4027773255
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.569
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.213
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.77534510582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.569
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2383560654
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.213
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.581
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.051
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7877236367
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.581
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0619570976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.051
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.79184981366
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3058667815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -68.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.599
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.77784182601
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -7.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5697821997
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.599
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.549
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.297
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.72316998131
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -7.549
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2409396794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.297
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.52
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.238
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.69325519836
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -7.52
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1766956109
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.238
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.534
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.102
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.70769681771
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -7.534
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0286075885
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.102
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.234
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.514
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.42977778552
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.234
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.743874738
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.514
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.22
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.867
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.41533616616
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.22
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.1282502667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.867
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.274
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.938
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.47103955511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.274
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2055609254
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.938
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.262
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.995
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.45866102423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.262
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2676272289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -70.995
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.973
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.348
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.12900049923
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.973
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2964726878
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.348
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.985
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.0062467347
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.901208334
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.985
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.025
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.0165621771
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9447636347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.829
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.89
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.98045812871
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -6.829
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7977644948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.89
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.062
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1136916165
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9850522878
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.062
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.966
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.928
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.184868169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8391420305
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.928
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.031
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.065
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2519185446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.031
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9883189354
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.983
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2024044211
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.983
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.932785927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -67.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -15.233
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.6495559193
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -15.233
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.092287757
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -110.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -15.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -114.864
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.9868708857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -15.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -126.035814386
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -114.864
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -15.654
operation: col(d(18_16)Mean)
inst: col
dD_actual | -116.219
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.0838360443
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -15.654
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -127.511250198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -116.219
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -15.666
operation: col(d(18_16)Mean)
inst: col
dD_actual | -116.484
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.0962145751
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -15.666
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -127.799804065
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -116.484
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.898
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.216
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.02009044153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.898
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.2862073293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -46.216
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.075
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.987
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.17112953224
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.3257930933
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.292
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.11439459905
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.4801372262
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.292
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.123
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.806
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.22064365574
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.123
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.9509403227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.806
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.654
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.32895580091
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.7854301801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.654
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.262
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.444
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.36402830506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.262
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.5567648514
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.444
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.267
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.08
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.36918602626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.267
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.1604116151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.08
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.271
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.033
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.37331220322
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.271
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.1092341367
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.033
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.134
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.252
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.23199064238
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.2588168906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -34.252
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.526
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.13605702808
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.468288183
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -33.526
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.726
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.20826512487
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.6860646865
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -33.726
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.952
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.20517049215
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.9321521354
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -33.952
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.349
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.45377265392
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.349
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.6101200948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.365
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.942
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.47027736175
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.365
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.0990283451
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.942
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.307
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.62
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.41044779585
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.307
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.7484081745
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.62
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.35
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.45480419816
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.35
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.8474964836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.406
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.51257067558
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.406
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.1817834164
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.399
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.50534986591
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.399
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.0010289185
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.433
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.54042237006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.433
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.0728951647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.419
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.949
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5259807507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.419
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.1066505227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.949
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.48
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.283
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58890494932
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.48
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.4703372835
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.283
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.389
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.57446332997
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.5857588304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.389
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.495
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.527
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.60437811292
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.495
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.7360246178
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.527
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.473
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.407
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58168413964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.473
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6053587157
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.407
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.269
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.428
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.37124911474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.269
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.5393427311
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.428
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.286
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.193
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.38878536682
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.286
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.2834553395
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.193
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.254
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.245
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.35577595115
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.254
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.3400772304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.232
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.254
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.33308197787
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.232
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.3498771731
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.254
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.684
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.256
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.89397269339
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.684
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.751882909
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -62.256
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.92
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.068
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.137417134
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.92
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.813820548
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.068
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.982
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.851
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2013728768
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.982
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6664155592
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.851
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.018
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.192
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2385084695
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -9.018
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0377244976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -66.192
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.679
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.386
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.79418225303
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.679
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.9380222526
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -40.386
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.504
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.577
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.61366201108
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.504
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.8793512611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.577
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.467
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.57549487421
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.467
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.0648671381
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.433
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.49
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.54042237006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.433
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6957359647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.49
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.59
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56930560877
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.8046242164
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.59
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.432
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.61985127651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6325807786
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.468
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.413
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.57652641845
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.468
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6118920108
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.413
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.483
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58374722812
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.688113787
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.483
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.142
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56930560877
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.3168048486
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.142
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.535
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.476
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.64563988251
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.535
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6804916094
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.476
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.49
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.409
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.59922039172
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.49
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6075364807
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.409
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.472
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5806525954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.472
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.5378479996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.46
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56827406453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.46
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.4736039311
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.486
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.39
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.59509421476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.486
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.5868477129
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.39
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.418
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.52494920646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.418
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.2590940752
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.281
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56414788757
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.4681595185
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.281
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.523
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.343
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.63326135163
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.523
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.5356702346
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.343
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.192
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62088282075
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.3712489745
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.192
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.503
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.251
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.61263046684
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.503
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.435493043
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.251
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.559
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.446
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.67039694426
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6478251339
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.446
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.345
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.681
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.44964647696
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.345
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.814830008
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.681
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.339
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.44345721152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.339
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.8921406668
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.349
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.781
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.45377265392
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.349
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.9237182598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.781
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.49606596775
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.8474964836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -35.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.847
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.418
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.9674816853
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.847
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.7951014583
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.418
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.763
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.96645014106
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.1707659268
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.763
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.676
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.99636492401
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.0760331478
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.676
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.855
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.876
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.97573403922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.855
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.2938096513
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.876
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.485
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.7570466604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.7791740695
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.636
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.544
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.74982585072
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.843418138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.544
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.468
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.73228959864
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.7606630667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.468
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.66
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.489
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.77458291247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.66
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.7835295996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.489
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.754
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.574
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.96618079017
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.754
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.0981475495
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -62.574
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.958
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.065
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1766158151
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.958
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8105539005
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.987
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.726
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.206530598
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.987
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5303052445
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.726
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.949
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2137514077
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.7731260459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.949
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.716
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.798
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.8323493899
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.716
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.3866418498
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -40.798
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.259
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.70959562537
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.621969138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.259
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.517
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.518
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62707208619
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.517
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.8151071926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.518
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.52
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.295
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.63016671891
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.52
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.5722863912
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.295
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.70546944841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.3392655325
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.164
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.71475334657
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.4296427814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.164
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.042
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6765862097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.2967991143
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.042
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.945
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.70546944841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.1911775101
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.945
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.589
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.354
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.70134327145
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.589
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.6365304597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.354
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.609
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.486
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.72197415625
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.780262952
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.486
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.69928018297
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.5831752164
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.575
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.74363658528
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.8771734961
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.437
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.72300570049
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.7269077087
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.437
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.574
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.236
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.68587010786
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.574
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.5080423227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.236
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.563
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.375
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.67452312122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.563
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.6593969926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.252
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.69721709449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.5254644429
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.252
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.766
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.251
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.88392660188
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.766
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.6132580779
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.251
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.718
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.98
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.83441247838
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.718
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.3181709156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.98
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.059
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.85916954013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.4041926345
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.059
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.687
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.828
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.80243460695
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.687
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.152660773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.828
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.70340635993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.1693998597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.523
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.798
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.63326135163
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.523
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.03111178
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.798
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.552
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.03
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.66317613458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.552
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.2837325241
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.03
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.883
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.69205937329
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.123666794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.883
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.08
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.84679100925
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.4270591674
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.08
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.786
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.497
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.90455748668
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.786
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.8811231772
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.497
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.818
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.722
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.93756690235
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.818
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.1261217436
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.722
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.818
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.455
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.93756690235
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.818
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.8353901114
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -38.455
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.586
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.69824863873
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.586
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.432909429
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.70546944841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.308776822
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.582
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.06
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.69412246177
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.582
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.3163989996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -37.06
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.628
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.977
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.7415734968
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -5.628
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.2260217506
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -36.977
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:20:40
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.752
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.96411770169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.752
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.8052381523
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -62.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.904
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.092
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1209124261
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.904
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8399537284
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.092
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.98
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.738
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1993097884
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.98
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5433718347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.738
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7784
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.5064
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.667
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.916
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4482
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.4025
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.988
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2075621423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.907418683174
measure_array: [ 2.7784 -14.4482]
inst: fit_type1
actual_mean: -8.988
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969420371412
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8210368767
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.883853756382
measure_array: [ 17.5064 -108.4025]
inst: fit_type1
actual_mean: -65.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918372720642
tarray: []
source_col: d(D_H)Mean