## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
220216_21-067_test_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4226 | -75.5619 |  |  |  | 0.20492 | 2.7313
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.22273 | 3.6171
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.33519 | 5.5672
4 |  | EV | -10.3456 | -74.5727 |  |  |  | 0.094208 | 1.8703
5 |  | CPER3-deep-2 | -8.8831 | -94.5392 |  |  |  | 0.052277 | 1.0268
6 |  | CPER3-E-2 | 2.3132 | -57.9838 |  |  |  | 0.26764 | 2.2353
7 |  | CPER3-D-5 | 5.9794 | -57.7918 | -5.2853 | 24.139 | 20.9866 | 0.055194 | 0.67968 | -15.3217 | 11.5307
8 | * | CPER3-010-1 | -0.58027 | -46.8631 |  |  |  | 2.5669 | 35.4253
9 |  | CPER3-010-3 | 2.2222 | -72.159 |  |  |  | 0.057731 | 0.21414
10 |  | CPER3-D-2 | 5.2541 | -61.813 | -2.9385 | 26.8835 |  | 0.01292 | 0.35898 | -8.8577
11 |  | CPER3-3040-1 | -5.5767 | -87.5651 |  |  |  | 0.19336 | 0.98884
12 |  | CPER3-B-2 | 1.3703 | -64.8383 |  |  |  | 0.186 | 1.3422
13 |  | EV | -9.9169 | -73.2309 |  |  |  | 0.29226 | 0.66285
14 |  | TEST | -15.3453 | -117.7163 |  |  |  | 0.087789 | 2.4446
15 |  | CPER3-B-3 | 1.0259 | -64.5153 |  | 62.9945 |  | 0.39857 | 3.2175
16 |  | CPER3-010-2 | -3.2188 | -83.4215 |  |  |  | 0.12299 | 1.2197
17 |  | CPER3-2030-2 | -2.3521 | -80.157 |  |  |  | 0.064211 | 0.35718
18 |  | CPER3-B-4 | -0.8115 | -67.1731 |  | 74.5783 |  | 0.029286 | 0.86164
19 |  | CPER3-E-3 | -1.0756 | -70.4967 |  |  |  | 0.057521 | 0.24255
20 | * | CPER3-E-4 | -0.61906 | -40.7157 |  |  |  | 2.5219 | 27.8115
21 |  | CPER3-2030-1 | -3.1142 | -80.5542 |  |  |  | 0.015966 | 0.60241
22 |  | EV | -10.0865 | -74.0584 |  |  |  | 0.20377 | 0.34073
23 |  | TEST | -15.4688 | -118.357 |  |  |  | 0.15998 | 2.5266
24 |  | CPER3-C-1 | -6.3338 | -84.5894 |  |  |  | 0.27901 | 2.2148
25 |  | CPER3-D-3 | 3.3771 | -65.4052 | -3.4138 |  | 7.351 | 0.18569 | 0.86374 | -9.4547 | 6.9597
26 |  | CPER3-E-5 | 4.4597 | -56.9387 |  | 50.9165 |  | 0.074536 | 0.88998
27 |  | CPER3-D-1 | 1.9974 | -70.3737 | -2.422 | 53.0773 |  | 0.1239 | 1.1614 | -8.304
28 |  | CPER3-A-1 | 1.197 | -67.2467 |  |  |  | 0.03509 | 0.62409
29 |  | CPER3-D-4 | 0.61619 | -67.5942 |  |  |  | 0.059867 | 0.29488
30 |  | CPER3-C-2 | -5.3522 | -84.3552 |  |  |  | 0.13126 | 0.85824
31 |  | EV | -10.1024 | -74.1385 |  |  |  | 0.098831 | 0.73882
32 |  | TEST | -15.481 | -118.3404 |  |  |  | 0.14644 | 2.4873
33 |  | CPER3-A-2 | -2.2138 | -77.0093 |  |  |  | 0.3314 | 2.2411
34 |  | CPER3-A-3 | -2.3571 | -77.0475 |  |  |  | 0.021142 | 0.22971
35 | * | CPER3-A-4 | -2.8558 | -60.4146 |  |  |  | 0.044249 | 30.1432
36 | * | CPER3-A-5 | -4.29 | -48.6887 |  |  |  | 1.7401 | 37.6865
37 |  | CPER3-B-1 | -2.8016 | -75.7102 |  |  |  | 0.062974 | 0.24964
38 |  | CPER3-B-5 | -0.31 | -70.5617 |  | 208.7035 |  | 0.10256 | 0.66311 | -6.5747
39 | * | CPER3-E-1 | -2.7912 | -18.2607 |  |  |  | 0 | 0
40 |  | EV | -10.1114 | -73.667 |  |  |  | 0.14521 | 0.24292
41 |  | TEST | -15.4624 | -118.1708 |  |  |  | 0.13786 | 2.5065
42 |  | CPER3-deep-3 | -8.9758 | -91.2125 |  |  |  | 0.10849 | 1.4514
43 |  | CPER3-3040-3 | -8.6611 | -90.7995 |  |  |  | 0.047244 | 0.12942
44 |  | CPER3-2030-3 | -6.7955 | -90.6156 |  |  |  | 0.044925 | 0.12316
45 |  | CPER3-C-5 | -6.2202 | -86.4342 |  |  |  | 0.025264 | 0.23104
46 |  | CPER3-C-4 | -6.2182 | -86.4181 |  | 83.1503 |  | 0.0080156 | 0.070477
47 |  | CPER3-1020-3 | -2.0331 | -80.7518 |  |  |  | 0.12898 | 0.40511
48 |  | CPER3-3040-2 | -8.0002 | -95.4302 |  |  |  | 0.14852 | 0.77958
49 |  | EV | -10.1822 | -74.7937 |  |  |  | 0.082156 | 1.0724
50 |  | TEST | -10.2422 | -73.9943 |  |  |  | 0.021266 | 0.17175
51 |  | CPER3-C-3 | -1.8595 | -72.0245 |  |  |  | 0.23098 | 0.22879
52 |  | CPER3-deep-1 | -9.1417 | -95.382 |  |  |  | 0.14687 | 1.2293
53 |  | CPER3-1020-2 | -2.0502 | -81.8072 |  |  |  | 0.19482 | 0.85696
54 |  | TEST | -15.3137 | -118.6397 |  |  |  | 0.29051 | 2.0172
55 |  | TEST | -15.5862 | -120.3519 |  |  |  | 0.025936 | 0.11017
56 |  | TEST | -15.6285 | -120.5628 |  |  |  | 0.07766 | 0.22991
57 |  | EV | -10.3463 | -76.1183 |  |  |  | 0.15208 | 2.5464

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -7.974 | -67.537 | -10.7218 | -81.2257 | 0.31 | 1330.593 | 2022/02/16 12:44:09
1 |  | EV | -8.064 | -65.143 | -10.8113 | -78.9937 | 0.246 | 1477.55 | 2022/02/16 12:53:06
1 |  | EV | -7.8 | -62.018 | -10.5487 | -76.0803 | 0.566 | 1337.615 | 2022/02/16 13:02:01
1 |  | EV | -7.615 | -60.512 | -10.3647 | -74.6762 | 0.516 | 1312.617 | 2022/02/16 13:10:58
1 |  | EV | -7.657 | -60.602 | -10.4065 | -74.7601 | 0.517 | 1319.453 | 2022/02/16 13:19:55
1 |  | EV | -7.568 | -60.179 | -10.318 | -74.3658 | 0.134 | 1311.544 | 2022/02/16 13:28:51
1 |  | EV | -7.537 | -59.83 | -10.2872 | -74.0404 | 0.273 | 1311.205 | 2022/02/16 13:37:47
1 |  | EV | -7.531 | -59.674 | -10.2812 | -73.895 | 0.516 | 1310.249 | 2022/02/16 13:46:44
1 |  | EV | -7.483 | -59.496 | -10.2335 | -73.729 | 0.373 | 1309.059 | 2022/02/16 13:55:41
1 |  | EV | -7.503 | -59.629 | -10.2534 | -73.853 | 0.174 | 1307.62 | 2022/02/16 14:04:37
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4226 | -75.5619 |  |  |  | 0.20492 | 2.7313
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 4.171 | 29.283 | 1.3571 | 9.04 | 0.51 | 1308.338 | 2022/02/16 14:14:25
2 |  | PZ | 4.625 | 37.152 | 1.8087 | 16.3763 | 0.393 | 1305.39 | 2022/02/16 14:23:22
2 |  | PZ | 4.765 | 39.253 | 1.9479 | 18.335 | 0.356 | 1307.494 | 2022/02/16 14:32:19
2 |  | PZ | 4.746 | 39.861 | 1.929 | 18.9019 | 0.122 | 1304.607 | 2022/02/16 14:41:16
2 |  | PZ | 4.846 | 40.325 | 2.0285 | 19.3345 | 0.537 | 1303.238 | 2022/02/16 14:50:12
2 |  | PZ | 4.842 | 40.524 | 2.0245 | 19.52 | 0.181 | 1304.289 | 2022/02/16 14:59:10
2 |  | PZ | 4.924 | 41.233 | 2.106 | 20.181 | 0.126 | 1306.379 | 2022/02/16 15:08:06
2 |  | PZ | 4.937 | 41.205 | 2.119 | 20.1549 | 0.314 | 1306.031 | 2022/02/16 15:17:03
2 |  | PZ | 4.862 | 40.859 | 2.0444 | 19.8323 | 0.344 | 1308.995 | 2022/02/16 15:26:00
2 |  | PZ | 4.752 | 40.314 | 1.935 | 19.3242 | 0.16 | 1368.675 | 2022/02/16 15:34:57
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.22273 | 3.6171
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -12.211 | -92.898 | -14.9357 | -104.8698 | 0.832 | 1317.615 | 2022/02/16 15:44:45
3 |  | UT2 | -13.006 | -105.884 | -15.7264 | -116.9767 | 0.322 | 1306.122 | 2022/02/16 15:53:42
3 |  | UT2 | -13.128 | -108.28 | -15.8478 | -119.2105 | 0.249 | 1305.727 | 2022/02/16 16:02:38
3 |  | UT2 | -13.168 | -109.384 | -15.8875 | -120.2398 | 0.094 | 1303.787 | 2022/02/16 16:11:35
3 |  | UT2 | -13.195 | -109.88 | -15.9144 | -120.7022 | 0.3 | 1305.497 | 2022/02/16 16:20:32
3 |  | UT2 | -13.248 | -110.377 | -15.9671 | -121.1656 | 0.485 | 1303.776 | 2022/02/16 16:29:28
3 |  | UT2 | -13.301 | -110.615 | -16.0198 | -121.3874 | 0.729 | 1302.092 | 2022/02/16 16:38:24
3 |  | UT2 | -13.299 | -110.91 | -16.0178 | -121.6625 | 0.746 | 1302.07 | 2022/02/16 16:47:21
3 |  | UT2 | -13.311 | -111.102 | -16.0298 | -121.8415 | 0.409 | 1303.021 | 2022/02/16 16:56:17
3 |  | UT2 | -13.335 | -111.212 | -16.0536 | -121.944 | 0.364 | 1302.412 | 2022/02/16 17:05:13
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.33519 | 5.5672
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -7.854 | -65.534 | -10.6024 | -79.3583 | 0.291 | 1302.011 | 2022/02/16 17:15:03
4 |  | EV | -7.579 | -60.996 | -10.3289 | -75.1275 | 0.368 | 1304.398 | 2022/02/16 17:23:59
4 |  | EV | -7.593 | -60.379 | -10.3429 | -74.5522 | 0.331 | 1304.206 | 2022/02/16 17:32:56
4 |  | EV | -7.566 | -59.861 | -10.316 | -74.0693 | 0.686 | 1303.613 | 2022/02/16 17:41:53
4 |  | EV | -7.601 | -59.717 | -10.3508 | -73.9351 | 0.558 | 1304.001 | 2022/02/16 17:50:49
4 |  | EV | -7.555 | -59.525 | -10.3051 | -73.7561 | 0.463 | 1303.641 | 2022/02/16 17:59:46
4 |  | EV | -7.536 | -59.513 | -10.2862 | -73.7449 | 0.332 | 1303.528 | 2022/02/16 18:08:43
4 |  | EV | -7.525 | -59.474 | -10.2752 | -73.7085 | 0.375 | 1304.275 | 2022/02/16 18:17:43
4 |  | EV | -7.594 | -59.541 | -10.3439 | -73.771 | 0.55 | 1303.666 | 2022/02/16 18:26:40
4 |  | EV | -7.554 | -59.47 | -10.3041 | -73.7048 | 0.347 | 1303.858 | 2022/02/16 18:35:37
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3456 | -74.5727 |  |  |  | 0.094208 | 1.8703
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | CPER3-deep-2 | -6.196 | -80.319 | -8.9535 | -93.1424 | 0.374 | 1295.988 | 2022/02/16 18:45:27
5 |  | CPER3-deep-2 | -6.132 | -82.009 | -8.8898 | -94.718 | 0.859 | 1297.212 | 2022/02/16 18:54:23
5 |  | CPER3-deep-2 | -6.095 | -82.352 | -8.853 | -95.0377 | 0.355 | 1297.045 | 2022/02/16 19:03:20
5 |  | CPER3-deep-2 | -6.078 | -82.589 | -8.8361 | -95.2587 | 0.608 | 1296.596 | 2022/02/16 19:12:17
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | CPER3-deep-2 | -8.8831 | -94.5392 |  |  |  | 0.052277 | 1.0268
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | CPER3-E-2 | 4.744 | -45.865 | 1.927 | -61.0208 | 0.072 | 1302.163 | 2022/02/16 19:22:06
6 |  | CPER3-E-2 | 5.174 | -42.202 | 2.3547 | -57.6058 | 0.404 | 1303.136 | 2022/02/16 19:31:03
6 |  | CPER3-E-2 | 5.271 | -41.453 | 2.4511 | -56.9075 | 0.509 | 1305.246 | 2022/02/16 19:40:00
6 |  | CPER3-E-2 | 5.34 | -40.91 | 2.5198 | -56.4012 | 0.36 | 1306.645 | 2022/02/16 19:48:57
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | CPER3-E-2 | 2.3132 | -57.9838 |  |  |  | 0.26764 | 2.2353
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | CPER3-D-5 | 8.89 | -41.445 | 6.0505 | -56.9 | -4.797 | 1333.215 | 2022/02/16 19:58:47
7 |  | CPER3-D-5 | 8.776 | -42.437 | 5.9371 | -57.8248 | -5.654 | 1333.013 | 2022/02/16 20:07:43
7 |  | CPER3-D-5 | 8.774 | -42.711 | 5.9351 | -58.0803 | -5.445 | 1333.479 | 2022/02/16 20:16:40
7 |  | CPER3-D-5 | 8.834 | -43.013 | 5.9948 | -58.3619 | -5.245 | 1331.96 | 2022/02/16 20:25:37
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | CPER3-D-5 | 5.9794 | -57.7918 | -5.2853 | 24.139 | 20.9866 | 0.055194 | 0.67968 | -15.3217 | 11.5307
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | CPER3-010-1 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/16 20:35:24
8 |  | CPER3-010-1 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/16 20:44:21
8 |  | CPER3-010-1 | 4.45 | -61.353 | 1.6346 | -75.4603 | 0.623 | 1306.382 | 2022/02/16 20:53:17
8 |  | CPER3-010-1 | 4.442 | -61.364 | 1.6267 | -75.4706 | 0.307 | 1306.143 | 2022/02/16 21:02:14
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 | * | CPER3-010-1 | -0.58027 | -46.8631 |  |  |  | 2.5669 | 35.4253
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | CPER3-010-3 | 4.984 | -58.028 | 2.1657 | -72.3604 | 0.13 | 1302.247 | 2022/02/16 21:12:01
9 |  | CPER3-010-3 | 5 | -57.964 | 2.1816 | -72.3007 | 0.266 | 1301.854 | 2022/02/16 21:20:58
9 |  | CPER3-010-3 | 5.076 | -57.635 | 2.2572 | -71.994 | 0.438 | 1303.73 | 2022/02/16 21:29:55
9 |  | CPER3-010-3 | 5.103 | -57.621 | 2.2841 | -71.9809 | 0.264 | 1304.685 | 2022/02/16 21:38:51
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | CPER3-010-3 | 2.2222 | -72.159 |  |  |  | 0.057731 | 0.21414
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | CPER3-D-2 | 8.089 | -47.177 | 5.2538 | -62.244 | -2.861 | 1332.618 | 2022/02/16 21:48:39
10 |  | CPER3-D-2 | 8.103 | -46.651 | 5.2677 | -61.7536 | -2.822 | 1334.359 | 2022/02/16 21:57:36
10 |  | CPER3-D-2 | 8.072 | -46.727 | 5.2369 | -61.8244 | -2.775 | 1335.805 | 2022/02/16 22:06:32
10 |  | CPER3-D-2 | 8.093 | -46.304 | 5.2578 | -61.4301 | -3.296 | 1339.863 | 2022/02/16 22:15:29
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | CPER3-D-2 | 5.2541 | -61.813 | -2.9385 | 26.8835 |  | 0.01292 | 0.35898 | -8.8577
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | CPER3-3040-1 | -2.532 | -72.923 | -5.3094 | -86.2471 | 0.288 | 1299.832 | 2022/02/16 22:25:18
11 |  | CPER3-3040-1 | -2.791 | -74.412 | -5.567 | -87.6353 | 0.254 | 1302.111 | 2022/02/16 22:34:15
11 |  | CPER3-3040-1 | -2.917 | -74.877 | -5.6923 | -88.0688 | 0.446 | 1301.978 | 2022/02/16 22:43:12
11 |  | CPER3-3040-1 | -2.963 | -75.135 | -5.7381 | -88.3093 | 0.407 | 1301.024 | 2022/02/16 22:52:09
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | CPER3-3040-1 | -5.5767 | -87.5651 |  |  |  | 0.19336 | 0.98884
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | CPER3-B-2 | 3.914 | -51.903 | 1.1015 | -66.65 | -0.067 | 1330.852 | 2022/02/16 23:01:58
12 |  | CPER3-B-2 | 4.211 | -49.732 | 1.3969 | -64.626 | 0.107 | 1330.882 | 2022/02/16 23:10:55
12 |  | CPER3-B-2 | 4.292 | -49.331 | 1.4775 | -64.2521 | 0.555 | 1331.319 | 2022/02/16 23:19:51
12 |  | CPER3-B-2 | 4.32 | -48.873 | 1.5053 | -63.8252 | 0.2 | 1335.401 | 2022/02/16 23:28:48
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | CPER3-B-2 | 1.3703 | -64.8383 |  |  |  | 0.186 | 1.3422
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -6.743 | -57.984 | -9.4975 | -72.3194 | 0.743 | 1304.47 | 2022/02/16 23:38:37
13 |  | EV | -7.196 | -59.147 | -9.948 | -73.4036 | 0.349 | 1300.958 | 2022/02/16 23:47:33
13 |  | EV | -7.336 | -59.275 | -10.0873 | -73.523 | 0.39 | 1301.381 | 2022/02/16 23:56:30
13 |  | EV | -7.384 | -59.441 | -10.135 | -73.6777 | 0.548 | 1300.483 | 2022/02/17 00:05:27
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -9.9169 | -73.2309 |  |  |  | 0.29226 | 0.66285
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | TEST | -12.499 | -103.112 | -15.2222 | -114.3924 | 0.476 | 1301.393 | 2022/02/17 00:15:16
14 |  | TEST | -12.625 | -107.072 | -15.3475 | -118.0843 | 0.333 | 1300.948 | 2022/02/17 00:24:13
14 |  | TEST | -12.669 | -108.13 | -15.3913 | -119.0707 | 0.476 | 1301.556 | 2022/02/17 00:33:10
14 |  | TEST | -12.698 | -108.395 | -15.4201 | -119.3177 | 0.222 | 1301.706 | 2022/02/17 00:42:06
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | TEST | -15.3453 | -117.7163 |  |  |  | 0.087789 | 2.4446
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | CPER3-B-3 | 3.265 | -54.242 | 0.45606 | -68.8307 | 0.282 | 1368.736 | 2022/02/17 00:51:52
15 |  | CPER3-B-3 | 3.882 | -49.308 | 1.0697 | -64.2307 | 0.031 | 1367.753 | 2022/02/17 01:00:49
15 |  | CPER3-B-3 | 4.046 | -47.761 | 1.2328 | -62.7884 | 0.487 | 1374.789 | 2022/02/17 01:09:46
15 |  | CPER3-B-3 | 4.159 | -47.142 | 1.3452 | -62.2113 | 0.214 | 1375.811 | 2022/02/17 01:18:42
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | CPER3-B-3 | 1.0259 | -64.5153 |  | 62.9945 |  | 0.39857 | 3.2175
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | CPER3-010-2 | -0.247 | -68.126 | -3.0368 | -81.7748 | 0.489 | 1305.577 | 2022/02/17 01:28:29
16 |  | CPER3-010-2 | -0.493 | -70.047 | -3.2815 | -83.5658 | 0.088 | 1302.189 | 2022/02/17 01:37:26
16 |  | CPER3-010-2 | -0.509 | -70.595 | -3.2974 | -84.0767 | 0.579 | 1301.063 | 2022/02/17 01:46:22
16 |  | CPER3-010-2 | -0.471 | -70.801 | -3.2596 | -84.2687 | 0.682 | 1300.665 | 2022/02/17 01:55:19
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | CPER3-010-2 | -3.2188 | -83.4215 |  |  |  | 0.12299 | 1.2197
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | CPER3-2030-2 | 0.346 | -66.918 | -2.4471 | -80.6486 | 0.25 | 1300.422 | 2022/02/17 02:05:06
17 |  | CPER3-2030-2 | 0.469 | -66.186 | -2.3247 | -79.9661 | 0.223 | 1302.378 | 2022/02/17 02:14:02
17 |  | CPER3-2030-2 | 0.466 | -66.156 | -2.3277 | -79.9382 | 0.357 | 1301.195 | 2022/02/17 02:22:59
17 |  | CPER3-2030-2 | 0.485 | -66.303 | -2.3088 | -80.0752 | 0.443 | 1301.187 | 2022/02/17 02:31:55
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | CPER3-2030-2 | -2.3521 | -80.157 |  |  |  | 0.064211 | 0.35718
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | CPER3-B-4 | 1.948 | -53.718 | -0.85377 | -68.3422 | 0.119 | 1380.841 | 2022/02/17 02:41:43
18 |  | CPER3-B-4 | 2.013 | -52.301 | -0.78913 | -67.0211 | 0.075 | 1383.721 | 2022/02/17 02:50:39
18 |  | CPER3-B-4 | 2.006 | -52.047 | -0.79609 | -66.7843 | 0.406 | 1384.152 | 2022/02/17 02:59:35
18 |  | CPER3-B-4 | 1.995 | -51.79 | -0.80703 | -66.5447 | 0.362 | 1384.71 | 2022/02/17 03:08:31
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | CPER3-B-4 | -0.8115 | -67.1731 |  | 74.5783 |  | 0.029286 | 0.86164
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | CPER3-E-3 | 1.725 | -55.926 | -1.0756 | -70.4007 | 0.1 | 1300.52 | 2022/02/17 03:18:19
19 |  | CPER3-E-3 | 1.803 | -55.874 | -0.99798 | -70.3522 | 0.491 | 1302.852 | 2022/02/17 03:27:14
19 |  | CPER3-E-3 | 1.666 | -56.391 | -1.1342 | -70.8342 | 0.397 | 1297.514 | 2022/02/17 03:36:10
19 |  | CPER3-E-3 | 1.706 | -55.925 | -1.0945 | -70.3998 | 0.486 | 1303.019 | 2022/02/17 03:45:07
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | CPER3-E-3 | -1.0756 | -70.4967 |  |  |  | 0.057521 | 0.24255
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | CPER3-E-4 | 4.368 | -48.171 | 1.5531 | -63.1707 | 0.464 | 1333.108 | 2022/02/17 03:54:55
20 |  | CPER3-E-4 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 04:03:51
20 |  | CPER3-E-4 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 04:13:50
20 |  | CPER3-E-4 | 4.368 | -48.171 | 1.5531 | -63.1707 | 0.464 | 1333.108 | 2022/02/17 04:23:48
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 | * | CPER3-E-4 | -0.61906 | -40.7157 |  |  |  | 2.5219 | 27.8115
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | CPER3-2030-1 | -0.325 | -65.92 | -3.1144 | -79.7181 | 0.266 | 1295.177 | 2022/02/17 04:34:39
21 |  | CPER3-2030-1 | -0.33 | -67.019 | -3.1194 | -80.7427 | 0.428 | 1296.323 | 2022/02/17 04:43:35
21 |  | CPER3-2030-1 | -0.303 | -67.199 | -3.0925 | -80.9105 | 0.19 | 1298.545 | 2022/02/17 04:52:31
21 |  | CPER3-2030-1 | -0.341 | -67.129 | -3.1303 | -80.8453 | 0.52 | 1297.895 | 2022/02/17 05:01:28
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | CPER3-2030-1 | -3.1142 | -80.5542 |  |  |  | 0.015966 | 0.60241
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -7.033 | -60.341 | -9.7859 | -74.5168 | 0.48 | 1300.13 | 2022/02/17 05:11:14
22 |  | EV | -7.397 | -59.75 | -10.1479 | -73.9658 | 0.468 | 1298.01 | 2022/02/17 05:20:10
22 |  | EV | -7.44 | -59.554 | -10.1907 | -73.7831 | 0.581 | 1299.647 | 2022/02/17 05:29:05
22 |  | EV | -7.471 | -59.752 | -10.2215 | -73.9677 | 0.642 | 1298.882 | 2022/02/17 05:38:02
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.0865 | -74.0584 |  |  |  | 0.20377 | 0.34073
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | TEST | -12.519 | -103.642 | -15.2421 | -114.8865 | 0.173 | 1297.732 | 2022/02/17 05:47:48
23 |  | TEST | -12.764 | -107.936 | -15.4857 | -118.8898 | 0.617 | 1296.775 | 2022/02/17 05:56:44
23 |  | TEST | -12.886 | -108.899 | -15.6071 | -119.7876 | 0.564 | 1296.742 | 2022/02/17 06:05:40
23 |  | TEST | -12.819 | -108.981 | -15.5404 | -119.8641 | 0.392 | 1297.276 | 2022/02/17 06:14:36
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | TEST | -15.4688 | -118.357 |  |  |  | 0.15998 | 2.5266
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | CPER3-C-1 | -3.972 | -74.414 | -6.7416 | -87.6371 | 0.219 | 1297.737 | 2022/02/17 06:24:23
24 |  | CPER3-C-1 | -3.504 | -70.529 | -6.2761 | -84.0151 | 0.612 | 1301.664 | 2022/02/17 06:33:19
24 |  | CPER3-C-1 | -3.394 | -70.075 | -6.1667 | -83.5919 | 0.577 | 1302.209 | 2022/02/17 06:42:15
24 |  | CPER3-C-1 | -3.378 | -69.562 | -6.1508 | -83.1136 | 0.261 | 1302.683 | 2022/02/17 06:51:12
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | CPER3-C-1 | -6.3338 | -84.5894 |  |  |  | 0.27901 | 2.2148
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | CPER3-D-3 | 5.929 | -51.829 | 3.1056 | -66.581 | -3.081 | 1318.072 | 2022/02/17 07:00:59
25 |  | CPER3-D-3 | 6.241 | -50.424 | 3.4159 | -65.2712 | -3.345 | 1322.448 | 2022/02/17 07:09:55
25 |  | CPER3-D-3 | 6.322 | -50.043 | 3.4964 | -64.9159 | -3.692 | 1322.32 | 2022/02/17 07:18:51
25 |  | CPER3-D-3 | 6.316 | -49.975 | 3.4905 | -64.8526 | -3.537 | 1322.139 | 2022/02/17 07:27:48
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | CPER3-D-3 | 3.3771 | -65.4052 | -3.4138 |  | 7.351 | 0.18569 | 0.86374 | -9.4547 | 6.9597
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | CPER3-E-5 | 7.184 | -42.786 | 4.3537 | -58.1502 | 0.43 | 1358.193 | 2022/02/17 07:37:35
26 |  | CPER3-E-5 | 7.312 | -41.321 | 4.481 | -56.7844 | 0.458 | 1360.865 | 2022/02/17 07:46:30
26 |  | CPER3-E-5 | 7.358 | -40.827 | 4.5268 | -56.3238 | 0.486 | 1360.413 | 2022/02/17 07:55:26
26 |  | CPER3-E-5 | 7.308 | -41.012 | 4.4771 | -56.4963 | 0.554 | 1359.306 | 2022/02/17 08:04:23
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | CPER3-E-5 | 4.4597 | -56.9387 |  | 50.9165 |  | 0.074536 | 0.88998
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | CPER3-D-1 | 5 | -54.199 | 2.1816 | -68.7906 | -2.45 | 1361.341 | 2022/02/17 08:14:10
27 |  | CPER3-D-1 | 4.742 | -56.096 | 1.925 | -70.5592 | -2.303 | 1362.153 | 2022/02/17 08:23:06
27 |  | CPER3-D-1 | 4.766 | -56.634 | 1.9489 | -71.0608 | -2.271 | 1361.924 | 2022/02/17 08:32:03
27 |  | CPER3-D-1 | 4.751 | -56.659 | 1.934 | -71.0841 | -2.664 | 1362.002 | 2022/02/17 08:40:59
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | CPER3-D-1 | 1.9974 | -70.3737 | -2.422 | 53.0773 |  | 0.1239 | 1.1614 | -8.304
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | CPER3-A-1 | 3.967 | -53.401 | 1.1542 | -68.0466 | 0.476 | 1305.276 | 2022/02/17 08:50:47
28 |  | CPER3-A-1 | 4.02 | -52.534 | 1.207 | -67.2383 | 0.48 | 1305.183 | 2022/02/17 08:59:43
28 |  | CPER3-A-1 | 4.051 | -51.929 | 1.2378 | -66.6743 | 0.627 | 1305.097 | 2022/02/17 09:08:39
28 |  | CPER3-A-1 | 4.002 | -52.308 | 1.1891 | -67.0276 | 0.473 | 1303.735 | 2022/02/17 09:17:35
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | CPER3-A-1 | 1.197 | -67.2467 |  |  |  | 0.03509 | 0.62409
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | CPER3-D-4 | 3.506 | -52.566 | 0.69575 | -67.2682 | -0.686 | 1307.988 | 2022/02/17 09:27:19
29 |  | CPER3-D-4 | 3.362 | -53.279 | 0.55253 | -67.9329 | -0.748 | 1304.252 | 2022/02/17 09:36:15
29 |  | CPER3-D-4 | 3.41 | -52.852 | 0.60027 | -67.5348 | -0.562 | 1308.183 | 2022/02/17 09:45:11
29 |  | CPER3-D-4 | 3.426 | -52.966 | 0.61619 | -67.6411 | -0.493 | 1308.136 | 2022/02/17 09:54:07
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | CPER3-D-4 | 0.61619 | -67.5942 |  |  |  | 0.059867 | 0.29488
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | CPER3-C-2 | -2.383 | -69.658 | -5.1612 | -83.2031 | 0.288 | 1299.871 | 2022/02/17 10:03:52
30 |  | CPER3-C-2 | -2.602 | -71.006 | -5.379 | -84.4598 | 0.306 | 1300.536 | 2022/02/17 10:12:49
30 |  | CPER3-C-2 | -2.642 | -71.316 | -5.4188 | -84.7488 | 0.243 | 1298.318 | 2022/02/17 10:21:44
30 |  | CPER3-C-2 | -2.673 | -71.595 | -5.4496 | -85.009 | 0.553 | 1298.131 | 2022/02/17 10:30:41
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | CPER3-C-2 | -5.3522 | -84.3552 |  |  |  | 0.13126 | 0.85824
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -7.204 | -61.032 | -9.956 | -75.161 | 0.417 | 1299.242 | 2022/02/17 10:40:26
31 |  | EV | -7.389 | -59.685 | -10.14 | -73.9052 | 0.533 | 1301.933 | 2022/02/17 10:49:22
31 |  | EV | -7.396 | -59.595 | -10.1469 | -73.8213 | 0.398 | 1301.434 | 2022/02/17 10:58:18
31 |  | EV | -7.416 | -59.429 | -10.1668 | -73.6666 | 0.634 | 1301.501 | 2022/02/17 11:07:14
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.1024 | -74.1385 |  |  |  | 0.098831 | 0.73882
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | TEST | -12.545 | -103.697 | -15.2679 | -114.9378 | 0.273 | 1297.432 | 2022/02/17 11:17:00
32 |  | TEST | -12.785 | -107.872 | -15.5066 | -118.8301 | 0.496 | 1298.327 | 2022/02/17 11:25:56
32 |  | TEST | -12.853 | -108.706 | -15.5743 | -119.6077 | 0.322 | 1297.627 | 2022/02/17 11:34:52
32 |  | TEST | -12.854 | -109.112 | -15.5752 | -119.9862 | 0.563 | 1298.301 | 2022/02/17 11:43:48
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | TEST | -15.481 | -118.3404 |  |  |  | 0.14644 | 2.4873
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | CPER3-A-2 | 0.103 | -66.249 | -2.6887 | -80.0249 | 0.645 | 1342.643 | 2022/02/17 11:53:34
33 |  | CPER3-A-2 | 0.612 | -62.783 | -2.1825 | -76.7935 | 0.182 | 1337.706 | 2022/02/17 12:02:30
33 |  | CPER3-A-2 | 0.781 | -61.66 | -2.0144 | -75.7465 | 0.671 | 1338.798 | 2022/02/17 12:11:26
33 |  | CPER3-A-2 | 0.826 | -61.366 | -1.9697 | -75.4724 | 0.237 | 1339.976 | 2022/02/17 12:20:23
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | CPER3-A-2 | -2.2138 | -77.0093 |  |  |  | 0.3314 | 2.2411
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | CPER3-A-3 | 0.42 | -63.033 | -2.3735 | -77.0266 | 0.276 | 1346.712 | 2022/02/17 12:30:09
34 |  | CPER3-A-3 | 0.417 | -63.324 | -2.3764 | -77.2979 | 0.346 | 1344.161 | 2022/02/17 12:39:05
34 |  | CPER3-A-3 | 0.459 | -63.099 | -2.3347 | -77.0881 | 0.331 | 1347.871 | 2022/02/17 12:48:01
34 |  | CPER3-A-3 | 0.45 | -62.766 | -2.3436 | -76.7776 | 0.312 | 1348.455 | 2022/02/17 12:56:57
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | CPER3-A-3 | -2.3571 | -77.0475 |  |  |  | 0.021142 | 0.22971
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | CPER3-A-4 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 13:06:44
35 |  | CPER3-A-4 | -0.074 | -60.203 | -2.8648 | -74.3882 | 0.466 | 1300.14 | 2022/02/17 13:15:41
35 |  | CPER3-A-4 | -0.093 | -60.328 | -2.8837 | -74.5047 | 0.398 | 1299.163 | 2022/02/17 13:24:38
35 |  | CPER3-A-4 | -0.093 | -60.328 | -2.8837 | -74.5047 | 0.398 | 1299.163 | 2022/02/17 13:33:34
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 | * | CPER3-A-4 | -2.8558 | -60.4146 |  |  |  | 0.044249 | 30.1432
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | CPER3-A-5 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 13:44:18
36 |  | CPER3-A-5 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 13:53:15
36 |  | CPER3-A-5 | -3.007 | -65.253 | -5.7818 | -79.0963 | 0.282 | 1306.035 | 2022/02/17 14:02:10
36 |  | CPER3-A-5 | -3.021 | -65.297 | -5.7957 | -79.1373 | 0.536 | 1308.784 | 2022/02/17 14:11:07
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 | * | CPER3-A-5 | -4.29 | -48.6887 |  |  |  | 1.7401 | 37.6865
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | CPER3-B-1 | -0.104 | -61.974 | -2.8946 | -76.0393 | 0.311 | 1321.896 | 2022/02/17 14:20:50
37 |  | CPER3-B-1 | 0.012 | -61.4 | -2.7792 | -75.5041 | 0.212 | 1322.05 | 2022/02/17 14:29:46
37 |  | CPER3-B-1 | 0.017 | -61.604 | -2.7743 | -75.6943 | 0.463 | 1319.478 | 2022/02/17 14:38:41
37 |  | CPER3-B-1 | 0.033 | -61.506 | -2.7584 | -75.6029 | 0.508 | 1320.56 | 2022/02/17 14:47:37
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | CPER3-B-1 | -2.8016 | -75.7102 |  |  |  | 0.062974 | 0.24964
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | CPER3-B-5 | 2.405 | -56.983 | -0.39926 | -71.3861 | 0 | 1504.964 | 2022/02/17 14:57:21
38 |  | CPER3-B-5 | 2.642 | -55.378 | -0.16355 | -69.8898 | 0.076 | 1530.553 | 2022/02/17 15:06:17
38 |  | CPER3-B-5 | 2.457 | -55.97 | -0.34754 | -70.4417 | 0.17 | 1517.874 | 2022/02/17 15:15:13
38 |  | CPER3-B-5 | 2.475 | -56.064 | -0.32964 | -70.5293 | 0.372 | 1516.534 | 2022/02/17 15:24:09
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | CPER3-B-5 | -0.31 | -70.5617 |  | 208.7035 |  | 0.10256 | 0.66311 | -6.5747
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | CPER3-E-1 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 15:33:53
39 |  | CPER3-E-1 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 15:43:51
39 |  | CPER3-E-1 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 15:53:49
39 |  | CPER3-E-1 |  |  | -2.7912 | -18.2607 |  |  | 2022/02/17 16:03:47
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 | * | CPER3-E-1 | -2.7912 | -18.2607 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.169 | -59.151 | -9.9212 | -73.4074 | 0.441 | 1305.392 | 2022/02/17 16:14:33
40 |  | EV | -7.331 | -59.429 | -10.0823 | -73.6666 | 0.336 | 1301.297 | 2022/02/17 16:23:29
40 |  | EV | -7.443 | -59.743 | -10.1937 | -73.9593 | 0.178 | 1301.153 | 2022/02/17 16:32:25
40 |  | EV | -7.498 | -59.395 | -10.2484 | -73.6349 | 0.495 | 1300.535 | 2022/02/17 16:41:21
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.1114 | -73.667 |  |  |  | 0.14521 | 0.24292
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | TEST | -12.573 | -103.567 | -15.2958 | -114.8166 | 0.291 | 1297.925 | 2022/02/17 16:51:05
41 |  | TEST | -12.682 | -107.336 | -15.4042 | -118.3304 | 0.41 | 1300.201 | 2022/02/17 17:00:01
41 |  | TEST | -12.853 | -108.836 | -15.5743 | -119.7289 | 0.342 | 1298.629 | 2022/02/17 17:08:57
41 |  | TEST | -12.854 | -108.92 | -15.5752 | -119.8072 | 0.616 | 1299.183 | 2022/02/17 17:17:52
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | TEST | -15.4624 | -118.1708 |  |  |  | 0.13786 | 2.5065
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | CPER3-deep-3 | -6.366 | -80.386 | -9.1225 | -93.2048 | 0.544 | 1300.28 | 2022/02/17 17:27:37
42 |  | CPER3-deep-3 | -6.223 | -77.926 | -8.9803 | -90.9114 | 0.519 | 1297.658 | 2022/02/17 17:36:34
42 |  | CPER3-deep-3 | -6.111 | -77.384 | -8.8689 | -90.4061 | 0.491 | 1298.895 | 2022/02/17 17:45:30
42 |  | CPER3-deep-3 | -6.174 | -77.3 | -8.9316 | -90.3277 | 0.515 | 1296.678 | 2022/02/17 17:54:25
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | CPER3-deep-3 | -8.9758 | -91.2125 |  |  |  | 0.10849 | 1.4514
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | CPER3-3040-3 | -5.932 | -77.765 | -8.6909 | -90.7613 | 0.601 | 1299.101 | 2022/02/17 18:04:09
43 |  | CPER3-3040-3 | -5.95 | -77.989 | -8.7088 | -90.9701 | 0.549 | 1299.66 | 2022/02/17 18:13:04
43 |  | CPER3-3040-3 | -5.848 | -77.685 | -8.6074 | -90.6867 | 0.409 | 1301.053 | 2022/02/17 18:22:00
43 |  | CPER3-3040-3 | -5.878 | -77.785 | -8.6372 | -90.7799 | 0.571 | 1300.309 | 2022/02/17 18:30:56
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | CPER3-3040-3 | -8.6611 | -90.7995 |  |  |  | 0.047244 | 0.12942
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | CPER3-2030-3 | -4.083 | -77.709 | -6.852 | -90.7091 | 0.397 | 1299.076 | 2022/02/17 18:40:40
44 |  | CPER3-2030-3 | -3.996 | -77.567 | -6.7654 | -90.5767 | 0.251 | 1299.174 | 2022/02/17 18:49:36
44 |  | CPER3-2030-3 | -4.041 | -77.706 | -6.8102 | -90.7063 | 0.579 | 1300.075 | 2022/02/17 18:58:31
44 |  | CPER3-2030-3 | -3.985 | -77.453 | -6.7545 | -90.4704 | 0.385 | 1299.32 | 2022/02/17 19:07:27
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | CPER3-2030-3 | -6.7955 | -90.6156 |  |  |  | 0.044925 | 0.12316
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | CPER3-C-5 | -3.449 | -73.41 | -6.2214 | -86.7011 | 0.479 | 1300.104 | 2022/02/17 19:17:12
45 |  | CPER3-C-5 | -3.483 | -73.147 | -6.2552 | -86.4559 | 0.461 | 1301.549 | 2022/02/17 19:26:08
45 |  | CPER3-C-5 | -3.428 | -73.091 | -6.2005 | -86.4037 | 0.633 | 1301.251 | 2022/02/17 19:35:04
45 |  | CPER3-C-5 | -3.431 | -72.847 | -6.2035 | -86.1762 | 0.466 | 1301.069 | 2022/02/17 19:43:59
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | CPER3-C-5 | -6.2202 | -86.4342 |  |  |  | 0.025264 | 0.23104
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | CPER3-C-4 | -3.44 | -73.008 | -6.2125 | -86.3263 | -0.107 | 1385.94 | 2022/02/17 19:53:45
46 |  | CPER3-C-4 | -3.44 | -73.113 | -6.2125 | -86.4242 | 0.039 | 1392.356 | 2022/02/17 20:02:41
46 |  | CPER3-C-4 | -3.457 | -73.174 | -6.2294 | -86.4811 | 0.13 | 1396.616 | 2022/02/17 20:11:36
46 |  | CPER3-C-4 | -3.446 | -73.131 | -6.2184 | -86.441 | 0.096 | 1392.8 | 2022/02/17 20:20:32
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | CPER3-C-4 | -6.2182 | -86.4181 |  | 83.1503 |  | 0.0080156 | 0.070477
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | CPER3-1020-3 | 0.573 | -67.626 | -2.2213 | -81.3086 | 0.467 | 1302.374 | 2022/02/17 20:30:18
47 |  | CPER3-1020-3 | 0.81 | -66.93 | -1.9856 | -80.6598 | 0.55 | 1302.065 | 2022/02/17 20:39:13
47 |  | CPER3-1020-3 | 0.803 | -66.808 | -1.9925 | -80.546 | 0.459 | 1301.832 | 2022/02/17 20:48:09
47 |  | CPER3-1020-3 | 0.863 | -66.751 | -1.9329 | -80.4929 | 0.455 | 1301.51 | 2022/02/17 20:57:05
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | CPER3-1020-3 | -2.0331 | -80.7518 |  |  |  | 0.12898 | 0.40511
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | CPER3-3040-2 | -5.02 | -81.611 | -7.7839 | -94.3469 | 0.474 | 1294.189 | 2022/02/17 21:06:52
48 |  | CPER3-3040-2 | -5.282 | -83.054 | -8.0444 | -95.6922 | 0.247 | 1294.786 | 2022/02/17 21:15:48
48 |  | CPER3-3040-2 | -5.293 | -83.268 | -8.0554 | -95.8917 | 0.388 | 1295.264 | 2022/02/17 21:24:44
48 |  | CPER3-3040-2 | -5.355 | -83.159 | -8.117 | -95.7901 | 0.303 | 1294.924 | 2022/02/17 21:33:40
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | CPER3-3040-2 | -8.0002 | -95.4302 |  |  |  | 0.14852 | 0.77958
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -7.31 | -62.167 | -10.0614 | -76.2192 | 0.486 | 1301.54 | 2022/02/17 21:43:26
49 |  | EV | -7.459 | -60.596 | -10.2096 | -74.7545 | 0.34 | 1300.085 | 2022/02/17 21:52:22
49 |  | EV | -7.466 | -59.844 | -10.2166 | -74.0535 | 0.444 | 1299.017 | 2022/02/17 22:01:18
49 |  | EV | -7.491 | -59.945 | -10.2414 | -74.1476 | 0.507 | 1300.17 | 2022/02/17 22:10:13
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.1822 | -74.7937 |  |  |  | 0.082156 | 1.0724
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | TEST | -7.494 | -60.037 | -10.2444 | -74.2334 | 0.435 | 1300.005 | 2022/02/17 22:19:58
50 |  | TEST | -7.461 | -59.711 | -10.2116 | -73.9295 | 0.412 | 1299.994 | 2022/02/17 22:28:54
50 |  | TEST | -7.506 | -59.701 | -10.2563 | -73.9201 | 0.295 | 1299.184 | 2022/02/17 22:37:50
50 |  | TEST | -7.506 | -59.673 | -10.2563 | -73.894 | 0.397 | 1301.181 | 2022/02/17 22:46:46
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 |  | TEST | -10.2422 | -73.9943 |  |  |  | 0.021266 | 0.17175
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | CPER3-C-3 | 0.595 | -57.996 | -2.1994 | -72.3306 | 0.616 | 1299.031 | 2022/02/17 22:56:31
51 |  | CPER3-C-3 | 1.003 | -57.566 | -1.7936 | -71.9297 | 0.324 | 1301.248 | 2022/02/17 23:05:27
51 |  | CPER3-C-3 | 1.053 | -57.636 | -1.7439 | -71.9949 | 0.159 | 1302.591 | 2022/02/17 23:14:23
51 |  | CPER3-C-3 | 1.096 | -57.473 | -1.7011 | -71.843 | 0.653 | 1301.618 | 2022/02/17 23:23:18
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 |  | CPER3-C-3 | -1.8595 | -72.0245 |  |  |  | 0.23098 | 0.22879
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
52 |  | CPER3-deep-1 | -6.204 | -81.014 | -8.9614 | -93.7903 | 0.386 | 1298.152 | 2022/02/17 23:33:04
52 |  | CPER3-deep-1 | -6.33 | -82.754 | -9.0867 | -95.4125 | 0.326 | 1300.641 | 2022/02/17 23:42:00
52 |  | CPER3-deep-1 | -6.483 | -83.23 | -9.2389 | -95.8563 | 0.46 | 1297.843 | 2022/02/17 23:50:56
52 |  | CPER3-deep-1 | -6.524 | -83.887 | -9.2797 | -96.4688 | 0.355 | 1299.581 | 2022/02/17 23:59:51
Sample 52 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
52 |  | CPER3-deep-1 | -9.1417 | -95.382 |  |  |  | 0.14687 | 1.2293
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
53 |  | CPER3-1020-2 | 0.462 | -69.442 | -2.3317 | -83.0017 | 0.446 | 1298.159 | 2022/02/18 00:09:37
53 |  | CPER3-1020-2 | 0.806 | -67.807 | -1.9896 | -81.4774 | 0.682 | 1298.926 | 2022/02/18 00:18:32
53 |  | CPER3-1020-2 | 0.804 | -67.753 | -1.9915 | -81.427 | 0.556 | 1298.101 | 2022/02/18 00:27:27
53 |  | CPER3-1020-2 | 0.908 | -67.641 | -1.8881 | -81.3226 | 0.52 | 1299.038 | 2022/02/18 00:36:23
Sample 53 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
53 |  | CPER3-1020-2 | -2.0502 | -81.8072 |  |  |  | 0.19482 | 0.85696
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
54 |  | TEST | -12.169 | -104.701 | -14.894 | -115.8738 | 0.573 | 1298.264 | 2022/02/18 00:46:10
54 |  | TEST | -12.632 | -108.126 | -15.3545 | -119.0669 | 0.394 | 1299.762 | 2022/02/18 00:55:05
54 |  | TEST | -12.761 | -108.764 | -15.4828 | -119.6617 | 0.332 | 1299.007 | 2022/02/18 01:04:01
54 |  | TEST | -12.802 | -109.08 | -15.5235 | -119.9564 | 0.428 | 1299.555 | 2022/02/18 01:12:57
Sample 54 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
54 |  | TEST | -15.3137 | -118.6397 |  |  |  | 0.29051 | 2.0172
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
55 |  | TEST | -12.848 | -109.339 | -15.5693 | -120.1978 | 0.507 | 1298.491 | 2022/02/18 01:22:44
55 |  | TEST | -12.898 | -109.558 | -15.619 | -120.402 | 0.646 | 1299.299 | 2022/02/18 01:31:40
55 |  | TEST | -12.873 | -109.561 | -15.5941 | -120.4048 | 0.613 | 1299.18 | 2022/02/18 01:40:36
55 |  | TEST | -12.841 | -109.559 | -15.5623 | -120.4029 | 0.424 | 1298.554 | 2022/02/18 01:49:30
Sample 55 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
55 |  | TEST | -15.5862 | -120.3519 |  |  |  | 0.025936 | 0.11017
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
56 |  | TEST | -13 | -110.015 | -15.7205 | -120.8281 | 0.322 | 1296.877 | 2022/02/18 01:59:18
56 |  | TEST | -12.943 | -109.662 | -15.6638 | -120.499 | 0.615 | 1297.444 | 2022/02/18 02:08:14
56 |  | TEST | -12.837 | -109.465 | -15.5583 | -120.3153 | 0.348 | 1302.591 | 2022/02/18 02:17:09
56 |  | TEST | -12.85 | -109.78 | -15.5713 | -120.609 | 0.597 | 1299.206 | 2022/02/18 02:26:04
Sample 56 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
56 |  | TEST | -15.6285 | -120.5628 |  |  |  | 0.07766 | 0.22991
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
57 |  | EV | -7.811 | -65.817 | -10.5597 | -79.6221 | 0.396 | 1299.929 | 2022/02/18 02:35:49
57 |  | EV | -7.592 | -61.441 | -10.3419 | -75.5423 | 0.487 | 1299.817 | 2022/02/18 02:44:45
57 |  | EV | -7.517 | -60.574 | -10.2673 | -74.734 | 0.158 | 1300.798 | 2022/02/18 02:53:40
57 |  | EV | -7.466 | -60.403 | -10.2166 | -74.5746 | 0.234 | 1300.424 | 2022/02/18 03:02:36
Sample 57 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
57 |  | EV | -10.3463 | -76.1183 |  |  |  | 0.15208 | 2.5464

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-20130 | 2022/02/16 12:44:09 | 1-01 | 1 | -7.974 | -67.537 | 20990 | -1 | 1 | EV | standard | H2O | 1645040956 | 0.313 | 0.773 | 82.349 | 0 | 0.003 | 1.933 | 1330.593 | 2.467 | 3.377 | 126.8 | 0.869 | 0.31 | 0 | 0 | 1 | 46.688 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-20130 | 2022/02/16 12:53:06 | 1-01 | 2 | -8.064 | -65.143 | 22286 | -1 | 1 | EV | standard | H2O | 1645041493 | 0.329 | 0.656 | 82.498 | 0.001 | 0.001 | 1.933 | 1477.55 | -3.449 | 3.607 | 135.391 | 0.871 | 0.246 | 0 | 0 | 1 | 46.625 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-20130 | 2022/02/16 13:02:01 | 1-01 | 3 | -7.8 | -62.018 | 19126 | -1 | 1 | EV | standard | H2O | 1645042028 | 0.328 | 0.739 | 81.19 | 0.001 | 0.003 | 1.918 | 1337.615 | 0.693 | 3.075 | 117.873 | 0.865 | 0.566 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-20130 | 2022/02/16 13:10:58 | 1-01 | 4 | -7.615 | -60.512 | 18982 | 0 | 1 | EV | standard | H2O | 1645042565 | 0.322 | 0.729 | 78.915 | 0.002 | 0.003 | 1.853 | 1312.617 | 1.979 | 3.042 | 116.047 | 0.864 | 0.516 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-20130 | 2022/02/16 13:19:55 | 1-01 | 5 | -7.657 | -60.602 | 19159 | 0 | 1 | EV | standard | H2O | 1645043101 | 0.325 | 0.74 | 79.269 | 0 | 0.001 | 1.876 | 1319.453 | 1.842 | 3.11 | 117.151 | 0.864 | 0.517 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-20130 | 2022/02/16 13:28:51 | 1-01 | 6 | -7.568 | -60.179 | 18964 | 0 | 1 | EV | standard | H2O | 1645043637 | 0.316 | 0.72 | 77.86 | 0.001 | 0.001 | 1.854 | 1311.544 | 1.59 | 3.027 | 115.341 | 0.864 | 0.134 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-20130 | 2022/02/16 13:37:47 | 1-01 | 7 | -7.537 | -59.83 | 19176 | 0 | 1 | EV | standard | H2O | 1645044174 | 0.306 | 0.735 | 80.35 | 0.001 | 0.003 | 1.895 | 1311.205 | 2.251 | 3.018 | 116.163 | 0.866 | 0.273 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-20130 | 2022/02/16 13:46:44 | 1-01 | 8 | -7.531 | -59.674 | 19275 | 0 | 1 | EV | standard | H2O | 1645044710 | 0.278 | 0.692 | 77.559 | 0.001 | 0.001 | 1.825 | 1310.249 | 2.107 | 3.083 | 117.632 | 0.863 | 0.516 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-20130 | 2022/02/16 13:55:41 | 1-01 | 9 | -7.483 | -59.496 | 19258 | 0 | 1 | EV | standard | H2O | 1645045248 | 0.3 | 0.721 | 79.715 | 0 | 0.001 | 1.888 | 1309.059 | 2.269 | 3.124 | 117.689 | 0.867 | 0.373 | 0 | 0 | 1 | 46.562 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-20130 | 2022/02/16 14:04:37 | 1-01 | 10 | -7.503 | -59.629 | 19075 | 0 | 1 | EV | standard | H2O | 1645045784 | 0.328 | 0.699 | 79.423 | 0 | 0.002 | 1.885 | 1307.62 | 2.025 | 3.034 | 116.002 | 0.867 | 0.174 | 0 | 0 | 1 | 46.625 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-20131 | 2022/02/16 14:14:25 | 1-02 | 1 | 4.171 | 29.283 | 19416 | -1 | 1 | PZ | standard | H2O | 1645046372 | 0.339 | 0.739 | 80.99 | -0.001 | -0.005 | 1.914 | 1308.338 | 1.611 | 3.152 | 120.215 | 0.865 | 0.51 | 0 | 0 | 1 | 46.688 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-20131 | 2022/02/16 14:23:22 | 1-02 | 2 | 4.625 | 37.152 | 18900 | -1 | 1 | PZ | standard | H2O | 1645046909 | 0.325 | 0.816 | 77.716 | -0.001 | -0.006 | 1.84 | 1305.39 | 1.914 | 3.102 | 116.411 | 0.863 | 0.393 | 0 | 0 | 1 | 46.625 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-20131 | 2022/02/16 14:32:19 | 1-02 | 3 | 4.765 | 39.253 | 19179 | -1 | 1 | PZ | standard | H2O | 1645047446 | 0.359 | 0.755 | 79.898 | 0 | -0.003 | 1.885 | 1307.494 | 1.706 | 3.097 | 117.247 | 0.867 | 0.356 | 0 | 0 | 1 | 46.562 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-20131 | 2022/02/16 14:41:16 | 1-02 | 4 | 4.746 | 39.861 | 18825 | 0 | 1 | PZ | standard | H2O | 1645047984 | 0.339 | 0.754 | 77.177 | 0.001 | -0.002 | 1.831 | 1304.607 | 2.064 | 2.964 | 113.446 | 0.862 | 0.122 | 0 | 0 | 1 | 46.5 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-20131 | 2022/02/16 14:50:12 | 1-02 | 5 | 4.846 | 40.325 | 18738 | 0 | 1 | PZ | standard | H2O | 1645048519 | 0.31 | 0.678 | 77.742 | 0.001 | -0.004 | 1.843 | 1303.238 | 2.017 | 3.066 | 115.143 | 0.862 | 0.537 | 0 | 0 | 1 | 46.375 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-20131 | 2022/02/16 14:59:10 | 1-02 | 6 | 4.842 | 40.524 | 18774 | 0 | 1 | PZ | standard | H2O | 1645049057 | 0.313 | 0.784 | 76.867 | 0.001 | -0.003 | 1.823 | 1304.289 | 1.756 | 2.977 | 114.263 | 0.862 | 0.181 | 0 | 0 | 1 | 46.312 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-20131 | 2022/02/16 15:08:06 | 1-02 | 7 | 4.924 | 41.233 | 19082 | 0 | 1 | PZ | standard | H2O | 1645049594 | 0.329 | 0.714 | 75.716 | 0.001 | -0.002 | 1.802 | 1306.379 | 1.587 | 3.085 | 116.297 | 0.863 | 0.126 | 0 | 0 | 1 | 46.312 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-20131 | 2022/02/16 15:17:03 | 1-02 | 8 | 4.937 | 41.205 | 19095 | 0 | 1 | PZ | standard | H2O | 1645050130 | 0.322 | 0.697 | 77.306 | 0 | 0 | 1.833 | 1306.031 | 1.822 | 3.112 | 117.066 | 0.862 | 0.314 | 0 | 0 | 1 | 46.312 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-20131 | 2022/02/16 15:26:00 | 1-02 | 9 | 4.862 | 40.859 | 18986 | 0 | 1 | PZ | standard | H2O | 1645050667 | 0.334 | 0.729 | 77.604 | 0.001 | 0 | 1.84 | 1308.995 | 1.542 | 3.084 | 116.177 | 0.862 | 0.344 | 0 | 0 | 1 | 46.25 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-20131 | 2022/02/16 15:34:57 | 1-02 | 10 | 4.752 | 40.314 | 21101 | 0 | 1 | PZ | standard | H2O | 1645051204 | 0.301 | 0.776 | 83.022 | 0 | -0.003 | 1.95 | 1368.675 | 1.842 | 3.427 | 126.802 | 0.865 | 0.16 | 0 | 0 | 1 | 46.312 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-20132 | 2022/02/16 15:44:45 | 1-03 | 1 | -12.211 | -92.898 | 20340 | -1 | 1 | UT2 | standard | H2O | 1645051792 | 0.329 | 0.725 | 81.657 | 0.003 | 0.007 | 1.92 | 1317.615 | 2.678 | 3.299 | 124.875 | 0.864 | 0.832 | 0 | 0 | 1 | 46.312 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-20132 | 2022/02/16 15:53:42 | 1-03 | 2 | -13.006 | -105.884 | 19062 | -1 | 1 | UT2 | standard | H2O | 1645052329 | 0.348 | 0.919 | 77.373 | 0.003 | 0.012 | 1.84 | 1306.122 | 2.628 | 3.024 | 115.507 | 0.865 | 0.322 | 0 | 0 | 1 | 46.812 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-20132 | 2022/02/16 16:02:38 | 1-03 | 3 | -13.128 | -108.28 | 19066 | -1 | 1 | UT2 | standard | H2O | 1645052865 | 0.323 | 0.848 | 80.09 | 0.002 | 0.009 | 1.898 | 1305.727 | 2.371 | 3.028 | 115.639 | 0.864 | 0.249 | 0 | 0 | 1 | 46.875 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-20132 | 2022/02/16 16:11:35 | 1-03 | 4 | -13.168 | -109.384 | 18863 | 0 | 1 | UT2 | standard | H2O | 1645053402 | 0.31 | 0.804 | 77.133 | 0.001 | 0.007 | 1.837 | 1303.787 | 3.135 | 2.932 | 112.908 | 0.863 | 0.094 | 0 | 0 | 1 | 46.562 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-20132 | 2022/02/16 16:20:32 | 1-03 | 5 | -13.195 | -109.88 | 19087 | 0 | 1 | UT2 | standard | H2O | 1645053939 | 0.276 | 0.695 | 76.734 | 0 | 0.007 | 1.814 | 1305.497 | 2.658 | 2.989 | 115.422 | 0.865 | 0.3 | 0 | 0 | 1 | 46.5 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-20132 | 2022/02/16 16:29:28 | 1-03 | 6 | -13.248 | -110.377 | 18907 | 0 | 1 | UT2 | standard | H2O | 1645054475 | 0.313 | 0.721 | 78.252 | 0.001 | 0.004 | 1.852 | 1303.776 | 2.608 | 3.042 | 115.613 | 0.863 | 0.485 | 0 | 0 | 1 | 46.5 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-20132 | 2022/02/16 16:38:24 | 1-03 | 7 | -13.301 | -110.615 | 18743 | 0 | 1 | UT2 | standard | H2O | 1645055011 | 0.298 | 0.756 | 77.373 | 0.001 | 0.008 | 1.845 | 1302.092 | 2.815 | 3.014 | 114.668 | 0.863 | 0.729 | 0 | 0 | 1 | 46.5 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-20132 | 2022/02/16 16:47:21 | 1-03 | 8 | -13.299 | -110.91 | 18806 | 0 | 1 | UT2 | standard | H2O | 1645055548 | 0.341 | 0.742 | 79.277 | 0.001 | 0.006 | 1.872 | 1302.07 | 2.98 | 3.05 | 115.818 | 0.865 | 0.746 | 0 | 0 | 1 | 46.562 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-20132 | 2022/02/16 16:56:17 | 1-03 | 9 | -13.311 | -111.102 | 18901 | 0 | 1 | UT2 | standard | H2O | 1645056083 | 0.302 | 0.744 | 77.87 | 0.001 | 0.004 | 1.845 | 1303.021 | 2.596 | 3.016 | 115.695 | 0.862 | 0.409 | 0 | 0 | 1 | 46.5 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-20132 | 2022/02/16 17:05:13 | 1-03 | 10 | -13.335 | -111.212 | 18815 | 0 | 1 | UT2 | standard | H2O | 1645056620 | 0.339 | 0.787 | 77.689 | 0.001 | 0.006 | 1.838 | 1302.412 | 2.559 | 2.96 | 113.968 | 0.864 | 0.364 | 0 | 0 | 1 | 46.5 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-20133 | 2022/02/16 17:15:03 | 1-04 | 1 | -7.854 | -65.534 | 18803 | -1 | 1 | EV | standard | H2O | 1645057209 | 0.283 | 0.753 | 77.108 | 0 | 0.003 | 1.837 | 1302.011 | 3.003 | 2.968 | 113.251 | 0.862 | 0.291 | 0 | 0 | 1 | 46.438 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-20133 | 2022/02/16 17:23:59 | 1-04 | 2 | -7.579 | -60.996 | 19141 | -1 | 1 | EV | standard | H2O | 1645057747 | 0.304 | 0.674 | 78.05 | 0.001 | 0.001 | 1.851 | 1304.398 | 2.697 | 3.037 | 116.137 | 0.862 | 0.368 | 0 | 0 | 1 | 46.375 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-20133 | 2022/02/16 17:32:56 | 1-04 | 3 | -7.593 | -60.379 | 19105 | -1 | 1 | EV | standard | H2O | 1645058283 | 0.326 | 0.698 | 76.241 | 0.001 | 0.001 | 1.808 | 1304.206 | 2.164 | 2.997 | 116.433 | 0.862 | 0.331 | 0 | 0 | 1 | 46.25 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-20133 | 2022/02/16 17:41:53 | 1-04 | 4 | -7.566 | -59.861 | 19065 | 0 | 1 | EV | standard | H2O | 1645058820 | 0.34 | 0.744 | 77.608 | 0.001 | 0.001 | 1.85 | 1303.613 | 2.207 | 3.087 | 117.656 | 0.863 | 0.686 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-20133 | 2022/02/16 17:50:49 | 1-04 | 5 | -7.601 | -59.717 | 19125 | 0 | 1 | EV | standard | H2O | 1645059356 | 0.304 | 0.689 | 76.073 | 0.001 | 0.003 | 1.803 | 1304.001 | 2.276 | 3.129 | 117.41 | 0.862 | 0.558 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-20133 | 2022/02/16 17:59:46 | 1-04 | 6 | -7.555 | -59.525 | 19046 | 0 | 1 | EV | standard | H2O | 1645059893 | 0.308 | 0.675 | 79.074 | 0.001 | 0.002 | 1.875 | 1303.641 | 1.932 | 3.068 | 117.372 | 0.862 | 0.463 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-20133 | 2022/02/16 18:08:43 | 1-04 | 7 | -7.536 | -59.513 | 19054 | 0 | 1 | EV | standard | H2O | 1645060430 | 0.314 | 0.761 | 77.504 | 0.001 | 0.002 | 1.838 | 1303.528 | 2.915 | 3.017 | 114.752 | 0.861 | 0.332 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-20133 | 2022/02/16 18:17:43 | 1-04 | 8 | -7.525 | -59.474 | 19167 | 0 | 1 | EV | standard | H2O | 1645060970 | 0.343 | 0.685 | 81.622 | 0.001 | 0.002 | 1.935 | 1304.275 | 2.476 | 3.067 | 117.059 | 0.862 | 0.375 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-20133 | 2022/02/16 18:26:40 | 1-04 | 9 | -7.594 | -59.541 | 19116 | 0 | 1 | EV | standard | H2O | 1645061507 | 0.331 | 0.694 | 77.331 | 0.001 | 0.001 | 1.839 | 1303.666 | 2.596 | 3.094 | 117.409 | 0.864 | 0.55 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-20133 | 2022/02/16 18:35:37 | 1-04 | 10 | -7.554 | -59.47 | 19078 | 0 | 1 | EV | standard | H2O | 1645062044 | 0.311 | 0.772 | 80.173 | 0.001 | 0.003 | 1.899 | 1303.858 | 2.276 | 3.102 | 115.68 | 0.863 | 0.347 | 0 | 0 | 1 | 46.188 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-20134 | 2022/02/16 18:45:27 | 1-05 | 1 | -6.196 | -80.319 | 18013 | -1 | 1 | CPER3-deep-2 | 21-067 | H2O | 1645062634 | 0.325 | 0.794 | 75.237 | 0 | 0.003 | 1.799 | 1295.988 | 2.172 | 2.868 | 111.11 | 0.86 | 0.374 | 0 | 0 | 1 | 46.312 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-20134 | 2022/02/16 18:54:23 | 1-05 | 2 | -6.132 | -82.009 | 18243 | -1 | 1 | CPER3-deep-2 | 21-067 | H2O | 1645063170 | 0.318 | 0.715 | 75.312 | 0.001 | 0.002 | 1.795 | 1297.212 | 2.425 | 2.961 | 113.367 | 0.862 | 0.859 | 0 | 0 | 1 | 46.375 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-20134 | 2022/02/16 19:03:20 | 1-05 | 3 | -6.095 | -82.352 | 18175 | -1 | 1 | CPER3-deep-2 | 21-067 | H2O | 1645063707 | 0.309 | 0.728 | 76.65 | 0.002 | 0.005 | 1.821 | 1297.045 | 2.441 | 2.891 | 111.321 | 0.866 | 0.355 | 0 | 0 | 1 | 46.438 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-20134 | 2022/02/16 19:12:17 | 1-05 | 4 | -6.078 | -82.589 | 18135 | 0 | 1 | CPER3-deep-2 | 21-067 | H2O | 1645064244 | 0.292 | 0.789 | 76.231 | 0 | 0.002 | 1.814 | 1296.596 | 2.499 | 2.908 | 111.739 | 0.864 | 0.608 | 0 | 0 | 1 | 46.5 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-20135 | 2022/02/16 19:22:06 | 1-06 | 1 | 4.744 | -45.865 | 18428 | -1 | 1 | CPER3-E-2 | 21-067 | H2O | 1645064833 | 0.328 | 0.743 | 75.74 | 0 | 0 | 1.801 | 1302.163 | 1.55 | 2.913 | 112.434 | 0.864 | 0.072 | 0 | 0 | 1 | 46.562 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-20135 | 2022/02/16 19:31:03 | 1-06 | 2 | 5.174 | -42.202 | 18600 | -1 | 1 | CPER3-E-2 | 21-067 | H2O | 1645065370 | 0.322 | 0.754 | 75.22 | 0 | 0 | 1.783 | 1303.136 | 1.76 | 3.009 | 114.456 | 0.863 | 0.404 | 0 | 0 | 1 | 46.562 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-20135 | 2022/02/16 19:40:00 | 1-06 | 3 | 5.271 | -41.453 | 18891 | -1 | 1 | CPER3-E-2 | 21-067 | H2O | 1645065907 | 0.306 | 0.684 | 79.585 | 0 | 0.003 | 1.888 | 1305.246 | 1.69 | 3.058 | 117.005 | 0.867 | 0.509 | 0 | 0 | 1 | 46.562 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-20135 | 2022/02/16 19:48:57 | 1-06 | 4 | 5.34 | -40.91 | 19064 | 0 | 1 | CPER3-E-2 | 21-067 | H2O | 1645066443 | 0.299 | 0.712 | 77.283 | 0 | 0.002 | 1.839 | 1306.645 | 1.61 | 3.071 | 117.385 | 0.869 | 0.36 | 0 | 0 | 1 | 46.562 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-20136 | 2022/02/16 19:58:47 | 1-07 | 1 | 8.89 | -41.445 | 18886 | -1 | 1 | CPER3-D-5 | 21-067 | H2O | 1645067034 | 0.329 | 0.921 | 76.458 | 0 | 0.001 | 1.811 | 1333.215 | -15.623 | 3.783 | 131.119 | 0.863 | -4.797 | 0 | 0 | 1 | 46.625 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-20136 | 2022/02/16 20:07:43 | 1-07 | 2 | 8.776 | -42.437 | 18787 | -1 | 1 | CPER3-D-5 | 21-067 | H2O | 1645067570 | 0.32 | 0.762 | 79.315 | -0.001 | 0.001 | 1.878 | 1333.013 | -14.979 | 3.636 | 126.233 | 0.863 | -5.654 | 0 | 0 | 1 | 46.625 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-20136 | 2022/02/16 20:16:40 | 1-07 | 3 | 8.774 | -42.711 | 18831 | -1 | 1 | CPER3-D-5 | 21-067 | H2O | 1645068107 | 0.302 | 0.765 | 79.272 | 0 | 0.003 | 1.88 | 1333.479 | -15.82 | 3.744 | 128.097 | 0.862 | -5.445 | 0 | 0 | 1 | 46.562 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-20136 | 2022/02/16 20:25:37 | 1-07 | 4 | 8.834 | -43.013 | 18692 | 0 | 1 | CPER3-D-5 | 21-067 | H2O | 1645068644 | 0.328 | 0.896 | 78.618 | 0.001 | 0.001 | 1.87 | 1331.96 | -14.865 | 3.699 | 127.31 | 0.865 | -5.245 | 0 | 0 | 1 | 46.5 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-20137 | 2022/02/16 20:35:24 | 1-08 | 1 | [5 empty] | CPER3-010-1 | 21-067 | [18 empty] | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-20137 | 2022/02/16 20:44:21 | 1-08 | 2 | [5 empty] | CPER3-010-1 | 21-067 | [18 empty] | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-20137 | 2022/02/16 20:53:17 | 1-08 | 3 | 4.45 | -61.353 | 19432 | -1 | 1 | CPER3-010-1 | 21-067 | H2O | 1645070304 | 0.319 | 0.708 | 80.109 | 0 | 0.004 | 1.906 | 1306.382 | 1.516 | 3.126 | 121.299 | 0.869 | 0.623 | 0 | 0 | 1 | 46.188 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-20137 | 2022/02/16 21:02:14 | 1-08 | 4 | 4.442 | -61.364 | 19379 | 0 | 1 | CPER3-010-1 | 21-067 | H2O | 1645070841 | 0.322 | 0.738 | 81.012 | 0 | 0.004 | 1.919 | 1306.143 | 1.992 | 3.114 | 118.414 | 0.863 | 0.307 | 0 | 0 | 1 | 46.188 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-20138 | 2022/02/16 21:12:01 | 1-09 | 1 | 4.984 | -58.028 | 18852 | -1 | 1 | CPER3-010-3 | 21-067 | H2O | 1645071428 | 0.325 | 0.768 | 77.657 | 0.002 | 0.004 | 1.847 | 1302.247 | 2.1 | 2.969 | 114.436 | 0.864 | 0.13 | 0 | 0 | 1 | 46.188 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-20138 | 2022/02/16 21:20:58 | 1-09 | 2 | 5 | -57.964 | 18832 | -1 | 1 | CPER3-010-3 | 21-067 | H2O | 1645071966 | 0.322 | 0.677 | 78.602 | 0.001 | -0.001 | 1.871 | 1301.854 | 2.331 | 3.022 | 114.975 | 0.863 | 0.266 | 0 | 0 | 1 | 46.25 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-20138 | 2022/02/16 21:29:55 | 1-09 | 3 | 5.076 | -57.635 | 19077 | -1 | 1 | CPER3-010-3 | 21-067 | H2O | 1645072502 | 0.296 | 0.745 | 78.339 | 0 | -0.001 | 1.846 | 1303.73 | 1.953 | 3.067 | 116.855 | 0.867 | 0.438 | 0 | 0 | 1 | 46.25 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-20138 | 2022/02/16 21:38:51 | 1-09 | 4 | 5.103 | -57.621 | 19193 | 0 | 1 | CPER3-010-3 | 21-067 | H2O | 1645073038 | 0.306 | 0.723 | 77.792 | 0.001 | 0.003 | 1.842 | 1304.685 | 2.343 | 3.033 | 116.186 | 0.868 | 0.264 | 0 | 0 | 1 | 46.25 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-20139 | 2022/02/16 21:48:39 | 1-10 | 1 | 8.089 | -47.177 | 18389 | -1 | 1 | CPER3-D-2 | 21-067 | H2O | 1645073626 | 0.33 | 0.802 | 75.623 | 0 | 0.001 | 1.808 | 1332.618 | -8.746 | 3.177 | 120.442 | 0.865 | -2.861 | 0 | 0 | 1 | 46.25 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-20139 | 2022/02/16 21:57:36 | 1-10 | 2 | 8.103 | -46.651 | 18533 | -1 | 1 | CPER3-D-2 | 21-067 | H2O | 1645074162 | 0.321 | 0.82 | 76.634 | 0.001 | 0.003 | 1.814 | 1334.359 | -8.605 | 3.215 | 120.563 | 0.868 | -2.822 | 0 | 0 | 1 | 46.25 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-20139 | 2022/02/16 22:06:32 | 1-10 | 3 | 8.072 | -46.727 | 18700 | -1 | 1 | CPER3-D-2 | 21-067 | H2O | 1645074699 | 0.329 | 0.779 | 77.874 | 0 | 0.001 | 1.844 | 1335.805 | -8.84 | 3.24 | 123.406 | 0.868 | -2.775 | 0 | 0 | 1 | 46.25 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-20139 | 2022/02/16 22:15:29 | 1-10 | 4 | 8.093 | -46.304 | 19078 | 0 | 1 | CPER3-D-2 | 21-067 | H2O | 1645075236 | 0.291 | 0.827 | 80.269 | 0.001 | 0.002 | 1.887 | 1339.863 | -9.24 | 3.273 | 123.54 | 0.864 | -3.296 | 0 | 0 | 1 | 46.25 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-20140 | 2022/02/16 22:25:18 | 1-11 | 1 | -2.532 | -72.923 | 18413 | -1 | 1 | CPER3-3040-1 | 21-067 | H2O | 1645075825 | 0.299 | 0.719 | 77.946 | 0.001 | 0.003 | 1.847 | 1299.832 | 2.384 | 2.891 | 112.018 | 0.868 | 0.288 | 0 | 0 | 1 | 46.25 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-20140 | 2022/02/16 22:34:15 | 1-11 | 2 | -2.791 | -74.412 | 18805 | -1 | 1 | CPER3-3040-1 | 21-067 | H2O | 1645076362 | 0.308 | 0.697 | 74.964 | 0.001 | 0.003 | 1.781 | 1302.111 | 2.341 | 2.96 | 115.039 | 0.865 | 0.254 | 0 | 0 | 1 | 46.25 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-20140 | 2022/02/16 22:43:12 | 1-11 | 3 | -2.917 | -74.877 | 18788 | -1 | 1 | CPER3-3040-1 | 21-067 | H2O | 1645076899 | 0.347 | 0.699 | 77.071 | 0 | 0.002 | 1.829 | 1301.978 | 1.998 | 3.037 | 115.171 | 0.867 | 0.446 | 0 | 0 | 1 | 46.188 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-20140 | 2022/02/16 22:52:09 | 1-11 | 4 | -2.963 | -75.135 | 18663 | 0 | 1 | CPER3-3040-1 | 21-067 | H2O | 1645077436 | 0.307 | 0.769 | 78.596 | 0.002 | 0.004 | 1.855 | 1301.024 | 1.846 | 3.016 | 114.721 | 0.866 | 0.407 | 0 | 0 | 1 | 46.25 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-20141 | 2022/02/16 23:01:58 | 1-12 | 1 | 3.914 | -51.903 | 18741 | -1 | 1 | CPER3-B-2 | 21-067 | H2O | 1645078024 | 0.341 | 0.687 | 77.05 | 0.002 | 0.002 | 1.821 | 1330.852 | 0.621 | 2.985 | 114.019 | 0.864 | -0.067 | 0 | 0 | 1 | 46.25 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-20141 | 2022/02/16 23:10:55 | 1-12 | 2 | 4.211 | -49.732 | 18740 | -1 | 1 | CPER3-B-2 | 21-067 | H2O | 1645078562 | 0.312 | 0.767 | 77.968 | 0 | 0.003 | 1.849 | 1330.882 | 1.352 | 2.971 | 113.667 | 0.867 | 0.107 | 0 | 0 | 1 | 46.312 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-20141 | 2022/02/16 23:19:51 | 1-12 | 3 | 4.292 | -49.331 | 18779 | -1 | 1 | CPER3-B-2 | 21-067 | H2O | 1645079098 | 0.27 | 0.8 | 77.409 | 0.001 | 0.001 | 1.84 | 1331.319 | 0.609 | 3.037 | 116.373 | 0.865 | 0.555 | 0 | 0 | 1 | 46.312 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-20141 | 2022/02/16 23:28:48 | 1-12 | 4 | 4.32 | -48.873 | 19181 | 0 | 1 | CPER3-B-2 | 21-067 | H2O | 1645079635 | 0.317 | 0.734 | 79.156 | 0 | 0.002 | 1.881 | 1335.401 | 0.494 | 3.068 | 116.734 | 0.875 | 0.2 | 0 | 0 | 1 | 46.375 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-20142 | 2022/02/16 23:38:37 | 1-13 | 1 | -6.743 | -57.984 | 19102 | -1 | 1 | EV | standard | H2O | 1645080224 | 0.29 | 0.677 | 76.84 | 0.002 | 0 | 1.816 | 1304.47 | 2.049 | 3.164 | 118.378 | 0.875 | 0.743 | 0 | 0 | 1 | 46.438 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-20142 | 2022/02/16 23:47:33 | 1-13 | 2 | -7.196 | -59.147 | 18662 | -1 | 1 | EV | standard | H2O | 1645080760 | 0.291 | 0.712 | 76.337 | 0.002 | 0.003 | 1.812 | 1300.958 | 2.418 | 3.028 | 113.544 | 0.872 | 0.349 | 0 | 0 | 1 | 46.438 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-20142 | 2022/02/16 23:56:30 | 1-13 | 3 | -7.336 | -59.275 | 18741 | -1 | 1 | EV | standard | H2O | 1645081297 | 0.31 | 0.77 | 79.018 | 0.001 | 0.003 | 1.881 | 1301.381 | 2.587 | 2.997 | 113.552 | 0.869 | 0.39 | 0 | 0 | 1 | 46.5 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-20142 | 2022/02/17 00:05:27 | 1-13 | 4 | -7.384 | -59.441 | 18633 | 0 | 1 | EV | standard | H2O | 1645081834 | 0.329 | 0.734 | 78.112 | 0 | 0.003 | 1.853 | 1300.483 | 2.312 | 3.052 | 114.719 | 0.871 | 0.548 | 0 | 0 | 1 | 46.5 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-20143 | 2022/02/17 00:15:16 | 1-14 | 1 | -12.499 | -103.112 | 18787 | -1 | 1 | TEST | TEST | H2O | 1645082422 | 0.295 | 0.705 | 78.429 | 0.001 | 0.002 | 1.865 | 1301.393 | 2.713 | 3.018 | 114.323 | 0.865 | 0.476 | 0 | 0 | 1 | 46.562 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-20143 | 2022/02/17 00:24:13 | 1-14 | 2 | -12.625 | -107.072 | 18706 | -1 | 1 | TEST | TEST | H2O | 1645082960 | 0.303 | 0.735 | 79.176 | 0 | 0.006 | 1.872 | 1300.948 | 2.894 | 2.938 | 113.282 | 0.87 | 0.333 | 0 | 0 | 1 | 46.562 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-20143 | 2022/02/17 00:33:10 | 1-14 | 3 | -12.669 | -108.13 | 18795 | -1 | 1 | TEST | TEST | H2O | 1645083497 | 0.326 | 0.737 | 75.963 | 0.001 | 0.004 | 1.8 | 1301.556 | 2.209 | 3.021 | 115.623 | 0.868 | 0.476 | 0 | 0 | 1 | 46.625 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-20143 | 2022/02/17 00:42:06 | 1-14 | 4 | -12.698 | -108.395 | 18805 | 0 | 1 | TEST | TEST | H2O | 1645084033 | 0.318 | 0.736 | 78.328 | 0.002 | 0.005 | 1.866 | 1301.706 | 2.745 | 2.916 | 113.405 | 0.868 | 0.222 | 0 | 0 | 1 | 46.625 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-20144 | 2022/02/17 00:51:52 | 1-15 | 1 | 3.265 | -54.242 | 18605 | -1 | 1 | CPER3-B-3 | 21-067 | H2O | 1645084619 | 0.29 | 0.755 | 76.707 | 0 | 0.001 | 1.815 | 1368.736 | -0.566 | 2.987 | 113.284 | 0.867 | 0.282 | 0 | 0 | 1 | 46.688 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-20144 | 2022/02/17 01:00:49 | 1-15 | 2 | 3.882 | -49.308 | 18404 | -1 | 1 | CPER3-B-3 | 21-067 | H2O | 1645085156 | 0.321 | 0.808 | 78.251 | -0.001 | 0.002 | 1.862 | 1367.753 | -0.673 | 2.936 | 111.449 | 0.869 | 0.031 | 0 | 0 | 1 | 46.688 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-20144 | 2022/02/17 01:09:46 | 1-15 | 3 | 4.046 | -47.761 | 19017 | -1 | 1 | CPER3-B-3 | 21-067 | H2O | 1645085693 | 0.321 | 0.785 | 78.507 | -0.001 | 0 | 1.861 | 1374.789 | -1.343 | 3.072 | 117.584 | 0.873 | 0.487 | 0 | 0 | 1 | 46.688 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-20144 | 2022/02/17 01:18:42 | 1-15 | 4 | 4.159 | -47.142 | 19072 | 0 | 1 | CPER3-B-3 | 21-067 | H2O | 1645086229 | 0.329 | 0.679 | 78.058 | 0.001 | -0.001 | 1.851 | 1375.811 | -1.156 | 3.116 | 116.969 | 0.872 | 0.214 | 0 | 0 | 1 | 46.688 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-20145 | 2022/02/17 01:28:29 | 1-16 | 1 | -0.247 | -68.126 | 19027 | -1 | 1 | CPER3-010-2 | 21-067 | H2O | 1645086816 | 0.312 | 0.694 | 75.872 | 0.001 | 0.001 | 1.807 | 1305.577 | 1.738 | 3.067 | 117.135 | 0.883 | 0.489 | 0 | 0 | 1 | 46.75 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-20145 | 2022/02/17 01:37:26 | 1-16 | 2 | -0.493 | -70.047 | 18699 | -1 | 1 | CPER3-010-2 | 21-067 | H2O | 1645087352 | 0.3 | 0.808 | 77.412 | 0.001 | 0.006 | 1.83 | 1302.189 | 2.007 | 2.96 | 113.085 | 0.865 | 0.088 | 0 | 0 | 1 | 46.75 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-20145 | 2022/02/17 01:46:22 | 1-16 | 3 | -0.509 | -70.595 | 18648 | -1 | 1 | CPER3-010-2 | 21-067 | H2O | 1645087889 | 0.314 | 0.812 | 77.347 | -0.001 | 0.004 | 1.837 | 1301.063 | 2.071 | 2.992 | 115.353 | 0.866 | 0.579 | 0 | 0 | 1 | 46.75 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-20145 | 2022/02/17 01:55:19 | 1-16 | 4 | -0.471 | -70.801 | 18615 | 0 | 1 | CPER3-010-2 | 21-067 | H2O | 1645088426 | 0.319 | 0.795 | 76.574 | 0.001 | 0.006 | 1.816 | 1300.665 | 2.055 | 3.02 | 115.514 | 0.872 | 0.682 | 0 | 0 | 1 | 46.812 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-20146 | 2022/02/17 02:05:06 | 1-17 | 1 | 0.346 | -66.918 | 18563 | -1 | 1 | CPER3-2030-2 | 21-067 | H2O | 1645089012 | 0.32 | 0.707 | 74.857 | 0.001 | 0.006 | 1.777 | 1300.422 | 2.084 | 2.95 | 112.912 | 0.871 | 0.25 | 0 | 0 | 1 | 46.812 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-20146 | 2022/02/17 02:14:02 | 1-17 | 2 | 0.469 | -66.186 | 18811 | -1 | 1 | CPER3-2030-2 | 21-067 | H2O | 1645089548 | 0.324 | 0.735 | 78.125 | 0 | 0.003 | 1.853 | 1302.378 | 1.92 | 2.982 | 114.443 | 0.878 | 0.223 | 0 | 0 | 1 | 46.75 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-20146 | 2022/02/17 02:22:59 | 1-17 | 3 | 0.466 | -66.156 | 18694 | -1 | 1 | CPER3-2030-2 | 21-067 | H2O | 1645090086 | 0.318 | 0.689 | 76.877 | 0.001 | 0.002 | 1.827 | 1301.195 | 2.153 | 2.981 | 114.427 | 0.866 | 0.357 | 0 | 0 | 1 | 46.75 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-20146 | 2022/02/17 02:31:55 | 1-17 | 4 | 0.485 | -66.303 | 18693 | 0 | 1 | CPER3-2030-2 | 21-067 | H2O | 1645090621 | 0.345 | 0.691 | 75.579 | 0.001 | 0.001 | 1.799 | 1301.187 | 1.814 | 3.051 | 115.709 | 0.873 | 0.443 | 0 | 0 | 1 | 46.812 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-20147 | 2022/02/17 02:41:43 | 1-18 | 1 | 1.948 | -53.718 | 18564 | -1 | 1 | CPER3-B-4 | 21-067 | H2O | 1645091210 | 0.325 | 0.775 | 78.467 | 0.001 | 0.003 | 1.86 | 1380.841 | -0.694 | 2.954 | 112.231 | 0.867 | 0.119 | 0 | 0 | 1 | 46.812 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-20147 | 2022/02/17 02:50:39 | 1-18 | 2 | 2.013 | -52.301 | 18691 | -1 | 1 | CPER3-B-4 | 21-067 | H2O | 1645091746 | 0.314 | 0.776 | 76.695 | 0.001 | 0.003 | 1.813 | 1383.721 | -0.88 | 2.995 | 113.258 | 0.869 | 0.075 | 0 | 0 | 1 | 46.75 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-20147 | 2022/02/17 02:59:35 | 1-18 | 3 | 2.006 | -52.047 | 18713 | -1 | 1 | CPER3-B-4 | 21-067 | H2O | 1645092282 | 0.312 | 0.8 | 78.129 | 0.001 | 0.004 | 1.861 | 1384.152 | -1.094 | 3.044 | 115.071 | 0.89 | 0.406 | 0 | 0 | 1 | 46.75 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-20147 | 2022/02/17 03:08:31 | 1-18 | 4 | 1.995 | -51.79 | 18757 | 0 | 1 | CPER3-B-4 | 21-067 | H2O | 1645092818 | 0.329 | 0.853 | 77.429 | 0.001 | 0 | 1.832 | 1384.71 | -0.894 | 3.007 | 114.665 | 0.879 | 0.362 | 0 | 0 | 1 | 46.812 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-20148 | 2022/02/17 03:18:19 | 1-19 | 1 | 1.725 | -55.926 | 18288 | -1 | 1 | CPER3-E-3 | 21-067 | H2O | 1645093406 | 0.338 | 0.703 | 73.759 | 0 | 0.002 | 1.764 | 1300.52 | 1.766 | 2.904 | 111.537 | 0.88 | 0.1 | 0 | 0 | 1 | 46.75 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-20148 | 2022/02/17 03:27:14 | 1-19 | 2 | 1.803 | -55.874 | 18858 | -1 | 1 | CPER3-E-3 | 21-067 | H2O | 1645093941 | 0.323 | 0.734 | 78.929 | 0 | 0.002 | 1.873 | 1302.852 | 1.944 | 3.045 | 115.962 | 0.872 | 0.491 | 0 | 0 | 1 | 46.75 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-20148 | 2022/02/17 03:36:10 | 1-19 | 3 | 1.666 | -56.391 | 18176 | -1 | 1 | CPER3-E-3 | 21-067 | H2O | 1645094478 | 0.282 | 0.784 | 76.51 | 0.001 | 0.004 | 1.824 | 1297.514 | 2.122 | 2.904 | 111.533 | 0.873 | 0.397 | 0 | 0 | 1 | 46.75 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-20148 | 2022/02/17 03:45:07 | 1-19 | 4 | 1.706 | -55.925 | 18965 | 0 | 1 | CPER3-E-3 | 21-067 | H2O | 1645095014 | 0.319 | 0.659 | 77.407 | 0.001 | 0.004 | 1.836 | 1303.019 | 2.586 | 3.089 | 115.487 | 0.872 | 0.486 | 0 | 0 | 1 | 46.688 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-20149 | 2022/02/17 03:54:55 | 1-20 | 1 | 4.368 | -48.171 | 18065 | -1 | 1 | CPER3-E-4 | 21-067 | H2O | 1645095602 | 0.307 | 0.817 | 77.377 | 0 | 0.002 | 1.84 | 1333.108 | 0.514 | 2.951 | 111.852 | 0.865 | 0.464 | 0 | 0 | 1 | 46.688 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-20149 | 2022/02/17 04:03:51 | 1-20 | 2 | [5 empty] | CPER3-E-4 | 21-067 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-20149 | 2022/02/17 04:13:50 | 1-20 | 3 | [5 empty] | CPER3-E-4 | 21-067 | H2O | [17 empty] | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-20149 | 2022/02/17 04:23:48 | 1-20 | 4 | 4.368 | -48.171 | 18065 | -1 | 1 | CPER3-E-4 | 21-067 | H2O | 1645095602 | 0.307 | 0.817 | 77.377 | 0 | 0.002 | 1.84 | 1333.108 | 0.514 | 2.951 | 111.852 | 0.865 | 0.464 | 0 | 0 | 1 | 46.688 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-20150 | 2022/02/17 04:34:39 | 1-21 | 1 | -0.325 | -65.92 | 17755 | -1 | 1 | CPER3-2030-1 | 21-067 | H2O | 1645097986 | 0.296 | 0.704 | 67.179 | 0.001 | 0 | 1.601 | 1295.177 | 1.839 | 2.838 | 108.223 | 0.874 | 0.266 | 0 | 0 | 1 | 46.938 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-20150 | 2022/02/17 04:43:35 | 1-21 | 2 | -0.33 | -67.019 | 18045 | -1 | 1 | CPER3-2030-1 | 21-067 | H2O | 1645098521 | 0.311 | 0.687 | 71.011 | 0.001 | 0.001 | 1.693 | 1296.323 | 2.232 | 2.883 | 110.465 | 0.887 | 0.428 | 0 | 0 | 1 | 46.75 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-20150 | 2022/02/17 04:52:31 | 1-21 | 3 | -0.303 | -67.199 | 18380 | -1 | 1 | CPER3-2030-1 | 21-067 | H2O | 1645099058 | 0.311 | 0.771 | 71.797 | 0.001 | 0.002 | 1.707 | 1298.545 | 3.006 | 2.923 | 110.654 | 0.875 | 0.19 | 0 | 0 | 1 | 46.688 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-20150 | 2022/02/17 05:01:28 | 1-21 | 4 | -0.341 | -67.129 | 18288 | 0 | 1 | CPER3-2030-1 | 21-067 | H2O | 1645099594 | 0.335 | 0.725 | 72.181 | 0.001 | 0.002 | 1.71 | 1297.895 | 2.128 | 2.925 | 112.248 | 0.868 | 0.52 | 0 | 0 | 1 | 46.75 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-20151 | 2022/02/17 05:11:14 | 1-22 | 1 | -7.033 | -60.341 | 18629 | -1 | 1 | EV | standard | H2O | 1645100180 | 0.305 | 0.841 | 73.387 | 0.001 | 0.003 | 1.739 | 1300.13 | 2.495 | 3.005 | 114.013 | 0.876 | 0.48 | 0 | 0 | 1 | 46.688 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-20151 | 2022/02/17 05:20:10 | 1-22 | 2 | -7.397 | -59.75 | 18343 | -1 | 1 | EV | standard | H2O | 1645100717 | 0.342 | 0.719 | 75.847 | 0.002 | 0.003 | 1.809 | 1298.01 | 2.446 | 2.993 | 111.834 | 0.877 | 0.468 | 0 | 0 | 1 | 46.688 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-20151 | 2022/02/17 05:29:05 | 1-22 | 3 | -7.44 | -59.554 | 18576 | -1 | 1 | EV | standard | H2O | 1645101252 | 0.318 | 0.782 | 76.41 | 0.001 | 0.003 | 1.816 | 1299.647 | 2.571 | 2.947 | 113.488 | 0.869 | 0.581 | 0 | 0 | 1 | 46.75 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-20151 | 2022/02/17 05:38:02 | 1-22 | 4 | -7.471 | -59.752 | 18485 | 0 | 1 | EV | standard | H2O | 1645101789 | 0.344 | 0.709 | 74.893 | 0.001 | 0 | 1.792 | 1298.882 | 2.232 | 2.998 | 114.485 | 0.869 | 0.642 | 0 | 0 | 1 | 46.688 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-20152 | 2022/02/17 05:47:48 | 1-23 | 1 | -12.519 | -103.642 | 18277 | -1 | 1 | TEST | TEST | H2O | 1645102374 | 0.288 | 0.743 | 77.948 | 0.002 | 0.005 | 1.854 | 1297.732 | 2.773 | 2.855 | 110.425 | 0.87 | 0.173 | 0 | 0 | 1 | 46.625 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-20152 | 2022/02/17 05:56:44 | 1-23 | 2 | -12.764 | -107.936 | 18208 | -1 | 1 | TEST | TEST | H2O | 1645102911 | 0.29 | 0.75 | 72.666 | 0.001 | 0.006 | 1.736 | 1296.775 | 2.793 | 2.961 | 112.003 | 0.876 | 0.617 | 0 | 0 | 1 | 46.625 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-20152 | 2022/02/17 06:05:40 | 1-23 | 3 | -12.886 | -108.899 | 18195 | -1 | 1 | TEST | TEST | H2O | 1645103446 | 0.308 | 0.714 | 73.738 | 0.001 | 0.002 | 1.754 | 1296.742 | 2.939 | 2.898 | 111.271 | 0.869 | 0.564 | 0 | 0 | 1 | 46.5 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-20152 | 2022/02/17 06:14:36 | 1-23 | 4 | -12.819 | -108.981 | 18244 | 0 | 1 | TEST | TEST | H2O | 1645103983 | 0.313 | 0.772 | 74.841 | 0.001 | 0.007 | 1.781 | 1297.276 | 3.041 | 2.887 | 110.187 | 0.87 | 0.392 | 0 | 0 | 1 | 46.375 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-20153 | 2022/02/17 06:24:23 | 1-24 | 1 | -3.972 | -74.414 | 18233 | -1 | 1 | CPER3-C-1 | 21-067 | H2O | 1645104570 | 0.283 | 0.711 | 76.553 | 0 | 0.003 | 1.81 | 1297.737 | 2.17 | 2.864 | 110.873 | 0.864 | 0.219 | 0 | 0 | 1 | 46.25 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-20153 | 2022/02/17 06:33:19 | 1-24 | 2 | -3.504 | -70.529 | 18823 | -1 | 1 | CPER3-C-1 | 21-067 | H2O | 1645105106 | 0.302 | 0.732 | 76.676 | 0.001 | 0.004 | 1.824 | 1301.664 | 2.287 | 3.005 | 116.057 | 0.864 | 0.612 | 0 | 0 | 1 | 46.188 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-20153 | 2022/02/17 06:42:15 | 1-24 | 3 | -3.394 | -70.075 | 18891 | -1 | 1 | CPER3-C-1 | 21-067 | H2O | 1645105641 | 0.298 | 0.725 | 76.252 | 0.001 | 0.002 | 1.824 | 1302.209 | 2.372 | 3.099 | 116.134 | 0.864 | 0.577 | 0 | 0 | 1 | 46.188 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-20153 | 2022/02/17 06:51:12 | 1-24 | 4 | -3.378 | -69.562 | 18902 | 0 | 1 | CPER3-C-1 | 21-067 | H2O | 1645106178 | 0.314 | 0.731 | 75.577 | 0 | 0.003 | 1.793 | 1302.683 | 1.982 | 3.04 | 115.172 | 0.87 | 0.261 | 0 | 0 | 1 | 45.875 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-20154 | 2022/02/17 07:00:59 | 1-25 | 1 | 5.929 | -51.829 | 18430 | -1 | 1 | CPER3-D-3 | 21-067 | H2O | 1645106766 | 0.288 | 0.837 | 76.685 | 0 | 0.001 | 1.825 | 1318.072 | -9.033 | 3.261 | 122.722 | 0.865 | -3.081 | 0 | 0 | 1 | 45.938 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-20154 | 2022/02/17 07:09:55 | 1-25 | 2 | 6.241 | -50.424 | 18883 | -1 | 1 | CPER3-D-3 | 21-067 | H2O | 1645107302 | 0.328 | 0.835 | 74.275 | 0 | 0 | 1.764 | 1322.448 | -9.588 | 3.338 | 124.692 | 0.866 | -3.345 | 0 | 0 | 1 | 45.875 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-20154 | 2022/02/17 07:18:51 | 1-25 | 3 | 6.322 | -50.043 | 18841 | -1 | 1 | CPER3-D-3 | 21-067 | H2O | 1645107838 | 0.333 | 0.734 | 78.33 | 0 | 0.002 | 1.862 | 1322.32 | -9.533 | 3.273 | 123.378 | 0.871 | -3.692 | 0 | 0 | 1 | 45.875 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-20154 | 2022/02/17 07:27:48 | 1-25 | 4 | 6.316 | -49.975 | 18822 | 0 | 1 | CPER3-D-3 | 21-067 | H2O | 1645108374 | 0.304 | 0.773 | 78.123 | 0 | 0.004 | 1.847 | 1322.139 | -9.665 | 3.315 | 123.683 | 0.87 | -3.537 | 0 | 0 | 1 | 45.938 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-20155 | 2022/02/17 07:37:35 | 1-26 | 1 | 7.184 | -42.786 | 18555 | -1 | 1 | CPER3-E-5 | 21-067 | H2O | 1645108962 | 0.311 | 0.749 | 75.946 | 0.001 | 0.001 | 1.806 | 1358.193 | -0.533 | 2.991 | 114.794 | 0.874 | 0.43 | 0 | 0 | 1 | 45.938 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-20155 | 2022/02/17 07:46:30 | 1-26 | 2 | 7.312 | -41.321 | 18744 | -1 | 1 | CPER3-E-5 | 21-067 | H2O | 1645109497 | 0.338 | 0.73 | 75.08 | 0 | 0 | 1.779 | 1360.865 | -0.314 | 3.04 | 115.67 | 0.873 | 0.458 | 0 | 0 | 1 | 46 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-20155 | 2022/02/17 07:55:26 | 1-26 | 3 | 7.358 | -40.827 | 18676 | -1 | 1 | CPER3-E-5 | 21-067 | H2O | 1645110033 | 0.317 | 0.676 | 78.503 | 0.001 | 0 | 1.863 | 1360.413 | -0.484 | 3 | 115.155 | 0.867 | 0.486 | 0 | 0 | 1 | 46 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-20155 | 2022/02/17 08:04:23 | 1-26 | 4 | 7.308 | -41.012 | 18596 | 0 | 1 | CPER3-E-5 | 21-067 | H2O | 1645110569 | 0.367 | 0.679 | 75.497 | 0.001 | 0.001 | 1.785 | 1359.306 | -0.592 | 3.024 | 116.068 | 0.869 | 0.554 | 0 | 0 | 1 | 46.125 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-20156 | 2022/02/17 08:14:10 | 1-27 | 1 | 5 | -54.199 | 18710 | -1 | 1 | CPER3-D-1 | 21-067 | H2O | 1645111157 | 0.338 | 0.747 | 75.115 | 0.001 | 0.003 | 1.776 | 1361.341 | -8.05 | 3.089 | 120.446 | 0.869 | -2.45 | 0 | 0 | 1 | 46.188 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-20156 | 2022/02/17 08:23:06 | 1-27 | 2 | 4.742 | -56.096 | 18804 | -1 | 1 | CPER3-D-1 | 21-067 | H2O | 1645111693 | 0.305 | 0.791 | 77.336 | 0 | 0.003 | 1.829 | 1362.153 | -8.721 | 3.127 | 123.497 | 0.87 | -2.303 | 0 | 0 | 1 | 46.25 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-20156 | 2022/02/17 08:32:03 | 1-27 | 3 | 4.766 | -56.634 | 18785 | -1 | 1 | CPER3-D-1 | 21-067 | H2O | 1645112230 | 0.347 | 0.781 | 78.94 | -0.001 | 0.002 | 1.872 | 1361.924 | -8.43 | 3.214 | 123.299 | 0.868 | -2.271 | 0 | 0 | 1 | 46.312 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-20156 | 2022/02/17 08:40:59 | 1-27 | 4 | 4.751 | -56.659 | 18768 | 0 | 1 | CPER3-D-1 | 21-067 | H2O | 1645112765 | 0.303 | 0.757 | 75.658 | 0 | 0.003 | 1.814 | 1362.002 | -8.015 | 3.121 | 120.597 | 0.873 | -2.664 | 0 | 0 | 1 | 46.375 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-20157 | 2022/02/17 08:50:47 | 1-28 | 1 | 3.967 | -53.401 | 18481 | -1 | 1 | CPER3-A-1 | 21-067 | H2O | 1645113354 | 0.313 | 0.811 | 77.305 | 0.001 | 0.005 | 1.845 | 1305.276 | 1.595 | 2.974 | 114.087 | 0.872 | 0.476 | 0 | 0 | 1 | 46.375 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-20157 | 2022/02/17 08:59:43 | 1-28 | 2 | 4.02 | -52.534 | 18620 | -1 | 1 | CPER3-A-1 | 21-067 | H2O | 1645113890 | 0.311 | 0.762 | 73.936 | 0 | -0.001 | 1.761 | 1305.183 | 1.783 | 2.979 | 114.736 | 0.864 | 0.48 | 0 | 0 | 1 | 46.375 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-20157 | 2022/02/17 09:08:39 | 1-28 | 3 | 4.051 | -51.929 | 18617 | -1 | 1 | CPER3-A-1 | 21-067 | H2O | 1645114426 | 0.327 | 0.74 | 76.535 | 0 | 0.001 | 1.815 | 1305.097 | 1.539 | 3.043 | 115.618 | 0.872 | 0.627 | 0 | 0 | 1 | 46.25 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-20157 | 2022/02/17 09:17:35 | 1-28 | 4 | 4.002 | -52.308 | 18451 | 0 | 1 | CPER3-A-1 | 21-067 | H2O | 1645114961 | 0.327 | 0.771 | 74.217 | 0.001 | -0.002 | 1.759 | 1303.735 | 1.674 | 2.972 | 114.012 | 0.867 | 0.473 | 0 | 0 | 1 | 46.188 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-20158 | 2022/02/17 09:27:19 | 1-29 | 1 | 3.506 | -52.566 | 18623 | -1 | 1 | CPER3-D-4 | 21-067 | H2O | 1645115546 | 0.277 | 0.83 | 73.713 | 0 | -0.002 | 1.744 | 1307.988 | -0.861 | 2.948 | 115.544 | 0.864 | -0.686 | 0 | 0 | 1 | 46.062 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-20158 | 2022/02/17 09:36:15 | 1-29 | 2 | 3.362 | -53.279 | 18120 | -1 | 1 | CPER3-D-4 | 21-067 | H2O | 1645116082 | 0.319 | 0.693 | 77.637 | 0 | 0.001 | 1.849 | 1304.252 | -0.915 | 2.897 | 112.707 | 0.864 | -0.748 | 0 | 0 | 1 | 46 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-20158 | 2022/02/17 09:45:11 | 1-29 | 3 | 3.41 | -52.852 | 18649 | -1 | 1 | CPER3-D-4 | 21-067 | H2O | 1645116618 | 0.325 | 0.773 | 78.557 | 0.001 | 0.002 | 1.866 | 1308.183 | -1.361 | 2.953 | 117.166 | 0.871 | -0.562 | 0 | 0 | 1 | 46 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-20158 | 2022/02/17 09:54:07 | 1-29 | 4 | 3.426 | -52.966 | 18673 | 0 | 1 | CPER3-D-4 | 21-067 | H2O | 1645117154 | 0.328 | 0.751 | 75.728 | 0 | 0.003 | 1.803 | 1308.136 | -0.605 | 2.988 | 116.641 | 0.864 | -0.493 | 0 | 0 | 1 | 46 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-20159 | 2022/02/17 10:03:52 | 1-30 | 1 | -2.383 | -69.658 | 18510 | -1 | 1 | CPER3-C-2 | 21-067 | H2O | 1645117739 | 0.297 | 0.726 | 74.84 | 0.001 | 0.003 | 1.771 | 1299.871 | 2.433 | 2.892 | 112.307 | 0.863 | 0.288 | 0 | 0 | 1 | 45.938 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-20159 | 2022/02/17 10:12:49 | 1-30 | 2 | -2.602 | -71.006 | 18631 | -1 | 1 | CPER3-C-2 | 21-067 | H2O | 1645118276 | 0.328 | 0.732 | 78.16 | 0.001 | 0.004 | 1.846 | 1300.536 | 2.455 | 2.961 | 113.226 | 0.865 | 0.306 | 0 | 0 | 1 | 45.938 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-20159 | 2022/02/17 10:21:44 | 1-30 | 3 | -2.642 | -71.316 | 18313 | -1 | 1 | CPER3-C-2 | 21-067 | H2O | 1645118811 | 0.307 | 0.74 | 75.656 | 0.001 | 0.004 | 1.795 | 1298.318 | 1.984 | 2.917 | 112.22 | 0.871 | 0.243 | 0 | 0 | 1 | 45.938 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-20159 | 2022/02/17 10:30:41 | 1-30 | 4 | -2.673 | -71.595 | 18357 | 0 | 1 | CPER3-C-2 | 21-067 | H2O | 1645119347 | 0.31 | 0.732 | 77.186 | 0.001 | 0.003 | 1.826 | 1298.131 | 2.698 | 2.98 | 112.78 | 0.864 | 0.553 | 0 | 0 | 1 | 46 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-20160 | 2022/02/17 10:40:26 | 1-31 | 1 | -7.204 | -61.032 | 18479 | -1 | 1 | EV | standard | H2O | 1645119932 | 0.324 | 0.71 | 76.992 | 0.002 | 0.006 | 1.83 | 1299.242 | 2.098 | 2.954 | 113.019 | 0.87 | 0.417 | 0 | 0 | 1 | 46.062 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-20160 | 2022/02/17 10:49:22 | 1-31 | 2 | -7.389 | -59.685 | 18884 | -1 | 1 | EV | standard | H2O | 1645120469 | 0.352 | 0.712 | 76.689 | 0 | 0.001 | 1.813 | 1301.933 | 2.177 | 3.067 | 116.544 | 0.868 | 0.533 | 0 | 0 | 1 | 46.062 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-20160 | 2022/02/17 10:58:18 | 1-31 | 3 | -7.396 | -59.595 | 18796 | -1 | 1 | EV | standard | H2O | 1645121004 | 0.323 | 0.712 | 74.718 | 0.001 | 0.001 | 1.775 | 1301.434 | 2.02 | 2.994 | 115.185 | 0.871 | 0.398 | 0 | 0 | 1 | 46.188 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-20160 | 2022/02/17 11:07:14 | 1-31 | 4 | -7.416 | -59.429 | 18823 | 0 | 1 | EV | standard | H2O | 1645121541 | 0.317 | 0.752 | 77.316 | 0 | 0 | 1.844 | 1301.501 | 2.277 | 3.028 | 115.445 | 0.867 | 0.634 | 0 | 0 | 1 | 46.25 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-20161 | 2022/02/17 11:17:00 | 1-32 | 1 | -12.545 | -103.697 | 18253 | -1 | 1 | TEST | TEST | H2O | 1645122126 | 0.329 | 0.739 | 74.612 | 0.002 | 0.004 | 1.78 | 1297.432 | 2.547 | 2.888 | 110.727 | 0.878 | 0.273 | 0 | 0 | 1 | 46.375 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-20161 | 2022/02/17 11:25:56 | 1-32 | 2 | -12.785 | -107.872 | 18390 | -1 | 1 | TEST | TEST | H2O | 1645122663 | 0.307 | 0.803 | 75.226 | 0.001 | 0.005 | 1.778 | 1298.327 | 2.611 | 2.929 | 112.015 | 0.865 | 0.496 | 0 | 0 | 1 | 46.562 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-20161 | 2022/02/17 11:34:52 | 1-32 | 3 | -12.853 | -108.706 | 18296 | -1 | 1 | TEST | TEST | H2O | 1645123199 | 0.299 | 0.76 | 77.888 | 0 | 0.006 | 1.853 | 1297.627 | 2.688 | 2.932 | 111.36 | 0.869 | 0.322 | 0 | 0 | 1 | 46.562 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-20161 | 2022/02/17 11:43:48 | 1-32 | 4 | -12.854 | -109.112 | 18387 | 0 | 1 | TEST | TEST | H2O | 1645123735 | 0.298 | 0.768 | 75.846 | 0 | 0.005 | 1.805 | 1298.301 | 2.554 | 2.925 | 112.269 | 0.864 | 0.563 | 0 | 0 | 1 | 46.562 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-20162 | 2022/02/17 11:53:34 | 1-33 | 1 | 0.103 | -66.249 | 18722 | -1 | 1 | CPER3-A-2 | 21-067 | H2O | 1645124320 | 0.353 | 0.771 | 77.463 | 0.001 | 0.002 | 1.834 | 1342.643 | 0.496 | 3.039 | 116.06 | 0.868 | 0.645 | 0 | 0 | 1 | 46.5 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-20162 | 2022/02/17 12:02:30 | 1-33 | 2 | 0.612 | -62.783 | 18076 | -1 | 1 | CPER3-A-2 | 21-067 | H2O | 1645124856 | 0.343 | 0.731 | 75.066 | 0.001 | 0 | 1.785 | 1337.706 | 0.769 | 2.874 | 109.976 | 0.864 | 0.182 | 0 | 0 | 1 | 46.375 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-20162 | 2022/02/17 12:11:26 | 1-33 | 3 | 0.781 | -61.66 | 18197 | -1 | 1 | CPER3-A-2 | 21-067 | H2O | 1645125393 | 0.289 | 0.714 | 78.483 | 0 | 0.001 | 1.856 | 1338.798 | 0.27 | 2.982 | 112.671 | 0.864 | 0.671 | 0 | 0 | 1 | 46.375 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-20162 | 2022/02/17 12:20:23 | 1-33 | 4 | 0.826 | -61.366 | 18326 | 0 | 1 | CPER3-A-2 | 21-067 | H2O | 1645125929 | 0.31 | 0.786 | 75.873 | 0 | 0 | 1.814 | 1339.976 | 0.629 | 2.878 | 112.231 | 0.865 | 0.237 | 0 | 0 | 1 | 46.375 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-20163 | 2022/02/17 12:30:09 | 1-34 | 1 | 0.42 | -63.033 | 18366 | -1 | 1 | CPER3-A-3 | 21-067 | H2O | 1645126516 | 0.313 | 0.746 | 77.178 | 0 | 0.001 | 1.827 | 1346.712 | 0.945 | 2.919 | 111.815 | 0.863 | 0.276 | 0 | 0 | 1 | 46.312 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-20163 | 2022/02/17 12:39:05 | 1-34 | 2 | 0.417 | -63.324 | 18091 | -1 | 1 | CPER3-A-3 | 21-067 | H2O | 1645127052 | 0.337 | 0.693 | 74.89 | 0.001 | 0.003 | 1.781 | 1344.161 | 0.242 | 2.879 | 111.636 | 0.864 | 0.346 | 0 | 0 | 1 | 46.312 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-20163 | 2022/02/17 12:48:01 | 1-34 | 3 | 0.459 | -63.099 | 18442 | -1 | 1 | CPER3-A-3 | 21-067 | H2O | 1645127588 | 0.318 | 0.73 | 74.881 | 0.001 | 0.004 | 1.775 | 1347.871 | 0.103 | 2.944 | 112.885 | 0.877 | 0.331 | 0 | 0 | 1 | 46.375 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-20163 | 2022/02/17 12:56:57 | 1-34 | 4 | 0.45 | -62.766 | 18520 | 0 | 1 | CPER3-A-3 | 21-067 | H2O | 1645128124 | 0.316 | 0.719 | 77.655 | 0.002 | 0.001 | 1.831 | 1348.455 | -0.006 | 3.001 | 114.161 | 0.866 | 0.312 | 0 | 0 | 1 | 46.375 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-20164 | 2022/02/17 13:06:44 | 1-35 | 1 | [5 empty] | CPER3-A-4 | 21-067 | [18 empty] | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-20164 | 2022/02/17 13:15:41 | 1-35 | 2 | -0.074 | -60.203 | 18506 | -1 | 1 | CPER3-A-4 | 21-067 | H2O | 1645129248 | 0.324 | 0.735 | 75.803 | 0.001 | 0.003 | 1.805 | 1300.14 | 2.205 | 2.993 | 113.983 | 0.864 | 0.466 | 0 | 0 | 1 | 46.562 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-20164 | 2022/02/17 13:24:38 | 1-35 | 3 | -0.093 | -60.328 | 18371 | -1 | 1 | CPER3-A-4 | 21-067 | H2O | 1645129784 | 0.318 | 0.753 | 75.353 | 0 | 0 | 1.793 | 1299.163 | 2.394 | 2.902 | 111.882 | 0.864 | 0.398 | 0 | 0 | 1 | 46.562 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-20164 | 2022/02/17 13:33:34 | 1-35 | 4 | -0.093 | -60.328 | 18371 | -1 | 1 | CPER3-A-4 | 21-067 | H2O | 1645129784 | 0.318 | 0.753 | 75.353 | 0 | 0 | 1.793 | 1299.163 | 2.394 | 2.902 | 111.882 | 0.864 | 0.398 | 0 | 0 | 1 | 46.562 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-20165 | 2022/02/17 13:44:18 | 1-36 | 1 | [5 empty] | CPER3-A-5 | 21-067 | [18 empty] | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-20165 | 2022/02/17 13:53:15 | 1-36 | 2 | [5 empty] | CPER3-A-5 | 21-067 | [18 empty] | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-20165 | 2022/02/17 14:02:10 | 1-36 | 3 | -3.007 | -65.253 | 18741 | -1 | 1 | CPER3-A-5 | 21-067 | H2O | 1645132037 | 0.293 | 0.663 | 75.746 | 0.001 | 0.003 | 1.798 | 1306.035 | 1.294 | 2.995 | 115.1 | 0.872 | 0.282 | 0 | 0 | 1 | 46.688 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-20165 | 2022/02/17 14:11:07 | 1-36 | 4 | -3.021 | -65.297 | 19175 | 0 | 1 | CPER3-A-5 | 21-067 | H2O | 1645132573 | 0.31 | 0.708 | 74.05 | 0.001 | 0.002 | 1.753 | 1308.784 | 1.844 | 3.115 | 119.042 | 0.868 | 0.536 | 0 | 0 | 1 | 46.688 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-20166 | 2022/02/17 14:20:50 | 1-37 | 1 | -0.104 | -61.974 | 18650 | -1 | 1 | CPER3-B-1 | 21-067 | H2O | 1645133156 | 0.309 | 0.694 | 76.659 | 0.002 | 0.002 | 1.811 | 1321.896 | 1.008 | 2.969 | 114.735 | 0.869 | 0.311 | 0 | 0 | 1 | 46.75 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-20166 | 2022/02/17 14:29:46 | 1-37 | 2 | 0.012 | -61.4 | 18620 | -1 | 1 | CPER3-B-1 | 21-067 | H2O | 1645133692 | 0.297 | 0.654 | 76.452 | 0.001 | 0.002 | 1.82 | 1322.05 | 1.442 | 2.952 | 112.998 | 0.868 | 0.212 | 0 | 0 | 1 | 46.75 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-20166 | 2022/02/17 14:38:41 | 1-37 | 3 | 0.017 | -61.604 | 18334 | -1 | 1 | CPER3-B-1 | 21-067 | H2O | 1645134227 | 0.338 | 0.702 | 74.58 | 0 | 0.003 | 1.777 | 1319.478 | 1.33 | 2.973 | 112.811 | 0.873 | 0.463 | 0 | 0 | 1 | 46.688 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-20166 | 2022/02/17 14:47:37 | 1-37 | 4 | 0.033 | -61.506 | 18472 | 0 | 1 | CPER3-B-1 | 21-067 | H2O | 1645134764 | 0.318 | 0.733 | 76.779 | 0.001 | 0.003 | 1.814 | 1320.56 | 1.632 | 2.981 | 113.107 | 0.865 | 0.508 | 0 | 0 | 1 | 46.75 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-20167 | 2022/02/17 14:57:21 | 1-38 | 1 | 2.405 | -56.983 | 17983 | -1 | 1 | CPER3-B-5 | 21-067 | H2O | 1645135348 | 0.368 | 0.866 | 75.267 | 0.001 | 0.006 | 1.786 | 1504.964 | -6.538 | 2.871 | 109.811 | 0.881 | 0 | 0 | 0 | 1 | 46.812 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-20167 | 2022/02/17 15:06:17 | 1-38 | 2 | 2.642 | -55.378 | 19225 | -1 | 1 | CPER3-B-5 | 21-067 | H2O | 1645135883 | 0.318 | 0.731 | 75.736 | 0 | 0 | 1.795 | 1530.553 | -6.387 | 3.111 | 116.319 | 0.884 | 0.076 | 0 | 0 | 1 | 46.75 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-20167 | 2022/02/17 15:15:13 | 1-38 | 3 | 2.457 | -55.97 | 18437 | -1 | 1 | CPER3-B-5 | 21-067 | H2O | 1645136419 | 0.367 | 0.798 | 73.99 | 0.001 | 0.003 | 1.765 | 1517.874 | -7.242 | 2.999 | 114.211 | 0.867 | 0.17 | 0 | 0 | 1 | 46.75 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-20167 | 2022/02/17 15:24:09 | 1-38 | 4 | 2.475 | -56.064 | 18362 | 0 | 1 | CPER3-B-5 | 21-067 | H2O | 1645136955 | 0.318 | 0.83 | 75.658 | 0.002 | 0.005 | 1.805 | 1516.534 | -6.132 | 2.932 | 113.106 | 0.872 | 0.372 | 0 | 0 | 1 | 46.625 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-20168 | 2022/02/17 15:33:53 | 1-39 | 1 | [5 empty] | CPER3-E-1 | 21-067 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-20168 | 2022/02/17 15:43:51 | 1-39 | 2 | [5 empty] | CPER3-E-1 | 21-067 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-20168 | 2022/02/17 15:53:49 | 1-39 | 3 | [5 empty] | CPER3-E-1 | 21-067 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-20168 | 2022/02/17 16:03:47 | 1-39 | 4 | [5 empty] | CPER3-E-1 | 21-067 | H2O | [17 empty] | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-20169 | 2022/02/17 16:14:33 | 1-40 | 1 | -7.169 | -59.151 | 18650 | -1 | 1 | EV | standard | H2O | 1645139980 | 0.307 | 0.759 | 68.426 | 0.001 | 0.001 | 1.623 | 1305.392 | 2.151 | 3.003 | 114.547 | 0.866 | 0.441 | 0 | 0 | 1 | 46.812 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-20169 | 2022/02/17 16:23:29 | 1-40 | 2 | -7.331 | -59.429 | 18638 | -1 | 1 | EV | standard | H2O | 1645140516 | 0.338 | 0.779 | 71.586 | 0.001 | 0 | 1.696 | 1301.297 | 3.134 | 2.974 | 112.552 | 0.876 | 0.336 | 0 | 0 | 1 | 46.5 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-20169 | 2022/02/17 16:32:25 | 1-40 | 3 | -7.443 | -59.743 | 18642 | -1 | 1 | EV | standard | H2O | 1645141051 | 0.328 | 0.795 | 77.336 | 0.002 | 0.005 | 1.841 | 1301.153 | 2.33 | 2.943 | 113.461 | 0.874 | 0.178 | 0 | 0 | 1 | 46.375 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-20169 | 2022/02/17 16:41:21 | 1-40 | 4 | -7.498 | -59.395 | 18615 | 0 | 1 | EV | standard | H2O | 1645141588 | 0.302 | 0.625 | 71.64 | 0.001 | 0.001 | 1.702 | 1300.535 | 2.298 | 2.983 | 114.623 | 0.865 | 0.495 | 0 | 0 | 1 | 46.312 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-20170 | 2022/02/17 16:51:05 | 1-41 | 1 | -12.573 | -103.567 | 18271 | -1 | 1 | TEST | TEST | H2O | 1645142172 | 0.299 | 0.746 | 73.153 | 0.002 | 0.006 | 1.745 | 1297.925 | 2.661 | 2.885 | 111.031 | 0.872 | 0.291 | 0 | 0 | 1 | 46.25 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-20170 | 2022/02/17 17:00:01 | 1-41 | 2 | -12.682 | -107.336 | 18588 | -1 | 1 | TEST | TEST | H2O | 1645142707 | 0.315 | 0.736 | 73.517 | 0.002 | 0.006 | 1.747 | 1300.201 | 2.875 | 2.951 | 112.91 | 0.869 | 0.41 | 0 | 0 | 1 | 46.375 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-20170 | 2022/02/17 17:08:57 | 1-41 | 3 | -12.853 | -108.836 | 18385 | -1 | 1 | TEST | TEST | H2O | 1645143244 | 0.323 | 0.796 | 74.325 | 0.001 | 0.009 | 1.757 | 1298.629 | 2.731 | 2.902 | 111.981 | 0.884 | 0.342 | 0 | 0 | 1 | 46.438 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-20170 | 2022/02/17 17:17:52 | 1-41 | 4 | -12.854 | -108.92 | 18487 | 0 | 1 | TEST | TEST | H2O | 1645143780 | 0.3 | 0.708 | 73.389 | 0 | 0.003 | 1.752 | 1299.183 | 2.344 | 2.98 | 114.131 | 0.871 | 0.616 | 0 | 0 | 1 | 46.438 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-20171 | 2022/02/17 17:27:37 | 1-42 | 1 | -6.366 | -80.386 | 18600 | -1 | 1 | CPER3-deep-3 | 21-067 | H2O | 1645144364 | 0.319 | 0.729 | 74.666 | 0 | 0.005 | 1.778 | 1300.28 | 2.443 | 2.971 | 114.372 | 0.873 | 0.544 | 0 | 0 | 1 | 46.562 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-20171 | 2022/02/17 17:36:34 | 1-42 | 2 | -6.223 | -77.926 | 18247 | -1 | 1 | CPER3-deep-3 | 21-067 | H2O | 1645144901 | 0.304 | 0.771 | 75.383 | 0 | 0.004 | 1.796 | 1297.658 | 2.427 | 2.883 | 112.053 | 0.869 | 0.519 | 0 | 0 | 1 | 46.625 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-20171 | 2022/02/17 17:45:30 | 1-42 | 3 | -6.111 | -77.384 | 18409 | -1 | 1 | CPER3-deep-3 | 21-067 | H2O | 1645145437 | 0.295 | 0.716 | 75.688 | 0.001 | 0.004 | 1.793 | 1298.895 | 2.366 | 2.912 | 112.828 | 0.868 | 0.491 | 0 | 0 | 1 | 46.625 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-20171 | 2022/02/17 17:54:25 | 1-42 | 4 | -6.174 | -77.3 | 18140 | 0 | 1 | CPER3-deep-3 | 21-067 | H2O | 1645145971 | 0.321 | 0.686 | 72.377 | 0.001 | 0 | 1.726 | 1296.678 | 2.358 | 2.929 | 111.839 | 0.867 | 0.515 | 0 | 0 | 1 | 46.625 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-20172 | 2022/02/17 18:04:09 | 1-43 | 1 | -5.932 | -77.765 | 18492 | -1 | 1 | CPER3-3040-3 | 21-067 | H2O | 1645146555 | 0.29 | 0.804 | 77.839 | 0 | 0.006 | 1.844 | 1299.101 | 2.352 | 2.963 | 114.067 | 0.868 | 0.601 | 0 | 0 | 1 | 46.562 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-20172 | 2022/02/17 18:13:04 | 1-43 | 2 | -5.95 | -77.989 | 18566 | -1 | 1 | CPER3-3040-3 | 21-067 | H2O | 1645147091 | 0.287 | 0.738 | 76.109 | 0.001 | 0.005 | 1.796 | 1299.66 | 2.4 | 2.985 | 114.092 | 0.866 | 0.549 | 0 | 0 | 1 | 46.438 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-20172 | 2022/02/17 18:22:00 | 1-43 | 3 | -5.848 | -77.685 | 18732 | -1 | 1 | CPER3-3040-3 | 21-067 | H2O | 1645147627 | 0.317 | 0.669 | 75.559 | 0.001 | 0.003 | 1.796 | 1301.053 | 2.414 | 3.009 | 114.7 | 0.865 | 0.409 | 0 | 0 | 1 | 46.375 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-20172 | 2022/02/17 18:30:56 | 1-43 | 4 | -5.878 | -77.785 | 18669 | 0 | 1 | CPER3-3040-3 | 21-067 | H2O | 1645148163 | 0.3 | 0.737 | 79.271 | 0.001 | 0.004 | 1.869 | 1300.309 | 2.563 | 3.013 | 114.841 | 0.868 | 0.571 | 0 | 0 | 1 | 46.375 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-20173 | 2022/02/17 18:40:40 | 1-44 | 1 | -4.083 | -77.709 | 18474 | -1 | 1 | CPER3-2030-3 | 21-067 | H2O | 1645148747 | 0.298 | 0.718 | 78.784 | 0.001 | 0.003 | 1.87 | 1299.076 | 2.226 | 2.952 | 113.483 | 0.867 | 0.397 | 0 | 0 | 1 | 46.438 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-20173 | 2022/02/17 18:49:36 | 1-44 | 2 | -3.996 | -77.567 | 18501 | -1 | 1 | CPER3-2030-3 | 21-067 | H2O | 1645149283 | 0.319 | 0.8 | 75.75 | 0.001 | 0.003 | 1.804 | 1299.174 | 2.823 | 2.939 | 112.116 | 0.864 | 0.251 | 0 | 0 | 1 | 46.438 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-20173 | 2022/02/17 18:58:31 | 1-44 | 3 | -4.041 | -77.706 | 18659 | -1 | 1 | CPER3-2030-3 | 21-067 | H2O | 1645149818 | 0.32 | 0.671 | 75.656 | 0.001 | 0.001 | 1.8 | 1300.075 | 2.563 | 3.002 | 114.699 | 0.869 | 0.579 | 0 | 0 | 1 | 46.438 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-20173 | 2022/02/17 19:07:27 | 1-44 | 4 | -3.985 | -77.453 | 18533 | 0 | 1 | CPER3-2030-3 | 21-067 | H2O | 1645150354 | 0.317 | 0.759 | 75.893 | 0.001 | 0.003 | 1.803 | 1299.32 | 2.377 | 2.942 | 113.381 | 0.865 | 0.385 | 0 | 0 | 1 | 46.375 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-20174 | 2022/02/17 19:17:12 | 1-45 | 1 | -3.449 | -73.41 | 18573 | -1 | 1 | CPER3-C-5 | 21-067 | H2O | 1645150939 | 0.317 | 0.686 | 76.638 | 0.001 | 0.002 | 1.817 | 1300.104 | 2.092 | 2.985 | 114.467 | 0.877 | 0.479 | 0 | 0 | 1 | 46.375 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-20174 | 2022/02/17 19:26:08 | 1-45 | 2 | -3.483 | -73.147 | 18776 | -1 | 1 | CPER3-C-5 | 21-067 | H2O | 1645151475 | 0.288 | 0.753 | 76.158 | 0.001 | 0.003 | 1.808 | 1301.549 | 2.047 | 2.989 | 115.191 | 0.866 | 0.461 | 0 | 0 | 1 | 46.438 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-20174 | 2022/02/17 19:35:04 | 1-45 | 3 | -3.428 | -73.091 | 18778 | -1 | 1 | CPER3-C-5 | 21-067 | H2O | 1645152010 | 0.316 | 0.775 | 75.718 | 0.001 | 0.005 | 1.79 | 1301.251 | 2.805 | 3.026 | 114.866 | 0.864 | 0.633 | 0 | 0 | 1 | 46.5 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-20174 | 2022/02/17 19:43:59 | 1-45 | 4 | -3.431 | -72.847 | 18696 | 0 | 1 | CPER3-C-5 | 21-067 | H2O | 1645152546 | 0.296 | 0.644 | 75.204 | 0.001 | 0.002 | 1.788 | 1301.069 | 2.302 | 2.941 | 113.914 | 0.87 | 0.466 | 0 | 0 | 1 | 46.5 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-20175 | 2022/02/17 19:53:45 | 1-46 | 1 | -3.44 | -73.008 | 18299 | -1 | 1 | CPER3-C-4 | 21-067 | H2O | 1645153132 | 0.316 | 0.779 | 75.764 | 0 | 0.001 | 1.791 | 1385.94 | -1.321 | 2.872 | 110.321 | 0.87 | -0.107 | 0 | 0 | 1 | 46.438 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-20175 | 2022/02/17 20:02:41 | 1-46 | 2 | -3.44 | -73.113 | 18655 | -1 | 1 | CPER3-C-4 | 21-067 | H2O | 1645153668 | 0.34 | 0.786 | 75.249 | 0.001 | 0.002 | 1.797 | 1392.356 | -1.793 | 2.99 | 113.939 | 0.871 | 0.039 | 0 | 0 | 1 | 46.5 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-20175 | 2022/02/17 20:11:36 | 1-46 | 3 | -3.457 | -73.174 | 19018 | -1 | 1 | CPER3-C-4 | 21-067 | H2O | 1645154203 | 0.332 | 0.819 | 78.795 | 0.001 | 0.004 | 1.877 | 1396.616 | -0.879 | 3.034 | 115.29 | 0.878 | 0.13 | 0 | 0 | 1 | 46.5 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-20175 | 2022/02/17 20:20:32 | 1-46 | 4 | -3.446 | -73.131 | 18657 | 0 | 1 | CPER3-C-4 | 21-067 | H2O | 1645154739 | 0.316 | 0.759 | 76.189 | 0.001 | 0 | 1.822 | 1392.8 | -1.592 | 2.963 | 113.292 | 0.866 | 0.096 | 0 | 0 | 1 | 46.562 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-20176 | 2022/02/17 20:30:18 | 1-47 | 1 | 0.573 | -67.626 | 18560 | -1 | 1 | CPER3-1020-3 | 21-067 | H2O | 1645155324 | 0.303 | 0.692 | 77.703 | 0 | 0.002 | 1.849 | 1302.374 | 2.082 | 2.983 | 114.183 | 0.869 | 0.467 | 0 | 0 | 1 | 46.562 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-20176 | 2022/02/17 20:39:13 | 1-47 | 2 | 0.81 | -66.93 | 18811 | -1 | 1 | CPER3-1020-3 | 21-067 | H2O | 1645155860 | 0.296 | 0.769 | 75.881 | 0.001 | 0.001 | 1.795 | 1302.065 | 2.391 | 3.028 | 115.53 | 0.864 | 0.55 | 0 | 0 | 1 | 46.625 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-20176 | 2022/02/17 20:48:09 | 1-47 | 3 | 0.803 | -66.808 | 18799 | -1 | 1 | CPER3-1020-3 | 21-067 | H2O | 1645156396 | 0.303 | 0.744 | 76.781 | 0.001 | 0 | 1.821 | 1301.832 | 2.043 | 3.031 | 115.351 | 0.871 | 0.459 | 0 | 0 | 1 | 46.562 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-20176 | 2022/02/17 20:57:05 | 1-47 | 4 | 0.863 | -66.751 | 18766 | 0 | 1 | CPER3-1020-3 | 21-067 | H2O | 1645156931 | 0.317 | 0.676 | 73.813 | 0 | 0.002 | 1.752 | 1301.51 | 2.363 | 2.999 | 114.203 | 0.876 | 0.455 | 0 | 0 | 1 | 46.5 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-20177 | 2022/02/17 21:06:52 | 1-48 | 1 | -5.02 | -81.611 | 17808 | -1 | 1 | CPER3-3040-2 | 21-067 | H2O | 1645157519 | 0.33 | 0.765 | 73.248 | 0.003 | 0.002 | 1.736 | 1294.189 | 2.807 | 2.82 | 108.542 | 0.862 | 0.474 | 0 | 0 | 1 | 46.438 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-20177 | 2022/02/17 21:15:48 | 1-48 | 2 | -5.282 | -83.054 | 17868 | -1 | 1 | CPER3-3040-2 | 21-067 | H2O | 1645158054 | 0.328 | 0.732 | 74.279 | 0 | 0.006 | 1.757 | 1294.786 | 2.701 | 2.822 | 108.06 | 0.864 | 0.247 | 0 | 0 | 1 | 46.375 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-20177 | 2022/02/17 21:24:44 | 1-48 | 3 | -5.293 | -83.268 | 17934 | -1 | 1 | CPER3-3040-2 | 21-067 | H2O | 1645158591 | 0.296 | 0.715 | 73.938 | 0.001 | 0.004 | 1.762 | 1295.264 | 2.516 | 2.831 | 109.448 | 0.863 | 0.388 | 0 | 0 | 1 | 46.312 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-20177 | 2022/02/17 21:33:40 | 1-48 | 4 | -5.355 | -83.159 | 17882 | 0 | 1 | CPER3-3040-2 | 21-067 | H2O | 1645159127 | 0.314 | 0.713 | 74.146 | 0 | 0.005 | 1.77 | 1294.924 | 2.643 | 2.843 | 108.241 | 0.861 | 0.303 | 0 | 0 | 1 | 46.25 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-20178 | 2022/02/17 21:43:26 | 1-49 | 1 | -7.31 | -62.167 | 18853 | -1 | 1 | EV | standard | H2O | 1645159713 | 0.295 | 0.695 | 77.276 | 0.002 | 0.003 | 1.838 | 1301.54 | 3.028 | 3.027 | 114.943 | 0.869 | 0.486 | 0 | 0 | 1 | 46.312 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-20178 | 2022/02/17 21:52:22 | 1-49 | 2 | -7.459 | -60.596 | 18606 | -1 | 1 | EV | standard | H2O | 1645160248 | 0.305 | 0.778 | 76.608 | 0.001 | 0.002 | 1.818 | 1300.085 | 2.356 | 2.971 | 113.283 | 0.866 | 0.34 | 0 | 0 | 1 | 46.25 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-20178 | 2022/02/17 22:01:18 | 1-49 | 3 | -7.466 | -59.844 | 18486 | -1 | 1 | EV | standard | H2O | 1645160785 | 0.281 | 0.73 | 75.98 | 0.001 | 0.003 | 1.796 | 1299.017 | 2.87 | 2.934 | 112.205 | 0.863 | 0.444 | 0 | 0 | 1 | 46.125 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-20178 | 2022/02/17 22:10:13 | 1-49 | 4 | -7.491 | -59.945 | 18660 | 0 | 1 | EV | standard | H2O | 1645161320 | 0.329 | 0.712 | 74.8 | 0 | -0.001 | 1.784 | 1300.17 | 2.9 | 2.946 | 113.31 | 0.862 | 0.507 | 0 | 0 | 1 | 46.062 | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-20179 | 2022/02/17 22:19:58 | 1-50 | 1 | -7.494 | -60.037 | 18631 | -1 | 1 | TEST | TEST | H2O | 1645161905 | 0.305 | 0.67 | 76.558 | 0.002 | 0.002 | 1.809 | 1300.005 | 2.692 | 2.967 | 113.522 | 0.864 | 0.435 | 0 | 0 | 1 | 46 | 1 | 50 | 1 | GJB_Rinse | 0
222 | P-20179 | 2022/02/17 22:28:54 | 1-50 | 2 | -7.461 | -59.711 | 18630 | -1 | 1 | TEST | TEST | H2O | 1645162441 | 0.319 | 0.752 | 76.358 | 0.002 | 0.001 | 1.808 | 1299.994 | 2.689 | 2.966 | 113.589 | 0.863 | 0.412 | 0 | 0 | 1 | 46.062 | 1 | 50 | 1 | GJB_Rinse | 0
223 | P-20179 | 2022/02/17 22:37:50 | 1-50 | 3 | -7.506 | -59.701 | 18479 | -1 | 1 | TEST | TEST | H2O | 1645162977 | 0.32 | 0.772 | 76.985 | 0.001 | 0.004 | 1.819 | 1299.184 | 2.295 | 2.947 | 112.617 | 0.864 | 0.295 | 0 | 0 | 1 | 46.062 | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-20179 | 2022/02/17 22:46:46 | 1-50 | 4 | -7.506 | -59.673 | 18751 | 0 | 1 | TEST | TEST | H2O | 1645163512 | 0.269 | 0.726 | 75.001 | 0 | 0.002 | 1.773 | 1301.181 | 2.106 | 3.02 | 114.418 | 0.868 | 0.397 | 0 | 0 | 1 | 46.125 | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-20180 | 2022/02/17 22:56:31 | 1-51 | 1 | 0.595 | -57.996 | 18444 | -1 | 1 | CPER3-C-3 | 21-067 | H2O | 1645164098 | 0.335 | 0.742 | 73.3 | 0.001 | 0 | 1.729 | 1299.031 | 2.076 | 2.992 | 113.922 | 0.866 | 0.616 | 0 | 0 | 1 | 46.188 | 1 | 51 | 1 | GJB_Rinse | 0
226 | P-20180 | 2022/02/17 23:05:27 | 1-51 | 2 | 1.003 | -57.566 | 18720 | -1 | 1 | CPER3-C-3 | 21-067 | H2O | 1645164634 | 0.294 | 0.803 | 78.079 | 0 | 0.002 | 1.845 | 1301.248 | 1.804 | 2.977 | 115.407 | 0.864 | 0.324 | 0 | 0 | 1 | 46.125 | 1 | 51 | 1 | GJB_Rinse | 0
227 | P-20180 | 2022/02/17 23:14:23 | 1-51 | 3 | 1.053 | -57.636 | 18893 | -1 | 1 | CPER3-C-3 | 21-067 | H2O | 1645165169 | 0.328 | 0.72 | 75.915 | 0 | 0.003 | 1.801 | 1302.591 | 2.162 | 2.933 | 114.418 | 0.866 | 0.159 | 0 | 0 | 1 | 46.125 | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-20180 | 2022/02/17 23:23:18 | 1-51 | 4 | 1.096 | -57.473 | 18800 | 0 | 1 | CPER3-C-3 | 21-067 | H2O | 1645165705 | 0.324 | 0.712 | 78.041 | 0 | 0.002 | 1.857 | 1301.618 | 1.91 | 3.035 | 116.668 | 0.864 | 0.653 | 0 | 0 | 1 | 46.188 | 1 | 51 | 1 | GJB_Rinse | 0
229 | P-20181 | 2022/02/17 23:33:04 | 1-52 | 1 | -6.204 | -81.014 | 18340 | -1 | 1 | CPER3-deep-1 | 21-067 | H2O | 1645166291 | 0.31 | 0.738 | 76.937 | 0 | 0.001 | 1.822 | 1298.152 | 2.569 | 2.887 | 111.067 | 0.873 | 0.386 | 0 | 0 | 1 | 46.25 | 1 | 52 | 1 | GJB_Rinse | 0
230 | P-20181 | 2022/02/17 23:42:00 | 1-52 | 2 | -6.33 | -82.754 | 18710 | -1 | 1 | CPER3-deep-1 | 21-067 | H2O | 1645166827 | 0.343 | 0.76 | 78.326 | 0.002 | 0.005 | 1.849 | 1300.641 | 2.432 | 2.975 | 114.174 | 0.866 | 0.326 | 0 | 0 | 1 | 46.312 | 1 | 52 | 1 | GJB_Rinse | 0
231 | P-20181 | 2022/02/17 23:50:56 | 1-52 | 3 | -6.483 | -83.23 | 18339 | -1 | 1 | CPER3-deep-1 | 21-067 | H2O | 1645167362 | 0.306 | 0.74 | 76.403 | 0.001 | 0.003 | 1.811 | 1297.843 | 2.824 | 2.931 | 111.809 | 0.872 | 0.46 | 0 | 0 | 1 | 46.375 | 1 | 52 | 1 | GJB_Rinse | 0
232 | P-20181 | 2022/02/17 23:59:51 | 1-52 | 4 | -6.524 | -83.887 | 18544 | 0 | 1 | CPER3-deep-1 | 21-067 | H2O | 1645167898 | 0.312 | 0.705 | 75.301 | 0 | 0.001 | 1.789 | 1299.581 | 2.307 | 2.938 | 113.148 | 0.869 | 0.355 | 0 | 0 | 1 | 46.375 | 1 | 52 | 1 | GJB_Rinse | 0
233 | P-20182 | 2022/02/18 00:09:37 | 1-53 | 1 | 0.462 | -69.442 | 18352 | -1 | 1 | CPER3-1020-2 | 21-067 | H2O | 1645168484 | 0.301 | 0.724 | 76.012 | 0 | 0.004 | 1.801 | 1298.159 | 1.692 | 2.934 | 113.35 | 0.873 | 0.446 | 0 | 0 | 1 | 46.438 | 1 | 53 | 1 | GJB_Rinse | 0
234 | P-20182 | 2022/02/18 00:18:32 | 1-53 | 2 | 0.806 | -67.807 | 18512 | -1 | 1 | CPER3-1020-2 | 21-067 | H2O | 1645169019 | 0.279 | 0.686 | 77.305 | 0 | 0 | 1.829 | 1298.926 | 2.381 | 2.979 | 114.031 | 0.869 | 0.682 | 0 | 0 | 1 | 46.5 | 1 | 53 | 1 | GJB_Rinse | 0
235 | P-20182 | 2022/02/18 00:27:27 | 1-53 | 3 | 0.804 | -67.753 | 18410 | -1 | 1 | CPER3-1020-2 | 21-067 | H2O | 1645169554 | 0.329 | 0.66 | 74.309 | 0 | 0.002 | 1.77 | 1298.101 | 2.47 | 2.958 | 113.677 | 0.868 | 0.556 | 0 | 0 | 1 | 46.5 | 1 | 53 | 1 | GJB_Rinse | 0
236 | P-20182 | 2022/02/18 00:36:23 | 1-53 | 4 | 0.908 | -67.641 | 18512 | 0 | 1 | CPER3-1020-2 | 21-067 | H2O | 1645170089 | 0.343 | 0.784 | 74.871 | 0.001 | 0.003 | 1.785 | 1299.038 | 2.053 | 3.01 | 114.389 | 0.879 | 0.52 | 0 | 0 | 1 | 46.5 | 1 | 53 | 1 | GJB_Rinse | 0
237 | P-20183 | 2022/02/18 00:46:10 | 1-54 | 1 | -12.169 | -104.701 | 18395 | -1 | 1 | TEST | TEST | H2O | 1645170676 | 0.3 | 0.792 | 75.785 | 0.002 | 0.006 | 1.79 | 1298.264 | 2.176 | 2.931 | 113.394 | 0.865 | 0.573 | 0 | 0 | 1 | 46.562 | 1 | 54 | 1 | GJB_Rinse | 0
238 | P-20183 | 2022/02/18 00:55:05 | 1-54 | 2 | -12.632 | -108.126 | 18594 | -1 | 1 | TEST | TEST | H2O | 1645171212 | 0.307 | 0.733 | 77.247 | 0.002 | 0.006 | 1.83 | 1299.762 | 2.704 | 2.955 | 113.035 | 0.868 | 0.394 | 0 | 0 | 1 | 46.562 | 1 | 54 | 1 | GJB_Rinse | 0
239 | P-20183 | 2022/02/18 01:04:01 | 1-54 | 3 | -12.761 | -108.764 | 18489 | -1 | 1 | TEST | TEST | H2O | 1645171747 | 0.306 | 0.737 | 76.103 | 0.002 | 0.005 | 1.804 | 1299.007 | 2.581 | 2.961 | 112.889 | 0.872 | 0.332 | 0 | 0 | 1 | 46.625 | 1 | 54 | 1 | GJB_Rinse | 0
240 | P-20183 | 2022/02/18 01:12:57 | 1-54 | 4 | -12.802 | -109.08 | 18548 | 0 | 1 | TEST | TEST | H2O | 1645172283 | 0.312 | 0.697 | 75.903 | 0.001 | 0.004 | 1.807 | 1299.555 | 2.763 | 2.978 | 112.914 | 0.875 | 0.428 | 0 | 0 | 1 | 46.688 | 1 | 54 | 1 | GJB_Rinse | 0
241 | P-20184 | 2022/02/18 01:22:44 | 1-55 | 1 | -12.848 | -109.339 | 18429 | -1 | 1 | TEST | TEST | H2O | 1645172871 | 0.315 | 0.697 | 74.552 | 0.001 | 0.002 | 1.78 | 1298.491 | 2.754 | 2.913 | 112.553 | 0.872 | 0.507 | 0 | 0 | 1 | 46.688 | 1 | 55 | 1 | GJB_Rinse | 0
242 | P-20184 | 2022/02/18 01:31:40 | 1-55 | 2 | -12.898 | -109.558 | 18568 | -1 | 1 | TEST | TEST | H2O | 1645173407 | 0.314 | 0.667 | 76.628 | 0.001 | 0.005 | 1.83 | 1299.299 | 2.844 | 2.958 | 113.933 | 0.877 | 0.646 | 0 | 0 | 1 | 46.688 | 1 | 55 | 1 | GJB_Rinse | 0
243 | P-20184 | 2022/02/18 01:40:36 | 1-55 | 3 | -12.873 | -109.561 | 18558 | -1 | 1 | TEST | TEST | H2O | 1645173942 | 0.309 | 0.749 | 78.011 | 0.001 | 0.004 | 1.846 | 1299.18 | 3.148 | 2.973 | 113.058 | 0.874 | 0.613 | 0 | 0 | 1 | 46.688 | 1 | 55 | 1 | GJB_Rinse | 0
244 | P-20184 | 2022/02/18 01:49:30 | 1-55 | 4 | -12.841 | -109.559 | 18435 | 0 | 1 | TEST | TEST | H2O | 1645174477 | 0.303 | 0.699 | 75.755 | 0.001 | 0.005 | 1.795 | 1298.554 | 2.858 | 2.95 | 112.354 | 0.867 | 0.424 | 0 | 0 | 1 | 46.688 | 1 | 55 | 1 | GJB_Rinse | 0
245 | P-20185 | 2022/02/18 01:59:18 | 1-56 | 1 | -13 | -110.015 | 18194 | -1 | 1 | TEST | TEST | H2O | 1645175065 | 0.297 | 0.798 | 76.789 | 0.001 | 0.004 | 1.823 | 1296.877 | 2.541 | 2.892 | 111.015 | 0.866 | 0.322 | 0 | 0 | 1 | 46.688 | 1 | 56 | 1 | GJB_Rinse | 0
246 | P-20185 | 2022/02/18 02:08:14 | 1-56 | 2 | -12.943 | -109.662 | 18289 | -1 | 1 | TEST | TEST | H2O | 1645175601 | 0.326 | 0.758 | 75.287 | 0.001 | 0.007 | 1.79 | 1297.444 | 2.256 | 2.961 | 113.335 | 0.87 | 0.615 | 0 | 0 | 1 | 46.75 | 1 | 56 | 1 | GJB_Rinse | 0
247 | P-20185 | 2022/02/18 02:17:09 | 1-56 | 3 | -12.837 | -109.465 | 18983 | -1 | 1 | TEST | TEST | H2O | 1645176136 | 0.335 | 0.735 | 78.815 | 0 | 0.004 | 1.862 | 1302.591 | 2.582 | 3.034 | 115.665 | 0.876 | 0.348 | 0 | 0 | 1 | 46.75 | 1 | 56 | 1 | GJB_Rinse | 0
248 | P-20185 | 2022/02/18 02:26:04 | 1-56 | 4 | -12.85 | -109.78 | 18540 | 0 | 1 | TEST | TEST | H2O | 1645176671 | 0.302 | 0.698 | 78.849 | 0 | 0.005 | 1.871 | 1299.206 | 2.822 | 2.932 | 113.21 | 0.877 | 0.597 | 0 | 0 | 1 | 46.688 | 1 | 56 | 1 | GJB_Rinse | 0
249 | P-20186 | 2022/02/18 02:35:49 | 1-57 | 1 | -7.811 | -65.817 | 18617 | -1 | 1 | EV | standard | H2O | 1645177256 | 0.281 | 0.749 | 76.073 | 0.001 | 0.001 | 1.814 | 1299.929 | 2.55 | 2.983 | 113.338 | 0.873 | 0.396 | 0 | 0 | 1 | 46.75 | 1 | 57 | 1 | GJB_Rinse | 0
250 | P-20186 | 2022/02/18 02:44:45 | 1-57 | 2 | -7.592 | -61.441 | 18602 | -1 | 1 | EV | standard | H2O | 1645177792 | 0.295 | 0.69 | 76.293 | 0.001 | 0 | 1.809 | 1299.817 | 2.48 | 2.972 | 114.024 | 0.868 | 0.487 | 0 | 0 | 1 | 46.75 | 1 | 57 | 1 | GJB_Rinse | 0
251 | P-20186 | 2022/02/18 02:53:40 | 1-57 | 3 | -7.517 | -60.574 | 18699 | -1 | 1 | EV | standard | H2O | 1645178327 | 0.287 | 0.706 | 76.463 | 0 | 0.002 | 1.813 | 1300.798 | 2.465 | 2.984 | 112.745 | 0.874 | 0.158 | 0 | 0 | 1 | 46.688 | 1 | 57 | 1 | GJB_Rinse | 0
252 | P-20186 | 2022/02/18 03:02:36 | 1-57 | 4 | -7.466 | -60.403 | 18645 | 0 | 1 | EV | standard | H2O | 1645178862 | 0.311 | 0.694 | 76.941 | 0.001 | 0.003 | 1.82 | 1300.424 | 2.39 | 2.974 | 112.615 | 0.876 | 0.234 | 0 | 0 | 1 | 46.75 | 1 | 57 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1308.77775
varray: [1308.338, 1305.39, 1307.494, 1304.607, 1303.238, 1304.289, 1306.379, 1306.031, 1308.995, 1368.675, 1317.615, 1306.122, 1305.727, 1303.787, 1305.497, 1303.776, 1302.092, 1302.07, 1303.021, 1302.412]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
baseline_shift_STD | 14.5114819979
varray: [1308.338, 1305.39, 1307.494, 1304.607, 1303.238, 1304.289, 1306.379, 1306.031, 1308.995, 1368.675, 1317.615, 1306.122, 1305.727, 1303.787, 1305.497, 1303.776, 1302.092, 1302.07, 1303.021, 1302.412]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.37865
varray: [0.51, 0.393, 0.356, 0.122, 0.537, 0.181, 0.126, 0.314, 0.344, 0.16, 0.832, 0.322, 0.249, 0.094, 0.3, 0.485, 0.729, 0.746, 0.409, 0.364]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.21124873715
varray: [0.51, 0.393, 0.356, 0.122, 0.537, 0.181, 0.126, 0.314, 0.344, 0.16, 0.832, 0.322, 0.249, 0.094, 0.3, 0.485, 0.729, 0.746, 0.409, 0.364]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
residuals_MEAN | 3.071
varray: [3.152, 3.102, 3.097, 2.964, 3.066, 2.977, 3.085, 3.112, 3.084, 3.427, 3.299, 3.024, 3.028, 2.932, 2.989, 3.042, 3.014, 3.05, 3.016, 2.96]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
residuals_STD | 0.116811589187
varray: [3.152, 3.102, 3.097, 2.964, 3.066, 2.977, 3.085, 3.112, 3.084, 3.427, 3.299, 3.024, 3.028, 2.932, 2.989, 3.042, 3.014, 3.05, 3.016, 2.96]
inst: std_standards
operation: std_standards(residuals)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
slope_shift_MEAN | 2.24445
varray: [1.611, 1.914, 1.706, 2.064, 2.017, 1.756, 1.587, 1.822, 1.542, 1.842, 2.678, 2.628, 2.371, 3.135, 2.658, 2.608, 2.815, 2.98, 2.596, 2.559]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
slope_shift_STD | 0.509169755066
varray: [1.611, 1.914, 1.706, 2.064, 2.017, 1.756, 1.587, 1.822, 1.542, 1.842, 2.678, 2.628, 2.371, 3.135, 2.658, 2.608, 2.815, 2.98, 2.596, 2.559]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
baseline_curvature_MEAN | 116.659
varray: [120.215, 116.411, 117.247, 113.446, 115.143, 114.263, 116.297, 117.066, 116.177, 126.802, 124.875, 115.507, 115.639, 112.908, 115.422, 115.613, 114.668, 115.818, 115.695, 113.968]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
baseline_curvature_STD | 3.51111675142
varray: [120.215, 116.411, 117.247, 113.446, 115.143, 114.263, 116.297, 117.066, 116.177, 126.802, 124.875, 115.507, 115.639, 112.908, 115.422, 115.613, 114.668, 115.818, 115.695, 113.968]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 3.071
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1308.77775
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.37865
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 2.24445
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.175217383781
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 3.24621738378
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.233623178374
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.30462317837
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.350434767561
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.42143476756
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.3071
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.633746211449
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.01239621145
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 7.25574099897
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1316.033491
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1323.289232
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 43.5344459938
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1352.31219599
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 261.206675963
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1569.98442596
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.54584877533
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.79029877533
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 10.5333502543
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 127.192350254
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -5.58173305195
actual_array: [ 1.93 -15.84]
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503, 4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752, -12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335, -7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554, -6.196, -6.132, -6.095, -6.078, 4.744, 5.174, 5.271, 5.34, 8.89, 8.776, 8.774, 8.834, 0.0, 0.0, 4.45, 4.442, 4.984, 5.0, 5.076, 5.103, 8.089, 8.103, 8.072, 8.093, -2.532, -2.791, -2.917, -2.963, 3.914, 4.211, 4.292, 4.32, -6.743, -7.196, -7.336, -7.384, -12.499, -12.625, -12.669, -12.698, 3.265, 3.882, 4.046, 4.159, -0.247, -0.493, -0.509, -0.471, 0.346, 0.469, 0.466, 0.485, 1.948, 2.013, 2.006, 1.995, 1.725, 1.803, 1.666, 1.706, 4.368, 0.0, 0.0, 4.368, -0.325, -0.33, -0.303, -0.341, -7.033, -7.397, -7.44, -7.471, -12.519, -12.764, -12.886, -12.819, -3.972, -3.504, -3.394, -3.378, 5.929, 6.241, 6.322, 6.316, 7.184, 7.312, 7.358, 7.308, 5.0, 4.742, 4.766, 4.751, 3.967, 4.02, 4.051, 4.002, 3.506, 3.362, 3.41, 3.426, -2.383, -2.602, -2.642, -2.673, -7.204, -7.389, -7.396, -7.416, -12.545, -12.785, -12.853, -12.854, 0.103, 0.612, 0.781, 0.826, 0.42, 0.417, 0.459, 0.45, 0.0, -0.074, -0.093, -0.093, 0.0, 0.0, -3.007, -3.021, -0.104, 0.012, 0.017, 0.033, 2.405, 2.642, 2.457, 2.475, 0.0, 0.0, 0.0, 0.0, -7.169, -7.331, -7.443, -7.498, -12.573, -12.682, -12.853, -12.854, -6.366, -6.223, -6.111, -6.174, -5.932, -5.95, -5.848, -5.878, -4.083, -3.996, -4.041, -3.985, -3.449, -3.483, -3.428, -3.431, -3.44, -3.44, -3.457, -3.446, 0.573, 0.81, 0.803, 0.863, -5.02, -5.282, -5.293, -5.355, -7.31, -7.459, -7.466, -7.491, -7.494, -7.461, -7.506, -7.506, 0.595, 1.003, 1.053, 1.096, -6.204, -6.33, -6.483, -6.524, 0.462, 0.806, 0.804, 0.908, -12.169, -12.632, -12.761, -12.802, -12.848, -12.898, -12.873, -12.841, -13.0, -12.943, -12.837, -12.85, -7.811, -7.592, -7.517, -7.466]
fit_rsq: 1.0
fit_intercept: 2.80644310636
except: True
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.80582142857
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0, 5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0, 16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0, 21679.0, 22215.0, 22752.0, 23289.0, 23878.0, 24415.0, 24952.0, 25489.0, 26079.0, 26615.0, 27152.0, 27689.0, 28276.0, 28813.0, 29349.0, 29886.0, 30473.0, 31010.0, 31547.0, 32083.0, 32671.0, 33208.0, 33744.0, 34281.0, 34870.0, 35407.0, 35944.0, 36481.0, 37070.0, 37607.0, 38143.0, 38680.0, 39269.0, 39805.0, 40342.0, 40879.0, 41468.0, 42005.0, 42542.0, 43078.0, 43664.0, 44201.0, 44738.0, 45274.0, 45861.0, 46398.0, 46934.0, 47471.0, 48058.0, 48594.0, 49131.0, 49667.0, 50255.0, 50791.0, 51327.0, 51863.0, 52451.0, 52986.0, 53522.0, 54059.0, 54647.0, 55183.0, 55782.0, 56380.0, 57031.0, 57567.0, 58103.0, 58640.0, 59226.0, 59762.0, 60297.0, 60834.0, 61420.0, 61956.0, 62492.0, 63028.0, 63615.0, 64151.0, 64687.0, 65224.0, 65811.0, 66347.0, 66883.0, 67420.0, 68007.0, 68542.0, 69078.0, 69615.0, 70202.0, 70738.0, 71275.0, 71811.0, 72399.0, 72935.0, 73471.0, 74007.0, 74591.0, 75127.0, 75663.0, 76199.0, 76784.0, 77321.0, 77856.0, 78393.0, 78978.0, 79514.0, 80050.0, 80586.0, 81172.0, 81708.0, 82244.0, 82780.0, 83366.0, 83902.0, 84438.0, 84975.0, 85561.0, 86097.0, 86633.0, 87169.0, 87756.0, 88293.0, 88830.0, 89366.0, 90010.0, 90547.0, 91082.0, 91619.0, 92202.0, 92738.0, 93273.0, 93809.0, 94393.0, 94929.0, 95465.0, 96001.0, 96585.0, 97183.0, 97781.0, 98379.0, 99025.0, 99561.0, 100097.0, 100633.0, 101217.0, 101753.0, 102289.0, 102824.0, 103409.0, 103946.0, 104482.0, 105017.0, 105601.0, 106136.0, 106672.0, 107208.0, 107792.0, 108328.0, 108863.0, 109399.0, 109984.0, 110520.0, 111056.0, 111591.0, 112177.0, 112713.0, 113248.0, 113784.0, 114370.0, 114905.0, 115441.0, 115977.0, 116564.0, 117100.0, 117636.0, 118172.0, 118758.0, 119294.0, 119830.0, 120365.0, 120950.0, 121486.0, 122022.0, 122558.0, 123143.0, 123679.0, 124215.0, 124750.0, 125336.0, 125872.0, 126408.0, 126943.0, 127529.0, 128064.0, 128599.0, 129135.0, 129722.0, 130257.0, 130793.0, 131329.0, 131916.0, 132452.0, 132988.0, 133522.0, 134110.0, 134646.0, 135181.0, 135716.0, 136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -76.2759501063
actual_array: [ 18.1 -119. ]
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629, 29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314, -92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212, -65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47, -80.319, -82.009, -82.352, -82.589, -45.865, -42.202, -41.453, -40.91, -41.445, -42.437, -42.711, -43.013, 0.0, 0.0, -61.353, -61.364, -58.028, -57.964, -57.635, -57.621, -47.177, -46.651, -46.727, -46.304, -72.923, -74.412, -74.877, -75.135, -51.903, -49.732, -49.331, -48.873, -57.984, -59.147, -59.275, -59.441, -103.112, -107.072, -108.13, -108.395, -54.242, -49.308, -47.761, -47.142, -68.126, -70.047, -70.595, -70.801, -66.918, -66.186, -66.156, -66.303, -53.718, -52.301, -52.047, -51.79, -55.926, -55.874, -56.391, -55.925, -48.171, 0.0, 0.0, -48.171, -65.92, -67.019, -67.199, -67.129, -60.341, -59.75, -59.554, -59.752, -103.642, -107.936, -108.899, -108.981, -74.414, -70.529, -70.075, -69.562, -51.829, -50.424, -50.043, -49.975, -42.786, -41.321, -40.827, -41.012, -54.199, -56.096, -56.634, -56.659, -53.401, -52.534, -51.929, -52.308, -52.566, -53.279, -52.852, -52.966, -69.658, -71.006, -71.316, -71.595, -61.032, -59.685, -59.595, -59.429, -103.697, -107.872, -108.706, -109.112, -66.249, -62.783, -61.66, -61.366, -63.033, -63.324, -63.099, -62.766, 0.0, -60.203, -60.328, -60.328, 0.0, 0.0, -65.253, -65.297, -61.974, -61.4, -61.604, -61.506, -56.983, -55.378, -55.97, -56.064, 0.0, 0.0, 0.0, 0.0, -59.151, -59.429, -59.743, -59.395, -103.567, -107.336, -108.836, -108.92, -80.386, -77.926, -77.384, -77.3, -77.765, -77.989, -77.685, -77.785, -77.709, -77.567, -77.706, -77.453, -73.41, -73.147, -73.091, -72.847, -73.008, -73.113, -73.174, -73.131, -67.626, -66.93, -66.808, -66.751, -81.611, -83.054, -83.268, -83.159, -62.167, -60.596, -59.844, -59.945, -60.037, -59.711, -59.701, -59.673, -57.996, -57.566, -57.636, -57.473, -81.014, -82.754, -83.23, -83.887, -69.442, -67.807, -67.753, -67.641, -104.701, -108.126, -108.764, -109.08, -109.339, -109.558, -109.561, -109.559, -110.015, -109.662, -109.465, -109.78, -65.817, -61.441, -60.574, -60.403]
fit_rsq: 1.0
fit_intercept: 19.5866234865
except: True
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -62.2278730159
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0, 5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0, 10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0, 16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0, 21679.0, 22215.0, 22752.0, 23289.0, 23878.0, 24415.0, 24952.0, 25489.0, 26079.0, 26615.0, 27152.0, 27689.0, 28276.0, 28813.0, 29349.0, 29886.0, 30473.0, 31010.0, 31547.0, 32083.0, 32671.0, 33208.0, 33744.0, 34281.0, 34870.0, 35407.0, 35944.0, 36481.0, 37070.0, 37607.0, 38143.0, 38680.0, 39269.0, 39805.0, 40342.0, 40879.0, 41468.0, 42005.0, 42542.0, 43078.0, 43664.0, 44201.0, 44738.0, 45274.0, 45861.0, 46398.0, 46934.0, 47471.0, 48058.0, 48594.0, 49131.0, 49667.0, 50255.0, 50791.0, 51327.0, 51863.0, 52451.0, 52986.0, 53522.0, 54059.0, 54647.0, 55183.0, 55782.0, 56380.0, 57031.0, 57567.0, 58103.0, 58640.0, 59226.0, 59762.0, 60297.0, 60834.0, 61420.0, 61956.0, 62492.0, 63028.0, 63615.0, 64151.0, 64687.0, 65224.0, 65811.0, 66347.0, 66883.0, 67420.0, 68007.0, 68542.0, 69078.0, 69615.0, 70202.0, 70738.0, 71275.0, 71811.0, 72399.0, 72935.0, 73471.0, 74007.0, 74591.0, 75127.0, 75663.0, 76199.0, 76784.0, 77321.0, 77856.0, 78393.0, 78978.0, 79514.0, 80050.0, 80586.0, 81172.0, 81708.0, 82244.0, 82780.0, 83366.0, 83902.0, 84438.0, 84975.0, 85561.0, 86097.0, 86633.0, 87169.0, 87756.0, 88293.0, 88830.0, 89366.0, 90010.0, 90547.0, 91082.0, 91619.0, 92202.0, 92738.0, 93273.0, 93809.0, 94393.0, 94929.0, 95465.0, 96001.0, 96585.0, 97183.0, 97781.0, 98379.0, 99025.0, 99561.0, 100097.0, 100633.0, 101217.0, 101753.0, 102289.0, 102824.0, 103409.0, 103946.0, 104482.0, 105017.0, 105601.0, 106136.0, 106672.0, 107208.0, 107792.0, 108328.0, 108863.0, 109399.0, 109984.0, 110520.0, 111056.0, 111591.0, 112177.0, 112713.0, 113248.0, 113784.0, 114370.0, 114905.0, 115441.0, 115977.0, 116564.0, 117100.0, 117636.0, 118172.0, 118758.0, 119294.0, 119830.0, 120365.0, 120950.0, 121486.0, 122022.0, 122558.0, 123143.0, 123679.0, 124215.0, 124750.0, 125336.0, 125872.0, 126408.0, 126943.0, 127529.0, 128064.0, 128599.0, 129135.0, 129722.0, 130257.0, 130793.0, 131329.0, 131916.0, 132452.0, 132988.0, 133522.0, 134110.0, 134646.0, 135181.0, 135716.0, 136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.1497
varray: [3.377, 3.607, 3.075, 3.042, 3.11, 3.027, 3.018, 3.083, 3.124, 3.034]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: residuals
SAMP_residuals_STD | 0.191613064969
varray: [3.377, 3.607, 3.075, 3.042, 3.11, 3.027, 3.018, 3.083, 3.124, 3.034]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3625
varray: [0.31, 0.246, 0.566, 0.516, 0.517, 0.134, 0.273, 0.516, 0.373, 0.174]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.157973450653
varray: [0.31, 0.246, 0.566, 0.516, 0.517, 0.134, 0.273, 0.516, 0.373, 0.174]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1332.7505
varray: [1330.593, 1477.55, 1337.615, 1312.617, 1319.453, 1311.544, 1311.205, 1310.249, 1309.059, 1307.62]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 51.8442813867
varray: [1330.593, 1477.55, 1337.615, 1312.617, 1319.453, 1311.544, 1311.205, 1310.249, 1309.059, 1307.62]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.3774
varray: [2.467, -3.449, 0.693, 1.979, 1.842, 1.59, 2.251, 2.107, 2.269, 2.025]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.7664330789
varray: [2.467, -3.449, 0.693, 1.979, 1.842, 1.59, 2.251, 2.107, 2.269, 2.025]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: slope_shift
SAMP_d18O_STD | 0.204916134607
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.73126189802
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 119.6089
varray: [126.8, 135.391, 117.873, 116.047, 117.151, 115.341, 116.163, 117.632, 117.689, 116.002]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.43964653015
varray: [126.8, 135.391, 117.873, 116.047, 117.151, 115.341, 116.163, 117.632, 117.689, 116.002]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.1497
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1332.7505
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.3774
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0787
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0787
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.56268316509
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.9499
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6732
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
dD_mean | -61.462
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
d18O_stdev | 0.204916134607
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
dD_stdev | 2.73126189802
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6732
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.462
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.204916134607
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.73126189802
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4226324214
actual_array: [ 1.93 -15.84]
varray: [-7.974, -8.064, -7.8, -7.615, -7.657, -7.568, -7.537, -7.531, -7.483, -7.503]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.6732
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.5619239319
actual_array: [ 18.1 -119. ]
varray: [-67.537, -65.143, -62.018, -60.512, -60.602, -60.179, -59.83, -59.674, -59.496, -59.629]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.462
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [1.0, 538.0, 1073.0, 1610.0, 2147.0, 2683.0, 3219.0, 3756.0, 4293.0, 4829.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.62805
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 23.97275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.1066
varray: [3.152, 3.102, 3.097, 2.964, 3.066, 2.977, 3.085, 3.112, 3.084, 3.427]
inst: mean_all
operation: mean_all(residuals)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: residuals
SAMP_residuals_STD | 0.126761543415
varray: [3.152, 3.102, 3.097, 2.964, 3.066, 2.977, 3.085, 3.112, 3.084, 3.427]
inst: std_all
operation: std_all(residuals)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3043
varray: [0.51, 0.393, 0.356, 0.122, 0.537, 0.181, 0.126, 0.314, 0.344, 0.16]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.152651418452
varray: [0.51, 0.393, 0.356, 0.122, 0.537, 0.181, 0.126, 0.314, 0.344, 0.16]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1312.3436
varray: [1308.338, 1305.39, 1307.494, 1304.607, 1303.238, 1304.289, 1306.379, 1306.031, 1308.995, 1368.675]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 19.8757938665
varray: [1308.338, 1305.39, 1307.494, 1304.607, 1303.238, 1304.289, 1306.379, 1306.031, 1308.995, 1368.675]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7861
varray: [1.611, 1.914, 1.706, 2.064, 2.017, 1.756, 1.587, 1.822, 1.542, 1.842]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.179047758992
varray: [1.611, 1.914, 1.706, 2.064, 2.017, 1.756, 1.587, 1.822, 1.542, 1.842]
inst: std_all
operation: std_all(slope_shift)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: slope_shift
SAMP_d18O_STD | 0.222733024044
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.61708149541
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 117.3067
varray: [120.215, 116.411, 117.247, 113.446, 115.143, 114.263, 116.297, 117.066, 116.177, 126.802]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.80977016075
varray: [120.215, 116.411, 117.247, 113.446, 115.143, 114.263, 116.297, 117.066, 116.177, 126.802]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.1066
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3043
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1312.3436
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7861
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0356
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0356
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.15923152068
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.6477
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.747
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 39.0009
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.747
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
dD_mean | 39.0009
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
d18O_stdev | 0.222733024044
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
dD_stdev | 3.61708149541
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.747
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
measured_dD_mean | 39.0009
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.222733024044
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.61708149541
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [4.171, 4.625, 4.765, 4.746, 4.846, 4.842, 4.924, 4.937, 4.862, 4.752]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.747
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [29.283, 37.152, 39.253, 39.861, 40.325, 40.524, 41.233, 41.205, 40.859, 40.314]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 39.0009
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [5417.0, 5954.0, 6491.0, 7028.0, 7564.0, 8102.0, 8638.0, 9175.0, 9712.0, 10249.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.67115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.56585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0354
varray: [3.299, 3.024, 3.028, 2.932, 2.989, 3.042, 3.014, 3.05, 3.016, 2.96]
inst: mean_all
operation: mean_all(residuals)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
SAMP_residuals_STD | 0.0996049976435
varray: [3.299, 3.024, 3.028, 2.932, 2.989, 3.042, 3.014, 3.05, 3.016, 2.96]
inst: std_all
operation: std_all(residuals)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.453
varray: [0.832, 0.322, 0.249, 0.094, 0.3, 0.485, 0.729, 0.746, 0.409, 0.364]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.242123476305
varray: [0.832, 0.322, 0.249, 0.094, 0.3, 0.485, 0.729, 0.746, 0.409, 0.364]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1305.2119
varray: [1317.615, 1306.122, 1305.727, 1303.787, 1305.497, 1303.776, 1302.092, 1302.07, 1303.021, 1302.412]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.61096187724
varray: [1317.615, 1306.122, 1305.727, 1303.787, 1305.497, 1303.776, 1302.092, 1302.07, 1303.021, 1302.412]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.7028
varray: [2.678, 2.628, 2.371, 3.135, 2.658, 2.608, 2.815, 2.98, 2.596, 2.559]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.220001414137
varray: [2.678, 2.628, 2.371, 3.135, 2.658, 2.608, 2.815, 2.98, 2.596, 2.559]
inst: std_all
operation: std_all(slope_shift)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
SAMP_d18O_STD | 0.335188437881
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.56719100924
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 116.0113
varray: [124.875, 115.507, 115.639, 112.908, 115.422, 115.613, 114.668, 115.818, 115.695, 113.968]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.25254485151
varray: [124.875, 115.507, 115.639, 112.908, 115.422, 115.613, 114.668, 115.818, 115.695, 113.968]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0354
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.453
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1305.2119
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.7028
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0356
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0356
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.15923152068
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.6477
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.1202
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
dD_mean | -108.0542
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
d18O_stdev | 0.335188437881
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
dD_stdev | 5.56719100924
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.1202
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.0542
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.335188437881
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.56719100924
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-12.211, -13.006, -13.128, -13.168, -13.195, -13.248, -13.301, -13.299, -13.311, -13.335]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.1202
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-92.898, -105.884, -108.28, -109.384, -109.88, -110.377, -110.615, -110.91, -111.102, -111.212]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.0542
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [10837.0, 11374.0, 11910.0, 12447.0, 12984.0, 13520.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.74235
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.56585
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0566
varray: [2.968, 3.037, 2.997, 3.087, 3.129, 3.068, 3.017, 3.067, 3.094, 3.102]
inst: mean_all
operation: mean_all(residuals)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: residuals
SAMP_residuals_STD | 0.0507919503684
varray: [2.968, 3.037, 2.997, 3.087, 3.129, 3.068, 3.017, 3.067, 3.094, 3.102]
inst: std_all
operation: std_all(residuals)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4301
varray: [0.291, 0.368, 0.331, 0.686, 0.558, 0.463, 0.332, 0.375, 0.55, 0.347]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129053348142
varray: [0.291, 0.368, 0.331, 0.686, 0.558, 0.463, 0.332, 0.375, 0.55, 0.347]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.7197
varray: [1302.011, 1304.398, 1304.206, 1303.613, 1304.001, 1303.641, 1303.528, 1304.275, 1303.666, 1303.858]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.673305288195
varray: [1302.011, 1304.398, 1304.206, 1303.613, 1304.001, 1303.641, 1303.528, 1304.275, 1303.666, 1303.858]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.4542
varray: [3.003, 2.697, 2.164, 2.207, 2.276, 1.932, 2.915, 2.476, 2.596, 2.276]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.345539770986
varray: [3.003, 2.697, 2.164, 2.207, 2.276, 1.932, 2.915, 2.476, 2.596, 2.276]
inst: std_all
operation: std_all(slope_shift)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0942078670931
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.87033164023
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 116.3159
varray: [113.251, 116.137, 116.433, 117.656, 117.41, 117.372, 114.752, 117.059, 117.409, 115.68]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.41989111085
varray: [113.251, 116.137, 116.433, 117.656, 117.41, 117.372, 114.752, 117.059, 117.409, 115.68]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0566
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4301
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.7197
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.4542
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0144
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0144
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.468902637577
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3431
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.5957
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
dD_mean | -60.401
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0942078670931
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
dD_stdev | 1.87033164023
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.5957
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.401
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0942078670931
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.87033164023
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.345554032
actual_array: [ 1.93 -15.84]
varray: [-7.854, -7.579, -7.593, -7.566, -7.601, -7.555, -7.536, -7.525, -7.594, -7.554]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.5957
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.572749806
actual_array: [ 18.1 -119. ]
varray: [-65.534, -60.996, -60.379, -59.861, -59.717, -59.525, -59.513, -59.474, -59.541, -59.47]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [16255.0, 16791.0, 17328.0, 17865.0, 18401.0, 18938.0, 19475.0, 20015.0, 20552.0, 21089.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.72115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.05805
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.907
varray: [2.868, 2.961, 2.891, 2.908]
inst: mean_all
operation: mean_all(residuals)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: residuals
SAMP_residuals_STD | 0.0395558676642
varray: [2.868, 2.961, 2.891, 2.908]
inst: std_all
operation: std_all(residuals)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.549
varray: [0.374, 0.859, 0.355, 0.608]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.236531886504
varray: [0.374, 0.859, 0.355, 0.608]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.71025
varray: [1295.988, 1297.212, 1297.045, 1296.596]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.547268596456
varray: [1295.988, 1297.212, 1297.045, 1296.596]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.38425
varray: [2.172, 2.425, 2.441, 2.499]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.145027296281
varray: [2.172, 2.425, 2.441, 2.499]
inst: std_all
operation: std_all(slope_shift)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0522773054649
varray: [-6.196, -6.132, -6.095, -6.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.0268198073
varray: [-80.319, -82.009, -82.352, -82.589]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.88425
varray: [111.11, 113.367, 111.321, 111.739]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.0224739198
varray: [111.11, 113.367, 111.321, 111.739]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.907
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.549
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.71025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.38425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.164
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.164
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.34028003908
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.77475
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-6.196, -6.132, -6.095, -6.078]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-80.319, -82.009, -82.352, -82.589]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.12525
varray: [-6.196, -6.132, -6.095, -6.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
dD_mean | -81.81725
varray: [-80.319, -82.009, -82.352, -82.589]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0522773054649
varray: [-6.196, -6.132, -6.095, -6.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
dD_stdev | 1.0268198073
varray: [-80.319, -82.009, -82.352, -82.589]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.12525
varray: [-6.196, -6.132, -6.095, -6.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
measured_dD_mean | -81.81725
varray: [-80.319, -82.009, -82.352, -82.589]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0522773054649
varray: [-6.196, -6.132, -6.095, -6.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.0268198073
varray: [-80.319, -82.009, -82.352, -82.589]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.88310348012
actual_array: [ 1.93 -15.84]
varray: [-6.196, -6.132, -6.095, -6.078]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.12525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -94.5391969065
actual_array: [ 18.1 -119. ]
varray: [-80.319, -82.009, -82.352, -82.589]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -81.81725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [21679.0, 22215.0, 22752.0, 23289.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.87075
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.0675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01275
varray: [2.913, 3.009, 3.058, 3.071]
inst: mean_all
operation: mean_all(residuals)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: residuals
SAMP_residuals_STD | 0.0716583328488
varray: [2.913, 3.009, 3.058, 3.071]
inst: std_all
operation: std_all(residuals)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33625
varray: [0.072, 0.404, 0.509, 0.36]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.18692667903
varray: [0.072, 0.404, 0.509, 0.36]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.2975
varray: [1302.163, 1303.136, 1305.246, 1306.645]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.02612775181
varray: [1302.163, 1303.136, 1305.246, 1306.645]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6525
varray: [1.55, 1.76, 1.69, 1.61]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0917877987534
varray: [1.55, 1.76, 1.69, 1.61]
inst: std_all
operation: std_all(slope_shift)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: slope_shift
SAMP_d18O_STD | 0.267639527474
varray: [4.744, 5.174, 5.271, 5.34]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.23533092256
varray: [-45.865, -42.202, -41.453, -40.91]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.32
varray: [112.434, 114.456, 117.005, 117.385]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.32227919654
varray: [112.434, 114.456, 117.005, 117.385]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.2975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.89677629437
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.339
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [4.744, 5.174, 5.271, 5.34]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-45.865, -42.202, -41.453, -40.91]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.13225
varray: [4.744, 5.174, 5.271, 5.34]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
dD_mean | -42.6075
varray: [-45.865, -42.202, -41.453, -40.91]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
d18O_stdev | 0.267639527474
varray: [4.744, 5.174, 5.271, 5.34]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
dD_stdev | 2.23533092256
varray: [-45.865, -42.202, -41.453, -40.91]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.13225
varray: [4.744, 5.174, 5.271, 5.34]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.6075
varray: [-45.865, -42.202, -41.453, -40.91]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.267639527474
varray: [4.744, 5.174, 5.271, 5.34]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.23533092256
varray: [-45.865, -42.202, -41.453, -40.91]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.31315418756
actual_array: [ 1.93 -15.84]
varray: [4.744, 5.174, 5.271, 5.34]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.13225
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.983805594
actual_array: [ 18.1 -119. ]
varray: [-45.865, -42.202, -41.453, -40.91]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -42.6075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [23878.0, 24415.0, 24952.0, 25489.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.765
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.48025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.7155
varray: [3.783, 3.636, 3.744, 3.699]
inst: mean_all
operation: mean_all(residuals)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: residuals
SAMP_residuals_STD | 0.0631426955395
varray: [3.783, 3.636, 3.744, 3.699]
inst: std_all
operation: std_all(residuals)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -5.28525
varray: [-4.797, -5.654, -5.445, -5.245]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.365834548214
varray: [-4.797, -5.654, -5.445, -5.245]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1332.91675
varray: [1333.215, 1333.013, 1333.479, 1331.96]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.665760905531
varray: [1333.215, 1333.013, 1333.479, 1331.96]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -15.32175
varray: [-15.623, -14.979, -15.82, -14.865]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.470851268095
varray: [-15.623, -14.979, -15.82, -14.865]
inst: std_all
operation: std_all(slope_shift)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0551935986626
varray: [8.89, 8.776, 8.774, 8.834]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.679675167513
varray: [-41.445, -42.437, -42.711, -43.013]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 128.18975
varray: [131.119, 126.233, 128.097, 127.31]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.09697709016
varray: [131.119, 126.233, 128.097, 127.31]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.7155
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 5.28525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1332.91675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 15.32175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.6445
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.6445
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 20.9866492999
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 11.53075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [8.89, 8.776, 8.774, 8.834]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-41.445, -42.437, -42.711, -43.013]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 8.8185
varray: [8.89, 8.776, 8.774, 8.834]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
dD_mean | -42.4015
varray: [-41.445, -42.437, -42.711, -43.013]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0551935986626
varray: [8.89, 8.776, 8.774, 8.834]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
dD_stdev | 0.679675167513
varray: [-41.445, -42.437, -42.711, -43.013]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
measured_d18O_mean | 8.8185
varray: [8.89, 8.776, 8.774, 8.834]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.4015
varray: [-41.445, -42.437, -42.711, -43.013]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0551935986626
varray: [8.89, 8.776, 8.774, 8.834]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.679675167513
varray: [-41.445, -42.437, -42.711, -43.013]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 5.97935048581
actual_array: [ 1.93 -15.84]
varray: [8.89, 8.776, 8.774, 8.834]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.8185
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.7917510511
actual_array: [ 18.1 -119. ]
varray: [-41.445, -42.437, -42.711, -43.013]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -42.4015
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [26079.0, 26615.0, 27152.0, 27689.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.06225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 24.139
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.56
varray: [0.0, 0.0, 3.126, 3.114]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: residuals
SAMP_residuals_STD | 1.80133950159
varray: [0.0, 0.0, 3.126, 3.114]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2325
varray: [0.0, 0.0, 0.623, 0.307]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.297855110191
varray: [0.0, 0.0, 0.623, 0.307]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 653.13125
varray: [0.0, 0.0, 1306.382, 1306.143]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 754.171012319
varray: [0.0, 0.0, 1306.382, 1306.143]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.877
varray: [0.0, 0.0, 1.516, 1.992]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.03114887383
varray: [0.0, 0.0, 1.516, 1.992]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: slope_shift
SAMP_d18O_STD | 2.56690137455
varray: [0.0, 0.0, 4.45, 4.442]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 35.4253467767
varray: [0.0, 0.0, -61.353, -61.364]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.92825
varray: [0.0, 0.0, 121.299, 118.414]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 69.2092050832
varray: [0.0, 0.0, 121.299, 118.414]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.56
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 653.13125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.877
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.511
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.511
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -49.2022142625
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -56.73075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.0, 0.0, 4.45, 4.442]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [0.0, 0.0, -61.353, -61.364]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.223
varray: [0.0, 0.0, 4.45, 4.442]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
dD_mean | -30.67925
varray: [0.0, 0.0, -61.353, -61.364]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
d18O_stdev | 2.56690137455
varray: [0.0, 0.0, 4.45, 4.442]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
dD_stdev | 35.4253467767
varray: [0.0, 0.0, -61.353, -61.364]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.223
varray: [0.0, 0.0, 4.45, 4.442]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
measured_dD_mean | -30.67925
varray: [0.0, 0.0, -61.353, -61.364]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.56690137455
varray: [0.0, 0.0, 4.45, 4.442]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
measured_dD_stdev | 35.4253467767
varray: [0.0, 0.0, -61.353, -61.364]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.580269096445
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 4.45, 4.442]
fit_rsq: 1.0
fit_intercept: 2.80644310636
except: True
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -46.8630551066
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -61.353, -61.364]
fit_rsq: 1.0
fit_intercept: 19.5866234865
except: True
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -30.67925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [28276.0, 28813.0, 29349.0, 29886.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1307.21775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -655.6465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.02275
varray: [2.969, 3.022, 3.067, 3.033]
inst: mean_all
operation: mean_all(residuals)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: residuals
SAMP_residuals_STD | 0.0406314738432
varray: [2.969, 3.022, 3.067, 3.033]
inst: std_all
operation: std_all(residuals)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2745
varray: [0.13, 0.266, 0.438, 0.264]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.126220706172
varray: [0.13, 0.266, 0.438, 0.264]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.129
varray: [1302.247, 1301.854, 1303.73, 1304.685]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31477323774
varray: [1302.247, 1301.854, 1303.73, 1304.685]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.18175
varray: [2.1, 2.331, 1.953, 2.343]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.189109095498
varray: [2.1, 2.331, 1.953, 2.343]
inst: std_all
operation: std_all(slope_shift)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: slope_shift
SAMP_d18O_STD | 0.057731418367
varray: [4.984, 5.0, 5.076, 5.103]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.214141697637
varray: [-58.028, -57.964, -57.635, -57.621]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.613
varray: [114.436, 114.975, 116.855, 116.186]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.10502880807
varray: [114.436, 114.975, 116.855, 116.186]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.02275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2745
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.129
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.18175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.57114946272
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.046
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [4.984, 5.0, 5.076, 5.103]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-58.028, -57.964, -57.635, -57.621]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.04075
varray: [4.984, 5.0, 5.076, 5.103]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
dD_mean | -57.812
varray: [-58.028, -57.964, -57.635, -57.621]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
d18O_stdev | 0.057731418367
varray: [4.984, 5.0, 5.076, 5.103]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
dD_stdev | 0.214141697637
varray: [-58.028, -57.964, -57.635, -57.621]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.04075
varray: [4.984, 5.0, 5.076, 5.103]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.812
varray: [-58.028, -57.964, -57.635, -57.621]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.057731418367
varray: [4.984, 5.0, 5.076, 5.103]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.214141697637
varray: [-58.028, -57.964, -57.635, -57.621]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.22215196002
actual_array: [ 1.93 -15.84]
varray: [4.984, 5.0, 5.076, 5.103]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.04075
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.1590157703
actual_array: [ 18.1 -119. ]
varray: [-58.028, -57.964, -57.635, -57.621]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.812
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [30473.0, 31010.0, 31547.0, 32083.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.755
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.64875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.22625
varray: [3.177, 3.215, 3.24, 3.273]
inst: mean_all
operation: mean_all(residuals)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: residuals
SAMP_residuals_STD | 0.0405246838359
varray: [3.177, 3.215, 3.24, 3.273]
inst: std_all
operation: std_all(residuals)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.9385
varray: [-2.861, -2.822, -2.775, -3.296]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.240912847312
varray: [-2.861, -2.822, -2.775, -3.296]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1335.66125
varray: [1332.618, 1334.359, 1335.805, 1339.863]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.08936847646
varray: [1332.618, 1334.359, 1335.805, 1339.863]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -8.85775
varray: [-8.746, -8.605, -8.84, -9.24]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.272519571163
varray: [-8.746, -8.605, -8.84, -9.24]
inst: std_all
operation: std_all(slope_shift)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0129196233175
varray: [8.089, 8.103, 8.072, 8.093]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.358982241901
varray: [-47.177, -46.651, -46.727, -46.304]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 121.98775
varray: [120.442, 120.563, 123.406, 123.54]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7166021428
varray: [120.442, 120.563, 123.406, 123.54]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.22625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.9385
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1335.66125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 8.85775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.15525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.05535656138
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.32875
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [8.089, 8.103, 8.072, 8.093]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-47.177, -46.651, -46.727, -46.304]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 8.08925
varray: [8.089, 8.103, 8.072, 8.093]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
dD_mean | -46.71475
varray: [-47.177, -46.651, -46.727, -46.304]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0129196233175
varray: [8.089, 8.103, 8.072, 8.093]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
dD_stdev | 0.358982241901
varray: [-47.177, -46.651, -46.727, -46.304]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
measured_d18O_mean | 8.08925
varray: [8.089, 8.103, 8.072, 8.093]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.71475
varray: [-47.177, -46.651, -46.727, -46.304]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0129196233175
varray: [8.089, 8.103, 8.072, 8.093]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.358982241901
varray: [-47.177, -46.651, -46.727, -46.304]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 5.25406770507
actual_array: [ 1.93 -15.84]
varray: [8.089, 8.103, 8.072, 8.093]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.08925
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.8130095794
actual_array: [ 18.1 -119. ]
varray: [-47.177, -46.651, -46.727, -46.304]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -46.71475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [32671.0, 33208.0, 33744.0, 34281.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.5515
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 26.8835
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.976
varray: [2.891, 2.96, 3.037, 3.016]
inst: mean_all
operation: mean_all(residuals)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: residuals
SAMP_residuals_STD | 0.0653248293785
varray: [2.891, 2.96, 3.037, 3.016]
inst: std_all
operation: std_all(residuals)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34875
varray: [0.288, 0.254, 0.446, 0.407]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0922293337285
varray: [0.288, 0.254, 0.446, 0.407]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.23625
varray: [1299.832, 1302.111, 1301.978, 1301.024]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.0539368656
varray: [1299.832, 1302.111, 1301.978, 1301.024]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.14225
varray: [2.384, 2.341, 1.998, 1.846]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.262371714685
varray: [2.384, 2.341, 1.998, 1.846]
inst: std_all
operation: std_all(slope_shift)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: slope_shift
SAMP_d18O_STD | 0.193356966946
varray: [-2.532, -2.791, -2.917, -2.963]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.988841873102
varray: [-72.923, -74.412, -74.877, -75.135]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.23725
varray: [112.018, 115.039, 115.171, 114.721]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.49150670464
varray: [112.018, 115.039, 115.171, 114.721]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.976
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.23625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.14225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.095
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.095
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.09345490068
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.42175
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-2.532, -2.791, -2.917, -2.963]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-72.923, -74.412, -74.877, -75.135]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.80075
varray: [-2.532, -2.791, -2.917, -2.963]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
dD_mean | -74.33675
varray: [-72.923, -74.412, -74.877, -75.135]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
d18O_stdev | 0.193356966946
varray: [-2.532, -2.791, -2.917, -2.963]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
dD_stdev | 0.988841873102
varray: [-72.923, -74.412, -74.877, -75.135]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.80075
varray: [-2.532, -2.791, -2.917, -2.963]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
measured_dD_mean | -74.33675
varray: [-72.923, -74.412, -74.877, -75.135]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.193356966946
varray: [-2.532, -2.791, -2.917, -2.963]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.988841873102
varray: [-72.923, -74.412, -74.877, -75.135]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.57668921264
actual_array: [ 1.93 -15.84]
varray: [-2.532, -2.791, -2.917, -2.963]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.80075
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.5650997823
actual_array: [ 18.1 -119. ]
varray: [-72.923, -74.412, -74.877, -75.135]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -74.33675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [34870.0, 35407.0, 35944.0, 36481.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.80175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.5415
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01525
varray: [2.985, 2.971, 3.037, 3.068]
inst: mean_all
operation: mean_all(residuals)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: residuals
SAMP_residuals_STD | 0.0451986356726
varray: [2.985, 2.971, 3.037, 3.068]
inst: std_all
operation: std_all(residuals)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.19875
varray: [-0.067, 0.107, 0.555, 0.2]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.262015743802
varray: [-0.067, 0.107, 0.555, 0.2]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1332.1135
varray: [1330.852, 1330.882, 1331.319, 1335.401]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.20203398399
varray: [1330.852, 1330.882, 1331.319, 1335.401]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.769
varray: [0.621, 1.352, 0.609, 0.494]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.392860450203
varray: [0.621, 1.352, 0.609, 0.494]
inst: std_all
operation: std_all(slope_shift)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: slope_shift
SAMP_d18O_STD | 0.186000672042
varray: [3.914, 4.211, 4.292, 4.32]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.34219257809
varray: [-51.903, -49.732, -49.331, -48.873]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.19825
varray: [114.019, 113.667, 116.373, 116.734]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.57838765306
varray: [114.019, 113.667, 116.373, 116.734]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.19875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1332.1135
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.769
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.81536958645
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.46075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [3.914, 4.211, 4.292, 4.32]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-51.903, -49.732, -49.331, -48.873]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.18425
varray: [3.914, 4.211, 4.292, 4.32]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
dD_mean | -49.95975
varray: [-51.903, -49.732, -49.331, -48.873]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
d18O_stdev | 0.186000672042
varray: [3.914, 4.211, 4.292, 4.32]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
dD_stdev | 1.34219257809
varray: [-51.903, -49.732, -49.331, -48.873]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.18425
varray: [3.914, 4.211, 4.292, 4.32]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.95975
varray: [-51.903, -49.732, -49.331, -48.873]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.186000672042
varray: [3.914, 4.211, 4.292, 4.32]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.34219257809
varray: [-51.903, -49.732, -49.331, -48.873]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.3703114366
actual_array: [ 1.93 -15.84]
varray: [3.914, 4.211, 4.292, 4.32]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.18425
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.8383347806
actual_array: [ 18.1 -119. ]
varray: [-51.903, -49.732, -49.331, -48.873]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -49.95975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [37070.0, 37607.0, 38143.0, 38680.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7625
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 23.33575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.06025
varray: [3.164, 3.028, 2.997, 3.052]
inst: mean_all
operation: mean_all(residuals)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: residuals
SAMP_residuals_STD | 0.0727386875512
varray: [3.164, 3.028, 2.997, 3.052]
inst: std_all
operation: std_all(residuals)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5075
varray: [0.743, 0.349, 0.39, 0.548]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.178912455315
varray: [0.743, 0.349, 0.39, 0.548]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.823
varray: [1304.47, 1300.958, 1301.381, 1300.483]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.80238712083
varray: [1304.47, 1300.958, 1301.381, 1300.483]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3415
varray: [2.049, 2.418, 2.587, 2.312]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.225498706574
varray: [2.049, 2.418, 2.587, 2.312]
inst: std_all
operation: std_all(slope_shift)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: slope_shift
SAMP_d18O_STD | 0.292259445242
varray: [-6.743, -7.196, -7.336, -7.384]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.662852107688
varray: [-57.984, -59.147, -59.275, -59.441]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.04825
varray: [118.378, 113.544, 113.552, 114.719]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.28744200291
varray: [118.378, 113.544, 113.552, 114.719]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.06025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.823
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3415
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.350048844025
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.61075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-6.743, -7.196, -7.336, -7.384]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-57.984, -59.147, -59.275, -59.441]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.16475
varray: [-6.743, -7.196, -7.336, -7.384]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
dD_mean | -58.96175
varray: [-57.984, -59.147, -59.275, -59.441]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
d18O_stdev | 0.292259445242
varray: [-6.743, -7.196, -7.336, -7.384]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
dD_stdev | 0.662852107688
varray: [-57.984, -59.147, -59.275, -59.441]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.16475
varray: [-6.743, -7.196, -7.336, -7.384]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.96175
varray: [-57.984, -59.147, -59.275, -59.441]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.292259445242
varray: [-6.743, -7.196, -7.336, -7.384]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.662852107688
varray: [-57.984, -59.147, -59.275, -59.441]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.91694845863
actual_array: [ 1.93 -15.84]
varray: [-6.743, -7.196, -7.336, -7.384]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.16475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.2309318412
actual_array: [ 18.1 -119. ]
varray: [-57.984, -59.147, -59.275, -59.441]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -58.96175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [39269.0, 39805.0, 40342.0, 40879.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.95475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97325
varray: [3.018, 2.938, 3.021, 2.916]
inst: mean_all
operation: mean_all(residuals)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: residuals
SAMP_residuals_STD | 0.0541687179099
varray: [3.018, 2.938, 3.021, 2.916]
inst: std_all
operation: std_all(residuals)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.37675
varray: [0.476, 0.333, 0.476, 0.222]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.123237913539
varray: [0.476, 0.333, 0.476, 0.222]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.40075
varray: [1301.393, 1300.948, 1301.556, 1301.706]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.327781812593
varray: [1301.393, 1300.948, 1301.556, 1301.706]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.64025
varray: [2.713, 2.894, 2.209, 2.745]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.298122318297
varray: [2.713, 2.894, 2.209, 2.745]
inst: std_all
operation: std_all(slope_shift)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0877890463934
varray: [-12.499, -12.625, -12.669, -12.698]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.44458467843
varray: [-103.112, -107.072, -108.13, -108.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.15825
varray: [114.323, 113.282, 115.623, 113.405]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.08133170212
varray: [114.323, 113.282, 115.623, 113.405]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.37675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.40075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.64025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.18300227939
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.50075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.499, -12.625, -12.669, -12.698]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-103.112, -107.072, -108.13, -108.395]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.62275
varray: [-12.499, -12.625, -12.669, -12.698]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
dD_mean | -106.67725
varray: [-103.112, -107.072, -108.13, -108.395]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0877890463934
varray: [-12.499, -12.625, -12.669, -12.698]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
dD_stdev | 2.44458467843
varray: [-103.112, -107.072, -108.13, -108.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.62275
varray: [-12.499, -12.625, -12.669, -12.698]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
measured_dD_mean | -106.67725
varray: [-103.112, -107.072, -108.13, -108.395]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0877890463934
varray: [-12.499, -12.625, -12.669, -12.698]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.44458467843
varray: [-103.112, -107.072, -108.13, -108.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.3452561957
actual_array: [ 1.93 -15.84]
varray: [-12.499, -12.625, -12.669, -12.698]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.62275
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -117.71626455
actual_array: [ 18.1 -119. ]
varray: [-103.112, -107.072, -108.13, -108.395]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -106.67725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [41468.0, 42005.0, 42542.0, 43078.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.8045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.377
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.02775
varray: [2.987, 2.936, 3.072, 3.116]
inst: mean_all
operation: mean_all(residuals)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: residuals
SAMP_residuals_STD | 0.081291143429
varray: [2.987, 2.936, 3.072, 3.116]
inst: std_all
operation: std_all(residuals)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2535
varray: [0.282, 0.031, 0.487, 0.214]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.188326843546
varray: [0.282, 0.031, 0.487, 0.214]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1371.77225
varray: [1368.736, 1367.753, 1374.789, 1375.811]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.11442449398
varray: [1368.736, 1367.753, 1374.789, 1375.811]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.9345
varray: [-0.566, -0.673, -1.343, -1.156]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.374214287987
varray: [-0.566, -0.673, -1.343, -1.156]
inst: std_all
operation: std_all(slope_shift)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: slope_shift
SAMP_d18O_STD | 0.398568271024
varray: [3.265, 3.882, 4.046, 4.159]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.21747368547
varray: [-54.242, -49.308, -47.761, -47.142]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.8215
varray: [113.284, 111.449, 117.584, 116.969]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.94283451794
varray: [113.284, 111.449, 117.584, 116.969]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.02775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2535
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1371.77225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.9345
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04325
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04325
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.40833604689
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8375
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [3.265, 3.882, 4.046, 4.159]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-54.242, -49.308, -47.761, -47.142]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.838
varray: [3.265, 3.882, 4.046, 4.159]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
dD_mean | -49.61325
varray: [-54.242, -49.308, -47.761, -47.142]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
d18O_stdev | 0.398568271024
varray: [3.265, 3.882, 4.046, 4.159]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
dD_stdev | 3.21747368547
varray: [-54.242, -49.308, -47.761, -47.142]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.838
varray: [3.265, 3.882, 4.046, 4.159]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.61325
varray: [-54.242, -49.308, -47.761, -47.142]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.398568271024
varray: [3.265, 3.882, 4.046, 4.159]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.21747368547
varray: [-54.242, -49.308, -47.761, -47.142]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.02594508373
actual_array: [ 1.93 -15.84]
varray: [3.265, 3.882, 4.046, 4.159]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.838
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.5152915812
actual_array: [ 18.1 -119. ]
varray: [-54.242, -49.308, -47.761, -47.142]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -49.61325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [43664.0, 44201.0, 44738.0, 45274.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.75
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 62.9945
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.00975
varray: [3.067, 2.96, 2.992, 3.02]
inst: mean_all
operation: mean_all(residuals)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: residuals
SAMP_residuals_STD | 0.0453605922948
varray: [3.067, 2.96, 2.992, 3.02]
inst: std_all
operation: std_all(residuals)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4595
varray: [0.489, 0.088, 0.579, 0.682]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.259916012076
varray: [0.489, 0.088, 0.579, 0.682]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.3735
varray: [1305.577, 1302.189, 1301.063, 1300.665]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.23105617739
varray: [1305.577, 1302.189, 1301.063, 1300.665]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.96775
varray: [1.738, 2.007, 2.071, 2.055]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.155562152638
varray: [1.738, 2.007, 2.071, 2.055]
inst: std_all
operation: std_all(slope_shift)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: slope_shift
SAMP_d18O_STD | 0.122990514539
varray: [-0.247, -0.493, -0.509, -0.471]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.21973641278
varray: [-68.126, -70.047, -70.595, -70.801]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.27175
varray: [117.135, 113.085, 115.353, 115.514]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.66521897959
varray: [117.135, 113.085, 115.353, 115.514]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.00975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4595
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.3735
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.96775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06125
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06125
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.99446434386
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.38725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-0.247, -0.493, -0.509, -0.471]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-68.126, -70.047, -70.595, -70.801]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.43
varray: [-0.247, -0.493, -0.509, -0.471]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
dD_mean | -69.89225
varray: [-68.126, -70.047, -70.595, -70.801]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
d18O_stdev | 0.122990514539
varray: [-0.247, -0.493, -0.509, -0.471]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
dD_stdev | 1.21973641278
varray: [-68.126, -70.047, -70.595, -70.801]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.43
varray: [-0.247, -0.493, -0.509, -0.471]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.89225
varray: [-68.126, -70.047, -70.595, -70.801]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.122990514539
varray: [-0.247, -0.493, -0.509, -0.471]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.21973641278
varray: [-68.126, -70.047, -70.595, -70.801]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.21883641533
actual_array: [ 1.93 -15.84]
varray: [-0.247, -0.493, -0.509, -0.471]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.43
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.4214764058
actual_array: [ 18.1 -119. ]
varray: [-68.126, -70.047, -70.595, -70.801]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -69.89225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [45861.0, 46398.0, 46934.0, 47471.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.768
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.40425
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.991
varray: [2.95, 2.982, 2.981, 3.051]
inst: mean_all
operation: mean_all(residuals)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: residuals
SAMP_residuals_STD | 0.0426692707539
varray: [2.95, 2.982, 2.981, 3.051]
inst: std_all
operation: std_all(residuals)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31825
varray: [0.25, 0.223, 0.357, 0.443]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.10131592504
varray: [0.25, 0.223, 0.357, 0.443]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.2955
varray: [1300.422, 1302.378, 1301.195, 1301.187]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.807605720633
varray: [1300.422, 1302.378, 1301.195, 1301.187]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.99275
varray: [2.084, 1.92, 2.153, 1.814]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.154111161179
varray: [2.084, 1.92, 2.153, 1.814]
inst: std_all
operation: std_all(slope_shift)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0642105910267
varray: [0.346, 0.469, 0.466, 0.485]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.357175377091
varray: [-66.918, -66.186, -66.156, -66.303]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.37275
varray: [112.912, 114.443, 114.427, 115.709]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14414928367
varray: [112.912, 114.443, 114.427, 115.709]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.991
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.2955
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.99275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.60501465321
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.28625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.346, 0.469, 0.466, 0.485]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-66.918, -66.186, -66.156, -66.303]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.4415
varray: [0.346, 0.469, 0.466, 0.485]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
dD_mean | -66.39075
varray: [-66.918, -66.186, -66.156, -66.303]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0642105910267
varray: [0.346, 0.469, 0.466, 0.485]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
dD_stdev | 0.357175377091
varray: [-66.918, -66.186, -66.156, -66.303]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.4415
varray: [0.346, 0.469, 0.466, 0.485]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.39075
varray: [-66.918, -66.186, -66.156, -66.303]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0642105910267
varray: [0.346, 0.469, 0.466, 0.485]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.357175377091
varray: [-66.918, -66.186, -66.156, -66.303]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.35207749396
actual_array: [ 1.93 -15.84]
varray: [0.346, 0.469, 0.466, 0.485]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.4415
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.1570153296
actual_array: [ 18.1 -119. ]
varray: [-66.918, -66.186, -66.156, -66.303]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.39075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [48058.0, 48594.0, 49131.0, 49667.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.78675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.48225
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0
varray: [2.954, 2.995, 3.044, 3.007]
inst: mean_all
operation: mean_all(residuals)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: residuals
SAMP_residuals_STD | 0.0370854868289
varray: [2.954, 2.995, 3.044, 3.007]
inst: std_all
operation: std_all(residuals)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2405
varray: [0.119, 0.075, 0.406, 0.362]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.167635517319
varray: [0.119, 0.075, 0.406, 0.362]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1383.356
varray: [1380.841, 1383.721, 1384.152, 1384.71]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.72485574276
varray: [1380.841, 1383.721, 1384.152, 1384.71]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.8905
varray: [-0.694, -0.88, -1.094, -0.894]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.163449278575
varray: [-0.694, -0.88, -1.094, -0.894]
inst: std_all
operation: std_all(slope_shift)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0292859465728
varray: [1.948, 2.013, 2.006, 1.995]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.861636040719
varray: [-53.718, -52.301, -52.047, -51.79]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.80625
varray: [112.231, 113.258, 115.071, 114.665]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.30627137941
varray: [112.231, 113.258, 115.071, 114.665]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2405
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1383.356
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.8905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.071
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.071
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.31195050472
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.85275
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [1.948, 2.013, 2.006, 1.995]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-53.718, -52.301, -52.047, -51.79]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.9905
varray: [1.948, 2.013, 2.006, 1.995]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
dD_mean | -52.464
varray: [-53.718, -52.301, -52.047, -51.79]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0292859465728
varray: [1.948, 2.013, 2.006, 1.995]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
dD_stdev | 0.861636040719
varray: [-53.718, -52.301, -52.047, -51.79]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.9905
varray: [1.948, 2.013, 2.006, 1.995]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.464
varray: [-53.718, -52.301, -52.047, -51.79]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0292859465728
varray: [1.948, 2.013, 2.006, 1.995]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.861636040719
varray: [-53.718, -52.301, -52.047, -51.79]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.811504264798
actual_array: [ 1.93 -15.84]
varray: [1.948, 2.013, 2.006, 1.995]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.9905
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.1730560858
actual_array: [ 18.1 -119. ]
varray: [-53.718, -52.301, -52.047, -51.79]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.464
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [50255.0, 50791.0, 51327.0, 51863.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.77775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 74.57825
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9855
varray: [2.904, 3.045, 2.904, 3.089]
inst: mean_all
operation: mean_all(residuals)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: residuals
SAMP_residuals_STD | 0.0958070978581
varray: [2.904, 3.045, 2.904, 3.089]
inst: std_all
operation: std_all(residuals)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3685
varray: [0.1, 0.491, 0.397, 0.486]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.184134914307
varray: [0.1, 0.491, 0.397, 0.486]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.97625
varray: [1300.52, 1302.852, 1297.514, 1303.019]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.5746588609
varray: [1300.52, 1302.852, 1297.514, 1303.019]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.1045
varray: [1.766, 1.944, 2.122, 2.586]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.352368651651
varray: [1.766, 1.944, 2.122, 2.586]
inst: std_all
operation: std_all(slope_shift)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0575210106541
varray: [1.725, 1.803, 1.666, 1.706]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.242551712699
varray: [-55.926, -55.874, -56.391, -55.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.62975
varray: [111.537, 115.962, 111.533, 115.487]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.42657033348
varray: [111.537, 115.962, 111.533, 115.487]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9855
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3685
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.97625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.1045
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0855
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0855
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.78410941062
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.02925
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [1.725, 1.803, 1.666, 1.706]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-55.926, -55.874, -56.391, -55.925]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.725
varray: [1.725, 1.803, 1.666, 1.706]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
dD_mean | -56.029
varray: [-55.926, -55.874, -56.391, -55.925]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0575210106541
varray: [1.725, 1.803, 1.666, 1.706]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
dD_stdev | 0.242551712699
varray: [-55.926, -55.874, -56.391, -55.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.725
varray: [1.725, 1.803, 1.666, 1.706]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.029
varray: [-55.926, -55.874, -56.391, -55.925]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0575210106541
varray: [1.725, 1.803, 1.666, 1.706]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.242551712699
varray: [-55.926, -55.874, -56.391, -55.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.07555990866
actual_array: [ 1.93 -15.84]
varray: [1.725, 1.803, 1.666, 1.706]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.725
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.4967184409
actual_array: [ 18.1 -119. ]
varray: [-55.926, -55.874, -56.391, -55.925]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.029
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [52451.0, 52986.0, 53522.0, 54059.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.79225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.8015
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.4755
varray: [2.951, 0.0, 0.0, 2.951]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: residuals
SAMP_residuals_STD | 1.70376064438
varray: [2.951, 0.0, 0.0, 2.951]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.232
varray: [0.464, 0.0, 0.0, 0.464]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.267890524904
varray: [0.464, 0.0, 0.0, 0.464]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 666.554
varray: [1333.108, 0.0, 0.0, 1333.108]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 769.670262659
varray: [1333.108, 0.0, 0.0, 1333.108]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.257
varray: [0.514, 0.0, 0.0, 0.514]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.296758038363
varray: [0.514, 0.0, 0.0, 0.514]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: slope_shift
SAMP_d18O_STD | 2.52186597582
varray: [4.368, 0.0, 0.0, 4.368]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 27.8115398171
varray: [-48.171, 0.0, 0.0, -48.171]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.926
varray: [111.852, 0.0, 0.0, 111.852]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 64.5777823094
varray: [111.852, 0.0, 0.0, 111.852]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.4755
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.232
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 666.554
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.257
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.5955
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.5955
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -51.9537609899
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -60.733
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [4.368, 0.0, 0.0, 4.368]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-48.171, 0.0, 0.0, -48.171]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.184
varray: [4.368, 0.0, 0.0, 4.368]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
dD_mean | -24.0855
varray: [-48.171, 0.0, 0.0, -48.171]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
d18O_stdev | 2.52186597582
varray: [4.368, 0.0, 0.0, 4.368]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
dD_stdev | 27.8115398171
varray: [-48.171, 0.0, 0.0, -48.171]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.184
varray: [4.368, 0.0, 0.0, 4.368]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.0855
varray: [-48.171, 0.0, 0.0, -48.171]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.52186597582
varray: [4.368, 0.0, 0.0, 4.368]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
measured_dD_stdev | 27.8115398171
varray: [-48.171, 0.0, 0.0, -48.171]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.619056931136
actual_array: [ 1.93 -15.84]
varray: [4.368, 0.0, 0.0, 4.368]
fit_rsq: 1.0
fit_intercept: 2.80644310636
except: True
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.184
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -40.715678205
actual_array: [ 18.1 -119. ]
varray: [-48.171, 0.0, 0.0, -48.171]
fit_rsq: 1.0
fit_intercept: 19.5866234865
except: True
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -24.0855
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [54647.0, 55183.0, 55782.0, 56380.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1307.30225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -642.22375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89225
varray: [2.838, 2.883, 2.923, 2.925]
inst: mean_all
operation: mean_all(residuals)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: residuals
SAMP_residuals_STD | 0.0410152410696
varray: [2.838, 2.883, 2.923, 2.925]
inst: std_all
operation: std_all(residuals)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.351
varray: [0.266, 0.428, 0.19, 0.52]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.150151035072
varray: [0.266, 0.428, 0.19, 0.52]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.985
varray: [1295.177, 1296.323, 1298.545, 1297.895]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.52411810566
varray: [1295.177, 1296.323, 1298.545, 1297.895]
inst: std_all
operation: std_all(baseline_shift)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.30125
varray: [1.839, 2.232, 3.006, 2.128]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.498383637372
varray: [1.839, 2.232, 3.006, 2.128]
inst: std_all
operation: std_all(slope_shift)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0159661099416
varray: [-0.325, -0.33, -0.303, -0.341]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.602406493436
varray: [-65.92, -67.019, -67.199, -67.129]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.3975
varray: [108.223, 110.465, 110.654, 112.248]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.65561156878
varray: [108.223, 110.465, 110.654, 112.248]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.351
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.985
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.30125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.17875
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17875
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.82057961576
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.2615
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-0.325, -0.33, -0.303, -0.341]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-65.92, -67.019, -67.199, -67.129]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.32475
varray: [-0.325, -0.33, -0.303, -0.341]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
dD_mean | -66.81675
varray: [-65.92, -67.019, -67.199, -67.129]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0159661099416
varray: [-0.325, -0.33, -0.303, -0.341]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
dD_stdev | 0.602406493436
varray: [-65.92, -67.019, -67.199, -67.129]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.32475
varray: [-0.325, -0.33, -0.303, -0.341]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.81675
varray: [-65.92, -67.019, -67.199, -67.129]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0159661099416
varray: [-0.325, -0.33, -0.303, -0.341]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.602406493436
varray: [-65.92, -67.019, -67.199, -67.129]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.11415898966
actual_array: [ 1.93 -15.84]
varray: [-0.325, -0.33, -0.303, -0.341]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.32475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.5541766658
actual_array: [ 18.1 -119. ]
varray: [-65.92, -67.019, -67.199, -67.129]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.81675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [57031.0, 57567.0, 58103.0, 58640.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.8855
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.79275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98575
varray: [3.005, 2.993, 2.947, 2.998]
inst: mean_all
operation: mean_all(residuals)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: residuals
SAMP_residuals_STD | 0.0262979720384
varray: [3.005, 2.993, 2.947, 2.998]
inst: std_all
operation: std_all(residuals)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.54275
varray: [0.48, 0.468, 0.581, 0.642]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0833441659626
varray: [0.48, 0.468, 0.581, 0.642]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.16725
varray: [1300.13, 1298.01, 1299.647, 1298.882]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.92693810473
varray: [1300.13, 1298.01, 1299.647, 1298.882]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.436
varray: [2.495, 2.446, 2.571, 2.232]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.145398303521
varray: [2.495, 2.446, 2.571, 2.232]
inst: std_all
operation: std_all(slope_shift)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: slope_shift
SAMP_d18O_STD | 0.203771726858
varray: [-7.033, -7.397, -7.44, -7.471]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.340733889323
varray: [-60.341, -59.75, -59.554, -59.752]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.455
varray: [114.013, 111.834, 113.488, 114.485]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.15484400101
varray: [114.013, 111.834, 113.488, 114.485]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.54275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.16725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.436
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.77596873982
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.204
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.033, -7.397, -7.44, -7.471]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-60.341, -59.75, -59.554, -59.752]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.33525
varray: [-7.033, -7.397, -7.44, -7.471]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
dD_mean | -59.84925
varray: [-60.341, -59.75, -59.554, -59.752]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
d18O_stdev | 0.203771726858
varray: [-7.033, -7.397, -7.44, -7.471]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
dD_stdev | 0.340733889323
varray: [-60.341, -59.75, -59.554, -59.752]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.33525
varray: [-7.033, -7.397, -7.44, -7.471]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.84925
varray: [-60.341, -59.75, -59.554, -59.752]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.203771726858
varray: [-7.033, -7.397, -7.44, -7.471]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.340733889323
varray: [-60.341, -59.75, -59.554, -59.752]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0865209154
actual_array: [ 1.93 -15.84]
varray: [-7.033, -7.397, -7.44, -7.471]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.33525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.0583512915
actual_array: [ 18.1 -119. ]
varray: [-60.341, -59.75, -59.554, -59.752]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.84925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [59226.0, 59762.0, 60297.0, 60834.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.792
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.6105
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.90025
varray: [2.855, 2.961, 2.898, 2.887]
inst: mean_all
operation: mean_all(residuals)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: residuals
SAMP_residuals_STD | 0.0444175265708
varray: [2.855, 2.961, 2.898, 2.887]
inst: std_all
operation: std_all(residuals)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4365
varray: [0.173, 0.617, 0.564, 0.392]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.200207392471
varray: [0.173, 0.617, 0.564, 0.392]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.13125
varray: [1297.732, 1296.775, 1296.742, 1297.276]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.469142036062
varray: [1297.732, 1296.775, 1296.742, 1297.276]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.8865
varray: [2.773, 2.793, 2.939, 3.041]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.126821396723
varray: [2.773, 2.793, 2.939, 3.041]
inst: std_all
operation: std_all(slope_shift)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: slope_shift
SAMP_d18O_STD | 0.159977081692
varray: [-12.519, -12.764, -12.886, -12.819]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.52661703997
varray: [-103.642, -107.936, -108.899, -108.981]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.9715
varray: [110.425, 112.003, 111.271, 110.187]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.830219850401
varray: [110.425, 112.003, 111.271, 110.187]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.90025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4365
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.13125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.8865
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.17075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.56007815044
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.6875
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.519, -12.764, -12.886, -12.819]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-103.642, -107.936, -108.899, -108.981]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.747
varray: [-12.519, -12.764, -12.886, -12.819]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
dD_mean | -107.3645
varray: [-103.642, -107.936, -108.899, -108.981]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
d18O_stdev | 0.159977081692
varray: [-12.519, -12.764, -12.886, -12.819]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
dD_stdev | 2.52661703997
varray: [-103.642, -107.936, -108.899, -108.981]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.747
varray: [-12.519, -12.764, -12.886, -12.819]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.3645
varray: [-103.642, -107.936, -108.899, -108.981]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.159977081692
varray: [-12.519, -12.764, -12.886, -12.819]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.52661703997
varray: [-103.642, -107.936, -108.899, -108.981]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.4688302588
actual_array: [ 1.93 -15.84]
varray: [-12.519, -12.764, -12.886, -12.819]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.747
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -118.356990203
actual_array: [ 18.1 -119. ]
varray: [-103.642, -107.936, -108.899, -108.981]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.3645
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [61420.0, 61956.0, 62492.0, 63028.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.8775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.6465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.002
varray: [2.864, 3.005, 3.099, 3.04]
inst: mean_all
operation: mean_all(residuals)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: residuals
SAMP_residuals_STD | 0.0998432104185
varray: [2.864, 3.005, 3.099, 3.04]
inst: std_all
operation: std_all(residuals)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41725
varray: [0.219, 0.612, 0.577, 0.261]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.205884069321
varray: [0.219, 0.612, 0.577, 0.261]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.07325
varray: [1297.737, 1301.664, 1302.209, 1302.683]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.26279861749
varray: [1297.737, 1301.664, 1302.209, 1302.683]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.20275
varray: [2.17, 2.287, 2.372, 1.982]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.168865577704
varray: [2.17, 2.287, 2.372, 1.982]
inst: std_all
operation: std_all(slope_shift)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: slope_shift
SAMP_d18O_STD | 0.279012544521
varray: [-3.972, -3.504, -3.394, -3.378]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.21484431357
varray: [-74.414, -70.529, -70.075, -69.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.559
varray: [110.873, 116.057, 116.134, 115.172]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.49579606539
varray: [110.873, 116.057, 116.134, 115.172]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.002
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.07325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.20275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.069
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.069
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.24682513839
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.1
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-3.972, -3.504, -3.394, -3.378]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-74.414, -70.529, -70.075, -69.562]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.562
varray: [-3.972, -3.504, -3.394, -3.378]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
dD_mean | -71.145
varray: [-74.414, -70.529, -70.075, -69.562]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
d18O_stdev | 0.279012544521
varray: [-3.972, -3.504, -3.394, -3.378]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
dD_stdev | 2.21484431357
varray: [-74.414, -70.529, -70.075, -69.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.562
varray: [-3.972, -3.504, -3.394, -3.378]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.145
varray: [-74.414, -70.529, -70.075, -69.562]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.279012544521
varray: [-3.972, -3.504, -3.394, -3.378]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.21484431357
varray: [-74.414, -70.529, -70.075, -69.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.33379790902
actual_array: [ 1.93 -15.84]
varray: [-3.972, -3.504, -3.394, -3.378]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -84.5894197481
actual_array: [ 18.1 -119. ]
varray: [-74.414, -70.529, -70.075, -69.562]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -71.145
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [63615.0, 64151.0, 64687.0, 65224.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.77575
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.7045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.29675
varray: [3.261, 3.338, 3.273, 3.315]
inst: mean_all
operation: mean_all(residuals)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: residuals
SAMP_residuals_STD | 0.035947878936
varray: [3.261, 3.338, 3.273, 3.315]
inst: std_all
operation: std_all(residuals)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.41375
varray: [-3.081, -3.345, -3.692, -3.537]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.263351950819
varray: [-3.081, -3.345, -3.692, -3.537]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1321.24475
varray: [1318.072, 1322.448, 1322.32, 1322.139]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.11896191172
varray: [1318.072, 1322.448, 1322.32, 1322.139]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -9.45475
varray: [-9.033, -9.588, -9.533, -9.665]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.286331247567
varray: [-9.033, -9.588, -9.533, -9.665]
inst: std_all
operation: std_all(slope_shift)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: slope_shift
SAMP_d18O_STD | 0.185693295517
varray: [5.929, 6.241, 6.322, 6.316]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.863738145119
varray: [-51.829, -50.424, -50.043, -49.975]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 123.61875
varray: [122.722, 124.692, 123.378, 123.683]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.820185903723
varray: [122.722, 124.692, 123.378, 123.683]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.29675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.41375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1321.24475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.45475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.22575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.22575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.35102572452
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.95975
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [5.929, 6.241, 6.322, 6.316]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-51.829, -50.424, -50.043, -49.975]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 6.202
varray: [5.929, 6.241, 6.322, 6.316]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
dD_mean | -50.56775
varray: [-51.829, -50.424, -50.043, -49.975]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
d18O_stdev | 0.185693295517
varray: [5.929, 6.241, 6.322, 6.316]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
dD_stdev | 0.863738145119
varray: [-51.829, -50.424, -50.043, -49.975]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
measured_d18O_mean | 6.202
varray: [5.929, 6.241, 6.322, 6.316]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.56775
varray: [-51.829, -50.424, -50.043, -49.975]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.185693295517
varray: [5.929, 6.241, 6.322, 6.316]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.863738145119
varray: [-51.829, -50.424, -50.043, -49.975]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.37708460195
actual_array: [ 1.93 -15.84]
varray: [5.929, 6.241, 6.322, 6.316]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 6.202
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.405175373
actual_array: [ 18.1 -119. ]
varray: [-51.829, -50.424, -50.043, -49.975]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -50.56775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [65811.0, 66347.0, 66883.0, 67420.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.481
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 12.467
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01375
varray: [2.991, 3.04, 3.0, 3.024]
inst: mean_all
operation: mean_all(residuals)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: residuals
SAMP_residuals_STD | 0.0223662692463
varray: [2.991, 3.04, 3.0, 3.024]
inst: std_all
operation: std_all(residuals)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.482
varray: [0.43, 0.458, 0.486, 0.554]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.053166405433
varray: [0.43, 0.458, 0.486, 0.554]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1359.69425
varray: [1358.193, 1360.865, 1360.413, 1359.306]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.19606950049
varray: [1358.193, 1360.865, 1360.413, 1359.306]
inst: std_all
operation: std_all(baseline_shift)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.48075
varray: [-0.533, -0.314, -0.484, -0.592]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.119614310738
varray: [-0.533, -0.314, -0.484, -0.592]
inst: std_all
operation: std_all(slope_shift)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0745363446023
varray: [7.184, 7.312, 7.358, 7.308]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.889977715077
varray: [-42.786, -41.321, -40.827, -41.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.42175
varray: [114.794, 115.67, 115.155, 116.068]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.561097956391
varray: [114.794, 115.67, 115.155, 116.068]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.482
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1359.69425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.48075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05725
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05725
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.8642136112
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.23725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [7.184, 7.312, 7.358, 7.308]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-42.786, -41.321, -40.827, -41.012]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 7.2905
varray: [7.184, 7.312, 7.358, 7.308]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
dD_mean | -41.4865
varray: [-42.786, -41.321, -40.827, -41.012]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0745363446023
varray: [7.184, 7.312, 7.358, 7.308]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
dD_stdev | 0.889977715077
varray: [-42.786, -41.321, -40.827, -41.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
measured_d18O_mean | 7.2905
varray: [7.184, 7.312, 7.358, 7.308]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.4865
varray: [-42.786, -41.321, -40.827, -41.012]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0745363446023
varray: [7.184, 7.312, 7.358, 7.308]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.889977715077
varray: [-42.786, -41.321, -40.827, -41.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.45966301379
actual_array: [ 1.93 -15.84]
varray: [7.184, 7.312, 7.358, 7.308]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 7.2905
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.9386932517
actual_array: [ 18.1 -119. ]
varray: [-42.786, -41.321, -40.827, -41.012]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -41.4865
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [68007.0, 68542.0, 69078.0, 69615.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.764
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 50.9165
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.13775
varray: [3.089, 3.127, 3.214, 3.121]
inst: mean_all
operation: mean_all(residuals)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: residuals
SAMP_residuals_STD | 0.0535
varray: [3.089, 3.127, 3.214, 3.121]
inst: std_all
operation: std_all(residuals)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.422
varray: [-2.45, -2.303, -2.271, -2.664]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.179174030856
varray: [-2.45, -2.303, -2.271, -2.664]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1361.855
varray: [1361.341, 1362.153, 1361.924, 1362.002]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.355607461489
varray: [1361.341, 1362.153, 1361.924, 1362.002]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -8.304
varray: [-8.05, -8.721, -8.43, -8.015]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.335560228076
varray: [-8.05, -8.721, -8.43, -8.015]
inst: std_all
operation: std_all(slope_shift)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: slope_shift
SAMP_d18O_STD | 0.123896125847
varray: [5.0, 4.742, 4.766, 4.751]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.16140977552
varray: [-54.199, -56.096, -56.634, -56.659]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 121.95975
varray: [120.446, 123.497, 123.299, 120.597]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.66385643912
varray: [120.446, 123.497, 123.299, 120.597]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.13775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.422
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1361.855
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 8.304
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.17355910127
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.30075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [5.0, 4.742, 4.766, 4.751]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-54.199, -56.096, -56.634, -56.659]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.81475
varray: [5.0, 4.742, 4.766, 4.751]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
dD_mean | -55.897
varray: [-54.199, -56.096, -56.634, -56.659]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
d18O_stdev | 0.123896125847
varray: [5.0, 4.742, 4.766, 4.751]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
dD_stdev | 1.16140977552
varray: [-54.199, -56.096, -56.634, -56.659]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.81475
varray: [5.0, 4.742, 4.766, 4.751]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.897
varray: [-54.199, -56.096, -56.634, -56.659]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.123896125847
varray: [5.0, 4.742, 4.766, 4.751]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.16140977552
varray: [-54.199, -56.096, -56.634, -56.659]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.99738143078
actual_array: [ 1.93 -15.84]
varray: [5.0, 4.742, 4.766, 4.751]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.81475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.3736543649
actual_array: [ 18.1 -119. ]
varray: [-54.199, -56.096, -56.634, -56.659]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -55.897
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [70202.0, 70738.0, 71275.0, 71811.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.64
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 53.07725
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.992
varray: [2.974, 2.979, 3.043, 2.972]
inst: mean_all
operation: mean_all(residuals)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: residuals
SAMP_residuals_STD | 0.034127212993
varray: [2.974, 2.979, 3.043, 2.972]
inst: std_all
operation: std_all(residuals)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.514
varray: [0.476, 0.48, 0.627, 0.473]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0753878858527
varray: [0.476, 0.48, 0.627, 0.473]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.82275
varray: [1305.276, 1305.183, 1305.097, 1303.735]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.728841260724
varray: [1305.276, 1305.183, 1305.097, 1303.735]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.64775
varray: [1.595, 1.783, 1.539, 1.674]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.105815484059
varray: [1.595, 1.783, 1.539, 1.674]
inst: std_all
operation: std_all(slope_shift)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0350903595498
varray: [3.967, 4.02, 4.051, 4.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.624090271248
varray: [-53.401, -52.534, -51.929, -52.308]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.61325
varray: [114.087, 114.736, 115.618, 114.012]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.744542532924
varray: [114.087, 114.736, 115.618, 114.012]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.992
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.514
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.82275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.64775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.079
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.079
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.57245197004
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.04575
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [3.967, 4.02, 4.051, 4.002]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-53.401, -52.534, -51.929, -52.308]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.01
varray: [3.967, 4.02, 4.051, 4.002]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
dD_mean | -52.543
varray: [-53.401, -52.534, -51.929, -52.308]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0350903595498
varray: [3.967, 4.02, 4.051, 4.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
dD_stdev | 0.624090271248
varray: [-53.401, -52.534, -51.929, -52.308]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.01
varray: [3.967, 4.02, 4.051, 4.002]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.543
varray: [-53.401, -52.534, -51.929, -52.308]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0350903595498
varray: [3.967, 4.02, 4.051, 4.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.624090271248
varray: [-53.401, -52.534, -51.929, -52.308]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.19700938032
actual_array: [ 1.93 -15.84]
varray: [3.967, 4.02, 4.051, 4.002]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.01
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.2467080706
actual_array: [ 18.1 -119. ]
varray: [-53.401, -52.534, -51.929, -52.308]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.543
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [72399.0, 72935.0, 73471.0, 74007.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.78575
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9465
varray: [2.948, 2.897, 2.953, 2.988]
inst: mean_all
operation: mean_all(residuals)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: residuals
SAMP_residuals_STD | 0.0374922214155
varray: [2.948, 2.897, 2.953, 2.988]
inst: std_all
operation: std_all(residuals)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.62225
varray: [-0.686, -0.748, -0.562, -0.493]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115776724777
varray: [-0.686, -0.748, -0.562, -0.493]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1307.13975
varray: [1307.988, 1304.252, 1308.183, 1308.136]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.92695898157
varray: [1307.988, 1304.252, 1308.183, 1308.136]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.9355
varray: [-0.861, -0.915, -1.361, -0.605]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.31424565762
varray: [-0.861, -0.915, -1.361, -0.605]
inst: std_all
operation: std_all(slope_shift)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0598665181884
varray: [3.506, 3.362, 3.41, 3.426]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.294881190764
varray: [-52.566, -53.279, -52.852, -52.966]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.5145
varray: [115.544, 112.707, 117.166, 116.641]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.98992303034
varray: [115.544, 112.707, 117.166, 116.641]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9465
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.62225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1307.13975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.9355
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1245
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1245
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.05405405405
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1445
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [3.506, 3.362, 3.41, 3.426]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-52.566, -53.279, -52.852, -52.966]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.426
varray: [3.506, 3.362, 3.41, 3.426]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
dD_mean | -52.91575
varray: [-52.566, -53.279, -52.852, -52.966]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0598665181884
varray: [3.506, 3.362, 3.41, 3.426]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
dD_stdev | 0.294881190764
varray: [-52.566, -53.279, -52.852, -52.966]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.426
varray: [3.506, 3.362, 3.41, 3.426]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.91575
varray: [-52.566, -53.279, -52.852, -52.966]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0598665181884
varray: [3.506, 3.362, 3.41, 3.426]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.294881190764
varray: [-52.566, -53.279, -52.852, -52.966]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.616186419808
actual_array: [ 1.93 -15.84]
varray: [3.506, 3.362, 3.41, 3.426]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.5942242398
actual_array: [ 18.1 -119. ]
varray: [-52.566, -53.279, -52.852, -52.966]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.91575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [74591.0, 75127.0, 75663.0, 76199.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.83125
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.638
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9375
varray: [2.892, 2.961, 2.917, 2.98]
inst: mean_all
operation: mean_all(residuals)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: residuals
SAMP_residuals_STD | 0.040203648259
varray: [2.892, 2.961, 2.917, 2.98]
inst: std_all
operation: std_all(residuals)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3475
varray: [0.288, 0.306, 0.243, 0.553]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.139538525146
varray: [0.288, 0.306, 0.243, 0.553]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.214
varray: [1299.871, 1300.536, 1298.318, 1298.131]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.17686561113
varray: [1299.871, 1300.536, 1298.318, 1298.131]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3925
varray: [2.433, 2.455, 1.984, 2.698]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.297628963644
varray: [2.433, 2.455, 1.984, 2.698]
inst: std_all
operation: std_all(slope_shift)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: slope_shift
SAMP_d18O_STD | 0.131258015628
varray: [-2.383, -2.602, -2.642, -2.673]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.85823942852
varray: [-69.658, -71.006, -71.316, -71.595]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.63325
varray: [112.307, 113.226, 112.22, 112.78]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.465511457073
varray: [112.307, 113.226, 112.22, 112.78]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.214
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1335
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1335
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.34711820254
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.02575
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-2.383, -2.602, -2.642, -2.673]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-69.658, -71.006, -71.316, -71.595]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.575
varray: [-2.383, -2.602, -2.642, -2.673]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
dD_mean | -70.89375
varray: [-69.658, -71.006, -71.316, -71.595]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
d18O_stdev | 0.131258015628
varray: [-2.383, -2.602, -2.642, -2.673]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
dD_stdev | 0.85823942852
varray: [-69.658, -71.006, -71.316, -71.595]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.575
varray: [-2.383, -2.602, -2.642, -2.673]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.89375
varray: [-69.658, -71.006, -71.316, -71.595]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.131258015628
varray: [-2.383, -2.602, -2.642, -2.673]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.85823942852
varray: [-69.658, -71.006, -71.316, -71.595]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.35216732336
actual_array: [ 1.93 -15.84]
varray: [-2.383, -2.602, -2.642, -2.673]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.575
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -84.3551784671
actual_array: [ 18.1 -119. ]
varray: [-69.658, -71.006, -71.316, -71.595]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -70.89375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [76784.0, 77321.0, 77856.0, 78393.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.84025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.56375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01075
varray: [2.954, 3.067, 2.994, 3.028]
inst: mean_all
operation: mean_all(residuals)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: residuals
SAMP_residuals_STD | 0.0481758930033
varray: [2.954, 3.067, 2.994, 3.028]
inst: std_all
operation: std_all(residuals)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4955
varray: [0.417, 0.533, 0.398, 0.634]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109934829179
varray: [0.417, 0.533, 0.398, 0.634]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.0275
varray: [1299.242, 1301.933, 1301.434, 1301.501]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.21070021062
varray: [1299.242, 1301.933, 1301.434, 1301.501]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.143
varray: [2.098, 2.177, 2.02, 2.277]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.10994847278
varray: [2.098, 2.177, 2.02, 2.277]
inst: std_all
operation: std_all(slope_shift)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0988310848536
varray: [-7.204, -7.389, -7.396, -7.416]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.738815437034
varray: [-61.032, -59.685, -59.595, -59.429]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.04825
varray: [113.019, 116.544, 115.185, 115.445]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.47549254488
varray: [113.019, 116.544, 115.185, 115.445]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4955
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.0275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.143
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06025
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06025
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.9619016607
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.61075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.204, -7.389, -7.396, -7.416]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-61.032, -59.685, -59.595, -59.429]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.35125
varray: [-7.204, -7.389, -7.396, -7.416]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
dD_mean | -59.93525
varray: [-61.032, -59.685, -59.595, -59.429]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0988310848536
varray: [-7.204, -7.389, -7.396, -7.416]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
dD_stdev | 0.738815437034
varray: [-61.032, -59.685, -59.595, -59.429]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.35125
varray: [-7.204, -7.389, -7.396, -7.416]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.93525
varray: [-61.032, -59.685, -59.595, -59.429]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0988310848536
varray: [-7.204, -7.389, -7.396, -7.416]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.738815437034
varray: [-61.032, -59.685, -59.595, -59.429]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1024338732
actual_array: [ 1.93 -15.84]
varray: [-7.204, -7.389, -7.396, -7.416]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.35125
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.1385294016
actual_array: [ 18.1 -119. ]
varray: [-61.032, -59.685, -59.595, -59.429]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.93525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [78978.0, 79514.0, 80050.0, 80586.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.767
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.75025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9185
varray: [2.888, 2.929, 2.932, 2.925]
inst: mean_all
operation: mean_all(residuals)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: residuals
SAMP_residuals_STD | 0.020534523775
varray: [2.888, 2.929, 2.932, 2.925]
inst: std_all
operation: std_all(residuals)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4135
varray: [0.273, 0.496, 0.322, 0.563]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.138165359865
varray: [0.273, 0.496, 0.322, 0.563]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.92175
varray: [1297.432, 1298.327, 1297.627, 1298.301]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.459996648539
varray: [1297.432, 1298.327, 1297.627, 1298.301]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.6
varray: [2.547, 2.611, 2.688, 2.554]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0652942059706
varray: [2.547, 2.611, 2.688, 2.554]
inst: std_all
operation: std_all(slope_shift)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: slope_shift
SAMP_d18O_STD | 0.146438553667
varray: [-12.545, -12.785, -12.853, -12.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.48731654265
varray: [-103.697, -107.872, -108.706, -109.112]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.59275
varray: [110.727, 112.015, 111.36, 112.269]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.692653051679
varray: [110.727, 112.015, 111.36, 112.269]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9185
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.92175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.6
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.96580918268
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.06625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.545, -12.785, -12.853, -12.854]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-103.697, -107.872, -108.706, -109.112]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.75925
varray: [-12.545, -12.785, -12.853, -12.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
dD_mean | -107.34675
varray: [-103.697, -107.872, -108.706, -109.112]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
d18O_stdev | 0.146438553667
varray: [-12.545, -12.785, -12.853, -12.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
dD_stdev | 2.48731654265
varray: [-103.697, -107.872, -108.706, -109.112]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.75925
varray: [-12.545, -12.785, -12.853, -12.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.34675
varray: [-103.697, -107.872, -108.706, -109.112]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.146438553667
varray: [-12.545, -12.785, -12.853, -12.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.48731654265
varray: [-103.697, -107.872, -108.706, -109.112]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.4810136171
actual_array: [ 1.93 -15.84]
varray: [-12.545, -12.785, -12.853, -12.854]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.75925
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -118.340441814
actual_array: [ 18.1 -119. ]
varray: [-103.697, -107.872, -108.706, -109.112]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.34675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [81172.0, 81708.0, 82244.0, 82780.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.85925
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.856
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94325
varray: [3.039, 2.874, 2.982, 2.878]
inst: mean_all
operation: mean_all(residuals)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: residuals
SAMP_residuals_STD | 0.0810817488711
varray: [3.039, 2.874, 2.982, 2.878]
inst: std_all
operation: std_all(residuals)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.43375
varray: [0.645, 0.182, 0.671, 0.237]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.260129935481
varray: [0.645, 0.182, 0.671, 0.237]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1339.78075
varray: [1342.643, 1337.706, 1338.798, 1339.976]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.12139755429
varray: [1342.643, 1337.706, 1338.798, 1339.976]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.541
varray: [0.496, 0.769, 0.27, 0.629]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.212284400432
varray: [0.496, 0.769, 0.27, 0.629]
inst: std_all
operation: std_all(slope_shift)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: slope_shift
SAMP_d18O_STD | 0.331395534068
varray: [0.103, 0.612, 0.781, 0.826]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.24111616537
varray: [-66.249, -62.783, -61.66, -61.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.7345
varray: [116.06, 109.976, 112.671, 112.231]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.51169378441
varray: [116.06, 109.976, 112.671, 112.231]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.43375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1339.78075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.541
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.15988277434
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.9245
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.103, 0.612, 0.781, 0.826]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-66.249, -62.783, -61.66, -61.366]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.5805
varray: [0.103, 0.612, 0.781, 0.826]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
dD_mean | -63.0145
varray: [-66.249, -62.783, -61.66, -61.366]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
d18O_stdev | 0.331395534068
varray: [0.103, 0.612, 0.781, 0.826]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
dD_stdev | 2.24111616537
varray: [-66.249, -62.783, -61.66, -61.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.5805
varray: [0.103, 0.612, 0.781, 0.826]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.0145
varray: [-66.249, -62.783, -61.66, -61.366]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.331395534068
varray: [0.103, 0.612, 0.781, 0.826]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.24111616537
varray: [-66.249, -62.783, -61.66, -61.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.21383367288
actual_array: [ 1.93 -15.84]
varray: [0.103, 0.612, 0.781, 0.826]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.5805
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.0093252801
actual_array: [ 18.1 -119. ]
varray: [-66.249, -62.783, -61.66, -61.366]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -63.0145
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [83366.0, 83902.0, 84438.0, 84975.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.8345
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 31.003
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.93575
varray: [2.919, 2.879, 2.944, 3.001]
inst: mean_all
operation: mean_all(residuals)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: residuals
SAMP_residuals_STD | 0.05107755541
varray: [2.919, 2.879, 2.944, 3.001]
inst: std_all
operation: std_all(residuals)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31625
varray: [0.276, 0.346, 0.331, 0.312]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0302255410759
varray: [0.276, 0.346, 0.331, 0.312]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1346.79975
varray: [1346.712, 1344.161, 1347.871, 1348.455]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.90246600933
varray: [1346.712, 1344.161, 1347.871, 1348.455]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.321
varray: [0.945, 0.242, 0.103, -0.006]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.428201665885
varray: [0.945, 0.242, 0.103, -0.006]
inst: std_all
operation: std_all(slope_shift)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0211423745119
varray: [0.42, 0.417, 0.459, 0.45]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.229710687605
varray: [-63.033, -63.324, -63.099, -62.766]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.62425
varray: [111.815, 111.636, 112.885, 114.161]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.16348825377
varray: [111.815, 111.636, 112.885, 114.161]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.93575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1346.79975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.321
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.40410289808
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.03475
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.42, 0.417, 0.459, 0.45]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-63.033, -63.324, -63.099, -62.766]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.4365
varray: [0.42, 0.417, 0.459, 0.45]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
dD_mean | -63.0555
varray: [-63.033, -63.324, -63.099, -62.766]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0211423745119
varray: [0.42, 0.417, 0.459, 0.45]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
dD_stdev | 0.229710687605
varray: [-63.033, -63.324, -63.099, -62.766]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.4365
varray: [0.42, 0.417, 0.459, 0.45]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.0555
varray: [-63.033, -63.324, -63.099, -62.766]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0211423745119
varray: [0.42, 0.417, 0.459, 0.45]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.229710687605
varray: [-63.033, -63.324, -63.099, -62.766]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.35705029327
actual_array: [ 1.93 -15.84]
varray: [0.42, 0.417, 0.459, 0.45]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.4365
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.047549728
actual_array: [ 18.1 -119. ]
varray: [-63.033, -63.324, -63.099, -62.766]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -63.0555
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [85561.0, 86097.0, 86633.0, 87169.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.842
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 38.022
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.19925
varray: [0.0, 2.993, 2.902, 2.902]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: residuals
SAMP_residuals_STD | 1.46679409484
varray: [0.0, 2.993, 2.902, 2.902]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3155
varray: [0.0, 0.466, 0.398, 0.398]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.212761995353
varray: [0.0, 0.466, 0.398, 0.398]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 974.6165
varray: [0.0, 1300.14, 1299.163, 1299.163]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 649.744496565
varray: [0.0, 1300.14, 1299.163, 1299.163]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.74825
varray: [0.0, 2.205, 2.394, 2.394]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.16890044486
varray: [0.0, 2.205, 2.394, 2.394]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0442492937797
varray: [0.0, -0.074, -0.093, -0.093]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 30.1432242621
varray: [0.0, -60.203, -60.328, -60.328]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 84.43675
varray: [0.0, 113.983, 111.882, 111.882]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 56.2998790252
varray: [0.0, 113.983, 111.882, 111.882]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.19925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3155
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 974.6165
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.74825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.87175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.87175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -28.3865190492
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -32.22225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.0, -0.074, -0.093, -0.093]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [0.0, -60.203, -60.328, -60.328]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.065
varray: [0.0, -0.074, -0.093, -0.093]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
dD_mean | -45.21475
varray: [0.0, -60.203, -60.328, -60.328]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0442492937797
varray: [0.0, -0.074, -0.093, -0.093]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
dD_stdev | 30.1432242621
varray: [0.0, -60.203, -60.328, -60.328]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.065
varray: [0.0, -0.074, -0.093, -0.093]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.21475
varray: [0.0, -60.203, -60.328, -60.328]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0442492937797
varray: [0.0, -0.074, -0.093, -0.093]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
measured_dD_stdev | 30.1432242621
varray: [0.0, -60.203, -60.328, -60.328]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.85582206501
actual_array: [ 1.93 -15.84]
varray: [0.0, -0.074, -0.093, -0.093]
fit_rsq: 1.0
fit_intercept: 2.80644310636
except: True
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.065
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.4145541705
actual_array: [ 18.1 -119. ]
varray: [0.0, -60.203, -60.328, -60.328]
fit_rsq: 1.0
fit_intercept: 19.5866234865
except: True
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -45.21475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [87756.0, 88293.0, 88830.0, 89366.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1306.5785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -334.16125
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.5275
varray: [0.0, 0.0, 2.995, 3.115]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: residuals
SAMP_residuals_STD | 1.7644852885
varray: [0.0, 0.0, 2.995, 3.115]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2045
varray: [0.0, 0.0, 0.282, 0.536]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.257901143852
varray: [0.0, 0.0, 0.282, 0.536]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 653.70475
varray: [0.0, 0.0, 1306.035, 1308.784]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 754.834061056
varray: [0.0, 0.0, 1306.035, 1308.784]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.7845
varray: [0.0, 0.0, 1.294, 1.844]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.933275772035
varray: [0.0, 0.0, 1.294, 1.844]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: slope_shift
SAMP_d18O_STD | 1.74014309756
varray: [0.0, 0.0, -3.007, -3.021]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 37.6865431023
varray: [0.0, 0.0, -65.253, -65.297]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.5355
varray: [0.0, 0.0, 115.1, 119.042]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 67.6101292288
varray: [0.0, 0.0, 115.1, 119.042]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.5275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2045
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 653.70475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.7845
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.5435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.5435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -50.2605014653
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -58.1235
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.0, 0.0, -3.007, -3.021]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [0.0, 0.0, -65.253, -65.297]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.507
varray: [0.0, 0.0, -3.007, -3.021]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
dD_mean | -32.6375
varray: [0.0, 0.0, -65.253, -65.297]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
d18O_stdev | 1.74014309756
varray: [0.0, 0.0, -3.007, -3.021]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
dD_stdev | 37.6865431023
varray: [0.0, 0.0, -65.253, -65.297]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.507
varray: [0.0, 0.0, -3.007, -3.021]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.6375
varray: [0.0, 0.0, -65.253, -65.297]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.74014309756
varray: [0.0, 0.0, -3.007, -3.021]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
measured_dD_stdev | 37.6865431023
varray: [0.0, 0.0, -65.253, -65.297]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.28997738873
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, -3.007, -3.021]
fit_rsq: 1.0
fit_intercept: 2.80644310636
except: True
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -1.507
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.6887386429
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, -65.253, -65.297]
fit_rsq: 1.0
fit_intercept: 19.5866234865
except: True
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -32.6375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [90010.0, 90547.0, 91082.0, 91619.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1307.25025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -655.073
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96875
varray: [2.969, 2.952, 2.973, 2.981]
inst: mean_all
operation: mean_all(residuals)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: residuals
SAMP_residuals_STD | 0.0122304265393
varray: [2.969, 2.952, 2.973, 2.981]
inst: std_all
operation: std_all(residuals)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3735
varray: [0.311, 0.212, 0.463, 0.508]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.136734535993
varray: [0.311, 0.212, 0.463, 0.508]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1320.996
varray: [1321.896, 1322.05, 1319.478, 1320.56]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.21316885332
varray: [1321.896, 1322.05, 1319.478, 1320.56]
inst: std_all
operation: std_all(baseline_shift)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.353
varray: [1.008, 1.442, 1.33, 1.632]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.261607849016
varray: [1.008, 1.442, 1.33, 1.632]
inst: std_all
operation: std_all(slope_shift)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0629735394167
varray: [-0.104, 0.012, 0.017, 0.033]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.249642410927
varray: [-61.974, -61.4, -61.604, -61.506]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.41275
varray: [114.735, 112.998, 112.811, 113.107]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.889934220416
varray: [114.735, 112.998, 112.811, 113.107]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3735
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1320.996
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.353
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.32953435363
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.24625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-0.104, 0.012, 0.017, 0.033]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-61.974, -61.4, -61.604, -61.506]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.0105
varray: [-0.104, 0.012, 0.017, 0.033]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
dD_mean | -61.621
varray: [-61.974, -61.4, -61.604, -61.506]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0629735394167
varray: [-0.104, 0.012, 0.017, 0.033]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
dD_stdev | 0.249642410927
varray: [-61.974, -61.4, -61.604, -61.506]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.0105
varray: [-0.104, 0.012, 0.017, 0.033]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.621
varray: [-61.974, -61.4, -61.604, -61.506]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0629735394167
varray: [-0.104, 0.012, 0.017, 0.033]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.249642410927
varray: [-61.974, -61.4, -61.604, -61.506]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.80161855243
actual_array: [ 1.93 -15.84]
varray: [-0.104, 0.012, 0.017, 0.033]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.0105
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.7101602053
actual_array: [ 18.1 -119. ]
varray: [-61.974, -61.4, -61.604, -61.506]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [92202.0, 92738.0, 93273.0, 93809.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.809
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 12.21825
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97825
varray: [2.871, 3.111, 2.999, 2.932]
inst: mean_all
operation: mean_all(residuals)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: residuals
SAMP_residuals_STD | 0.102785780469
varray: [2.871, 3.111, 2.999, 2.932]
inst: std_all
operation: std_all(residuals)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.1545
varray: [0.0, 0.076, 0.17, 0.372]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.160809411002
varray: [0.0, 0.076, 0.17, 0.372]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1517.48125
varray: [1504.964, 1530.553, 1517.874, 1516.534]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 10.4658767868
varray: [1504.964, 1530.553, 1517.874, 1516.534]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -6.57475
varray: [-6.538, -6.387, -7.242, -6.132]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.475342245124
varray: [-6.538, -6.387, -7.242, -6.132]
inst: std_all
operation: std_all(slope_shift)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: slope_shift
SAMP_d18O_STD | 0.102555269652
varray: [2.405, 2.642, 2.457, 2.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.663114557725
varray: [-56.983, -55.378, -55.97, -56.064]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.36175
varray: [109.811, 116.319, 114.211, 113.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.71660245834
varray: [109.811, 116.319, 114.211, 113.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.1545
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1517.48125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 6.57475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.02018886356
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.29725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [2.405, 2.642, 2.457, 2.475]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-56.983, -55.378, -55.97, -56.064]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.49475
varray: [2.405, 2.642, 2.457, 2.475]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
dD_mean | -56.09875
varray: [-56.983, -55.378, -55.97, -56.064]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
d18O_stdev | 0.102555269652
varray: [2.405, 2.642, 2.457, 2.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
dD_stdev | 0.663114557725
varray: [-56.983, -55.378, -55.97, -56.064]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.49475
varray: [2.405, 2.642, 2.457, 2.475]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.09875
varray: [-56.983, -55.378, -55.97, -56.064]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.102555269652
varray: [2.405, 2.642, 2.457, 2.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.663114557725
varray: [-56.983, -55.378, -55.97, -56.064]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.309997453434
actual_array: [ 1.93 -15.84]
varray: [2.405, 2.642, 2.457, 2.475]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.49475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.5617466174
actual_array: [ 18.1 -119. ]
varray: [-56.983, -55.378, -55.97, -56.064]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.09875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [94393.0, 94929.0, 95465.0, 96001.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7995
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 208.7035
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -3.071
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 3.071
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -116.659
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 2.80644310636
except: True
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 19.5866234865
except: True
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [96585.0, 97183.0, 97781.0, 98379.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1308.77775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1308.77775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97575
varray: [3.003, 2.974, 2.943, 2.983]
inst: mean_all
operation: mean_all(residuals)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: residuals
SAMP_residuals_STD | 0.0249716505929
varray: [3.003, 2.974, 2.943, 2.983]
inst: std_all
operation: std_all(residuals)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3625
varray: [0.441, 0.336, 0.178, 0.495]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.13959584521
varray: [0.441, 0.336, 0.178, 0.495]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.09425
varray: [1305.392, 1301.297, 1301.153, 1300.535]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.22320899827
varray: [1305.392, 1301.297, 1301.153, 1300.535]
inst: std_all
operation: std_all(baseline_shift)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.47825
varray: [2.151, 3.134, 2.33, 2.298]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.444060337492
varray: [2.151, 3.134, 2.33, 2.298]
inst: std_all
operation: std_all(slope_shift)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: slope_shift
SAMP_d18O_STD | 0.145206462207
varray: [-7.169, -7.331, -7.443, -7.498]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.242923170296
varray: [-59.151, -59.429, -59.743, -59.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.79575
varray: [114.547, 112.552, 113.461, 114.623]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.984494921267
varray: [114.547, 112.552, 113.461, 114.623]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.09425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.47825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.10159557148
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.86325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.169, -7.331, -7.443, -7.498]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-59.151, -59.429, -59.743, -59.395]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.36025
varray: [-7.169, -7.331, -7.443, -7.498]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
dD_mean | -59.4295
varray: [-59.151, -59.429, -59.743, -59.395]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
d18O_stdev | 0.145206462207
varray: [-7.169, -7.331, -7.443, -7.498]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
dD_stdev | 0.242923170296
varray: [-59.151, -59.429, -59.743, -59.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.36025
varray: [-7.169, -7.331, -7.443, -7.498]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.4295
varray: [-59.151, -59.429, -59.743, -59.395]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.145206462207
varray: [-7.169, -7.331, -7.443, -7.498]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.242923170296
varray: [-59.151, -59.429, -59.743, -59.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.111384912
actual_array: [ 1.93 -15.84]
varray: [-7.169, -7.331, -7.443, -7.498]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.36025
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.6670168529
actual_array: [ 18.1 -119. ]
varray: [-59.151, -59.429, -59.743, -59.395]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.4295
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [99025.0, 99561.0, 100097.0, 100633.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.802
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.6835
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9295
varray: [2.885, 2.951, 2.902, 2.98]
inst: mean_all
operation: mean_all(residuals)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: residuals
SAMP_residuals_STD | 0.0437759446881
varray: [2.885, 2.951, 2.902, 2.98]
inst: std_all
operation: std_all(residuals)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41475
varray: [0.291, 0.41, 0.342, 0.616]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.142747737869
varray: [0.291, 0.41, 0.342, 0.616]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.9845
varray: [1297.925, 1300.201, 1298.629, 1299.183]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.960589228894
varray: [1297.925, 1300.201, 1298.629, 1299.183]
inst: std_all
operation: std_all(baseline_shift)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.65275
varray: [2.661, 2.875, 2.731, 2.344]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.224286089627
varray: [2.661, 2.875, 2.731, 2.344]
inst: std_all
operation: std_all(slope_shift)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: slope_shift
SAMP_d18O_STD | 0.137861041149
varray: [-12.573, -12.682, -12.853, -12.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.50646568897
varray: [-103.567, -107.336, -108.836, -108.92]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.51325
varray: [111.031, 112.91, 111.981, 114.131]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.32349042938
varray: [111.031, 112.91, 111.981, 114.131]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9295
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.9845
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.65275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1415
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1415
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.60761966786
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.14575
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.573, -12.682, -12.853, -12.854]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-103.567, -107.336, -108.836, -108.92]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.7405
varray: [-12.573, -12.682, -12.853, -12.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
dD_mean | -107.16475
varray: [-103.567, -107.336, -108.836, -108.92]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
d18O_stdev | 0.137861041149
varray: [-12.573, -12.682, -12.853, -12.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
dD_stdev | 2.50646568897
varray: [-103.567, -107.336, -108.836, -108.92]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.7405
varray: [-12.573, -12.682, -12.853, -12.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.16475
varray: [-103.567, -107.336, -108.836, -108.92]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.137861041149
varray: [-12.573, -12.682, -12.853, -12.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.50646568897
varray: [-103.567, -107.336, -108.836, -108.92]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.4623656197
actual_array: [ 1.93 -15.84]
varray: [-12.573, -12.682, -12.853, -12.854]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.7405
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -118.170762558
actual_array: [ 18.1 -119. ]
varray: [-103.567, -107.336, -108.836, -108.92]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.16475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [101217.0, 101753.0, 102289.0, 102824.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.84825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.79325
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.92375
varray: [2.971, 2.883, 2.912, 2.929]
inst: mean_all
operation: mean_all(residuals)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: residuals
SAMP_residuals_STD | 0.036782015533
varray: [2.971, 2.883, 2.912, 2.929]
inst: std_all
operation: std_all(residuals)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.51725
varray: [0.544, 0.519, 0.491, 0.515]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0217006144306
varray: [0.544, 0.519, 0.491, 0.515]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.37775
varray: [1300.28, 1297.658, 1298.895, 1296.678]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.55919773709
varray: [1300.28, 1297.658, 1298.895, 1296.678]
inst: std_all
operation: std_all(baseline_shift)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3985
varray: [2.443, 2.427, 2.366, 2.358]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0427746030568
varray: [2.443, 2.427, 2.366, 2.358]
inst: std_all
operation: std_all(slope_shift)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: slope_shift
SAMP_d18O_STD | 0.108494239478
varray: [-6.366, -6.223, -6.111, -6.174]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.45142734346
varray: [-80.386, -77.926, -77.384, -77.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.773
varray: [114.372, 112.053, 112.828, 111.839]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14754549656
varray: [114.372, 112.053, 112.828, 111.839]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.92375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.51725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.37775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3985
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.14725
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14725
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.79485509606
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.886
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-6.366, -6.223, -6.111, -6.174]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-80.386, -77.926, -77.384, -77.3]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.2185
varray: [-6.366, -6.223, -6.111, -6.174]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
dD_mean | -78.249
varray: [-80.386, -77.926, -77.384, -77.3]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
d18O_stdev | 0.108494239478
varray: [-6.366, -6.223, -6.111, -6.174]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
dD_stdev | 1.45142734346
varray: [-80.386, -77.926, -77.384, -77.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.2185
varray: [-6.366, -6.223, -6.111, -6.174]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.249
varray: [-80.386, -77.926, -77.384, -77.3]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.108494239478
varray: [-6.366, -6.223, -6.111, -6.174]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.45142734346
varray: [-80.386, -77.926, -77.384, -77.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.97584618743
actual_array: [ 1.93 -15.84]
varray: [-6.366, -6.223, -6.111, -6.174]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.2185
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.2125045646
actual_array: [ 18.1 -119. ]
varray: [-80.386, -77.926, -77.384, -77.3]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -78.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [103409.0, 103946.0, 104482.0, 105017.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.854
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.4
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9925
varray: [2.963, 2.985, 3.009, 3.013]
inst: mean_all
operation: mean_all(residuals)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: residuals
SAMP_residuals_STD | 0.0232307267787
varray: [2.963, 2.985, 3.009, 3.013]
inst: std_all
operation: std_all(residuals)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5325
varray: [0.601, 0.549, 0.409, 0.571]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0850470458041
varray: [0.601, 0.549, 0.409, 0.571]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.03075
varray: [1299.101, 1299.66, 1301.053, 1300.309]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.841488512498
varray: [1299.101, 1299.66, 1301.053, 1300.309]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.43225
varray: [2.352, 2.4, 2.414, 2.563]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0911203416733
varray: [2.352, 2.4, 2.414, 2.563]
inst: std_all
operation: std_all(slope_shift)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0472440472441
varray: [-5.932, -5.95, -5.848, -5.878]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.129424366588
varray: [-77.765, -77.989, -77.685, -77.785]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.425
varray: [114.067, 114.092, 114.7, 114.841]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.403209622901
varray: [114.067, 114.092, 114.7, 114.841]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.03075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.43225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.55617062846
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.234
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-5.932, -5.95, -5.848, -5.878]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-77.765, -77.989, -77.685, -77.785]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.902
varray: [-5.932, -5.95, -5.848, -5.878]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
dD_mean | -77.806
varray: [-77.765, -77.989, -77.685, -77.785]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0472440472441
varray: [-5.932, -5.95, -5.848, -5.878]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
dD_stdev | 0.129424366588
varray: [-77.765, -77.989, -77.685, -77.785]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.902
varray: [-5.932, -5.95, -5.848, -5.878]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.806
varray: [-77.765, -77.989, -77.685, -77.785]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0472440472441
varray: [-5.932, -5.95, -5.848, -5.878]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.129424366588
varray: [-77.765, -77.989, -77.685, -77.785]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.66106799051
actual_array: [ 1.93 -15.84]
varray: [-5.932, -5.95, -5.848, -5.878]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.7994940672
actual_array: [ 18.1 -119. ]
varray: [-77.765, -77.989, -77.685, -77.785]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.806
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [105601.0, 106136.0, 106672.0, 107208.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.78525
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.747
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95875
varray: [2.952, 2.939, 3.002, 2.942]
inst: mean_all
operation: mean_all(residuals)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: residuals
SAMP_residuals_STD | 0.0293640937201
varray: [2.952, 2.939, 3.002, 2.942]
inst: std_all
operation: std_all(residuals)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.403
varray: [0.397, 0.251, 0.579, 0.385]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.134709564125
varray: [0.397, 0.251, 0.579, 0.385]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.41125
varray: [1299.076, 1299.174, 1300.075, 1299.32]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.453714576211
varray: [1299.076, 1299.174, 1300.075, 1299.32]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.49725
varray: [2.226, 2.823, 2.563, 2.377]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.257211164869
varray: [2.226, 2.823, 2.563, 2.377]
inst: std_all
operation: std_all(slope_shift)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0449249373956
varray: [-4.083, -3.996, -4.041, -3.985]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.123164862413
varray: [-77.709, -77.567, -77.706, -77.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.41975
varray: [113.483, 112.116, 114.699, 113.381]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.05542199301
varray: [113.483, 112.116, 114.699, 113.381]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.403
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.41125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.49725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.65516118528
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.23925
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-4.083, -3.996, -4.041, -3.985]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-77.709, -77.567, -77.706, -77.453]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.02625
varray: [-4.083, -3.996, -4.041, -3.985]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
dD_mean | -77.60875
varray: [-77.709, -77.567, -77.706, -77.453]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0449249373956
varray: [-4.083, -3.996, -4.041, -3.985]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
dD_stdev | 0.123164862413
varray: [-77.709, -77.567, -77.706, -77.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.02625
varray: [-4.083, -3.996, -4.041, -3.985]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.60875
varray: [-77.709, -77.567, -77.706, -77.453]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0449249373956
varray: [-4.083, -3.996, -4.041, -3.985]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.123164862413
varray: [-77.709, -77.567, -77.706, -77.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.79552232583
actual_array: [ 1.93 -15.84]
varray: [-4.083, -3.996, -4.041, -3.985]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -4.02625
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.6155971809
actual_array: [ 18.1 -119. ]
varray: [-77.709, -77.567, -77.706, -77.453]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.60875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [107792.0, 108328.0, 108863.0, 109399.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.819
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.3665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98525
varray: [2.985, 2.989, 3.026, 2.941]
inst: mean_all
operation: mean_all(residuals)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: residuals
SAMP_residuals_STD | 0.0347982279242
varray: [2.985, 2.989, 3.026, 2.941]
inst: std_all
operation: std_all(residuals)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.50975
varray: [0.479, 0.461, 0.633, 0.466]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0825161600335
varray: [0.479, 0.461, 0.633, 0.466]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.99325
varray: [1300.104, 1301.549, 1301.251, 1301.069]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.624979132985
varray: [1300.104, 1301.549, 1301.251, 1301.069]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.3115
varray: [2.092, 2.047, 2.805, 2.302]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.347262148816
varray: [2.092, 2.047, 2.805, 2.302]
inst: std_all
operation: std_all(slope_shift)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0252636101933
varray: [-3.449, -3.483, -3.428, -3.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.231043105646
varray: [-73.41, -73.147, -73.091, -72.847]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.6095
varray: [114.467, 115.191, 114.866, 113.914]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.550139679233
varray: [114.467, 115.191, 114.866, 113.914]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.50975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.99325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.3115
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.79225008141
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0495
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-3.449, -3.483, -3.428, -3.431]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-73.41, -73.147, -73.091, -72.847]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.44775
varray: [-3.449, -3.483, -3.428, -3.431]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
dD_mean | -73.12375
varray: [-73.41, -73.147, -73.091, -72.847]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0252636101933
varray: [-3.449, -3.483, -3.428, -3.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
dD_stdev | 0.231043105646
varray: [-73.41, -73.147, -73.091, -72.847]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.44775
varray: [-3.449, -3.483, -3.428, -3.431]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.12375
varray: [-73.41, -73.147, -73.091, -72.847]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0252636101933
varray: [-3.449, -3.483, -3.428, -3.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.231043105646
varray: [-73.41, -73.147, -73.091, -72.847]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.22016944457
actual_array: [ 1.93 -15.84]
varray: [-3.449, -3.483, -3.428, -3.431]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.44775
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.4342155083
actual_array: [ 18.1 -119. ]
varray: [-73.41, -73.147, -73.091, -72.847]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.12375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [109984.0, 110520.0, 111056.0, 111591.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7925
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.7845
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96475
varray: [2.872, 2.99, 3.034, 2.963]
inst: mean_all
operation: mean_all(residuals)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: residuals
SAMP_residuals_STD | 0.0684074800978
varray: [2.872, 2.99, 3.034, 2.963]
inst: std_all
operation: std_all(residuals)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0395
varray: [-0.107, 0.039, 0.13, 0.096]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.104634283738
varray: [-0.107, 0.039, 0.13, 0.096]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1391.928
varray: [1385.94, 1392.356, 1396.616, 1392.8]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.4263248864
varray: [1385.94, 1392.356, 1396.616, 1392.8]
inst: std_all
operation: std_all(baseline_shift)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.39625
varray: [-1.321, -1.793, -0.879, -1.592]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.395364283499
varray: [-1.321, -1.793, -0.879, -1.592]
inst: std_all
operation: std_all(slope_shift)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00801560977094
varray: [-3.44, -3.44, -3.457, -3.446]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0704769465854
varray: [-73.008, -73.113, -73.174, -73.131]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.2105
varray: [110.321, 113.939, 115.29, 113.292]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.09848238814
varray: [110.321, 113.939, 115.29, 113.292]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0395
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1391.928
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.39625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10625
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10625
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.45978508629
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4485
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-3.44, -3.44, -3.457, -3.446]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-73.008, -73.113, -73.174, -73.131]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.44575
varray: [-3.44, -3.44, -3.457, -3.446]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
dD_mean | -73.1065
varray: [-73.008, -73.113, -73.174, -73.131]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00801560977094
varray: [-3.44, -3.44, -3.457, -3.446]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
dD_stdev | 0.0704769465854
varray: [-73.008, -73.113, -73.174, -73.131]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.44575
varray: [-3.44, -3.44, -3.457, -3.446]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.1065
varray: [-73.008, -73.113, -73.174, -73.131]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00801560977094
varray: [-3.44, -3.44, -3.457, -3.446]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0704769465854
varray: [-73.008, -73.113, -73.174, -73.131]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.21818032484
actual_array: [ 1.93 -15.84]
varray: [-3.44, -3.44, -3.457, -3.446]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.44575
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.4181332711
actual_array: [ 18.1 -119. ]
varray: [-73.008, -73.113, -73.174, -73.131]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.1065
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [112177.0, 112713.0, 113248.0, 113784.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.813
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 83.15025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01025
varray: [2.983, 3.028, 3.031, 2.999]
inst: mean_all
operation: mean_all(residuals)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: residuals
SAMP_residuals_STD | 0.0232002155162
varray: [2.983, 3.028, 3.031, 2.999]
inst: std_all
operation: std_all(residuals)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.48275
varray: [0.467, 0.55, 0.459, 0.455]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.045110050617
varray: [0.467, 0.55, 0.459, 0.455]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.94525
varray: [1302.374, 1302.065, 1301.832, 1301.51]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.365346753099
varray: [1302.374, 1302.065, 1301.832, 1301.51]
inst: std_all
operation: std_all(baseline_shift)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.21975
varray: [2.082, 2.391, 2.043, 2.363]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.182631459502
varray: [2.082, 2.391, 2.043, 2.363]
inst: std_all
operation: std_all(slope_shift)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: slope_shift
SAMP_d18O_STD | 0.128979003459
varray: [0.573, 0.81, 0.803, 0.863]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.405106878902
varray: [-67.626, -66.93, -66.808, -66.751]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.81675
varray: [114.183, 115.53, 115.351, 114.203]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.723988202022
varray: [114.183, 115.53, 115.351, 114.203]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.48275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.94525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.21975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.97818300228
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.84225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.573, 0.81, 0.803, 0.863]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-67.626, -66.93, -66.808, -66.751]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.76225
varray: [0.573, 0.81, 0.803, 0.863]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
dD_mean | -67.02875
varray: [-67.626, -66.93, -66.808, -66.751]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
d18O_stdev | 0.128979003459
varray: [0.573, 0.81, 0.803, 0.863]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
dD_stdev | 0.405106878902
varray: [-67.626, -66.93, -66.808, -66.751]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.76225
varray: [0.573, 0.81, 0.803, 0.863]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.02875
varray: [-67.626, -66.93, -66.808, -66.751]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.128979003459
varray: [0.573, 0.81, 0.803, 0.863]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.405106878902
varray: [-67.626, -66.93, -66.808, -66.751]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.03307241761
actual_array: [ 1.93 -15.84]
varray: [0.573, 0.81, 0.803, 0.863]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.76225
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.7518250302
actual_array: [ 18.1 -119. ]
varray: [-67.626, -66.93, -66.808, -66.751]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.02875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [114370.0, 114905.0, 115441.0, 115977.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.8325
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.829
varray: [2.82, 2.822, 2.831, 2.843]
inst: mean_all
operation: mean_all(residuals)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: residuals
SAMP_residuals_STD | 0.0104880884817
varray: [2.82, 2.822, 2.831, 2.843]
inst: std_all
operation: std_all(residuals)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.353
varray: [0.474, 0.247, 0.388, 0.303]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0993344518953
varray: [0.474, 0.247, 0.388, 0.303]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.79075
varray: [1294.189, 1294.786, 1295.264, 1294.924]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.448644904128
varray: [1294.189, 1294.786, 1295.264, 1294.924]
inst: std_all
operation: std_all(baseline_shift)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.66675
varray: [2.807, 2.701, 2.516, 2.643]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.121288567749
varray: [2.807, 2.701, 2.516, 2.643]
inst: std_all
operation: std_all(slope_shift)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: slope_shift
SAMP_d18O_STD | 0.148518236815
varray: [-5.02, -5.282, -5.293, -5.355]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.779578091021
varray: [-81.611, -83.054, -83.268, -83.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.57275
varray: [108.542, 108.06, 109.448, 108.241]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.616435654928
varray: [108.542, 108.06, 109.448, 108.241]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.829
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.353
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.79075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.66675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.242
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.242
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.88016932595
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.08625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-5.02, -5.282, -5.293, -5.355]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-81.611, -83.054, -83.268, -83.159]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.2375
varray: [-5.02, -5.282, -5.293, -5.355]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
dD_mean | -82.773
varray: [-81.611, -83.054, -83.268, -83.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
d18O_stdev | 0.148518236815
varray: [-5.02, -5.282, -5.293, -5.355]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
dD_stdev | 0.779578091021
varray: [-81.611, -83.054, -83.268, -83.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.2375
varray: [-5.02, -5.282, -5.293, -5.355]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
measured_dD_mean | -82.773
varray: [-81.611, -83.054, -83.268, -83.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.148518236815
varray: [-5.02, -5.282, -5.293, -5.355]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.779578091021
varray: [-81.611, -83.054, -83.268, -83.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.00018296096
actual_array: [ 1.93 -15.84]
varray: [-5.02, -5.282, -5.293, -5.355]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.2375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -95.4302460778
actual_array: [ 18.1 -119. ]
varray: [-81.611, -83.054, -83.268, -83.159]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -82.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [116564.0, 117100.0, 117636.0, 118172.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.94875
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -13.987
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9695
varray: [3.027, 2.971, 2.934, 2.946]
inst: mean_all
operation: mean_all(residuals)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: residuals
SAMP_residuals_STD | 0.0413158565202
varray: [3.027, 2.971, 2.934, 2.946]
inst: std_all
operation: std_all(residuals)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.44425
varray: [0.486, 0.34, 0.444, 0.507]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0742714615448
varray: [0.486, 0.34, 0.444, 0.507]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.203
varray: [1301.54, 1300.085, 1299.017, 1300.17]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.0342755919
varray: [1301.54, 1300.085, 1299.017, 1300.17]
inst: std_all
operation: std_all(baseline_shift)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.7885
varray: [3.028, 2.356, 2.87, 2.9]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.296361828401
varray: [3.028, 2.356, 2.87, 2.9]
inst: std_all
operation: std_all(slope_shift)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0821563549987
varray: [-7.31, -7.459, -7.466, -7.491]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.07242559959
varray: [-62.167, -60.596, -59.844, -59.945]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.43525
varray: [114.943, 113.283, 112.205, 113.31]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.12926122605
varray: [114.943, 113.283, 112.205, 113.31]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9695
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.44425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.203
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.7885
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1015
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1015
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.30511234126
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.22375
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.31, -7.459, -7.466, -7.491]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-62.167, -60.596, -59.844, -59.945]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4315
varray: [-7.31, -7.459, -7.466, -7.491]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
dD_mean | -60.638
varray: [-62.167, -60.596, -59.844, -59.945]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0821563549987
varray: [-7.31, -7.459, -7.466, -7.491]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
dD_stdev | 1.07242559959
varray: [-62.167, -60.596, -59.844, -59.945]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4315
varray: [-7.31, -7.459, -7.466, -7.491]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.638
varray: [-62.167, -60.596, -59.844, -59.945]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0821563549987
varray: [-7.31, -7.459, -7.466, -7.491]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.07242559959
varray: [-62.167, -60.596, -59.844, -59.945]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1822473023
actual_array: [ 1.93 -15.84]
varray: [-7.31, -7.459, -7.466, -7.491]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.4315
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.7937057606
actual_array: [ 18.1 -119. ]
varray: [-62.167, -60.596, -59.844, -59.945]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.638
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [118758.0, 119294.0, 119830.0, 120365.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.80825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.57475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.975
varray: [2.967, 2.966, 2.947, 3.02]
inst: mean_all
operation: mean_all(residuals)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: residuals
SAMP_residuals_STD | 0.031379398762
varray: [2.967, 2.966, 2.947, 3.02]
inst: std_all
operation: std_all(residuals)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.38475
varray: [0.435, 0.412, 0.295, 0.397]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0618405206964
varray: [0.435, 0.412, 0.295, 0.397]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.091
varray: [1300.005, 1299.994, 1299.184, 1301.181]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.822101777965
varray: [1300.005, 1299.994, 1299.184, 1301.181]
inst: std_all
operation: std_all(baseline_shift)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.4455
varray: [2.692, 2.689, 2.295, 2.106]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.293237673796
varray: [2.692, 2.689, 2.295, 2.106]
inst: std_all
operation: std_all(slope_shift)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0212661703181
varray: [-7.494, -7.461, -7.506, -7.506]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.17175467
varray: [-60.037, -59.711, -59.701, -59.673]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.5365
varray: [113.522, 113.589, 112.617, 114.418]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.736090800558
varray: [113.522, 113.589, 112.617, 114.418]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.38475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.091
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.4455
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.096
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.096
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.12601758385
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.1225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.494, -7.461, -7.506, -7.506]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-60.037, -59.711, -59.701, -59.673]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.49175
varray: [-7.494, -7.461, -7.506, -7.506]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
dD_mean | -59.7805
varray: [-60.037, -59.711, -59.701, -59.673]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0212661703181
varray: [-7.494, -7.461, -7.506, -7.506]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
dD_stdev | 0.17175467
varray: [-60.037, -59.711, -59.701, -59.673]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.49175
varray: [-7.494, -7.461, -7.506, -7.506]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.7805
varray: [-60.037, -59.711, -59.701, -59.673]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0212661703181
varray: [-7.494, -7.461, -7.506, -7.506]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.17175467
varray: [-60.037, -59.711, -59.701, -59.673]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2421695341
actual_array: [ 1.93 -15.84]
varray: [-7.494, -7.461, -7.506, -7.506]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.49175
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.9942554185
actual_array: [ 18.1 -119. ]
varray: [-60.037, -59.711, -59.701, -59.673]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.7805
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [120950.0, 121486.0, 122022.0, 122558.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.80275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.68675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98425
varray: [2.992, 2.977, 2.933, 3.035]
inst: mean_all
operation: mean_all(residuals)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: residuals
SAMP_residuals_STD | 0.042090180961
varray: [2.992, 2.977, 2.933, 3.035]
inst: std_all
operation: std_all(residuals)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.438
varray: [0.616, 0.324, 0.159, 0.653]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.237168013582
varray: [0.616, 0.324, 0.159, 0.653]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.122
varray: [1299.031, 1301.248, 1302.591, 1301.618]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.50467427261
varray: [1299.031, 1301.248, 1302.591, 1301.618]
inst: std_all
operation: std_all(baseline_shift)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.988
varray: [2.076, 1.804, 2.162, 1.91]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.161203804752
varray: [2.076, 1.804, 2.162, 1.91]
inst: std_all
operation: std_all(slope_shift)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: slope_shift
SAMP_d18O_STD | 0.23098105983
varray: [0.595, 1.003, 1.053, 1.096]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.228791571814
varray: [-57.996, -57.566, -57.636, -57.473]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.10375
varray: [113.922, 115.407, 114.418, 116.668]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.21183397515
varray: [113.922, 115.407, 114.418, 116.668]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.438
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.122
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.988
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.82481276457
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.55525
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.595, 1.003, 1.053, 1.096]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-57.996, -57.566, -57.636, -57.473]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.93675
varray: [0.595, 1.003, 1.053, 1.096]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
dD_mean | -57.66775
varray: [-57.996, -57.566, -57.636, -57.473]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
d18O_stdev | 0.23098105983
varray: [0.595, 1.003, 1.053, 1.096]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
dD_stdev | 0.228791571814
varray: [-57.996, -57.566, -57.636, -57.473]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.93675
varray: [0.595, 1.003, 1.053, 1.096]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.66775
varray: [-57.996, -57.566, -57.636, -57.473]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.23098105983
varray: [0.595, 1.003, 1.053, 1.096]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.228791571814
varray: [-57.996, -57.566, -57.636, -57.473]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.85952172137
actual_array: [ 1.93 -15.84]
varray: [0.595, 1.003, 1.053, 1.096]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.93675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.0245309751
actual_array: [ 18.1 -119. ]
varray: [-57.996, -57.566, -57.636, -57.473]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.66775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [123143.0, 123679.0, 124215.0, 124750.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.65575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 52 - Sample 52 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 52
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.93275
varray: [2.887, 2.975, 2.931, 2.938]
inst: mean_all
operation: mean_all(residuals)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: residuals
SAMP_residuals_STD | 0.0360959369828
varray: [2.887, 2.975, 2.931, 2.938]
inst: std_all
operation: std_all(residuals)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.38175
varray: [0.386, 0.326, 0.46, 0.355]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0576331791014
varray: [0.386, 0.326, 0.46, 0.355]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.05425
varray: [1298.152, 1300.641, 1297.843, 1299.581]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.30082316887
varray: [1298.152, 1300.641, 1297.843, 1299.581]
inst: std_all
operation: std_all(baseline_shift)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.533
varray: [2.569, 2.432, 2.824, 2.307]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.221550596178
varray: [2.569, 2.432, 2.824, 2.307]
inst: std_all
operation: std_all(slope_shift)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: slope_shift
SAMP_d18O_STD | 0.146868138138
varray: [-6.204, -6.33, -6.483, -6.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.22930532551
varray: [-81.014, -82.754, -83.23, -83.887]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.5495
varray: [111.067, 114.174, 111.809, 113.148]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.38363639251
varray: [111.067, 114.174, 111.809, 113.148]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.93275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.38175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.05425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.533
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.50179094757
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.1095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-6.204, -6.33, -6.483, -6.524]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-81.014, -82.754, -83.23, -83.887]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.38525
varray: [-6.204, -6.33, -6.483, -6.524]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
dD_mean | -82.72125
varray: [-81.014, -82.754, -83.23, -83.887]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
d18O_stdev | 0.146868138138
varray: [-6.204, -6.33, -6.483, -6.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
dD_stdev | 1.22930532551
varray: [-81.014, -82.754, -83.23, -83.887]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.38525
varray: [-6.204, -6.33, -6.483, -6.524]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
measured_dD_mean | -82.72125
varray: [-81.014, -82.754, -83.23, -83.887]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.146868138138
varray: [-6.204, -6.33, -6.483, -6.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.22930532551
varray: [-81.014, -82.754, -83.23, -83.887]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.14168904473
actual_array: [ 1.93 -15.84]
varray: [-6.204, -6.33, -6.483, -6.524]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.38525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -95.3819993662
actual_array: [ 18.1 -119. ]
varray: [-81.014, -82.754, -83.23, -83.887]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -82.72125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [125336.0, 125872.0, 126408.0, 126943.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.845
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.7235
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 53 - Sample 53 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 53
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97025
varray: [2.934, 2.979, 2.958, 3.01]
inst: mean_all
operation: mean_all(residuals)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: residuals
SAMP_residuals_STD | 0.0322529068457
varray: [2.934, 2.979, 2.958, 3.01]
inst: std_all
operation: std_all(residuals)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.551
varray: [0.446, 0.682, 0.556, 0.52]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.098610344285
varray: [0.446, 0.682, 0.556, 0.52]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.556
varray: [1298.159, 1298.926, 1298.101, 1299.038]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.494590065947
varray: [1298.159, 1298.926, 1298.101, 1299.038]
inst: std_all
operation: std_all(baseline_shift)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.149
varray: [1.692, 2.381, 2.47, 2.053]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.353520390737
varray: [1.692, 2.381, 2.47, 2.053]
inst: std_all
operation: std_all(slope_shift)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: slope_shift
SAMP_d18O_STD | 0.19481615265
varray: [0.462, 0.806, 0.804, 0.908]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.856959888209
varray: [-69.442, -67.807, -67.753, -67.641]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.86175
varray: [113.35, 114.031, 113.677, 114.389]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.448203357863
varray: [113.35, 114.031, 113.677, 114.389]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.551
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.556
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.149
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.28069032888
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.79725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [0.462, 0.806, 0.804, 0.908]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-69.442, -67.807, -67.753, -67.641]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.745
varray: [0.462, 0.806, 0.804, 0.908]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
dD_mean | -68.16075
varray: [-69.442, -67.807, -67.753, -67.641]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
d18O_stdev | 0.19481615265
varray: [0.462, 0.806, 0.804, 0.908]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
dD_stdev | 0.856959888209
varray: [-69.442, -67.807, -67.753, -67.641]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.745
varray: [0.462, 0.806, 0.804, 0.908]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.16075
varray: [-69.442, -67.807, -67.753, -67.641]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.19481615265
varray: [0.462, 0.806, 0.804, 0.908]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.856959888209
varray: [-69.442, -67.807, -67.753, -67.641]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.05022857527
actual_array: [ 1.93 -15.84]
varray: [0.462, 0.806, 0.804, 0.908]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.745
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.8071927121
actual_array: [ 18.1 -119. ]
varray: [-69.442, -67.807, -67.753, -67.641]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -68.16075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [127529.0, 128064.0, 128599.0, 129135.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.8075
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.22175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 54 - Sample 54 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 54
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95625
varray: [2.931, 2.955, 2.961, 2.978]
inst: mean_all
operation: mean_all(residuals)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: residuals
SAMP_residuals_STD | 0.0194486503388
varray: [2.931, 2.955, 2.961, 2.978]
inst: std_all
operation: std_all(residuals)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.43175
varray: [0.573, 0.394, 0.332, 0.428]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.102210159312
varray: [0.573, 0.394, 0.332, 0.428]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.147
varray: [1298.264, 1299.762, 1299.007, 1299.555]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.669322543074
varray: [1298.264, 1299.762, 1299.007, 1299.555]
inst: std_all
operation: std_all(baseline_shift)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.556
varray: [2.176, 2.704, 2.581, 2.763]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.264435247272
varray: [2.176, 2.704, 2.581, 2.763]
inst: std_all
operation: std_all(slope_shift)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: slope_shift
SAMP_d18O_STD | 0.290508749151
varray: [-12.169, -12.632, -12.761, -12.802]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.01724339549
varray: [-104.701, -108.126, -108.764, -109.08]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.058
varray: [113.394, 113.035, 112.889, 112.914]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.232896257305
varray: [113.394, 113.035, 112.889, 112.914]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.43175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.147
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.556
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.73656789319
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.601
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.169, -12.632, -12.761, -12.802]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-104.701, -108.126, -108.764, -109.08]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.591
varray: [-12.169, -12.632, -12.761, -12.802]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
dD_mean | -107.66775
varray: [-104.701, -108.126, -108.764, -109.08]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
d18O_stdev | 0.290508749151
varray: [-12.169, -12.632, -12.761, -12.802]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
dD_stdev | 2.01724339549
varray: [-104.701, -108.126, -108.764, -109.08]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.591
varray: [-12.169, -12.632, -12.761, -12.802]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.66775
varray: [-104.701, -108.126, -108.764, -109.08]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.290508749151
varray: [-12.169, -12.632, -12.761, -12.802]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.01724339549
varray: [-104.701, -108.126, -108.764, -109.08]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.31367892
actual_array: [ 1.93 -15.84]
varray: [-12.169, -12.632, -12.761, -12.802]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -118.639711271
actual_array: [ 18.1 -119. ]
varray: [-104.701, -108.126, -108.764, -109.08]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.66775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [129722.0, 130257.0, 130793.0, 131329.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.8215
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.63075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 55 - Sample 55 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 55
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9485
varray: [2.913, 2.958, 2.973, 2.95]
inst: mean_all
operation: mean_all(residuals)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: residuals
SAMP_residuals_STD | 0.0255147016443
varray: [2.913, 2.958, 2.973, 2.95]
inst: std_all
operation: std_all(residuals)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5475
varray: [0.507, 0.646, 0.613, 0.424]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101464279429
varray: [0.507, 0.646, 0.613, 0.424]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.881
varray: [1298.491, 1299.299, 1299.18, 1298.554]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.417593901616
varray: [1298.491, 1299.299, 1299.18, 1298.554]
inst: std_all
operation: std_all(baseline_shift)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.901
varray: [2.754, 2.844, 3.148, 2.858]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.170993177252
varray: [2.754, 2.844, 3.148, 2.858]
inst: std_all
operation: std_all(slope_shift)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0259358182186
varray: [-12.848, -12.898, -12.873, -12.841]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.110173726451
varray: [-109.339, -109.558, -109.561, -109.559]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.9745
varray: [112.553, 113.933, 113.058, 112.354]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.704361886154
varray: [112.553, 113.933, 113.058, 112.354]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9485
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.881
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.901
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.98892868772
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.6845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-12.848, -12.898, -12.873, -12.841]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-109.339, -109.558, -109.561, -109.559]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.865
varray: [-12.848, -12.898, -12.873, -12.841]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
dD_mean | -109.50425
varray: [-109.339, -109.558, -109.561, -109.559]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0259358182186
varray: [-12.848, -12.898, -12.873, -12.841]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
dD_stdev | 0.110173726451
varray: [-109.339, -109.558, -109.561, -109.559]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.865
varray: [-12.848, -12.898, -12.873, -12.841]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
measured_dD_mean | -109.50425
varray: [-109.339, -109.558, -109.561, -109.559]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0259358182186
varray: [-12.848, -12.898, -12.873, -12.841]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.110173726451
varray: [-109.339, -109.558, -109.561, -109.559]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.5861883227
actual_array: [ 1.93 -15.84]
varray: [-12.848, -12.898, -12.873, -12.841]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.865
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -120.351886844
actual_array: [ 18.1 -119. ]
varray: [-109.339, -109.558, -109.561, -109.559]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.50425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [131916.0, 132452.0, 132988.0, 133522.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.82925
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.89675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 56 - Sample 56 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 56
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95475
varray: [2.892, 2.961, 3.034, 2.932]
inst: mean_all
operation: mean_all(residuals)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: residuals
SAMP_residuals_STD | 0.0599298200676
varray: [2.892, 2.961, 3.034, 2.932]
inst: std_all
operation: std_all(residuals)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4705
varray: [0.322, 0.615, 0.348, 0.597]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.156993630444
varray: [0.322, 0.615, 0.348, 0.597]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.0295
varray: [1296.877, 1297.444, 1302.591, 1299.206]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.57309806783
varray: [1296.877, 1297.444, 1302.591, 1299.206]
inst: std_all
operation: std_all(baseline_shift)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.55025
varray: [2.541, 2.256, 2.582, 2.822]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.232037892021
varray: [2.541, 2.256, 2.582, 2.822]
inst: std_all
operation: std_all(slope_shift)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0776595132614
varray: [-13.0, -12.943, -12.837, -12.85]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.229907952596
varray: [-110.015, -109.662, -109.465, -109.78]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.30625
varray: [111.015, 113.335, 115.665, 113.21]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.89944015875
varray: [111.015, 113.335, 115.665, 113.21]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4705
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.0295
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.55025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11625
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11625
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.78541191794
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.35275
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-13.0, -12.943, -12.837, -12.85]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-110.015, -109.662, -109.465, -109.78]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.9075
varray: [-13.0, -12.943, -12.837, -12.85]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
dD_mean | -109.7305
varray: [-110.015, -109.662, -109.465, -109.78]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0776595132614
varray: [-13.0, -12.943, -12.837, -12.85]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
dD_stdev | 0.229907952596
varray: [-110.015, -109.662, -109.465, -109.78]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.9075
varray: [-13.0, -12.943, -12.837, -12.85]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
measured_dD_mean | -109.7305
varray: [-110.015, -109.662, -109.465, -109.78]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0776595132614
varray: [-13.0, -12.943, -12.837, -12.85]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.229907952596
varray: [-110.015, -109.662, -109.465, -109.78]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.628457117
actual_array: [ 1.93 -15.84]
varray: [-13.0, -12.943, -12.837, -12.85]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.9075
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -120.562820535
actual_array: [ 18.1 -119. ]
varray: [-110.015, -109.662, -109.465, -109.78]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.7305
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [134110.0, 134646.0, 135181.0, 135716.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.823
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.74825
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 57 - Sample 57 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Sample Variables
Sample 57
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97825
varray: [2.983, 2.972, 2.984, 2.974]
inst: mean_all
operation: mean_all(residuals)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: residuals
SAMP_residuals_STD | 0.00613052471925
varray: [2.983, 2.972, 2.984, 2.974]
inst: std_all
operation: std_all(residuals)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31875
varray: [0.396, 0.487, 0.158, 0.234]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149776221967
varray: [0.396, 0.487, 0.158, 0.234]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.242
varray: [1299.929, 1299.817, 1300.798, 1300.424]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.454919040416
varray: [1299.929, 1299.817, 1300.798, 1300.424]
inst: std_all
operation: std_all(baseline_shift)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.47125
varray: [2.55, 2.48, 2.465, 2.39]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.065622023742
varray: [2.55, 2.48, 2.465, 2.39]
inst: std_all
operation: std_all(slope_shift)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: slope_shift
SAMP_d18O_STD | 0.152075639075
varray: [-7.811, -7.592, -7.517, -7.466]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.54637315608
varray: [-65.817, -61.441, -60.574, -60.403]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.1805
varray: [113.338, 114.024, 112.745, 112.615]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.644398686115
varray: [113.338, 114.024, 112.745, 112.615]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.242
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.47125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.02018886356
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4785
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1202
varray: [-7.811, -7.592, -7.517, -7.466]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.0542
varray: [-65.817, -61.441, -60.574, -60.403]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.5965
varray: [-7.811, -7.592, -7.517, -7.466]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
dD_mean | -62.05875
varray: [-65.817, -61.441, -60.574, -60.403]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
d18O_stdev | 0.152075639075
varray: [-7.811, -7.592, -7.517, -7.466]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
dD_stdev | 2.54637315608
varray: [-65.817, -61.441, -60.574, -60.403]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.5965
varray: [-7.811, -7.592, -7.517, -7.466]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.05875
varray: [-65.817, -61.441, -60.574, -60.403]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.152075639075
varray: [-7.811, -7.592, -7.517, -7.466]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.54637315608
varray: [-65.817, -61.441, -60.574, -60.403]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3463496799
actual_array: [ 1.93 -15.84]
varray: [-7.811, -7.592, -7.517, -7.466]
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.5965
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.1182761088
actual_array: [ 18.1 -119. ]
varray: [-65.817, -61.441, -60.574, -60.403]
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -62.05875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: [136301.0, 136837.0, 137372.0, 137908.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1305.7995
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.53575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.974
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.537
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7217960285
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.974
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.2256683379
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.537
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.064
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8113064162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -8.064
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.9937335053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -65.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.8
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5487426122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.8
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0802847368
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -62.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.615
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.512
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3647490373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.615
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6762355063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.512
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.657
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.602
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4065205516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.657
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7601428308
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.602
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.568
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.179
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3180047237
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.568
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3657784055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.179
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.537
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.83
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.287173368
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.537
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.040404447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.83
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.531
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.674
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2812060088
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.531
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8949650845
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.674
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.483
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.496
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2334671353
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7290150427
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.496
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.503
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.629
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2533583326
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.503
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8530114222
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.629
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.171
operation: col(d(18_16)Mean)
inst: col
dD_actual | 29.283
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.3571335184
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.171
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 9.03996678796
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 29.283
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | 37.152
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.80866369661
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 16.376263863
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 37.152
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.765
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94790207755
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.765
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.3350337391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 39.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.746
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.861
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.92900544014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.746
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.9018743315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 39.861
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.02846142652
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.3344632046
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 40.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.842
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.524
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.02448318707
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.842
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5199916222
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 40.524
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.924
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.233
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10603709591
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.924
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1809948788
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 41.233
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.937
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.11896637414
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.937
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1548903778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 41.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.862
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.859
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04437438435
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.8323133302
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 40.859
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.752
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.314
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.747
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 39.0009
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.752
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.314
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.93497279932
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.752
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.3242078649
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 40.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.211
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.9357461718
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.211
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.869820088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -92.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.006
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.884
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7264212635
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.006
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.976714714
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -105.884
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.128
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.28
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8477575669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.128
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.210514154
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.28
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.168
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.384
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8875399615
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.168
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.239777335
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.384
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.195
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.88
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9143930778
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.195
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.702199924
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.88
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.248
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.377
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9671047506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.248
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.165554816
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -110.377
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.301
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.615
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0198164234
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.301
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.387443074
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -110.615
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.299
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.91
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0178273037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.299
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.662472638
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -110.91
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.311
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.102
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.029762022
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.311
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.84147493
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -111.102
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.335
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.212
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0536314588
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.335
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.944028327
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -111.212
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.534
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6024488448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.3582642152
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -65.534
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.579
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.996
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3289448822
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.579
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1274704516
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.996
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.379
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3428687203
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5522391267
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.379
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.566
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.861
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.316015604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.566
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0693058588
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.861
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.601
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.717
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3508251992
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.601
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9350541396
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.717
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.555
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3050754455
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.555
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7560518472
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.513
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2861788081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.744864204
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.513
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.525
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.474
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2752386496
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7085043633
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.594
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3438632802
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.594
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7709687049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.554
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.47
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3040808856
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.554
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7047751489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.47
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.196
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.319
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.95346859049
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.196
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.1423730289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -80.319
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.132
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.009
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.8898167592
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.7179661229
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -82.009
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.095
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.352
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.85301804424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.095
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.0377462597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -82.352
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.078
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.589
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.83611052655
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.078
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.2587022143
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -82.589
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.744
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.92701632041
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.744
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.0207845903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -45.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.174
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.202
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.35467706188
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.174
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.6057564818
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -42.202
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.271
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.453
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.45114936868
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.271
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.9074610809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -41.453
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.34
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.91
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.51977399928
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.34
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.4012202229
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -40.91
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 6.05046151607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.9000026521
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -41.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.776
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.437
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.93708169159
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.776
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.8248478291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -42.437
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.774
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.93509257186
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.774
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.0802990172
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -42.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.834
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.013
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.9947661637
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.834
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.3618547062
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -43.013
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.45
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.353
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.63461572043
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.45
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4603028389
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.353
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.442
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.62665924152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.442
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4705581785
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.984
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.028
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.16571068774
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.984
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3603933492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -58.028
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.964
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.18162364556
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3007259184
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.964
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.076
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.635
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.25721019522
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.076
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.993998032
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.635
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.103
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.28406331154
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.103
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9809457815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.089
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.177
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.2538190651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.089
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.2439669212
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -47.177
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.103
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.651
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.2677429032
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.103
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.7535752245
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -46.651
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.072
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.727
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.23691154742
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.072
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.8244302986
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -46.727
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.093
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.304
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.25779730456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 8.093
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.4300658733
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -46.304
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.923
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.30940124922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.2470555594
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -72.923
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.791
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.56699225396
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.791
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.6352556287
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -74.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.917
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.69230679681
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.917
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.0687768054
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -74.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.963
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.135
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.73805655055
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.963
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.3093111358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -75.135
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.903
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.10153163338
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.6500337629
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -51.903
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.211
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.732
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.39691591296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.211
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.6260026344
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -49.732
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.292
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.331
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.47747526193
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.292
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.2521488884
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -49.331
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.32
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.873
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.50532293812
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.32
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.8251538369
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -48.873
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.743
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.984
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.49749283603
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.743
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3193719905
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.984
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.196
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.94802845437
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.196
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4036410842
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.336
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0872668353
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.336
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5229759458
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.384
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.441
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1350057088
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.384
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6777383443
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.441
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.499
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.112
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.2221794126
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.499
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.392369119
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -103.112
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.3474939554
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.084291398
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.669
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.13
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.3912545894
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.669
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.070668613
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.13
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.698
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.395
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.4200968255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.698
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.317729069
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.395
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.265
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.242
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.456062281723
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.265
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.8306918971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -54.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.882
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.308
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.06970571774
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.882
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.2307059055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -49.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.046
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.761
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.23281353542
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.046
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.7884322271
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -47.761
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.159
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.142
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.34519880004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.159
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.211336295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -47.142
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.247
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.126
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.03683196024
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.247
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.7747951618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -68.126
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.493
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.047
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.28149368676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.493
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5657503888
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -70.047
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.509
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.29740664458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.509
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.0766527648
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -70.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.471
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.801
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.25961336975
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.471
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.2687073077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -70.801
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.346
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.44705796096
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.346
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6485724059
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.186
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.3247270977
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9661261663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.186
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.156
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.32771077729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9381570581
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.156
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.485
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.303
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.30881413988
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.485
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.0752056882
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.303
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.948
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.718
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.853773059013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.948
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.3421648076
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -53.718
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.013
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.301
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.789126667861
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.013
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.021090598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.301
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.006
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.047
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.796088586908
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.006
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.7842854821
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.047
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.995
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.807028745411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.995
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.5446834554
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -51.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.725
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.926
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.07555990866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.725
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.4006911695
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -55.926
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.997984239276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.352211382
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -55.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.666
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.391
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.13423894063
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.666
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8342123463
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.391
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.706
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.09445654607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.706
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.3997588659
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -55.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.368
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.55306181159
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.368
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.1706767055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -48.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.368
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.55306181159
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.368
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.1706767055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -48.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.325
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.92
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.11440762962
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.325
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7181334071
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -65.92
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.33
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.019
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.11938042894
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.33
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.7427350701
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.019
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.303
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.199
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.09252731262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.303
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.9105497191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.199
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.341
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.129
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.13032058745
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.341
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.8452884667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.129
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.033
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.341
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78591519656
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.033
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5168115897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.341
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.397
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.75
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.147934987
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.397
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9658201586
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.75
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1907010612
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7830886518
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.471
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2215324169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.471
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9676847658
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.519
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.642
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.2420706098
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.519
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.88649003
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -103.642
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.764
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.936
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.4857377765
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.764
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.889801714
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.936
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.886
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.899
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6070740799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.886
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.787610086
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.899
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.819
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.981
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.540438569
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.819
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.864058982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.981
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.972
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.414
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.74156745321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.972
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.6371202359
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -74.414
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.504
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.529
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.27611343691
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.504
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.0151207269
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -70.529
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.394
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.075
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.16671185189
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.394
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5918548898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -70.075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.378
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.562
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.15079889406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.378
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.1135831399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -69.562
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.929
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.10556975911
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.929
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.581043296
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -51.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.241
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.424
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.41587243664
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 6.241
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.2711567297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -50.424
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.322
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.043
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.49643178562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 6.322
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.9159490558
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -50.043
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.316
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.975
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.49046442644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 6.316
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.8525524106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -49.975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.184
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.786
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.35374238829
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 7.184
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.1502217876
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -42.786
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.312
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.321
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.48104605086
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 7.312
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.7843970049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -41.321
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.358
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.827
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.5267958046
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 7.358
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.3238390236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -40.827
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.308
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.012
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.47706781141
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 7.308
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.4963151907
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -41.012
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.199
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.18162364556
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 5.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.7906028421
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -54.199
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.096
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.92502720068
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5591827825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.096
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.766
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.634
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94889663741
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.766
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0607621225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.634
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.751
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.659
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.93397823946
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0840697126
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.659
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.967
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.401
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.15424330617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.967
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.0466245645
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -53.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.534
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.20695497896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2383173382
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.534
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.051
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.929
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.23778633474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.051
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.6742736566
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -51.929
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.308
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.18905290141
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 4.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.0276167233
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.695751208919
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2681510536
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.362
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.279
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.55253458852
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.362
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.9328835246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -53.279
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.41
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.600273461986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.41
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5347898849
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.966
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.616186419808
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 3.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.641072496
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -52.966
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.383
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.658
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.1612118295
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.383
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.2030842861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -69.658
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.006
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.37902043969
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.4598295469
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -71.006
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.642
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.41880283424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.642
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.7488436647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -71.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.673
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.44963419002
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -2.673
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.0089563708
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -71.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.204
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.032
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.95598493329
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.204
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1610333814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.032
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.389
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1399785081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.389
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9052204242
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.396
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1469404272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.396
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8213130996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.416
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.429
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1668316244
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.416
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6665507011
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.429
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.545
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.697
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.2679291663
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.545
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.937766728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -103.697
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.785
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.872
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5066235336
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.785
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.830134283
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.872
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.706
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5742536044
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.60767549
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.706
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.112
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5752481642
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.986190754
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.112
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.103
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.249
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.68873600788
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.103
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.0248612935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.612
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.783
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.18250503716
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.612
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.7934969953
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -62.783
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.781
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.66
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.01442442017
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.781
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7465200459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.66
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.826
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.96966922629
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.826
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4724227857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.033
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.37346053103
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0265728968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -63.033
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.417
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.37644421062
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.417
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.2978732462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -63.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.459
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.099
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.33467269634
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0881049348
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -63.099
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.45
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.766
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.34362373511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.45
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.777647834
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -62.766
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.074
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.203
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.86477310379
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.074
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3881536921
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.203
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.093
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.8836697412
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.093
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5046916428
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.093
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.8836697412
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.093
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5046916428
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.007
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.78181718456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.007
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.096286902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -65.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.021
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.297
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.79574102266
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.021
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1373082606
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -65.297
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.104
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.974
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.8946098997
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -0.104
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0392633781
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.974
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.012
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.4
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.77924095549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.012
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.5041211083
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.4
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.017
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.604
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.77426815617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.017
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.694311044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.604
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.033
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.506
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.75835519835
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.033
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6029452906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.506
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.405
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.399259201218
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.405
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.386136081
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.642
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.378
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.163548513477
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.642
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8897887935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -55.378
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.457
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.347542088296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.457
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.4417125282
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -55.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.329640010746
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 2.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5293490671
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -56.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.79117567386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2606797044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.92117533805
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4073702986
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.331
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.429
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.082294036
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.331
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6665507011
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.429
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.443
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.743
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1936847408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.443
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9592940333
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.395
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2483855333
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6348523785
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.395
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.573
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.567
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.2957768425
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.573
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.816567259
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -103.567
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.682
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.4041838676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.682
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.33041955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -107.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.836
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5742536044
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.728874959
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.836
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.92
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5752481642
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.807188462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.92
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.366
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.386
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.12254376735
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.366
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.2048373705
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -80.386
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.223
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.926
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.98032170681
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9113704999
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.926
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.384
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.86893100206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.4060619455
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.384
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.174
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.3
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.93158827348
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.174
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.3277484426
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.3
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.765
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.69090478642
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7612696193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.765
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.95
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.989
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.70880686397
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.95
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9701056271
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.989
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.848
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.60736175786
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.848
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.6866853309
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.878
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.785
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.63719855377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.878
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7799156915
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.785
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.083
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.709
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.8519635981
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -4.083
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7090606174
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.709
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.996
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.567
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.76543688994
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.5766735054
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.567
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.041
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.706
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.81019208382
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -4.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7062637066
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.706
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.985
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.453
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.75449673144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.985
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.4703908943
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -77.453
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.449
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.41
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.2214126444
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.449
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.7010874155
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.41
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.483
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.25522767977
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4558915672
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.428
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.091
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.20052688726
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.428
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4036825652
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.091
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.431
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.847
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.20351056685
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.431
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.1762004854
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -72.847
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.21246160562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.326301366
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.113
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.21246160562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4241932446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.113
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.457
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.22936912331
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.457
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4810637645
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.131
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.21842896481
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -3.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4409747095
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -73.131
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.573
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.626
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.22129287185
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.573
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.3086433588
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.626
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.81
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.93
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.98558218411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.81
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6597600491
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.93
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.808
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.99254410316
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.5460190092
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.808
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.863
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.751
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.93287051133
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.863
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.4928777037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -66.751
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.611
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.78386619056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.3469092877
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -81.611
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.282
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.054
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.0444408749
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.6922233911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -83.054
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.293
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.268
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.0553810334
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.293
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.8917363628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -83.268
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.355
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.159
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.11704374496
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -5.355
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.7901152697
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -83.159
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.31
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0614082789
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.31
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2191979741
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -62.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.459
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.596
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2095976986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7545490092
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.596
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2165596176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0534566975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.491
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.945
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2414236142
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.491
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1476193617
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.945
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 50 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.037
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2444072938
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2333912935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.037
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2115868183
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9294603179
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2563420122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9201372819
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2563420122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8940327809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -59.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 51 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.595
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.996
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.19941255485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.595
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3305596338
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.996
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.003
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.79363213038
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.003
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9296690832
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.053
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74390413719
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.053
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9949303356
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.096
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.70113806304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 1.096
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8429648479
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -57.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 52 - Sample 52 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 52 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.204
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.9614250694
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.204
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.790324035
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -81.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 52 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.33
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.08673961225
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.33
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.4125323093
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -82.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 52 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.483
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.23
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.23890727142
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.8563088257
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -83.23
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 52 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.524
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.27968422584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -6.524
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.4688322948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -83.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 53 - Sample 53 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 53 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.462
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.442
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.33168901675
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.462
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.0017067072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -69.442
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 53 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.806
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.807
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.98956042357
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.806
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.4773903115
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.807
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 53 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.804
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.753
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.9915495433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.804
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.4270459168
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.753
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 53 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.908
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.641
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.88811531745
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: 0.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.3226279129
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -67.641
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 54 - Sample 54 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 54 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.8939746575
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.873799549
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -104.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 54 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.126
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.3544558745
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.066939399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.126
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 54 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.764
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.4827540969
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.661749099
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -108.764
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 54 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.08
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5235310513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.956357039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.08
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 55 - Sample 55 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 55 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.848
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5692808051
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.848
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.197823673
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 55 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.898
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.558
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6190087982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.898
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.401998163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.558
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 55 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.873
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5941448016
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.873
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.404795073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 55 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.841
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.559
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.562318886
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.841
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.402930466
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.559
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 56 - Sample 56 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 56 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7204539044
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -13.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.828060911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -110.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 56 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.662
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6637639921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.498957738
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.662
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 56 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.465
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5583406465
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.315293927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.465
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 56 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.85
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.78
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5712699248
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -12.85
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.608969563
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -109.78
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean

## 57 - Sample 57 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:25:18
Detail Variables
Detail 57 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.811
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.817
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5596827707
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.811
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.6221061357
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -65.817
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 57 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.592
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.441
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3418741605
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.592
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.5423455562
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -61.441
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 57 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.517
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.574
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2672821707
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.517
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7340383299
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.574
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean
Detail 57 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.747
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 39.0009
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.335
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.212
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1202
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.0542
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.403
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2165596176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.80644310636
measure_array: [ 4.747 -13.1202]
inst: fit_type1
actual_mean: -7.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00546989308
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5746144132
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.5866234865
measure_array: [ 39.0009 -108.0542]
inst: fit_type1
actual_mean: -60.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.07261196207
tarray: []
source_col: d(D_H)Mean