## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
220506_22-084_22-095_21-069_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.2631 | -74.6411 |  | 3.2158 |  | 0.055296 | 0.57775
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.18884 | 3.1561
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.28935 | 4.6024
4 |  | EV | -10.2801 | -74.7206 |  |  |  | 0.062823 | 1.4104
5 |  | A00000209752 | -9.2615 | -70.1155 |  |  |  | 0.042603 | 0.1838
6 |  | A00000209761 | -1.4425 | -3.1778 |  |  |  | 0.16996 | 2.9438
7 |  | A00000209762 | -2.2332 | -7.7736 |  |  |  | 0.062174 | 0.27227
8 |  | A00000319094 | -5.1461 | -33.2853 |  |  |  | 0.072709 | 1.0738
9 |  | A00000319080 | -5.4498 | -36.0876 |  |  |  | 0.030977 | 0.16078
10 |  | A00000319071 | -5.6976 | -36.7362 |  |  |  | 0.031337 | 0.04997
11 |  | A00000307888 | -14.4095 | -108.2873 |  |  |  | 0.18784 | 3.1714
12 |  | A00000307886 | -20.9073 | -173.2338 |  |  |  | 0.12777 | 2.8322
13 |  | EV | -10.4227 | -78.0988 |  |  |  | 0.21567 | 4.0086
14 |  | A00000307890 | -18.212 | -145.5774 |  |  |  | 0.16147 | 3.0078
15 |  | A00000307918 | -20.2499 | -183.0071 |  |  |  | 0.01103 | 1.5519
16 |  | A00000278712 | -12.3215 | -91.4925 |  |  |  | 0.13298 | 3.8113
17 |  | A00000278697 | -10.1485 | -69.794 |  |  |  | 0.062195 | 1.1646
18 |  | A00000278679 | -10.6818 | -75.9924 |  |  |  | 0.021119 | 0.42002
19 |  | A00000278680 | -10.2579 | -72.2678 |  |  |  | 0.040237 | 0.35716
20 |  | A00000280479 | -9.2041 | -66.0354 |  |  |  | 0.033867 | 0.22243
21 |  | A00000280480 | -9.6398 | -69.0198 |  |  |  | 0.03567 | 0.36051
22 |  | EV | -10.2526 | -73.4681 |  |  |  | 0.035246 | 0.13207
23 |  | A00000280481 | -9.7062 | -69.4374 |  |  |  | 0.0094296 | 0.35837
24 |  | A00000280478 | -9.7144 | -69.5541 |  |  |  | 0.022396 | 0.10687
25 |  | A00000280477 | -7.3014 | -55.085 |  |  |  | 0.063275 | 0.7241
26 |  | A00000280492 | -7.5857 | -55.43 |  |  |  | 0.021188 | 0.18506
27 |  | A00000279569 | -7.5485 | -55.1388 |  |  |  | 0.046307 | 0.1537
28 |  | EVIAN 1-2 - 3-31-22 | -10.1101 | -72.102 |  |  |  | 0.046743 | 0.74951
29 |  | FLUKE D-G 1491 | -0.74596 | -3.2834 |  |  |  | 0.1711 | 2.895
30 |  | Beta2 -3-16-22 | -0.43535 | 0.68539 |  |  |  | 0.055307 | 0.40103
31 |  | EV | -10.0181 | -69.5387 |  |  |  | 0.17407 | 3.0217
32 |  | FLUKE D-G 1490 | -0.27517 | -1.0671 |  |  |  | 0.18316 | 2.8989
33 |  | Beta1 -3-16-22 | -0.44534 | 1.0142 |  |  |  | 0.024118 | 0.080616
34 |  | EVIAN 1-1 3-31-22 | -9.936 | -68.993 |  |  |  | 0.21281 | 3.0237
35 |  | HARV2-C-4 | -3.6253 | -59.204 |  |  |  | 0.12137 | 0.31702
36 |  | EV | -10.086 | -72.2843 |  |  |  | 0.12372 | 0.59485

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.28 | -69.503 | -10.3979 | -76.0559 | -0.969 | 69.656 | 2022/05/06 09:38:49
1 |  | EV | -9.187 | -68.842 | -10.3025 | -75.3416 | -1.023 | 74.195 | 2022/05/06 09:47:45
1 |  | EV | -9.169 | -68.384 | -10.2841 | -74.8467 | -0.972 | 73.767 | 2022/05/06 09:56:40
1 |  | EV | -9.157 | -68.187 | -10.2718 | -74.6338 | -1.05 | 73.338 | 2022/05/06 10:05:36
1 |  | EV | -9.104 | -67.863 | -10.2174 | -74.2837 | -0.927 | 72.992 | 2022/05/06 10:14:33
1 |  | EV | -9.136 | -67.903 | -10.2502 | -74.3269 | -0.995 | 72.372 | 2022/05/06 10:23:29
1 |  | EV | -9.101 | -67.973 | -10.2144 | -74.4026 | -0.875 | 71.642 | 2022/05/06 10:32:27
1 |  | EV | -9.141 | -67.883 | -10.2554 | -74.3053 | -1.038 | 71.328 | 2022/05/06 10:41:22
1 |  | EV | -9.108 | -67.646 | -10.2215 | -74.0492 | -1.107 | 70.893 | 2022/05/06 10:50:19
1 |  | EV | -9.102 | -67.753 | -10.2154 | -74.1649 | -0.906 | 70.473 | 2022/05/06 10:59:16
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.2631 | -74.6411 |  | 3.2158 |  | 0.055296 | 0.57775
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.251 | 9.338 | 1.423 | 9.1406 | -1.123 | 68.605 | 2022/05/06 11:09:05
2 |  | PZ | 2.643 | 15.707 | 1.8248 | 16.0231 | -1.224 | 69.876 | 2022/05/06 11:18:02
2 |  | PZ | 2.704 | 17.493 | 1.8874 | 17.953 | -0.9 | 69.314 | 2022/05/06 11:26:58
2 |  | PZ | 2.775 | 18.145 | 1.9601 | 18.6576 | -0.976 | 69.331 | 2022/05/06 11:35:56
2 |  | PZ | 2.835 | 18.667 | 2.0216 | 19.2217 | -1.026 | 69.196 | 2022/05/06 11:44:51
2 |  | PZ | 2.821 | 18.973 | 2.0073 | 19.5523 | -1.231 | 69.166 | 2022/05/06 11:53:47
2 |  | PZ | 2.86 | 19.152 | 2.0473 | 19.7458 | -1.159 | 69.008 | 2022/05/06 12:02:44
2 |  | PZ | 2.863 | 19.669 | 2.0504 | 20.3044 | -1.119 | 68.776 | 2022/05/06 12:11:39
2 |  | PZ | 2.854 | 19.632 | 2.0411 | 20.2645 | -1.277 | 68.737 | 2022/05/06 12:20:36
2 |  | PZ | 2.85 | 19.514 | 2.037 | 20.137 | -1.117 | 68.82 | 2022/05/06 12:29:33
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.18884 | 3.1561
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.817 | -97.112 | -15.0489 | -105.8906 | -1.096 | 68.507 | 2022/05/06 12:39:21
3 |  | UT2 | -14.452 | -106.547 | -15.6999 | -116.0861 | -1.004 | 68.526 | 2022/05/06 12:48:18
3 |  | UT2 | -14.558 | -108.938 | -15.8085 | -118.6699 | -1.07 | 68.765 | 2022/05/06 12:57:14
3 |  | UT2 | -14.646 | -110.062 | -15.8987 | -119.8845 | -0.824 | 68.718 | 2022/05/06 13:06:10
3 |  | UT2 | -14.684 | -110.899 | -15.9377 | -120.789 | -0.948 | 68.699 | 2022/05/06 13:15:07
3 |  | UT2 | -14.723 | -111.157 | -15.9777 | -121.0678 | -1.007 | 68.752 | 2022/05/06 13:24:04
3 |  | UT2 | -14.755 | -111.676 | -16.0105 | -121.6286 | -1.025 | 68.653 | 2022/05/06 13:33:00
3 |  | UT2 | -14.729 | -111.8 | -15.9838 | -121.7626 | -1.108 | 68.618 | 2022/05/06 13:41:56
3 |  | UT2 | -14.783 | -112.1 | -16.0392 | -122.0868 | -0.951 | 68.589 | 2022/05/06 13:50:53
3 |  | UT2 | -14.74 | -112.144 | -15.9951 | -122.1343 | -0.841 | 68.34 | 2022/05/06 13:59:49
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.28935 | 4.6024
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.303 | -72.016 | -10.4214 | -78.7715 | -0.869 | 68.028 | 2022/05/06 14:09:39
4 |  | EV | -9.222 | -68.931 | -10.3384 | -75.4378 | -0.828 | 67.869 | 2022/05/06 14:18:36
4 |  | EV | -9.21 | -68.525 | -10.3261 | -74.9991 | -0.914 | 67.965 | 2022/05/06 14:27:32
4 |  | EV | -9.116 | -67.717 | -10.2297 | -74.1259 | -1.018 | 67.95 | 2022/05/06 14:36:27
4 |  | EV | -9.152 | -67.778 | -10.2666 | -74.1919 | -0.828 | 67.863 | 2022/05/06 14:45:24
4 |  | EV | -9.107 | -67.675 | -10.2205 | -74.0806 | -0.819 | 67.843 | 2022/05/06 14:54:20
4 |  | EV | -9.147 | -67.77 | -10.2615 | -74.1832 | -0.903 | 67.879 | 2022/05/06 15:03:16
4 |  | EV | -9.111 | -67.564 | -10.2246 | -73.9606 | -0.883 | 67.877 | 2022/05/06 15:12:12
4 |  | EV | -9.12 | -67.45 | -10.2338 | -73.8374 | -1.061 | 67.949 | 2022/05/06 15:21:09
4 |  | EV | -9.163 | -67.247 | -10.2779 | -73.6181 | -1.085 | 67.822 | 2022/05/06 15:30:06
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.2801 | -74.7206 |  |  |  | 0.062823 | 1.4104
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | A00000209752 | -8.222 | -64.153 | -9.3133 | -70.2746 | -1.097 | 68.113 | 2022/05/06 15:39:56
5 |  | A00000209752 | -8.161 | -63.977 | -9.2507 | -70.0845 | -1.114 | 67.983 | 2022/05/06 15:48:53
5 |  | A00000209752 | -8.183 | -64.136 | -9.2733 | -70.2563 | -0.964 | 67.922 | 2022/05/06 15:57:49
5 |  | A00000209752 | -8.12 | -63.757 | -9.2087 | -69.8467 | -1.2 | 67.913 | 2022/05/06 16:06:46
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | A00000209752 | -9.2615 | -70.1155 |  |  |  | 0.042603 | 0.1838
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | A00000209761 | -0.797 | -6.386 | -1.7016 | -7.8509 | -0.966 | 67.924 | 2022/05/06 16:16:37
6 |  | A00000209761 | -0.49 | -1.44 | -1.3869 | -2.5062 | -0.991 | 68.056 | 2022/05/06 16:25:33
6 |  | A00000209761 | -0.453 | -0.375 | -1.349 | -1.3553 | -1.118 | 68.012 | 2022/05/06 16:34:29
6 |  | A00000209761 | -0.437 | -0.045 | -1.3326 | -0.99874 | -1.051 | 67.946 | 2022/05/06 16:43:26
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | A00000209761 | -1.4425 | -3.1778 |  |  |  | 0.16996 | 2.9438
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | A00000209762 | -1.233 | -6.008 | -2.1486 | -7.4424 | -1.062 | 68.008 | 2022/05/06 16:53:16
7 |  | A00000209762 | -1.378 | -6.616 | -2.2972 | -8.0994 | -1.214 | 68.101 | 2022/05/06 17:02:12
7 |  | A00000209762 | -1.307 | -6.454 | -2.2245 | -7.9244 | -1.036 | 67.482 | 2022/05/06 17:11:08
7 |  | A00000209762 | -1.344 | -6.18 | -2.2624 | -7.6283 | -1.098 | 67.411 | 2022/05/06 17:20:05
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | A00000209762 | -2.2332 | -7.7736 |  |  |  | 0.062174 | 0.27227
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | A00000319094 | -4.058 | -28.361 | -5.0446 | -31.5974 | -0.987 | 67.673 | 2022/05/06 17:29:53
8 |  | A00000319094 | -4.184 | -30.226 | -5.1738 | -33.6127 | -0.94 | 67.292 | 2022/05/06 17:38:49
8 |  | A00000319094 | -4.156 | -30.293 | -5.1451 | -33.6851 | -1.05 | 67.465 | 2022/05/06 17:47:45
8 |  | A00000319094 | -4.23 | -30.812 | -5.2209 | -34.2459 | -1.214 | 67.445 | 2022/05/06 17:56:42
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | A00000319094 | -5.1461 | -33.2853 |  |  |  | 0.072709 | 1.0738
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | A00000319080 | -4.483 | -32.608 | -5.4803 | -36.1867 | -1.047 | 67.389 | 2022/05/06 18:06:30
9 |  | A00000319080 | -4.416 | -32.326 | -5.4116 | -35.882 | -0.914 | 67.606 | 2022/05/06 18:15:25
9 |  | A00000319080 | -4.474 | -32.684 | -5.4711 | -36.2689 | -1.076 | 67.832 | 2022/05/06 18:24:22
9 |  | A00000319080 | -4.44 | -32.447 | -5.4362 | -36.0127 | -1.146 | 68 | 2022/05/06 18:33:19
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | A00000319080 | -5.4498 | -36.0876 |  |  |  | 0.030977 | 0.16078
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | A00000319071 | -4.654 | -33.081 | -5.6556 | -36.6979 | -1.039 | 67.97 | 2022/05/06 18:43:08
10 |  | A00000319071 | -4.729 | -33.081 | -5.7325 | -36.6979 | -0.876 | 68.082 | 2022/05/06 18:52:03
10 |  | A00000319071 | -4.705 | -33.187 | -5.7079 | -36.8124 | -1.153 | 67.981 | 2022/05/06 19:01:00
10 |  | A00000319071 | -4.692 | -33.117 | -5.6945 | -36.7368 | -1.05 | 67.626 | 2022/05/06 19:09:56
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | A00000319071 | -5.6976 | -36.7362 |  |  |  | 0.031337 | 0.04997
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | A00000307888 | -12.914 | -94.713 | -14.1232 | -103.2982 | -1.035 | 68.276 | 2022/05/06 19:19:45
11 |  | A00000307888 | -13.251 | -99.801 | -14.4687 | -108.7963 | -1.212 | 68.11 | 2022/05/06 19:28:42
11 |  | A00000307888 | -13.302 | -101.26 | -14.521 | -110.3729 | -0.935 | 68.35 | 2022/05/06 19:37:37
11 |  | A00000307888 | -13.306 | -101.546 | -14.5251 | -110.682 | -0.947 | 67.632 | 2022/05/06 19:46:33
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | A00000307888 | -14.4095 | -108.2873 |  |  |  | 0.18784 | 3.1714
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | A00000307886 | -19.354 | -155.421 | -20.7251 | -168.9 | -1.011 | 68.517 | 2022/05/06 19:56:24
12 |  | A00000307886 | -19.532 | -159.526 | -20.9076 | -173.3359 | -0.982 | 67.918 | 2022/05/06 20:05:20
12 |  | A00000307886 | -19.592 | -161 | -20.9691 | -174.9287 | -0.852 | 68.144 | 2022/05/06 20:14:15
12 |  | A00000307886 | -19.649 | -161.779 | -21.0275 | -175.7705 | -0.862 | 67.812 | 2022/05/06 20:23:12
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | A00000307886 | -20.9073 | -173.2338 |  |  |  | 0.12777 | 2.8322
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -9.608 | -77.179 | -10.7341 | -84.3507 | -0.65 | 68.073 | 2022/05/06 20:33:01
13 |  | EV | -9.306 | -70.983 | -10.4245 | -77.6552 | -0.79 | 67.933 | 2022/05/06 20:41:58
13 |  | EV | -9.169 | -68.951 | -10.2841 | -75.4594 | -0.818 | 67.685 | 2022/05/06 20:50:55
13 |  | EV | -9.134 | -68.461 | -10.2482 | -74.9299 | -0.967 | 67.85 | 2022/05/06 20:59:52
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.4227 | -78.0988 |  |  |  | 0.21567 | 4.0086
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | A00000307890 | -16.663 | -129.461 | -17.9664 | -140.8473 | -0.915 | 67.771 | 2022/05/06 21:09:43
14 |  | A00000307890 | -16.953 | -134.313 | -18.2637 | -146.0904 | -0.796 | 68.147 | 2022/05/06 21:18:38
14 |  | A00000307890 | -16.982 | -135.525 | -18.2935 | -147.4001 | -1.023 | 68.15 | 2022/05/06 21:27:35
14 |  | A00000307890 | -17.012 | -136.054 | -18.3242 | -147.9718 | -1.028 | 68.093 | 2022/05/06 21:36:31
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | A00000307890 | -18.212 | -145.5774 |  |  |  | 0.16147 | 3.0078
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | A00000307918 | -18.877 | -166.247 | -20.2361 | -180.5987 | -1.281 | 68.012 | 2022/05/06 21:46:18
15 |  | A00000307918 | -18.89 | -168.692 | -20.2494 | -183.2408 | -1.054 | 67.86 | 2022/05/06 21:55:14
15 |  | A00000307918 | -18.891 | -169.177 | -20.2504 | -183.7649 | -0.987 | 67.62 | 2022/05/06 22:04:09
15 |  | A00000307918 | -18.904 | -169.787 | -20.2638 | -184.4241 | -0.907 | 67.627 | 2022/05/06 22:13:05
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | A00000307918 | -20.2499 | -183.0071 |  |  |  | 0.01103 | 1.5519
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | A00000278712 | -11.338 | -89.277 | -12.5076 | -97.424 | -0.946 | 68.277 | 2022/05/06 22:22:52
16 |  | A00000278712 | -11.136 | -83.297 | -12.3005 | -90.9619 | -1.013 | 67.821 | 2022/05/06 22:31:49
16 |  | A00000278712 | -11.134 | -81.888 | -12.2985 | -89.4393 | -1.127 | 67.897 | 2022/05/06 22:40:46
16 |  | A00000278712 | -11.018 | -80.69 | -12.1795 | -88.1447 | -1 | 67.821 | 2022/05/06 22:49:41
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | A00000278712 | -12.3215 | -91.4925 |  |  |  | 0.13298 | 3.8113
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | A00000278697 | -9.125 | -65.359 | -10.239 | -71.5779 | -1.065 | 68.239 | 2022/05/06 22:59:29
17 |  | A00000278697 | -9.011 | -63.647 | -10.1221 | -69.7279 | -1.042 | 67.829 | 2022/05/06 23:08:25
17 |  | A00000278697 | -9.03 | -63.11 | -10.1416 | -69.1476 | -0.975 | 68.078 | 2022/05/06 23:17:22
17 |  | A00000278697 | -8.981 | -62.717 | -10.0913 | -68.7229 | -0.944 | 68.293 | 2022/05/06 23:26:18
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | A00000278697 | -10.1485 | -69.794 |  |  |  | 0.062195 | 1.1646
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | A00000278679 | -9.53 | -68.95 | -10.6541 | -75.4583 | -0.695 | 68.163 | 2022/05/06 23:36:07
18 |  | A00000278679 | -9.58 | -69.742 | -10.7054 | -76.3142 | -0.538 | 67.454 | 2022/05/06 23:45:04
18 |  | A00000278679 | -9.565 | -69.84 | -10.69 | -76.4201 | -0.842 | 67.759 | 2022/05/06 23:53:59
18 |  | A00000278679 | -9.553 | -69.245 | -10.6777 | -75.7771 | -1.146 | 67.861 | 2022/05/07 00:02:56
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | A00000278679 | -10.6818 | -75.9924 |  |  |  | 0.021119 | 0.42002
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | A00000278680 | -9.195 | -66.475 | -10.3107 | -72.7838 | -1.208 | 68.425 | 2022/05/07 00:12:44
19 |  | A00000278680 | -9.156 | -66.057 | -10.2707 | -72.3321 | -1.186 | 68.244 | 2022/05/07 00:21:40
19 |  | A00000278680 | -9.111 | -65.788 | -10.2246 | -72.0414 | -0.773 | 67.96 | 2022/05/07 00:30:37
19 |  | A00000278680 | -9.112 | -65.67 | -10.2256 | -71.9139 | -1.107 | 68.045 | 2022/05/07 00:39:33
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | A00000278680 | -10.2579 | -72.2678 |  |  |  | 0.040237 | 0.35716
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | A00000280479 | -8.077 | -60.396 | -9.1646 | -66.2148 | -1.242 | 68.404 | 2022/05/07 00:49:21
20 |  | A00000280479 | -8.158 | -60.426 | -9.2477 | -66.2472 | -1.285 | 68.046 | 2022/05/07 00:58:17
20 |  | A00000280479 | -8.105 | -59.957 | -9.1933 | -65.7404 | -1.008 | 67.784 | 2022/05/07 01:07:13
20 |  | A00000280479 | -8.122 | -60.141 | -9.2108 | -65.9392 | -1.008 | 67.882 | 2022/05/07 01:16:08
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | A00000280479 | -9.2041 | -66.0354 |  |  |  | 0.033867 | 0.22243
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | A00000280480 | -8.508 | -62.634 | -9.6065 | -68.6332 | -0.922 | 67.65 | 2022/05/07 01:25:57
21 |  | A00000280480 | -8.588 | -63.492 | -9.6885 | -69.5604 | -1.236 | 67.534 | 2022/05/07 01:34:54
21 |  | A00000280480 | -8.519 | -62.893 | -9.6177 | -68.9131 | -0.881 | 67.362 | 2022/05/07 01:43:50
21 |  | A00000280480 | -8.547 | -62.948 | -9.6464 | -68.9725 | -0.89 | 68.093 | 2022/05/07 01:52:45
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | A00000280480 | -9.6398 | -69.0198 |  |  |  | 0.03567 | 0.36051
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -9.187 | -67.18 | -10.3025 | -73.5457 | -1.108 | 67.699 | 2022/05/07 02:02:31
22 |  | EV | -9.104 | -67.008 | -10.2174 | -73.3598 | -0.798 | 67.404 | 2022/05/07 02:11:27
22 |  | EV | -9.137 | -66.987 | -10.2513 | -73.3371 | -0.931 | 68.047 | 2022/05/07 02:20:23
22 |  | EV | -9.125 | -67.258 | -10.239 | -73.6299 | -1.176 | 67.525 | 2022/05/07 02:29:19
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.2526 | -73.4681 |  |  |  | 0.035246 | 0.13207
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | A00000280481 | -8.618 | -63.912 | -9.7192 | -70.0142 | -0.935 | 67.908 | 2022/05/07 02:39:05
23 |  | A00000280481 | -8.596 | -63.143 | -9.6967 | -69.1832 | -0.919 | 67.806 | 2022/05/07 02:48:01
23 |  | A00000280481 | -8.601 | -63.245 | -9.7018 | -69.2934 | -1.08 | 67.797 | 2022/05/07 02:56:58
23 |  | A00000280481 | -8.606 | -63.213 | -9.7069 | -69.2589 | -1.072 | 67.837 | 2022/05/07 03:05:55
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | A00000280481 | -9.7062 | -69.4374 |  |  |  | 0.0094296 | 0.35837
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | A00000280478 | -8.597 | -63.328 | -9.6977 | -69.3831 | -1.102 | 68.043 | 2022/05/07 03:15:42
24 |  | A00000280478 | -8.634 | -63.529 | -9.7356 | -69.6003 | -1.065 | 68.249 | 2022/05/07 03:24:37
24 |  | A00000280478 | -8.631 | -63.525 | -9.7326 | -69.596 | -1.372 | 68.321 | 2022/05/07 03:33:33
24 |  | A00000280478 | -8.591 | -63.563 | -9.6915 | -69.6371 | -1.214 | 67.76 | 2022/05/07 03:42:28
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | A00000280478 | -9.7144 | -69.5541 |  |  |  | 0.022396 | 0.10687
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | A00000280477 | -6.348 | -51.147 | -7.3922 | -56.2202 | -0.982 | 67.608 | 2022/05/07 03:52:15
25 |  | A00000280477 | -6.258 | -50.006 | -7.2999 | -54.9872 | -0.898 | 67.749 | 2022/05/07 04:01:11
25 |  | A00000280477 | -6.202 | -49.631 | -7.2425 | -54.582 | -1.009 | 67.841 | 2022/05/07 04:10:08
25 |  | A00000280477 | -6.23 | -49.602 | -7.2712 | -54.5507 | -0.92 | 67.809 | 2022/05/07 04:19:04
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | A00000280477 | -7.3014 | -55.085 |  |  |  | 0.063275 | 0.7241
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | A00000280492 | -6.518 | -50.459 | -7.5664 | -55.4767 | -0.984 | 68.288 | 2022/05/07 04:28:51
26 |  | A00000280492 | -6.52 | -50.143 | -7.5685 | -55.1353 | -0.873 | 67.666 | 2022/05/07 04:37:47
26 |  | A00000280492 | -6.561 | -50.522 | -7.6105 | -55.5448 | -0.807 | 67.762 | 2022/05/07 04:46:44
26 |  | A00000280492 | -6.548 | -50.539 | -7.5972 | -55.5632 | -1.049 | 67.843 | 2022/05/07 04:55:39
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | A00000280492 | -7.5857 | -55.43 |  |  |  | 0.021188 | 0.18506
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | A00000279569 | -6.558 | -50.273 | -7.6074 | -55.2757 | -0.916 | 67.268 | 2022/05/07 05:05:28
27 |  | A00000279569 | -6.518 | -50.246 | -7.5664 | -55.2466 | -0.85 | 67.223 | 2022/05/07 05:14:24
27 |  | A00000279569 | -6.461 | -50.131 | -7.508 | -55.1223 | -0.941 | 67.199 | 2022/05/07 05:23:20
27 |  | A00000279569 | -6.465 | -49.935 | -7.5121 | -54.9105 | -0.912 | 67.932 | 2022/05/07 05:32:15
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | A00000279569 | -7.5485 | -55.1388 |  |  |  | 0.046307 | 0.1537
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | EVIAN 1-2 - 3-31-22 | -8.94 | -64.819 | -10.0493 | -70.9943 | -1.019 | 68.694 | 2022/05/07 05:41:59
28 |  | EVIAN 1-2 - 3-31-22 | -8.99 | -65.756 | -10.1006 | -72.0069 | -1.094 | 67.967 | 2022/05/07 05:50:54
28 |  | EVIAN 1-2 - 3-31-22 | -9.051 | -66.462 | -10.1631 | -72.7698 | -1.024 | 67.482 | 2022/05/07 05:59:50
28 |  | EVIAN 1-2 - 3-31-22 | -9.016 | -66.339 | -10.1272 | -72.6369 | -1.251 | 67.571 | 2022/05/07 06:08:46
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | EVIAN 1-2 - 3-31-22 | -10.1101 | -72.102 |  |  |  | 0.046743 | 0.74951
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | FLUKE D-G 1491 | -0.117 | -6.387 | -1.0046 | -7.852 | -0.953 | 67.493 | 2022/05/07 06:18:31
29 |  | FLUKE D-G 1491 | 0.223 | -1.57 | -0.65601 | -2.6467 | -1.184 | 67.409 | 2022/05/07 06:27:26
29 |  | FLUKE D-G 1491 | 0.179 | -0.729 | -0.70111 | -1.7379 | -1.004 | 67.494 | 2022/05/07 06:36:22
29 |  | FLUKE D-G 1491 | 0.256 | 0.049 | -0.62218 | -0.89717 | -1.17 | 67.405 | 2022/05/07 06:45:17
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | FLUKE D-G 1491 | -0.74596 | -3.2834 |  |  |  | 0.1711 | 2.895
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | Beta2 -3-16-22 | 0.36 | 0.92 | -0.51556 | 0.044047 | -1.19 | 67.929 | 2022/05/07 06:55:03
30 |  | Beta2 -3-16-22 | 0.455 | 1.619 | -0.41817 | 0.7994 | -1.071 | 67.056 | 2022/05/07 07:03:58
30 |  | Beta2 -3-16-22 | 0.448 | 1.754 | -0.42535 | 0.94528 | -1.133 | 67.219 | 2022/05/07 07:12:54
30 |  | Beta2 -3-16-22 | 0.49 | 1.761 | -0.38229 | 0.95284 | -0.991 | 67.087 | 2022/05/07 07:21:50
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | Beta2 -3-16-22 | -0.43535 | 0.68539 |  |  |  | 0.055307 | 0.40103
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.651 | -59.091 | -9.7531 | -64.8046 | -1.034 | 67.787 | 2022/05/07 07:31:34
31 |  | EV | -8.966 | -63.889 | -10.076 | -69.9894 | -1.016 | 67.531 | 2022/05/07 07:40:30
31 |  | EV | -8.995 | -65.164 | -10.1057 | -71.3671 | -0.736 | 67.833 | 2022/05/07 07:49:25
31 |  | EV | -9.026 | -65.744 | -10.1375 | -71.9939 | -0.99 | 67.39 | 2022/05/07 07:58:22
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.0181 | -69.5387 |  |  |  | 0.17407 | 3.0217
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | FLUKE D-G 1490 | 0.327 | -4.323 | -0.54939 | -5.6216 | -1.362 | 68.723 | 2022/05/07 08:08:08
32 |  | FLUKE D-G 1490 | 0.635 | 0.38 | -0.23365 | -0.53948 | -0.97 | 67.563 | 2022/05/07 08:17:04
32 |  | FLUKE D-G 1490 | 0.679 | 1.391 | -0.18854 | 0.55302 | -1.056 | 67.783 | 2022/05/07 08:26:00
32 |  | FLUKE D-G 1490 | 0.737 | 2.119 | -0.12909 | 1.3397 | -1.148 | 67.725 | 2022/05/07 08:34:55
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | FLUKE D-G 1490 | -0.27517 | -1.0671 |  |  |  | 0.18316 | 2.8989
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | Beta1 -3-16-22 | 0.448 | 1.894 | -0.42535 | 1.0966 | -1.142 | 67.739 | 2022/05/07 08:44:42
33 |  | Beta1 -3-16-22 | 0.417 | 1.781 | -0.45713 | 0.97445 | -1.077 | 67.275 | 2022/05/07 08:53:38
33 |  | Beta1 -3-16-22 | 0.449 | 1.722 | -0.42433 | 0.9107 | -0.875 | 67.297 | 2022/05/07 09:02:33
33 |  | Beta1 -3-16-22 | 0.4 | 1.874 | -0.47456 | 1.075 | -0.99 | 67.319 | 2022/05/07 09:11:29
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | Beta1 -3-16-22 | -0.44534 | 1.0142 |  |  |  | 0.024118 | 0.080616
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | EVIAN 1-1 3-31-22 | -8.515 | -58.58 | -9.6136 | -64.2524 | -1.097 | 68.112 | 2022/05/07 09:21:15
34 |  | EVIAN 1-1 3-31-22 | -8.901 | -63.425 | -10.0093 | -69.488 | -0.763 | 67.734 | 2022/05/07 09:30:11
34 |  | EVIAN 1-1 3-31-22 | -8.917 | -64.575 | -10.0257 | -70.7307 | -0.931 | 67.855 | 2022/05/07 09:39:06
34 |  | EVIAN 1-1 3-31-22 | -8.985 | -65.288 | -10.0955 | -71.5011 | -0.967 | 67.695 | 2022/05/07 09:48:02
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | EVIAN 1-1 3-31-22 | -9.936 | -68.993 |  |  |  | 0.21281 | 3.0237
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | HARV2-C-4 | -2.853 | -54.364 | -3.8093 | -59.6965 | -1.498 | 64.944 | 2022/05/07 09:57:51
35 |  | HARV2-C-4 | -2.642 | -53.884 | -3.593 | -59.1778 | -1.44 | 64.98 | 2022/05/07 10:06:45
35 |  | HARV2-C-4 | -2.603 | -53.685 | -3.553 | -58.9628 | -1.24 | 65.669 | 2022/05/07 10:15:41
35 |  | HARV2-C-4 | -2.596 | -53.7 | -3.5459 | -58.979 | -1.333 | 66.966 | 2022/05/07 10:24:37
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | HARV2-C-4 | -3.6253 | -59.204 |  |  |  | 0.12137 | 0.31702
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | EV | -8.793 | -65.125 | -9.8986 | -71.325 | -0.983 | 67.388 | 2022/05/07 10:34:20
36 |  | EV | -9.015 | -66.288 | -10.1262 | -72.5818 | -0.93 | 67.137 | 2022/05/07 10:43:15
36 |  | EV | -9.029 | -66.248 | -10.1406 | -72.5385 | -1.124 | 67.243 | 2022/05/07 10:52:11
36 |  | EV | -9.066 | -66.39 | -10.1785 | -72.692 | -1.083 | 67.399 | 2022/05/07 11:01:07
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | EV | -10.086 | -72.2843 |  |  |  | 0.12372 | 0.59485

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-19703 | 2022/05/06 09:38:49 | 1-01 | 1 | -9.28 | -69.503 | 12774 | -1 | 1 | EV | standard | H2O | 1651851834.88 | 0.221 | 0.868 | 63.852 | 0 | -0.002 | 1.577 | 69.656 | 29.1 | 2.167 | 65.051 | 0.855 | -0.969 | 0 | 0 | 1 | 43.5 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-19703 | 2022/05/06 09:47:45 | 1-01 | 2 | -9.187 | -68.842 | 12873 | -1 | 1 | EV | standard | H2O | 1651852370.65 | 0.196 | 0.776 | 64.213 | 0.001 | -0.001 | 1.584 | 74.195 | 29.684 | 2.185 | 65.198 | 0.856 | -1.023 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-19703 | 2022/05/06 09:56:40 | 1-01 | 3 | -9.169 | -68.384 | 12854 | -1 | 1 | EV | standard | H2O | 1651852905.21 | 0.209 | 0.713 | 60.528 | 0.001 | -0.004 | 1.486 | 73.767 | 29.559 | 2.173 | 65.697 | 0.856 | -0.972 | 0 | 0 | 1 | 43 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-19703 | 2022/05/06 10:05:36 | 1-01 | 4 | -9.157 | -68.187 | 12868 | 0 | 1 | EV | standard | H2O | 1651853442 | 0.19 | 0.764 | 60.784 | 0.001 | -0.002 | 1.504 | 73.338 | 29.544 | 2.174 | 64.995 | 0.855 | -1.05 | 0 | 0 | 1 | 42.875 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-19703 | 2022/05/06 10:14:33 | 1-01 | 5 | -9.104 | -67.863 | 12915 | 0 | 1 | EV | standard | H2O | 1651853978.27 | 0.194 | 0.751 | 65.071 | 0.001 | -0.001 | 1.598 | 72.992 | 29.441 | 2.198 | 66.067 | 0.856 | -0.927 | 0 | 0 | 1 | 42.938 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-19703 | 2022/05/06 10:23:29 | 1-01 | 6 | -9.136 | -67.903 | 12882 | 0 | 1 | EV | standard | H2O | 1651854515.51 | 0.193 | 0.786 | 59.956 | 0.001 | -0.003 | 1.471 | 72.372 | 29.531 | 2.178 | 65.678 | 0.856 | -0.995 | 0 | 0 | 1 | 42.938 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-19703 | 2022/05/06 10:32:27 | 1-01 | 7 | -9.101 | -67.973 | 12877 | 0 | 1 | EV | standard | H2O | 1651855052.03 | 0.185 | 0.731 | 61.261 | 0.001 | 0.002 | 1.505 | 71.642 | 29.377 | 2.179 | 66.202 | 0.856 | -0.875 | 0 | 0 | 1 | 43.062 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-19703 | 2022/05/06 10:41:22 | 1-01 | 8 | -9.141 | -67.883 | 12899 | 0 | 1 | EV | standard | H2O | 1651855587.95 | 0.198 | 0.743 | 60.192 | 0.001 | 0 | 1.479 | 71.328 | 29.621 | 2.191 | 65.605 | 0.856 | -1.038 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-19703 | 2022/05/06 10:50:19 | 1-01 | 9 | -9.108 | -67.646 | 12902 | 0 | 1 | EV | standard | H2O | 1651856124.62 | 0.208 | 0.749 | 60.222 | 0.001 | -0.002 | 1.478 | 70.893 | 29.64 | 2.185 | 65.27 | 0.855 | -1.107 | 0 | 0 | 1 | 43.312 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-19703 | 2022/05/06 10:59:16 | 1-01 | 10 | -9.102 | -67.753 | 12908 | 0 | 1 | EV | standard | H2O | 1651856661.89 | 0.195 | 0.811 | 62.81 | 0.001 | -0.001 | 1.544 | 70.473 | 29.879 | 2.191 | 65.723 | 0.855 | -0.906 | 0 | 0 | 1 | 43.438 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-19704 | 2022/05/06 11:09:05 | 1-02 | 1 | 2.251 | 9.338 | 12811 | -1 | 1 | PZ | standard | H2O | 1651857250.54 | 0.209 | 0.823 | 61.98 | 0 | 0 | 1.523 | 68.605 | 29.847 | 2.279 | 66.591 | 0.856 | -1.123 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-19704 | 2022/05/06 11:18:02 | 1-02 | 2 | 2.643 | 15.707 | 12832 | -1 | 1 | PZ | standard | H2O | 1651857786.9 | 0.217 | 0.822 | 62.827 | 0 | -0.002 | 1.541 | 69.876 | 29.727 | 2.295 | 66.448 | 0.855 | -1.224 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-19704 | 2022/05/06 11:26:58 | 1-02 | 3 | 2.704 | 17.493 | 12893 | -1 | 1 | PZ | standard | H2O | 1651858323.76 | 0.206 | 0.882 | 63.068 | 0 | -0.004 | 1.551 | 69.314 | 29.867 | 2.317 | 68.577 | 0.855 | -0.9 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-19704 | 2022/05/06 11:35:56 | 1-02 | 4 | 2.775 | 18.145 | 12890 | 0 | 1 | PZ | standard | H2O | 1651858861.33 | 0.206 | 0.904 | 58.088 | 0 | -0.006 | 1.425 | 69.331 | 29.88 | 2.323 | 68.532 | 0.856 | -0.976 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-19704 | 2022/05/06 11:44:51 | 1-02 | 5 | 2.835 | 18.667 | 12888 | 0 | 1 | PZ | standard | H2O | 1651859397.01 | 0.199 | 0.855 | 58.285 | 0.001 | -0.003 | 1.428 | 69.196 | 30.086 | 2.315 | 67.496 | 0.856 | -1.026 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-19704 | 2022/05/06 11:53:47 | 1-02 | 6 | 2.821 | 18.973 | 12892 | 0 | 1 | PZ | standard | H2O | 1651859932.96 | 0.22 | 0.863 | 60.081 | 0.001 | -0.005 | 1.476 | 69.166 | 30.113 | 2.318 | 66.557 | 0.855 | -1.231 | 0 | 0 | 1 | 43.188 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-19704 | 2022/05/06 12:02:44 | 1-02 | 7 | 2.86 | 19.152 | 12863 | 0 | 1 | PZ | standard | H2O | 1651860469.62 | 0.204 | 0.758 | 56.348 | 0.001 | -0.004 | 1.39 | 69.008 | 29.835 | 2.313 | 68.145 | 0.856 | -1.159 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-19704 | 2022/05/06 12:11:39 | 1-02 | 8 | 2.863 | 19.669 | 12835 | 0 | 1 | PZ | standard | H2O | 1651861005.01 | 0.191 | 0.79 | 63.081 | 0.001 | -0.002 | 1.563 | 68.776 | 29.719 | 2.321 | 68.367 | 0.856 | -1.119 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-19704 | 2022/05/06 12:20:36 | 1-02 | 9 | 2.854 | 19.632 | 12823 | 0 | 1 | PZ | standard | H2O | 1651861540.88 | 0.196 | 0.762 | 60.931 | 0.001 | -0.004 | 1.504 | 68.737 | 30.09 | 2.302 | 66.538 | 0.856 | -1.277 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-19704 | 2022/05/06 12:29:33 | 1-02 | 10 | 2.85 | 19.514 | 12882 | 0 | 1 | PZ | standard | H2O | 1651862078.25 | 0.211 | 0.868 | 59.227 | 0.001 | -0.001 | 1.452 | 68.82 | 29.83 | 2.315 | 67.601 | 0.856 | -1.117 | 0 | 0 | 1 | 43.375 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-19705 | 2022/05/06 12:39:21 | 1-03 | 1 | -13.817 | -97.112 | 12801 | -1 | 1 | UT2 | standard | H2O | 1651862667.1 | 0.197 | 0.914 | 59.794 | 0.001 | -0.007 | 1.468 | 68.507 | 29.587 | 2.122 | 63.502 | 0.856 | -1.096 | 0 | 0 | 1 | 43.375 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-19705 | 2022/05/06 12:48:18 | 1-03 | 2 | -14.452 | -106.547 | 12822 | -1 | 1 | UT2 | standard | H2O | 1651863203.26 | 0.212 | 0.681 | 58.26 | 0.001 | 0.004 | 1.438 | 68.526 | 29.864 | 2.117 | 63.536 | 0.856 | -1.004 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-19705 | 2022/05/06 12:57:14 | 1-03 | 3 | -14.558 | -108.938 | 12921 | -1 | 1 | UT2 | standard | H2O | 1651863739.43 | 0.193 | 0.704 | 59.069 | 0.001 | 0.003 | 1.463 | 68.765 | 30.016 | 2.126 | 63.455 | 0.855 | -1.07 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-19705 | 2022/05/06 13:06:10 | 1-03 | 4 | -14.646 | -110.062 | 12913 | 0 | 1 | UT2 | standard | H2O | 1651864275.8 | 0.209 | 0.835 | 61.152 | 0.001 | 0.007 | 1.496 | 68.718 | 29.607 | 2.117 | 64.918 | 0.856 | -0.824 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-19705 | 2022/05/06 13:15:07 | 1-03 | 5 | -14.684 | -110.899 | 12915 | 0 | 1 | UT2 | standard | H2O | 1651864813.38 | 0.2 | 0.657 | 60.505 | 0.001 | 0.001 | 1.478 | 68.699 | 29.841 | 2.116 | 64.39 | 0.856 | -0.948 | 0 | 0 | 1 | 43.062 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-19705 | 2022/05/06 13:24:04 | 1-03 | 6 | -14.723 | -111.157 | 12890 | 0 | 1 | UT2 | standard | H2O | 1651865349.66 | 0.203 | 0.737 | 61.059 | 0.001 | 0.002 | 1.501 | 68.752 | 29.71 | 2.105 | 63.153 | 0.856 | -1.007 | 0 | 0 | 1 | 43.125 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-19705 | 2022/05/06 13:33:00 | 1-03 | 7 | -14.755 | -111.676 | 12861 | 0 | 1 | UT2 | standard | H2O | 1651865885.24 | 0.186 | 0.709 | 59.037 | 0.001 | 0.001 | 1.446 | 68.653 | 29.696 | 2.098 | 63.274 | 0.855 | -1.025 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-19705 | 2022/05/06 13:41:56 | 1-03 | 8 | -14.729 | -111.8 | 12840 | 0 | 1 | UT2 | standard | H2O | 1651866421.54 | 0.211 | 0.709 | 63.765 | 0.002 | 0.004 | 1.578 | 68.618 | 29.414 | 2.112 | 63.526 | 0.856 | -1.108 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-19705 | 2022/05/06 13:50:53 | 1-03 | 9 | -14.783 | -112.1 | 12844 | 0 | 1 | UT2 | standard | H2O | 1651866958.09 | 0.209 | 0.829 | 60.937 | 0.001 | 0.002 | 1.497 | 68.589 | 29.317 | 2.102 | 63.678 | 0.856 | -0.951 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-19705 | 2022/05/06 13:59:49 | 1-03 | 10 | -14.74 | -112.144 | 12797 | 0 | 1 | UT2 | standard | H2O | 1651867494.67 | 0.201 | 0.758 | 57.06 | 0.002 | 0.001 | 1.398 | 68.34 | 29.34 | 2.107 | 64.516 | 0.856 | -0.841 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-19706 | 2022/05/06 14:09:39 | 1-04 | 1 | -9.303 | -72.016 | 12694 | -1 | 1 | EV | standard | H2O | 1651868084.75 | 0.161 | 0.744 | 64.191 | 0 | 0.006 | 1.589 | 68.028 | 29.13 | 2.152 | 65.745 | 0.856 | -0.869 | 0 | 0 | 1 | 42.938 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-19706 | 2022/05/06 14:18:36 | 1-04 | 2 | -9.222 | -68.931 | 12586 | -1 | 1 | EV | standard | H2O | 1651868621.9 | 0.19 | 0.765 | 59.896 | 0.001 | 0.002 | 1.477 | 67.869 | 28.551 | 2.134 | 65.698 | 0.856 | -0.828 | 0 | 0 | 1 | 42.438 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-19706 | 2022/05/06 14:27:32 | 1-04 | 3 | -9.21 | -68.525 | 12657 | -1 | 1 | EV | standard | H2O | 1651869157.92 | 0.205 | 0.77 | 61.68 | 0.001 | -0.002 | 1.513 | 67.965 | 28.894 | 2.146 | 65.036 | 0.856 | -0.914 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-19706 | 2022/05/06 14:36:27 | 1-04 | 4 | -9.116 | -67.717 | 12556 | 0 | 1 | EV | standard | H2O | 1651869693.32 | 0.201 | 0.779 | 59.388 | 0.001 | -0.001 | 1.467 | 67.95 | 28.673 | 2.127 | 63.46 | 0.855 | -1.018 | 0 | 0 | 1 | 42.625 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-19706 | 2022/05/06 14:45:24 | 1-04 | 5 | -9.152 | -67.778 | 12624 | 0 | 1 | EV | standard | H2O | 1651870229.44 | 0.199 | 0.775 | 61.345 | 0.001 | -0.002 | 1.515 | 67.863 | 28.983 | 2.138 | 64.912 | 0.855 | -0.828 | 0 | 0 | 1 | 42.625 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-19706 | 2022/05/06 14:54:20 | 1-04 | 6 | -9.107 | -67.675 | 12597 | 0 | 1 | EV | standard | H2O | 1651870765.22 | 0.2 | 0.8 | 63.109 | 0.001 | 0 | 1.559 | 67.843 | 28.842 | 2.146 | 64.962 | 0.856 | -0.819 | 0 | 0 | 1 | 42.75 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-19706 | 2022/05/06 15:03:16 | 1-04 | 7 | -9.147 | -67.77 | 12569 | 0 | 1 | EV | standard | H2O | 1651871301.11 | 0.172 | 0.782 | 58.629 | 0.001 | 0 | 1.444 | 67.879 | 28.709 | 2.133 | 64.429 | 0.856 | -0.903 | 0 | 0 | 1 | 42.812 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-19706 | 2022/05/06 15:12:12 | 1-04 | 8 | -9.111 | -67.564 | 12540 | 0 | 1 | EV | standard | H2O | 1651871837.78 | 0.197 | 0.786 | 58.005 | 0.001 | 0.003 | 1.434 | 67.877 | 28.56 | 2.119 | 64.552 | 0.856 | -0.883 | 0 | 0 | 1 | 42.875 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-19706 | 2022/05/06 15:21:09 | 1-04 | 9 | -9.12 | -67.45 | 12595 | 0 | 1 | EV | standard | H2O | 1651872374.25 | 0.209 | 0.753 | 59.826 | 0.001 | 0.002 | 1.474 | 67.949 | 29.235 | 2.127 | 63.361 | 0.856 | -1.061 | 0 | 0 | 1 | 42.938 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-19706 | 2022/05/06 15:30:06 | 1-04 | 10 | -9.163 | -67.247 | 12569 | 0 | 1 | EV | standard | H2O | 1651872910.94 | 0.193 | 0.759 | 56.804 | 0.001 | -0.001 | 1.404 | 67.822 | 29.248 | 2.136 | 63.002 | 0.856 | -1.085 | 0 | 0 | 1 | 43.062 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-19707 | 2022/05/06 15:39:56 | 1-05 | 1 | -8.222 | -64.153 | 12673 | -1 | 1 | A00000209752 | 22-084 | H2O | 1651873500.97 | 0.204 | 0.784 | 61.316 | 0.001 | 0 | 1.511 | 68.113 | 29.081 | 2.137 | 63.895 | 0.856 | -1.097 | 0 | 0 | 1 | 43.125 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-19707 | 2022/05/06 15:48:53 | 1-05 | 2 | -8.161 | -63.977 | 12573 | -1 | 1 | A00000209752 | 22-084 | H2O | 1651874038.7 | 0.197 | 0.766 | 56.869 | 0.001 | 0 | 1.407 | 67.983 | 28.822 | 2.135 | 63.635 | 0.855 | -1.114 | 0 | 0 | 1 | 43.125 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-19707 | 2022/05/06 15:57:49 | 1-05 | 3 | -8.183 | -64.136 | 12595 | -1 | 1 | A00000209752 | 22-084 | H2O | 1651874575.38 | 0.204 | 0.813 | 64.456 | 0.001 | 0.002 | 1.596 | 67.922 | 28.98 | 2.149 | 64.252 | 0.856 | -0.964 | 0 | 0 | 1 | 43.125 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-19707 | 2022/05/06 16:06:46 | 1-05 | 4 | -8.12 | -63.757 | 12595 | 0 | 1 | A00000209752 | 22-084 | H2O | 1651875112.16 | 0.203 | 0.792 | 62.601 | 0.001 | 0.001 | 1.535 | 67.913 | 29.339 | 2.136 | 63.324 | 0.855 | -1.2 | 0 | 0 | 1 | 43.188 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-19708 | 2022/05/06 16:16:37 | 1-06 | 1 | -0.797 | -6.386 | 12539 | -1 | 1 | A00000209761 | 22-084 | H2O | 1651875702.06 | 0.213 | 0.893 | 58.424 | 0 | 0.003 | 1.443 | 67.924 | 28.81 | 2.219 | 65.474 | 0.855 | -0.966 | 0 | 0 | 1 | 43.25 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-19708 | 2022/05/06 16:25:33 | 1-06 | 2 | -0.49 | -1.44 | 12614 | -1 | 1 | A00000209761 | 22-084 | H2O | 1651876238.71 | 0.193 | 0.856 | 63.813 | 0 | -0.005 | 1.58 | 68.056 | 28.968 | 2.243 | 66.361 | 0.856 | -0.991 | 0 | 0 | 1 | 43.312 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-19708 | 2022/05/06 16:34:29 | 1-06 | 3 | -0.453 | -0.375 | 12586 | -1 | 1 | A00000209761 | 22-084 | H2O | 1651876774.48 | 0.209 | 0.853 | 60.006 | 0 | -0.002 | 1.484 | 68.012 | 29.025 | 2.251 | 65.753 | 0.856 | -1.118 | 0 | 0 | 1 | 43.25 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-19708 | 2022/05/06 16:43:26 | 1-06 | 4 | -0.437 | -0.045 | 12559 | 0 | 1 | A00000209761 | 22-084 | H2O | 1651877311.76 | 0.205 | 0.895 | 61.35 | 0.001 | -0.003 | 1.502 | 67.946 | 29.005 | 2.241 | 65.329 | 0.855 | -1.051 | 0 | 0 | 1 | 43.25 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-19709 | 2022/05/06 16:53:16 | 1-07 | 1 | -1.233 | -6.008 | 12644 | -1 | 1 | A00000209762 | 22-084 | H2O | 1651877901.74 | 0.201 | 0.861 | 59.227 | 0 | -0.003 | 1.46 | 68.008 | 29.262 | 2.24 | 65.838 | 0.856 | -1.062 | 0 | 0 | 1 | 43.312 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-19709 | 2022/05/06 17:02:12 | 1-07 | 2 | -1.378 | -6.616 | 12669 | -1 | 1 | A00000209762 | 22-084 | H2O | 1651878437.96 | 0.206 | 0.788 | 63.581 | 0.001 | -0.002 | 1.573 | 68.101 | 29.199 | 2.25 | 65.342 | 0.856 | -1.214 | 0 | 0 | 1 | 43.375 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-19709 | 2022/05/06 17:11:08 | 1-07 | 3 | -1.307 | -6.454 | 12397 | -1 | 1 | A00000209762 | 22-084 | H2O | 1651878973.77 | 0.204 | 0.775 | 58.089 | 0 | -0.002 | 1.434 | 67.482 | 28.533 | 2.203 | 64.843 | 0.856 | -1.036 | 0 | 0 | 1 | 43.312 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-19709 | 2022/05/06 17:20:05 | 1-07 | 4 | -1.344 | -6.18 | 12413 | 0 | 1 | A00000209762 | 22-084 | H2O | 1651879510.76 | 0.217 | 0.787 | 58.132 | 0.001 | -0.001 | 1.431 | 67.411 | 28.621 | 2.197 | 64.368 | 0.856 | -1.098 | 0 | 0 | 1 | 43.312 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-19710 | 2022/05/06 17:29:53 | 1-08 | 1 | -4.058 | -28.361 | 12511 | -1 | 1 | A00000319094 | 22-084 | H2O | 1651880098.61 | 0.207 | 0.869 | 63.919 | 0.001 | 0.001 | 1.579 | 67.673 | 28.741 | 2.181 | 65.041 | 0.856 | -0.987 | 0 | 0 | 1 | 43.375 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-19710 | 2022/05/06 17:38:49 | 1-08 | 2 | -4.184 | -30.226 | 12312 | -1 | 1 | A00000319094 | 22-084 | H2O | 1651880634.95 | 0.192 | 0.866 | 56.079 | 0.001 | 0.001 | 1.392 | 67.292 | 28.036 | 2.154 | 64.402 | 0.856 | -0.94 | 0 | 0 | 1 | 43.438 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-19710 | 2022/05/06 17:47:45 | 1-08 | 3 | -4.156 | -30.293 | 12421 | -1 | 1 | A00000319094 | 22-084 | H2O | 1651881171.38 | 0.205 | 0.756 | 59.701 | 0.001 | 0 | 1.475 | 67.465 | 28.495 | 2.165 | 64.397 | 0.855 | -1.05 | 0 | 0 | 1 | 43.438 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-19710 | 2022/05/06 17:56:42 | 1-08 | 4 | -4.23 | -30.812 | 12389 | 0 | 1 | A00000319094 | 22-084 | H2O | 1651881707.68 | 0.191 | 0.757 | 59.691 | 0.001 | 0.004 | 1.475 | 67.445 | 28.737 | 2.168 | 63.129 | 0.856 | -1.214 | 0 | 0 | 1 | 43.438 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-19711 | 2022/05/06 18:06:30 | 1-09 | 1 | -4.483 | -32.608 | 12389 | -1 | 1 | A00000319080 | 22-084 | H2O | 1651882296 | 0.198 | 0.802 | 59.735 | 0.001 | 0.001 | 1.485 | 67.389 | 28.525 | 2.144 | 63.882 | 0.856 | -1.047 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-19711 | 2022/05/06 18:15:25 | 1-09 | 2 | -4.416 | -32.326 | 12550 | -1 | 1 | A00000319080 | 22-084 | H2O | 1651882831.33 | 0.188 | 0.856 | 63.576 | 0.001 | -0.002 | 1.574 | 67.606 | 29.021 | 2.19 | 65.142 | 0.856 | -0.914 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-19711 | 2022/05/06 18:24:22 | 1-09 | 3 | -4.474 | -32.684 | 12604 | -1 | 1 | A00000319080 | 22-084 | H2O | 1651883367.06 | 0.195 | 0.848 | 57.739 | 0.001 | -0.002 | 1.422 | 67.832 | 29.139 | 2.194 | 64.844 | 0.855 | -1.076 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-19711 | 2022/05/06 18:33:19 | 1-09 | 4 | -4.44 | -32.447 | 12628 | 0 | 1 | A00000319080 | 22-084 | H2O | 1651883904.15 | 0.226 | 0.736 | 63.789 | 0.001 | 0 | 1.57 | 68 | 28.987 | 2.184 | 64.512 | 0.856 | -1.146 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-19712 | 2022/05/06 18:43:08 | 1-10 | 1 | -4.654 | -33.081 | 12643 | -1 | 1 | A00000319071 | 22-084 | H2O | 1651884492.93 | 0.229 | 0.865 | 63.536 | 0.001 | -0.003 | 1.57 | 67.97 | 29.139 | 2.193 | 65.2 | 0.856 | -1.039 | 0 | 0 | 1 | 43.5 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-19712 | 2022/05/06 18:52:03 | 1-10 | 2 | -4.729 | -33.081 | 12763 | -1 | 1 | A00000319071 | 22-084 | H2O | 1651885028.77 | 0.189 | 0.776 | 59.524 | 0 | -0.002 | 1.472 | 68.082 | 29.483 | 2.223 | 66.125 | 0.856 | -0.876 | 0 | 0 | 1 | 43.5 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-19712 | 2022/05/06 19:01:00 | 1-10 | 3 | -4.705 | -33.187 | 12643 | -1 | 1 | A00000319071 | 22-084 | H2O | 1651885565.33 | 0.196 | 0.822 | 57.515 | 0.001 | 0.003 | 1.419 | 67.981 | 29.291 | 2.191 | 64.556 | 0.855 | -1.153 | 0 | 0 | 1 | 43.5 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-19712 | 2022/05/06 19:09:56 | 1-10 | 4 | -4.692 | -33.117 | 12455 | 0 | 1 | A00000319071 | 22-084 | H2O | 1651886102.05 | 0.208 | 0.827 | 59.515 | 0.001 | 0 | 1.465 | 67.626 | 28.572 | 2.169 | 64.339 | 0.856 | -1.05 | 0 | 0 | 1 | 43.5 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-19713 | 2022/05/06 19:19:45 | 1-11 | 1 | -12.914 | -94.713 | 12751 | -1 | 1 | A00000307888 | 22-084 | H2O | 1651886690.67 | 0.181 | 0.752 | 52.119 | 0 | -0.004 | 1.283 | 68.276 | 29.168 | 2.125 | 63.912 | 0.856 | -1.035 | 0 | 0 | 1 | 43.5 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-19713 | 2022/05/06 19:28:42 | 1-11 | 2 | -13.251 | -99.801 | 12631 | -1 | 1 | A00000307888 | 22-084 | H2O | 1651887227.26 | 0.216 | 0.784 | 59.48 | 0.001 | 0 | 1.469 | 68.11 | 29.103 | 2.097 | 62.895 | 0.856 | -1.212 | 0 | 0 | 1 | 43.438 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-19713 | 2022/05/06 19:37:37 | 1-11 | 3 | -13.302 | -101.26 | 12781 | -1 | 1 | A00000307888 | 22-084 | H2O | 1651887763.09 | 0.206 | 0.827 | 59.911 | 0.001 | 0.004 | 1.48 | 68.35 | 29.148 | 2.126 | 64.559 | 0.856 | -0.935 | 0 | 0 | 1 | 43.438 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-19713 | 2022/05/06 19:46:33 | 1-11 | 4 | -13.306 | -101.546 | 12460 | 0 | 1 | A00000307888 | 22-084 | H2O | 1651888298.91 | 0.213 | 0.714 | 57.701 | 0.001 | 0.001 | 1.426 | 67.632 | 28.558 | 2.059 | 62.482 | 0.856 | -0.947 | 0 | 0 | 1 | 43.438 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-19714 | 2022/05/06 19:56:24 | 1-12 | 1 | -19.354 | -155.421 | 12849 | -1 | 1 | A00000307886 | 22-084 | H2O | 1651888889.28 | 0.211 | 0.694 | 57.727 | 0.001 | 0.003 | 1.426 | 68.517 | 29.58 | 2.047 | 62.774 | 0.856 | -1.011 | 0 | 0 | 1 | 43.5 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-19714 | 2022/05/06 20:05:20 | 1-12 | 2 | -19.532 | -159.526 | 12527 | -1 | 1 | A00000307886 | 22-084 | H2O | 1651889426.06 | 0.207 | 0.724 | 57.332 | 0.001 | 0.003 | 1.415 | 67.918 | 28.68 | 1.99 | 60.926 | 0.856 | -0.982 | 0 | 0 | 1 | 43.625 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-19714 | 2022/05/06 20:14:15 | 1-12 | 3 | -19.592 | -161 | 12619 | -1 | 1 | A00000307886 | 22-084 | H2O | 1651889960.92 | 0.185 | 0.748 | 60.212 | 0.001 | 0.005 | 1.478 | 68.144 | 28.539 | 2.012 | 62.306 | 0.856 | -0.852 | 0 | 0 | 1 | 43.75 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-19714 | 2022/05/06 20:23:12 | 1-12 | 4 | -19.649 | -161.779 | 12491 | 0 | 1 | A00000307886 | 22-084 | H2O | 1651890498.31 | 0.195 | 0.742 | 60.199 | 0.001 | 0.002 | 1.491 | 67.812 | 28.351 | 1.976 | 61.265 | 0.856 | -0.862 | 0 | 0 | 1 | 43.75 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-19715 | 2022/05/06 20:33:01 | 1-13 | 1 | -9.608 | -77.179 | 12749 | -1 | 1 | EV | standard | H2O | 1651891087.01 | 0.197 | 0.858 | 62.285 | 0 | 0.01 | 1.542 | 68.073 | 28.901 | 2.152 | 66.414 | 0.855 | -0.65 | 0 | 0 | 1 | 43.375 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-19715 | 2022/05/06 20:41:58 | 1-13 | 2 | -9.306 | -70.983 | 12622 | -1 | 1 | EV | standard | H2O | 1651891624.1 | 0.189 | 0.837 | 62.811 | 0.001 | -0.001 | 1.55 | 67.933 | 28.642 | 2.137 | 65.25 | 0.856 | -0.79 | 0 | 0 | 1 | 43.125 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-19715 | 2022/05/06 20:50:55 | 1-13 | 3 | -9.169 | -68.951 | 12527 | -1 | 1 | EV | standard | H2O | 1651892161.24 | 0.202 | 0.729 | 57.424 | 0.001 | -0.001 | 1.412 | 67.685 | 28.5 | 2.126 | 65.201 | 0.855 | -0.818 | 0 | 0 | 1 | 43.125 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-19715 | 2022/05/06 20:59:52 | 1-13 | 4 | -9.134 | -68.461 | 12584 | 0 | 1 | EV | standard | H2O | 1651892698.3 | 0.218 | 0.744 | 58.842 | 0.001 | 0.001 | 1.454 | 67.85 | 29.081 | 2.116 | 63.11 | 0.856 | -0.967 | 0 | 0 | 1 | 43.188 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-19716 | 2022/05/06 21:09:43 | 1-14 | 1 | -16.663 | -129.461 | 12441 | -1 | 1 | A00000307890 | 22-084 | H2O | 1651893288.47 | 0.211 | 0.718 | 60.471 | 0.001 | -0.002 | 1.498 | 67.771 | 28.035 | 2.02 | 61.775 | 0.856 | -0.915 | 0 | 0 | 1 | 43 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-19716 | 2022/05/06 21:18:38 | 1-14 | 2 | -16.953 | -134.313 | 12652 | -1 | 1 | A00000307890 | 22-084 | H2O | 1651893823.88 | 0.196 | 0.712 | 60.416 | 0.001 | 0.005 | 1.488 | 68.147 | 28.667 | 2.065 | 63.488 | 0.855 | -0.796 | 0 | 0 | 1 | 42.938 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-19716 | 2022/05/06 21:27:35 | 1-14 | 3 | -16.982 | -135.525 | 12620 | -1 | 1 | A00000307890 | 22-084 | H2O | 1651894360.55 | 0.187 | 0.759 | 57.706 | 0.002 | 0.004 | 1.424 | 68.15 | 28.821 | 2.047 | 62.304 | 0.856 | -1.023 | 0 | 0 | 1 | 43.062 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-19716 | 2022/05/06 21:36:31 | 1-14 | 4 | -17.012 | -136.054 | 12630 | 0 | 1 | A00000307890 | 22-084 | H2O | 1651894896.37 | 0.194 | 0.755 | 62.735 | 0.001 | 0.004 | 1.538 | 68.093 | 28.98 | 2.045 | 62.23 | 0.856 | -1.028 | 0 | 0 | 1 | 43 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-19717 | 2022/05/06 21:46:18 | 1-15 | 1 | -18.877 | -166.247 | 12526 | -1 | 1 | A00000307918 | 22-084 | H2O | 1651895483.21 | 0.205 | 0.755 | 60.237 | 0.002 | 0.002 | 1.498 | 68.012 | 28.839 | 1.982 | 59.659 | 0.855 | -1.281 | 0 | 0 | 1 | 43 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-19717 | 2022/05/06 21:55:14 | 1-15 | 2 | -18.89 | -168.692 | 12435 | -1 | 1 | A00000307918 | 22-084 | H2O | 1651896019.29 | 0.206 | 0.853 | 59.518 | 0.002 | 0.006 | 1.472 | 67.86 | 28.159 | 1.955 | 60.266 | 0.855 | -1.054 | 0 | 0 | 1 | 42.938 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-19717 | 2022/05/06 22:04:09 | 1-15 | 3 | -18.891 | -169.177 | 12374 | -1 | 1 | A00000307918 | 22-084 | H2O | 1651896554.22 | 0.19 | 0.823 | 61.371 | 0.001 | 0.006 | 1.517 | 67.62 | 28.3 | 1.959 | 59.099 | 0.855 | -0.987 | 0 | 0 | 1 | 43 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-19717 | 2022/05/06 22:13:05 | 1-15 | 4 | -18.904 | -169.787 | 12399 | 0 | 1 | A00000307918 | 22-084 | H2O | 1651897090.32 | 0.209 | 0.862 | 54.944 | 0.001 | 0.006 | 1.354 | 67.627 | 28.102 | 1.97 | 60.915 | 0.856 | -0.907 | 0 | 0 | 1 | 43 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-19718 | 2022/05/06 22:22:52 | 1-16 | 1 | -11.338 | -89.277 | 12778 | -1 | 1 | A00000278712 | 22-084 | H2O | 1651897676.96 | 0.183 | 0.801 | 64.126 | 0.001 | 0.007 | 1.583 | 68.277 | 29.472 | 2.147 | 64.292 | 0.856 | -0.946 | 0 | 0 | 1 | 43.062 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-19718 | 2022/05/06 22:31:49 | 1-16 | 2 | -11.136 | -83.297 | 12545 | -1 | 1 | A00000278712 | 22-084 | H2O | 1651898214.45 | 0.19 | 0.708 | 62.696 | 0.001 | -0.003 | 1.552 | 67.821 | 28.684 | 2.101 | 63.96 | 0.855 | -1.013 | 0 | 0 | 1 | 43.062 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-19718 | 2022/05/06 22:40:46 | 1-16 | 3 | -11.134 | -81.888 | 12544 | -1 | 1 | A00000278712 | 22-084 | H2O | 1651898751.24 | 0.2 | 0.776 | 62.174 | 0.001 | 0 | 1.531 | 67.897 | 28.909 | 2.107 | 62.518 | 0.855 | -1.127 | 0 | 0 | 1 | 43.125 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-19718 | 2022/05/06 22:49:41 | 1-16 | 4 | -11.018 | -80.69 | 12560 | 0 | 1 | A00000278712 | 22-084 | H2O | 1651899286.9 | 0.204 | 0.78 | 58.755 | 0.001 | -0.001 | 1.453 | 67.821 | 29.057 | 2.108 | 63.976 | 0.856 | -1 | 0 | 0 | 1 | 43.125 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-19719 | 2022/05/06 22:59:29 | 1-17 | 1 | -9.125 | -65.359 | 12690 | -1 | 1 | A00000278697 | 22-084 | H2O | 1651899875.2 | 0.194 | 0.89 | 63.386 | 0.001 | -0.003 | 1.555 | 68.239 | 29.021 | 2.171 | 64.849 | 0.855 | -1.065 | 0 | 0 | 1 | 43.125 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-19719 | 2022/05/06 23:08:25 | 1-17 | 2 | -9.011 | -63.647 | 12491 | -1 | 1 | A00000278697 | 22-084 | H2O | 1651900410.23 | 0.193 | 0.842 | 59.218 | 0.001 | -0.001 | 1.463 | 67.829 | 28.435 | 2.135 | 64.423 | 0.856 | -1.042 | 0 | 0 | 1 | 43.25 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-19719 | 2022/05/06 23:17:22 | 1-17 | 3 | -9.03 | -63.11 | 12610 | -1 | 1 | A00000278697 | 22-084 | H2O | 1651900947.35 | 0.212 | 0.652 | 58.996 | 0.002 | -0.001 | 1.468 | 68.078 | 28.744 | 2.162 | 64.824 | 0.856 | -0.975 | 0 | 0 | 1 | 43.375 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-19719 | 2022/05/06 23:26:18 | 1-17 | 4 | -8.981 | -62.717 | 12692 | 0 | 1 | A00000278697 | 22-084 | H2O | 1651901483.51 | 0.195 | 0.846 | 63.357 | 0.001 | 0.003 | 1.551 | 68.293 | 28.778 | 2.169 | 65.302 | 0.856 | -0.944 | 0 | 0 | 1 | 43.5 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-19720 | 2022/05/06 23:36:07 | 1-18 | 1 | -9.53 | -68.95 | 12845 | -1 | 1 | A00000278679 | 22-084 | H2O | 1651902072.31 | 0.202 | 0.747 | 60.446 | 0.001 | -0.001 | 1.476 | 68.163 | 29.364 | 2.175 | 66.983 | 0.856 | -0.695 | 0 | 0 | 1 | 43.438 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-19720 | 2022/05/06 23:45:04 | 1-18 | 2 | -9.58 | -69.742 | 12489 | -1 | 1 | A00000278679 | 22-084 | H2O | 1651902609.27 | 0.199 | 0.827 | 57.302 | 0.001 | -0.002 | 1.415 | 67.454 | 28.269 | 2.122 | 65.392 | 0.856 | -0.538 | 0 | 0 | 1 | 43 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-19720 | 2022/05/06 23:53:59 | 1-18 | 3 | -9.565 | -69.84 | 12582 | -1 | 1 | A00000278679 | 22-084 | H2O | 1651903145.06 | 0.205 | 0.905 | 60.768 | 0 | -0.002 | 1.496 | 67.759 | 28.562 | 2.127 | 64.946 | 0.855 | -0.842 | 0 | 0 | 1 | 42.812 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-19720 | 2022/05/07 00:02:56 | 1-18 | 4 | -9.553 | -69.245 | 12567 | 0 | 1 | A00000278679 | 22-084 | H2O | 1651903681.96 | 0.176 | 0.809 | 62.587 | 0 | 0.003 | 1.547 | 67.861 | 28.979 | 2.12 | 63.233 | 0.856 | -1.146 | 0 | 0 | 1 | 42.938 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-19721 | 2022/05/07 00:12:44 | 1-19 | 1 | -9.195 | -66.475 | 12762 | -1 | 1 | A00000278680 | 22-084 | H2O | 1651904269.89 | 0.217 | 0.823 | 60.146 | 0.001 | -0.002 | 1.481 | 68.425 | 29.374 | 2.162 | 64.134 | 0.856 | -1.208 | 0 | 0 | 1 | 43.062 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-19721 | 2022/05/07 00:21:40 | 1-19 | 2 | -9.156 | -66.057 | 12609 | -1 | 1 | A00000278680 | 22-084 | H2O | 1651904805.82 | 0.202 | 0.873 | 63.661 | 0.001 | 0.004 | 1.575 | 68.244 | 28.931 | 2.14 | 62.887 | 0.856 | -1.186 | 0 | 0 | 1 | 43.125 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-19721 | 2022/05/07 00:30:37 | 1-19 | 3 | -9.111 | -65.788 | 12592 | -1 | 1 | A00000278680 | 22-084 | H2O | 1651905342.7 | 0.21 | 0.856 | 61.14 | 0.001 | -0.001 | 1.508 | 67.96 | 28.536 | 2.16 | 65.459 | 0.856 | -0.773 | 0 | 0 | 1 | 43.188 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-19721 | 2022/05/07 00:39:33 | 1-19 | 4 | -9.112 | -65.67 | 12587 | 0 | 1 | A00000278680 | 22-084 | H2O | 1651905878.62 | 0.214 | 0.642 | 61.776 | 0.002 | 0 | 1.527 | 68.045 | 28.797 | 2.154 | 63.999 | 0.856 | -1.107 | 0 | 0 | 1 | 43.062 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-19722 | 2022/05/07 00:49:21 | 1-20 | 1 | -8.077 | -60.396 | 12819 | -1 | 1 | A00000280479 | 22-084 | H2O | 1651906466.32 | 0.207 | 0.812 | 63.162 | 0 | 0.001 | 1.541 | 68.404 | 29.886 | 2.169 | 63.758 | 0.856 | -1.242 | 0 | 0 | 1 | 43.125 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-19722 | 2022/05/07 00:58:17 | 1-20 | 2 | -8.158 | -60.426 | 12627 | -1 | 1 | A00000280479 | 22-084 | H2O | 1651907002.53 | 0.177 | 0.759 | 58.698 | 0.001 | -0.002 | 1.446 | 68.046 | 29.151 | 2.146 | 62.891 | 0.856 | -1.285 | 0 | 0 | 1 | 43.25 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-19722 | 2022/05/07 01:07:13 | 1-20 | 3 | -8.105 | -59.957 | 12603 | -1 | 1 | A00000280479 | 22-084 | H2O | 1651907538.39 | 0.213 | 0.822 | 65.756 | 0.001 | 0.001 | 1.626 | 67.784 | 29.164 | 2.146 | 63.465 | 0.856 | -1.008 | 0 | 0 | 1 | 43.125 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-19722 | 2022/05/07 01:16:08 | 1-20 | 4 | -8.122 | -60.141 | 12627 | 0 | 1 | A00000280479 | 22-084 | H2O | 1651908073.95 | 0.2 | 0.691 | 57.967 | 0.001 | -0.001 | 1.421 | 67.882 | 29.137 | 2.14 | 64.455 | 0.856 | -1.008 | 0 | 0 | 1 | 43.188 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-19723 | 2022/05/07 01:25:57 | 1-21 | 1 | -8.508 | -62.634 | 12534 | -1 | 1 | A00000280480 | 22-084 | H2O | 1651908662.86 | 0.214 | 0.822 | 56.462 | 0.001 | -0.001 | 1.396 | 67.65 | 28.487 | 2.141 | 64.688 | 0.855 | -0.922 | 0 | 0 | 1 | 43.25 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-19723 | 2022/05/07 01:34:54 | 1-21 | 2 | -8.588 | -63.492 | 12369 | -1 | 1 | A00000280480 | 22-084 | H2O | 1651909199.53 | 0.203 | 0.757 | 63.34 | 0.001 | 0.004 | 1.565 | 67.534 | 28.379 | 2.096 | 62.332 | 0.856 | -1.236 | 0 | 0 | 1 | 43.312 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-19723 | 2022/05/07 01:43:50 | 1-21 | 3 | -8.519 | -62.893 | 12401 | -1 | 1 | A00000280480 | 22-084 | H2O | 1651909735.78 | 0.194 | 0.788 | 55.021 | 0 | 0 | 1.352 | 67.362 | 28.413 | 2.111 | 63.928 | 0.855 | -0.881 | 0 | 0 | 1 | 43.375 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-19723 | 2022/05/07 01:52:45 | 1-21 | 4 | -8.547 | -62.948 | 12738 | 0 | 1 | A00000280480 | 22-084 | H2O | 1651910270.92 | 0.203 | 0.706 | 62.075 | 0.001 | 0.001 | 1.534 | 68.093 | 29.284 | 2.168 | 65.701 | 0.856 | -0.89 | 0 | 0 | 1 | 43.188 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-19724 | 2022/05/07 02:02:31 | 1-22 | 1 | -9.187 | -67.18 | 12511 | -1 | 1 | EV | standard | H2O | 1651910857.05 | 0.203 | 0.869 | 58.536 | 0.001 | -0.002 | 1.447 | 67.699 | 28.887 | 2.113 | 62.903 | 0.855 | -1.108 | 0 | 0 | 1 | 43.312 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-19724 | 2022/05/07 02:11:27 | 1-22 | 2 | -9.104 | -67.008 | 12402 | -1 | 1 | EV | standard | H2O | 1651911392.39 | 0.214 | 0.873 | 60.269 | 0.001 | -0.001 | 1.501 | 67.404 | 28.058 | 2.108 | 64.457 | 0.855 | -0.798 | 0 | 0 | 1 | 43.312 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-19724 | 2022/05/07 02:20:23 | 1-22 | 3 | -9.137 | -66.987 | 12697 | -1 | 1 | EV | standard | H2O | 1651911929.04 | 0.187 | 0.806 | 58.793 | 0.001 | -0.003 | 1.447 | 68.047 | 29.085 | 2.15 | 64.759 | 0.856 | -0.931 | 0 | 0 | 1 | 43.125 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-19724 | 2022/05/07 02:29:19 | 1-22 | 4 | -9.125 | -67.258 | 12407 | 0 | 1 | EV | standard | H2O | 1651912464.56 | 0.205 | 0.777 | 54.718 | 0.001 | 0.001 | 1.356 | 67.525 | 28.63 | 2.101 | 61.795 | 0.856 | -1.176 | 0 | 0 | 1 | 43.188 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-19725 | 2022/05/07 02:39:05 | 1-23 | 1 | -8.618 | -63.912 | 12614 | -1 | 1 | A00000280481 | 22-084 | H2O | 1651913050.79 | 0.198 | 0.772 | 60.118 | 0.001 | 0.001 | 1.487 | 67.908 | 28.761 | 2.148 | 64.71 | 0.856 | -0.935 | 0 | 0 | 1 | 43.188 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-19725 | 2022/05/07 02:48:01 | 1-23 | 2 | -8.596 | -63.143 | 12588 | -1 | 1 | A00000280481 | 22-084 | H2O | 1651913587.49 | 0.191 | 0.7 | 60.119 | 0.001 | 0 | 1.486 | 67.806 | 28.643 | 2.148 | 64.923 | 0.856 | -0.919 | 0 | 0 | 1 | 43.188 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-19725 | 2022/05/07 02:56:58 | 1-23 | 3 | -8.601 | -63.245 | 12633 | -1 | 1 | A00000280481 | 22-084 | H2O | 1651914124.33 | 0.199 | 0.784 | 57.551 | 0.001 | 0 | 1.422 | 67.797 | 29.203 | 2.155 | 64.396 | 0.856 | -1.08 | 0 | 0 | 1 | 43.25 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-19725 | 2022/05/07 03:05:55 | 1-23 | 4 | -8.606 | -63.213 | 12583 | 0 | 1 | A00000280481 | 22-084 | H2O | 1651914660.98 | 0.225 | 0.859 | 59.353 | 0.001 | 0 | 1.466 | 67.837 | 28.816 | 2.136 | 63.528 | 0.856 | -1.072 | 0 | 0 | 1 | 43.25 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-19726 | 2022/05/07 03:15:42 | 1-24 | 1 | -8.597 | -63.328 | 12585 | -1 | 1 | A00000280478 | 22-084 | H2O | 1651915247.21 | 0.195 | 0.751 | 61.388 | 0.001 | -0.001 | 1.513 | 68.043 | 28.632 | 2.164 | 64.022 | 0.855 | -1.102 | 0 | 0 | 1 | 43.25 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-19726 | 2022/05/07 03:24:37 | 1-24 | 2 | -8.634 | -63.529 | 12674 | -1 | 1 | A00000280478 | 22-084 | H2O | 1651915782.91 | 0.192 | 0.798 | 61.8 | 0.001 | 0.005 | 1.514 | 68.249 | 28.65 | 2.17 | 65.107 | 0.856 | -1.065 | 0 | 0 | 1 | 43.125 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-19726 | 2022/05/07 03:33:33 | 1-24 | 3 | -8.631 | -63.525 | 12613 | -1 | 1 | A00000280478 | 22-084 | H2O | 1651916318.58 | 0.17 | 0.773 | 58.607 | 0 | 0.002 | 1.456 | 68.321 | 28.662 | 2.155 | 63.487 | 0.856 | -1.372 | 0 | 0 | 1 | 43.188 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-19726 | 2022/05/07 03:42:28 | 1-24 | 4 | -8.591 | -63.563 | 12409 | 0 | 1 | A00000280478 | 22-084 | H2O | 1651916853.65 | 0.201 | 0.79 | 56.836 | 0 | 0.001 | 1.41 | 67.76 | 28.312 | 2.122 | 63.106 | 0.855 | -1.214 | 0 | 0 | 1 | 43.375 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-19727 | 2022/05/07 03:52:15 | 1-25 | 1 | -6.348 | -51.147 | 12485 | -1 | 1 | A00000280477 | 22-084 | H2O | 1651917441.15 | 0.202 | 0.889 | 64.672 | 0.002 | 0 | 1.6 | 67.608 | 28.61 | 2.149 | 64.077 | 0.856 | -0.982 | 0 | 0 | 1 | 43.312 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-19727 | 2022/05/07 04:01:11 | 1-25 | 2 | -6.258 | -50.006 | 12582 | -1 | 1 | A00000280477 | 22-084 | H2O | 1651917977.28 | 0.209 | 0.717 | 60.544 | 0.001 | 0 | 1.493 | 67.749 | 28.978 | 2.162 | 64.918 | 0.856 | -0.898 | 0 | 0 | 1 | 43.25 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-19727 | 2022/05/07 04:10:08 | 1-25 | 3 | -6.202 | -49.631 | 12596 | -1 | 1 | A00000280477 | 22-084 | H2O | 1651918513.84 | 0.187 | 0.682 | 59.482 | 0.001 | -0.001 | 1.471 | 67.841 | 28.867 | 2.158 | 64.533 | 0.855 | -1.009 | 0 | 0 | 1 | 43.312 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-19727 | 2022/05/07 04:19:04 | 1-25 | 4 | -6.23 | -49.602 | 12553 | 0 | 1 | A00000280477 | 22-084 | H2O | 1651919049.69 | 0.187 | 0.791 | 60.346 | 0.001 | -0.002 | 1.489 | 67.809 | 28.457 | 2.151 | 65.204 | 0.856 | -0.92 | 0 | 0 | 1 | 43.188 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-19728 | 2022/05/07 04:28:51 | 1-26 | 1 | -6.518 | -50.459 | 12879 | -1 | 1 | A00000280492 | 22-084 | H2O | 1651919636.56 | 0.218 | 0.889 | 58.984 | 0.001 | -0.001 | 1.45 | 68.288 | 30.189 | 2.216 | 64.971 | 0.856 | -0.984 | 0 | 0 | 1 | 43.375 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-19728 | 2022/05/07 04:37:47 | 1-26 | 2 | -6.52 | -50.143 | 12566 | -1 | 1 | A00000280492 | 22-084 | H2O | 1651920173.47 | 0.187 | 0.653 | 63.422 | 0.001 | 0.001 | 1.563 | 67.666 | 28.851 | 2.146 | 65.202 | 0.856 | -0.873 | 0 | 0 | 1 | 43.312 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-19728 | 2022/05/07 04:46:44 | 1-26 | 3 | -6.561 | -50.522 | 12602 | -1 | 1 | A00000280492 | 22-084 | H2O | 1651920710.05 | 0.196 | 0.745 | 60.657 | 0.001 | -0.002 | 1.498 | 67.762 | 28.864 | 2.173 | 65.33 | 0.855 | -0.807 | 0 | 0 | 1 | 43.25 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-19728 | 2022/05/07 04:55:39 | 1-26 | 4 | -6.548 | -50.539 | 12594 | 0 | 1 | A00000280492 | 22-084 | H2O | 1651921245.71 | 0.206 | 0.841 | 59.566 | 0 | -0.002 | 1.482 | 67.843 | 28.712 | 2.161 | 65.044 | 0.856 | -1.049 | 0 | 0 | 1 | 43.312 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-19729 | 2022/05/07 05:05:28 | 1-27 | 1 | -6.558 | -50.273 | 12371 | -1 | 1 | A00000279569 | 22-084 | H2O | 1651921833.41 | 0.209 | 0.851 | 57.912 | 0.002 | 0 | 1.431 | 67.268 | 28.365 | 2.126 | 63.371 | 0.856 | -0.916 | 0 | 0 | 1 | 43.25 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-19729 | 2022/05/07 05:14:24 | 1-27 | 2 | -6.518 | -50.246 | 12367 | -1 | 1 | A00000279569 | 22-084 | H2O | 1651922370 | 0.217 | 0.756 | 57.981 | 0.001 | 0.001 | 1.434 | 67.223 | 28.407 | 2.133 | 64.128 | 0.855 | -0.85 | 0 | 0 | 1 | 43.188 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-19729 | 2022/05/07 05:23:20 | 1-27 | 3 | -6.461 | -50.131 | 12277 | -1 | 1 | A00000279569 | 22-084 | H2O | 1651922905.19 | 0.208 | 0.969 | 59.101 | 0.001 | -0.002 | 1.461 | 67.199 | 27.788 | 2.116 | 64.254 | 0.856 | -0.941 | 0 | 0 | 1 | 43.25 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-19729 | 2022/05/07 05:32:15 | 1-27 | 4 | -6.465 | -49.935 | 12693 | 0 | 1 | A00000279569 | 22-084 | H2O | 1651923440.69 | 0.193 | 0.819 | 60.078 | 0.001 | 0 | 1.486 | 67.932 | 29.104 | 2.182 | 65.891 | 0.856 | -0.912 | 0 | 0 | 1 | 43.312 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-19730 | 2022/05/07 05:41:59 | 1-28 | 1 | -8.94 | -64.819 | 13028 | -1 | 1 | EVIAN 1-2 - 3-31-22 | 22-095 | H2O | 1651924024.31 | 0.197 | 0.809 | 62.023 | 0.001 | -0.001 | 1.521 | 68.694 | 30.063 | 2.201 | 66.347 | 0.856 | -1.019 | 0 | 0 | 1 | 43.312 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-19730 | 2022/05/07 05:50:54 | 1-28 | 2 | -8.99 | -65.756 | 12629 | -1 | 1 | EVIAN 1-2 - 3-31-22 | 22-095 | H2O | 1651924560.42 | 0.198 | 0.75 | 57.398 | 0.001 | -0.001 | 1.423 | 67.967 | 29.087 | 2.142 | 63.767 | 0.855 | -1.094 | 0 | 0 | 1 | 43.312 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-19730 | 2022/05/07 05:59:50 | 1-28 | 3 | -9.051 | -66.462 | 12443 | -1 | 1 | EVIAN 1-2 - 3-31-22 | 22-095 | H2O | 1651925096.24 | 0.183 | 0.84 | 56.164 | 0.001 | -0.004 | 1.39 | 67.482 | 28.695 | 2.114 | 62.994 | 0.855 | -1.024 | 0 | 0 | 1 | 43.25 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-19730 | 2022/05/07 06:08:46 | 1-28 | 4 | -9.016 | -66.339 | 12443 | 0 | 1 | EVIAN 1-2 - 3-31-22 | 22-095 | H2O | 1651925631.94 | 0.213 | 0.804 | 57.812 | 0.001 | 0.001 | 1.432 | 67.571 | 29.008 | 2.104 | 61.568 | 0.855 | -1.251 | 0 | 0 | 1 | 43.062 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-19731 | 2022/05/07 06:18:31 | 1-29 | 1 | -0.117 | -6.387 | 12487 | -1 | 1 | FLUKE D-G 1491 | 22-095 | H2O | 1651926216.72 | 0.205 | 0.895 | 59.592 | 0.001 | 0.004 | 1.473 | 67.493 | 28.57 | 2.227 | 66.043 | 0.856 | -0.953 | 0 | 0 | 1 | 43.062 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-19731 | 2022/05/07 06:27:26 | 1-29 | 2 | 0.223 | -1.57 | 12375 | -1 | 1 | FLUKE D-G 1491 | 22-095 | H2O | 1651926752.06 | 0.195 | 0.89 | 58.287 | 0 | -0.004 | 1.45 | 67.409 | 28.556 | 2.194 | 64.151 | 0.855 | -1.184 | 0 | 0 | 1 | 42.938 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-19731 | 2022/05/07 06:36:22 | 1-29 | 3 | 0.179 | -0.729 | 12455 | -1 | 1 | FLUKE D-G 1491 | 22-095 | H2O | 1651927287.45 | 0.205 | 0.748 | 60.512 | 0.001 | -0.004 | 1.508 | 67.494 | 28.6 | 2.222 | 65.52 | 0.856 | -1.004 | 0 | 0 | 1 | 42.938 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-19731 | 2022/05/07 06:45:17 | 1-29 | 4 | 0.256 | 0.049 | 12390 | 0 | 1 | FLUKE D-G 1491 | 22-095 | H2O | 1651927823.42 | 0.2 | 0.876 | 58.503 | 0.001 | -0.005 | 1.444 | 67.405 | 28.572 | 2.203 | 64.659 | 0.855 | -1.17 | 0 | 0 | 1 | 43.062 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-19732 | 2022/05/07 06:55:03 | 1-30 | 1 | 0.36 | 0.92 | 12638 | -1 | 1 | Beta2 -3-16-22 | 22-095 | H2O | 1651928408.44 | 0.207 | 0.751 | 63.096 | 0.001 | -0.001 | 1.571 | 67.929 | 29.28 | 2.248 | 65.184 | 0.856 | -1.19 | 0 | 0 | 1 | 43 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-19732 | 2022/05/07 07:03:58 | 1-30 | 2 | 0.455 | 1.619 | 12195 | -1 | 1 | Beta2 -3-16-22 | 22-095 | H2O | 1651928944.18 | 0.202 | 0.818 | 63.123 | 0.001 | -0.002 | 1.551 | 67.056 | 27.909 | 2.176 | 63.698 | 0.855 | -1.071 | 0 | 0 | 1 | 42.938 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-19732 | 2022/05/07 07:12:54 | 1-30 | 3 | 0.448 | 1.754 | 12309 | -1 | 1 | Beta2 -3-16-22 | 22-095 | H2O | 1651929480.16 | 0.208 | 0.88 | 54.278 | 0 | -0.002 | 1.346 | 67.219 | 28.218 | 2.201 | 64.873 | 0.856 | -1.133 | 0 | 0 | 1 | 43.125 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-19732 | 2022/05/07 07:21:50 | 1-30 | 4 | 0.49 | 1.761 | 12297 | 0 | 1 | Beta2 -3-16-22 | 22-095 | H2O | 1651930015.26 | 0.184 | 0.811 | 56.731 | 0 | -0.005 | 1.407 | 67.087 | 28.473 | 2.202 | 64.68 | 0.856 | -0.991 | 0 | 0 | 1 | 43.125 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-19733 | 2022/05/07 07:31:34 | 1-31 | 1 | -8.651 | -59.091 | 12592 | -1 | 1 | EV | standard | H2O | 1651930600.2 | 0.205 | 0.789 | 57.836 | 0.001 | -0.004 | 1.427 | 67.787 | 28.996 | 2.152 | 64.489 | 0.855 | -1.034 | 0 | 0 | 1 | 43 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-19733 | 2022/05/07 07:40:30 | 1-31 | 2 | -8.966 | -63.889 | 12461 | -1 | 1 | EV | standard | H2O | 1651931135.8 | 0.208 | 0.81 | 61.156 | 0.001 | 0.003 | 1.513 | 67.531 | 28.535 | 2.134 | 63.549 | 0.855 | -1.016 | 0 | 0 | 1 | 43.062 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-19733 | 2022/05/07 07:49:25 | 1-31 | 3 | -8.995 | -65.164 | 12654 | -1 | 1 | EV | standard | H2O | 1651931670.67 | 0.185 | 0.745 | 57.965 | 0.001 | 0.002 | 1.429 | 67.833 | 28.717 | 2.162 | 65.698 | 0.855 | -0.736 | 0 | 0 | 1 | 43.062 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-19733 | 2022/05/07 07:58:22 | 1-31 | 4 | -9.026 | -65.744 | 12402 | 0 | 1 | EV | standard | H2O | 1651932207.21 | 0.187 | 0.755 | 61.004 | 0.001 | -0.001 | 1.506 | 67.39 | 28.233 | 2.109 | 63.816 | 0.855 | -0.99 | 0 | 0 | 1 | 43 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-19734 | 2022/05/07 08:08:08 | 1-32 | 1 | 0.327 | -4.323 | 12966 | -1 | 1 | FLUKE D-G 1490 | 22-095 | H2O | 1651932793.3 | 0.21 | 0.837 | 60.374 | 0.001 | 0.002 | 1.484 | 68.723 | 30.144 | 2.297 | 66.139 | 0.856 | -1.362 | 0 | 0 | 1 | 42.938 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-19734 | 2022/05/07 08:17:04 | 1-32 | 2 | 0.635 | 0.38 | 12573 | -1 | 1 | FLUKE D-G 1490 | 22-095 | H2O | 1651933329.83 | 0.195 | 0.789 | 56.658 | 0 | -0.001 | 1.411 | 67.563 | 29.06 | 2.246 | 66.784 | 0.855 | -0.97 | 0 | 0 | 1 | 43.188 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-19734 | 2022/05/07 08:26:00 | 1-32 | 3 | 0.679 | 1.391 | 12550 | -1 | 1 | FLUKE D-G 1490 | 22-095 | H2O | 1651933866.09 | 0.2 | 0.816 | 63.009 | 0.002 | 0.001 | 1.56 | 67.783 | 28.581 | 2.238 | 66.005 | 0.856 | -1.056 | 0 | 0 | 1 | 43.062 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-19734 | 2022/05/07 08:34:55 | 1-32 | 4 | 0.737 | 2.119 | 12527 | 0 | 1 | FLUKE D-G 1490 | 22-095 | H2O | 1651934400.19 | 0.18 | 0.867 | 58.382 | 0 | -0.001 | 1.438 | 67.725 | 28.772 | 2.229 | 65.481 | 0.855 | -1.148 | 0 | 0 | 1 | 42.938 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-19735 | 2022/05/07 08:44:42 | 1-33 | 1 | 0.448 | 1.894 | 12545 | -1 | 1 | Beta1 -3-16-22 | 22-095 | H2O | 1651934987.73 | 0.205 | 0.796 | 58.405 | 0.001 | 0 | 1.432 | 67.739 | 28.874 | 2.244 | 65.563 | 0.855 | -1.142 | 0 | 0 | 1 | 42.875 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-19735 | 2022/05/07 08:53:38 | 1-33 | 2 | 0.417 | 1.781 | 12364 | -1 | 1 | Beta1 -3-16-22 | 22-095 | H2O | 1651935523.28 | 0.193 | 0.756 | 59.29 | 0 | -0.003 | 1.475 | 67.275 | 28.394 | 2.215 | 65.222 | 0.855 | -1.077 | 0 | 0 | 1 | 42.875 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-19735 | 2022/05/07 09:02:33 | 1-33 | 3 | 0.449 | 1.722 | 12395 | -1 | 1 | Beta1 -3-16-22 | 22-095 | H2O | 1651936059.51 | 0.187 | 0.822 | 59.835 | 0.001 | -0.001 | 1.476 | 67.297 | 28.269 | 2.211 | 66.234 | 0.856 | -0.875 | 0 | 0 | 1 | 42.875 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-19735 | 2022/05/07 09:11:29 | 1-33 | 4 | 0.4 | 1.874 | 12392 | 0 | 1 | Beta1 -3-16-22 | 22-095 | H2O | 1651936594.47 | 0.182 | 0.789 | 58.126 | 0.001 | -0.004 | 1.435 | 67.319 | 28.275 | 2.223 | 65.35 | 0.856 | -0.99 | 0 | 0 | 1 | 42.938 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-19736 | 2022/05/07 09:21:15 | 1-34 | 1 | -8.515 | -58.58 | 12689 | -1 | 1 | EVIAN 1-1 3-31-22 | 22-095 | H2O | 1651937181.03 | 0.223 | 0.819 | 58.944 | 0.001 | -0.007 | 1.451 | 68.112 | 29.186 | 2.168 | 63.922 | 0.856 | -1.097 | 0 | 0 | 1 | 43.188 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-19736 | 2022/05/07 09:30:11 | 1-34 | 2 | -8.901 | -63.425 | 12657 | -1 | 1 | EVIAN 1-1 3-31-22 | 22-095 | H2O | 1651937716.02 | 0.229 | 0.783 | 58.003 | 0.001 | 0.002 | 1.426 | 67.734 | 29.111 | 2.16 | 65.589 | 0.856 | -0.763 | 0 | 0 | 1 | 43.312 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-19736 | 2022/05/07 09:39:06 | 1-34 | 3 | -8.917 | -64.575 | 12628 | -1 | 1 | EVIAN 1-1 3-31-22 | 22-095 | H2O | 1651938252.1 | 0.205 | 0.727 | 55.149 | 0.001 | 0 | 1.361 | 67.855 | 28.704 | 2.158 | 65.33 | 0.855 | -0.931 | 0 | 0 | 1 | 42.938 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-19736 | 2022/05/07 09:48:02 | 1-34 | 4 | -8.985 | -65.288 | 12543 | 0 | 1 | EVIAN 1-1 3-31-22 | 22-095 | H2O | 1651938788.16 | 0.195 | 0.719 | 61.588 | 0.001 | 0 | 1.512 | 67.695 | 28.489 | 2.134 | 64.754 | 0.856 | -0.967 | 0 | 0 | 1 | 43 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-19737 | 2022/05/07 09:57:51 | 1-35 | 1 | -2.853 | -54.364 | 10038 | -1 | 1 | HARV2-C-4 | 21-069 | H2O | 1651939377.74 | 0.206 | 0.922 | 50.977 | 0.001 | 0.001 | 1.316 | 64.944 | 20.781 | 1.817 | 53.687 | 0.856 | -1.498 | 0 | 0 | 1 | 43.125 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-19737 | 2022/05/07 10:06:45 | 1-35 | 2 | -2.642 | -53.884 | 10057 | -1 | 1 | HARV2-C-4 | 21-069 | H2O | 1651939912.4 | 0.215 | 0.952 | 51.717 | 0.001 | 0.003 | 1.32 | 64.98 | 20.928 | 1.823 | 53.963 | 0.855 | -1.44 | 0 | 0 | 1 | 43.125 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-19737 | 2022/05/07 10:15:41 | 1-35 | 3 | -2.603 | -53.685 | 10408 | -1 | 1 | HARV2-C-4 | 21-069 | H2O | 1651940447.45 | 0.205 | 0.882 | 53.506 | 0.001 | 0.002 | 1.357 | 65.669 | 21.285 | 1.881 | 56.881 | 0.856 | -1.24 | 0 | 0 | 1 | 43.125 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-19737 | 2022/05/07 10:24:37 | 1-35 | 4 | -2.596 | -53.7 | 11006 | 0 | 1 | HARV2-C-4 | 21-069 | H2O | 1651940983.62 | 0.219 | 0.839 | 50.901 | 0.001 | -0.003 | 1.283 | 66.966 | 22.94 | 1.995 | 59.79 | 0.855 | -1.333 | 0 | 0 | 1 | 43.062 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-19738 | 2022/05/07 10:34:20 | 1-36 | 1 | -8.793 | -65.125 | 12406 | -1 | 1 | EV | standard | H2O | 1651941566 | 0.186 | 0.747 | 59.296 | 0.001 | 0 | 1.463 | 67.388 | 28.535 | 2.111 | 62.83 | 0.856 | -0.983 | 0 | 0 | 1 | 43 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-19738 | 2022/05/07 10:43:15 | 1-36 | 2 | -9.015 | -66.288 | 12307 | -1 | 1 | EV | standard | H2O | 1651942101.27 | 0.191 | 0.767 | 55.759 | 0.002 | 0.002 | 1.374 | 67.137 | 28.075 | 2.1 | 63.221 | 0.856 | -0.93 | 0 | 0 | 1 | 43.125 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-19738 | 2022/05/07 10:52:11 | 1-36 | 3 | -9.029 | -66.248 | 12310 | -1 | 1 | EV | standard | H2O | 1651942636.57 | 0.211 | 0.933 | 59.548 | 0.001 | 0.001 | 1.481 | 67.243 | 28.452 | 2.088 | 61.751 | 0.855 | -1.124 | 0 | 0 | 1 | 43.188 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-19738 | 2022/05/07 11:01:07 | 1-36 | 4 | -9.066 | -66.39 | 12326 | 0 | 1 | EV | standard | H2O | 1651943172.91 | 0.21 | 0.843 | 57.956 | 0.001 | -0.001 | 1.432 | 67.399 | 28.007 | 2.097 | 62.603 | 0.856 | -1.083 | 0 | 0 | 1 | 43.375 | 1 | 36 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 68.8498
varray: [68.605, 69.876, 69.314, 69.331, 69.196, 69.166, 69.008, 68.776, 68.737, 68.82, 68.507, 68.526, 68.765, 68.718, 68.699, 68.752, 68.653, 68.618, 68.589, 68.34]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_shift
baseline_shift_STD | 0.364444147768
varray: [68.605, 69.876, 69.314, 69.331, 69.196, 69.166, 69.008, 68.776, 68.737, 68.82, 68.507, 68.526, 68.765, 68.718, 68.699, 68.752, 68.653, 68.618, 68.589, 68.34]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_shift
ch4_ppm_MEAN | -1.0513
varray: [-1.123, -1.224, -0.9, -0.976, -1.026, -1.231, -1.159, -1.119, -1.277, -1.117, -1.096, -1.004, -1.07, -0.824, -0.948, -1.007, -1.025, -1.108, -0.951, -0.841]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.124599611387
varray: [-1.123, -1.224, -0.9, -0.976, -1.026, -1.231, -1.159, -1.119, -1.277, -1.117, -1.096, -1.004, -1.07, -0.824, -0.948, -1.007, -1.025, -1.108, -0.951, -0.841]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: ch4_ppm
residuals_MEAN | 2.211
varray: [2.279, 2.295, 2.317, 2.323, 2.315, 2.318, 2.313, 2.321, 2.302, 2.315, 2.122, 2.117, 2.126, 2.117, 2.116, 2.105, 2.098, 2.112, 2.102, 2.107]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: residuals
residuals_STD | 0.101998968003
varray: [2.279, 2.295, 2.317, 2.323, 2.315, 2.318, 2.313, 2.321, 2.302, 2.315, 2.122, 2.117, 2.126, 2.117, 2.116, 2.105, 2.098, 2.112, 2.102, 2.107]
inst: std_standards
operation: std_standards(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: residuals
slope_shift_MEAN | 29.7693
varray: [29.847, 29.727, 29.867, 29.88, 30.086, 30.113, 29.835, 29.719, 30.09, 29.83, 29.587, 29.864, 30.016, 29.607, 29.841, 29.71, 29.696, 29.414, 29.317, 29.34]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: slope_shift
slope_shift_STD | 0.231696284867
varray: [29.847, 29.727, 29.867, 29.88, 30.086, 30.113, 29.835, 29.719, 30.09, 29.83, 29.587, 29.864, 30.016, 29.607, 29.841, 29.71, 29.696, 29.414, 29.317, 29.34]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: slope_shift
baseline_curvature_MEAN | 65.64
varray: [66.591, 66.448, 68.577, 68.532, 67.496, 66.557, 68.145, 68.367, 66.538, 67.601, 63.502, 63.536, 63.455, 64.918, 64.39, 63.153, 63.274, 63.526, 63.678, 64.516]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_curvature
baseline_curvature_STD | 2.03155128697
varray: [66.591, 66.448, 68.577, 68.532, 67.496, 66.557, 68.145, 68.367, 66.538, 67.601, 63.502, 63.536, 63.455, 64.918, 64.39, 63.153, 63.274, 63.526, 63.678, 64.516]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.211
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 68.8498
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 1.0513
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 29.7693
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.152998452005
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.363998452
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.203997936006
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.41499793601
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.305996904009
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.51699690401
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.2211
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.373798834162
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.42509883416
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.182222073884
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 69.0320220739
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 69.2142441478
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 1.0933324433
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 69.9431324433
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 6.55999465982
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 75.4097946598
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 1.15848142433
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 30.9277814243
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 6.0946538609
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 71.7346538609
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -8.56201282387
actual_array: [ 1.93 -15.84]
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102, 2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85, -13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74, -9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163, -8.222, -8.161, -8.183, -8.12, -0.797, -0.49, -0.453, -0.437, -1.233, -1.378, -1.307, -1.344, -4.058, -4.184, -4.156, -4.23, -4.483, -4.416, -4.474, -4.44, -4.654, -4.729, -4.705, -4.692, -12.914, -13.251, -13.302, -13.306, -19.354, -19.532, -19.592, -19.649, -9.608, -9.306, -9.169, -9.134, -16.663, -16.953, -16.982, -17.012, -18.877, -18.89, -18.891, -18.904, -11.338, -11.136, -11.134, -11.018, -9.125, -9.011, -9.03, -8.981, -9.53, -9.58, -9.565, -9.553, -9.195, -9.156, -9.111, -9.112, -8.077, -8.158, -8.105, -8.122, -8.508, -8.588, -8.519, -8.547, -9.187, -9.104, -9.137, -9.125, -8.618, -8.596, -8.601, -8.606, -8.597, -8.634, -8.631, -8.591, -6.348, -6.258, -6.202, -6.23, -6.518, -6.52, -6.561, -6.548, -6.558, -6.518, -6.461, -6.465, -8.94, -8.99, -9.051, -9.016, -0.117, 0.223, 0.179, 0.256, 0.36, 0.455, 0.448, 0.49, -8.651, -8.966, -8.995, -9.026, 0.327, 0.635, 0.679, 0.737, 0.448, 0.417, 0.449, 0.4, -8.515, -8.901, -8.917, -8.985, -2.853, -2.642, -2.603, -2.596, -8.793, -9.015, -9.029, -9.066]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -7.48916071429
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0, 5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0, 16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0, 21668.0, 22205.0, 22741.0, 23278.0, 23869.0, 24405.0, 24941.0, 25478.0, 26068.0, 26604.0, 27140.0, 27677.0, 28265.0, 28801.0, 29337.0, 29874.0, 30462.0, 30997.0, 31534.0, 32071.0, 32660.0, 33195.0, 33732.0, 34268.0, 34857.0, 35394.0, 35929.0, 36465.0, 37056.0, 37592.0, 38127.0, 38664.0, 39253.0, 39790.0, 40327.0, 40864.0, 41455.0, 41990.0, 42527.0, 43063.0, 43650.0, 44186.0, 44721.0, 45257.0, 45844.0, 46381.0, 46918.0, 47453.0, 48041.0, 48577.0, 49114.0, 49650.0, 50239.0, 50776.0, 51311.0, 51848.0, 52436.0, 52972.0, 53509.0, 54045.0, 54633.0, 55169.0, 55705.0, 56240.0, 56829.0, 57366.0, 57902.0, 58437.0, 59023.0, 59559.0, 60095.0, 60631.0, 61217.0, 61753.0, 62290.0, 62827.0, 63414.0, 63949.0, 64485.0, 65020.0, 65607.0, 66143.0, 66680.0, 67216.0, 67803.0, 68339.0, 68876.0, 69411.0, 70000.0, 70536.0, 71072.0, 71607.0, 72191.0, 72726.0, 73262.0, 73798.0, 74383.0, 74918.0, 75454.0, 75989.0, 76575.0, 77110.0, 77646.0, 78182.0, 78766.0, 79302.0, 79837.0, 80374.0, 80960.0, 81496.0, 82032.0, 82567.0, 83154.0, 83690.0, 84225.0, 84761.0, 85347.0, 85883.0, 86418.0, 86954.0, 87543.0, 88077.0, 88613.0, 89149.0, 89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -63.5362679349
actual_array: [ 18.1 -119. ]
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753, 9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514, -97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144, -72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247, -64.153, -63.977, -64.136, -63.757, -6.386, -1.44, -0.375, -0.045, -6.008, -6.616, -6.454, -6.18, -28.361, -30.226, -30.293, -30.812, -32.608, -32.326, -32.684, -32.447, -33.081, -33.081, -33.187, -33.117, -94.713, -99.801, -101.26, -101.546, -155.421, -159.526, -161.0, -161.779, -77.179, -70.983, -68.951, -68.461, -129.461, -134.313, -135.525, -136.054, -166.247, -168.692, -169.177, -169.787, -89.277, -83.297, -81.888, -80.69, -65.359, -63.647, -63.11, -62.717, -68.95, -69.742, -69.84, -69.245, -66.475, -66.057, -65.788, -65.67, -60.396, -60.426, -59.957, -60.141, -62.634, -63.492, -62.893, -62.948, -67.18, -67.008, -66.987, -67.258, -63.912, -63.143, -63.245, -63.213, -63.328, -63.529, -63.525, -63.563, -51.147, -50.006, -49.631, -49.602, -50.459, -50.143, -50.522, -50.539, -50.273, -50.246, -50.131, -49.935, -64.819, -65.756, -66.462, -66.339, -6.387, -1.57, -0.729, 0.049, 0.92, 1.619, 1.754, 1.761, -59.091, -63.889, -65.164, -65.744, -4.323, 0.38, 1.391, 2.119, 1.894, 1.781, 1.722, 1.874, -58.58, -63.425, -64.575, -65.288, -54.364, -53.884, -53.685, -53.7, -65.125, -66.288, -66.248, -66.39]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -57.917297619
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0, 5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0, 10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0, 16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0, 21668.0, 22205.0, 22741.0, 23278.0, 23869.0, 24405.0, 24941.0, 25478.0, 26068.0, 26604.0, 27140.0, 27677.0, 28265.0, 28801.0, 29337.0, 29874.0, 30462.0, 30997.0, 31534.0, 32071.0, 32660.0, 33195.0, 33732.0, 34268.0, 34857.0, 35394.0, 35929.0, 36465.0, 37056.0, 37592.0, 38127.0, 38664.0, 39253.0, 39790.0, 40327.0, 40864.0, 41455.0, 41990.0, 42527.0, 43063.0, 43650.0, 44186.0, 44721.0, 45257.0, 45844.0, 46381.0, 46918.0, 47453.0, 48041.0, 48577.0, 49114.0, 49650.0, 50239.0, 50776.0, 51311.0, 51848.0, 52436.0, 52972.0, 53509.0, 54045.0, 54633.0, 55169.0, 55705.0, 56240.0, 56829.0, 57366.0, 57902.0, 58437.0, 59023.0, 59559.0, 60095.0, 60631.0, 61217.0, 61753.0, 62290.0, 62827.0, 63414.0, 63949.0, 64485.0, 65020.0, 65607.0, 66143.0, 66680.0, 67216.0, 67803.0, 68339.0, 68876.0, 69411.0, 70000.0, 70536.0, 71072.0, 71607.0, 72191.0, 72726.0, 73262.0, 73798.0, 74383.0, 74918.0, 75454.0, 75989.0, 76575.0, 77110.0, 77646.0, 78182.0, 78766.0, 79302.0, 79837.0, 80374.0, 80960.0, 81496.0, 82032.0, 82567.0, 83154.0, 83690.0, 84225.0, 84761.0, 85347.0, 85883.0, 86418.0, 86954.0, 87543.0, 88077.0, 88613.0, 89149.0, 89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1821
varray: [2.167, 2.185, 2.173, 2.174, 2.198, 2.178, 2.179, 2.191, 2.185, 2.191]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: residuals
SAMP_residuals_STD | 0.00960844535928
varray: [2.167, 2.185, 2.173, 2.174, 2.198, 2.178, 2.179, 2.191, 2.185, 2.191]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9862
varray: [-0.969, -1.023, -0.972, -1.05, -0.927, -0.995, -0.875, -1.038, -1.107, -0.906]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0711630365975
varray: [-0.969, -1.023, -0.972, -1.05, -0.927, -0.995, -0.875, -1.038, -1.107, -0.906]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 72.0656
varray: [69.656, 74.195, 73.767, 73.338, 72.992, 72.372, 71.642, 71.328, 70.893, 70.473]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.50706611224
varray: [69.656, 74.195, 73.767, 73.338, 72.992, 72.372, 71.642, 71.328, 70.893, 70.473]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.5376
varray: [29.1, 29.684, 29.559, 29.544, 29.441, 29.531, 29.377, 29.621, 29.64, 29.879]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.206179102293
varray: [29.1, 29.684, 29.559, 29.544, 29.441, 29.531, 29.377, 29.621, 29.64, 29.879]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0552956699129
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.577746782288
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.5486
varray: [65.051, 65.198, 65.697, 64.995, 66.067, 65.678, 66.202, 65.605, 65.27, 65.723]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.411319070525
varray: [65.051, 65.198, 65.697, 64.995, 66.067, 65.678, 66.202, 65.605, 65.27, 65.723]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1821
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9862
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 72.0656
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.5376
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0289
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0289
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.30710085934
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.0914
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1485
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
dD_mean | -68.1937
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0552956699129
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
dD_stdev | 0.577746782288
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1485
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.1937
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0552956699129
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.577746782288
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2630598294
actual_array: [ 1.93 -15.84]
varray: [-9.28, -9.187, -9.169, -9.157, -9.104, -9.136, -9.101, -9.141, -9.108, -9.102]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.1485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.6410760409
actual_array: [ 18.1 -119. ]
varray: [-69.503, -68.842, -68.384, -68.187, -67.863, -67.903, -67.973, -67.883, -67.646, -67.753]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.1937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [1.0, 537.0, 1072.0, 1608.0, 2145.0, 2681.0, 3219.0, 3754.0, 4291.0, 4828.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6677
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.2158
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3098
varray: [2.279, 2.295, 2.317, 2.323, 2.315, 2.318, 2.313, 2.321, 2.302, 2.315]
inst: mean_all
operation: mean_all(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: residuals
SAMP_residuals_STD | 0.0137904959221
varray: [2.279, 2.295, 2.317, 2.323, 2.315, 2.318, 2.313, 2.321, 2.302, 2.315]
inst: std_all
operation: std_all(residuals)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1152
varray: [-1.123, -1.224, -0.9, -0.976, -1.026, -1.231, -1.159, -1.119, -1.277, -1.117]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.118821808698
varray: [-1.123, -1.224, -0.9, -0.976, -1.026, -1.231, -1.159, -1.119, -1.277, -1.117]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 69.0829
varray: [68.605, 69.876, 69.314, 69.331, 69.196, 69.166, 69.008, 68.776, 68.737, 68.82]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.377214189082
varray: [68.605, 69.876, 69.314, 69.331, 69.196, 69.166, 69.008, 68.776, 68.737, 68.82]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.8994
varray: [29.847, 29.727, 29.867, 29.88, 30.086, 30.113, 29.835, 29.719, 30.09, 29.83]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.146125668899
varray: [29.847, 29.727, 29.867, 29.88, 30.086, 30.113, 29.835, 29.719, 30.09, 29.83]
inst: std_all
operation: std_all(slope_shift)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: slope_shift
SAMP_d18O_STD | 0.188841026616
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.15611681237
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 67.4852
varray: [66.591, 66.448, 68.577, 68.532, 67.496, 66.557, 68.145, 68.367, 66.538, 67.601]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.89121462685
varray: [66.591, 66.448, 68.577, 68.532, 67.496, 66.557, 68.145, 68.367, 66.538, 67.601]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3098
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1152
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 69.0829
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.8994
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0988
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0988
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.46856625961
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8452
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.7456
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.629
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7456
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
dD_mean | 17.629
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
d18O_stdev | 0.188841026616
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
dD_stdev | 3.15611681237
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7456
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.629
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.188841026616
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.15611681237
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.251, 2.643, 2.704, 2.775, 2.835, 2.821, 2.86, 2.863, 2.854, 2.85]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.7456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [9.338, 15.707, 17.493, 18.145, 18.667, 18.973, 19.152, 19.669, 19.632, 19.514]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 17.629
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [5417.0, 5954.0, 6490.0, 7028.0, 7563.0, 8099.0, 8636.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.54
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.2331
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1122
varray: [2.122, 2.117, 2.126, 2.117, 2.116, 2.105, 2.098, 2.112, 2.102, 2.107]
inst: mean_all
operation: mean_all(residuals)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: residuals
SAMP_residuals_STD | 0.00901603509803
varray: [2.122, 2.117, 2.126, 2.117, 2.116, 2.105, 2.098, 2.112, 2.102, 2.107]
inst: std_all
operation: std_all(residuals)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9874
varray: [-1.096, -1.004, -1.07, -0.824, -0.948, -1.007, -1.025, -1.108, -0.951, -0.841]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0978913229613
varray: [-1.096, -1.004, -1.07, -0.824, -0.948, -1.007, -1.025, -1.108, -0.951, -0.841]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.6167
varray: [68.507, 68.526, 68.765, 68.718, 68.699, 68.752, 68.653, 68.618, 68.589, 68.34]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.131759081492
varray: [68.507, 68.526, 68.765, 68.718, 68.699, 68.752, 68.653, 68.618, 68.589, 68.34]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.6392
varray: [29.587, 29.864, 30.016, 29.607, 29.841, 29.71, 29.696, 29.414, 29.317, 29.34]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.233163176052
varray: [29.587, 29.864, 30.016, 29.607, 29.841, 29.71, 29.696, 29.414, 29.317, 29.34]
inst: std_all
operation: std_all(slope_shift)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: slope_shift
SAMP_d18O_STD | 0.28935485711
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.60235544163
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.7948
varray: [63.502, 63.536, 63.455, 64.918, 64.39, 63.153, 63.274, 63.526, 63.678, 64.516]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.593780692774
varray: [63.502, 63.536, 63.455, 64.918, 64.39, 63.153, 63.274, 63.526, 63.678, 64.516]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1122
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9874
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.6167
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.6392
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0988
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0988
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.46856625961
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8452
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.5887
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
dD_mean | -109.2435
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
d18O_stdev | 0.28935485711
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
dD_stdev | 4.60235544163
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.5887
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
measured_dD_mean | -109.2435
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.28935485711
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.60235544163
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.817, -14.452, -14.558, -14.646, -14.684, -14.723, -14.755, -14.729, -14.783, -14.74]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.5887
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-97.112, -106.547, -108.938, -110.062, -110.899, -111.157, -111.676, -111.8, -112.1, -112.144]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -109.2435
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [10833.0, 11370.0, 11906.0, 12442.0, 12979.0, 13516.0, 14052.0, 14588.0, 15125.0, 15661.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.7376
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.2331
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1358
varray: [2.152, 2.134, 2.146, 2.127, 2.138, 2.146, 2.133, 2.119, 2.127, 2.136]
inst: mean_all
operation: mean_all(residuals)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: residuals
SAMP_residuals_STD | 0.0101302626927
varray: [2.152, 2.134, 2.146, 2.127, 2.138, 2.146, 2.133, 2.119, 2.127, 2.136]
inst: std_all
operation: std_all(residuals)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9208
varray: [-0.869, -0.828, -0.914, -1.018, -0.828, -0.819, -0.903, -0.883, -1.061, -1.085]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0989543104445
varray: [-0.869, -0.828, -0.914, -1.018, -0.828, -0.819, -0.903, -0.883, -1.061, -1.085]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.9045
varray: [68.028, 67.869, 67.965, 67.95, 67.863, 67.843, 67.879, 67.877, 67.949, 67.822]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0649277376097
varray: [68.028, 67.869, 67.965, 67.95, 67.863, 67.843, 67.879, 67.877, 67.949, 67.822]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.8825
varray: [29.13, 28.551, 28.894, 28.673, 28.983, 28.842, 28.709, 28.56, 29.235, 29.248]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.262226894968
varray: [29.13, 28.551, 28.894, 28.673, 28.983, 28.842, 28.709, 28.56, 29.235, 29.248]
inst: std_all
operation: std_all(slope_shift)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0628232971649
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.41042358098
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.5157
varray: [65.745, 65.698, 65.036, 63.46, 64.912, 64.962, 64.429, 64.552, 63.361, 63.002]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.959393216801
varray: [65.745, 65.698, 65.036, 63.46, 64.912, 64.962, 64.429, 64.552, 63.361, 63.002]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1358
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9208
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.9045
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.8825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0752
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0752
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.40117593849
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1243
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1651
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
dD_mean | -68.2673
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0628232971649
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
dD_stdev | 1.41042358098
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1651
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.2673
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0628232971649
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.41042358098
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2800770726
actual_array: [ 1.93 -15.84]
varray: [-9.303, -9.222, -9.21, -9.116, -9.152, -9.107, -9.147, -9.111, -9.12, -9.163]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.1651
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.7206091155
actual_array: [ 18.1 -119. ]
varray: [-72.016, -68.931, -68.525, -67.717, -67.778, -67.675, -67.77, -67.564, -67.45, -67.247]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.2673
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [16251.0, 16788.0, 17324.0, 17859.0, 18396.0, 18932.0, 19468.0, 20004.0, 20541.0, 21078.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.714
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9453
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13925
varray: [2.137, 2.135, 2.149, 2.136]
inst: mean_all
operation: mean_all(residuals)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: residuals
SAMP_residuals_STD | 0.00655108133568
varray: [2.137, 2.135, 2.149, 2.136]
inst: std_all
operation: std_all(residuals)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.09375
varray: [-1.097, -1.114, -0.964, -1.2]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0975444343193
varray: [-1.097, -1.114, -0.964, -1.2]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.98275
varray: [68.113, 67.983, 67.922, 67.913]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0922329478368
varray: [68.113, 67.983, 67.922, 67.913]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.0555
varray: [29.081, 28.822, 28.98, 29.339]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.216983102261
varray: [29.081, 28.822, 28.98, 29.339]
inst: std_all
operation: std_all(slope_shift)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0426028168083
varray: [-8.222, -8.161, -8.183, -8.12]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.183803110238
varray: [-64.153, -63.977, -64.136, -63.757]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.7765
varray: [63.895, 63.635, 64.252, 63.324]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.393666948913
varray: [63.895, 63.635, 64.252, 63.324]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.09375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.98275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.0555
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.24513794663
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.222, -8.161, -8.183, -8.12]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-64.153, -63.977, -64.136, -63.757]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.1715
varray: [-8.222, -8.161, -8.183, -8.12]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
dD_mean | -64.00575
varray: [-64.153, -63.977, -64.136, -63.757]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0426028168083
varray: [-8.222, -8.161, -8.183, -8.12]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
dD_stdev | 0.183803110238
varray: [-64.153, -63.977, -64.136, -63.757]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.1715
varray: [-8.222, -8.161, -8.183, -8.12]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.00575
varray: [-64.153, -63.977, -64.136, -63.757]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0426028168083
varray: [-8.222, -8.161, -8.183, -8.12]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.183803110238
varray: [-64.153, -63.977, -64.136, -63.757]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.26150280081
actual_array: [ 1.93 -15.84]
varray: [-8.222, -8.161, -8.183, -8.12]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.1715
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.115525232
actual_array: [ 18.1 -119. ]
varray: [-64.153, -63.977, -64.136, -63.757]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -64.00575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [21668.0, 22205.0, 22741.0, 23278.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.71055
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.86705
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2385
varray: [2.219, 2.243, 2.251, 2.241]
inst: mean_all
operation: mean_all(residuals)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: residuals
SAMP_residuals_STD | 0.013699148392
varray: [2.219, 2.243, 2.251, 2.241]
inst: std_all
operation: std_all(residuals)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0315
varray: [-0.966, -0.991, -1.118, -1.051]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0678060960878
varray: [-0.966, -0.991, -1.118, -1.051]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.9845
varray: [67.924, 68.056, 68.012, 67.946]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0605832760201
varray: [67.924, 68.056, 68.012, 67.946]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.952
varray: [28.81, 28.968, 29.025, 29.005]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0975670709478
varray: [28.81, 28.968, 29.025, 29.005]
inst: std_all
operation: std_all(slope_shift)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: slope_shift
SAMP_d18O_STD | 0.169955631465
varray: [-0.797, -0.49, -0.453, -0.437]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.94381368296
varray: [-6.386, -1.44, -0.375, -0.045]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.72925
varray: [65.474, 66.361, 65.753, 65.329]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.45644450192
varray: [65.474, 66.361, 65.753, 65.329]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2385
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0315
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.9845
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.952
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.24378109453
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.08925
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-0.797, -0.49, -0.453, -0.437]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-6.386, -1.44, -0.375, -0.045]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.54425
varray: [-0.797, -0.49, -0.453, -0.437]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
dD_mean | -2.0615
varray: [-6.386, -1.44, -0.375, -0.045]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
d18O_stdev | 0.169955631465
varray: [-0.797, -0.49, -0.453, -0.437]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
dD_stdev | 2.94381368296
varray: [-6.386, -1.44, -0.375, -0.045]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.54425
varray: [-0.797, -0.49, -0.453, -0.437]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.0615
varray: [-6.386, -1.44, -0.375, -0.045]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.169955631465
varray: [-0.797, -0.49, -0.453, -0.437]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.94381368296
varray: [-6.386, -1.44, -0.375, -0.045]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.44254082945
actual_array: [ 1.93 -15.84]
varray: [-0.797, -0.49, -0.453, -0.437]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -0.54425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.17779897141
actual_array: [ 18.1 -119. ]
varray: [-6.386, -1.44, -0.375, -0.045]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -2.0615
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [23869.0, 24405.0, 24941.0, 25478.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6113
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8653
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2225
varray: [2.24, 2.25, 2.203, 2.197]
inst: mean_all
operation: mean_all(residuals)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: residuals
SAMP_residuals_STD | 0.0264133804475
varray: [2.24, 2.25, 2.203, 2.197]
inst: std_all
operation: std_all(residuals)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1025
varray: [-1.062, -1.214, -1.036, -1.098]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0785599558724
varray: [-1.062, -1.214, -1.036, -1.098]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.7505
varray: [68.008, 68.101, 67.482, 67.411]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.354264025834
varray: [68.008, 68.101, 67.482, 67.411]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.90375
varray: [29.262, 29.199, 28.533, 28.621]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.379876624709
varray: [29.262, 29.199, 28.533, 28.621]
inst: std_all
operation: std_all(slope_shift)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0621744856566
varray: [-1.233, -1.378, -1.307, -1.344]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.272271310767
varray: [-6.008, -6.616, -6.454, -6.18]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.09775
varray: [65.838, 65.342, 64.843, 64.368]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.633787753642
varray: [65.838, 65.342, 64.843, 64.368]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.7505
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.90375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.52012663953
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.54225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-1.233, -1.378, -1.307, -1.344]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-6.008, -6.616, -6.454, -6.18]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.3155
varray: [-1.233, -1.378, -1.307, -1.344]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
dD_mean | -6.3145
varray: [-6.008, -6.616, -6.454, -6.18]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0621744856566
varray: [-1.233, -1.378, -1.307, -1.344]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
dD_stdev | 0.272271310767
varray: [-6.008, -6.616, -6.454, -6.18]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.3155
varray: [-1.233, -1.378, -1.307, -1.344]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
measured_dD_mean | -6.3145
varray: [-6.008, -6.616, -6.454, -6.18]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0621744856566
varray: [-1.233, -1.378, -1.307, -1.344]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.272271310767
varray: [-6.008, -6.616, -6.454, -6.18]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.23317630363
actual_array: [ 1.93 -15.84]
varray: [-1.233, -1.378, -1.307, -1.344]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -1.3155
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.77364361859
actual_array: [ 18.1 -119. ]
varray: [-6.008, -6.616, -6.454, -6.18]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.3145
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [26068.0, 26604.0, 27140.0, 27677.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6273
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0993
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.167
varray: [2.181, 2.154, 2.165, 2.168]
inst: mean_all
operation: mean_all(residuals)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: residuals
SAMP_residuals_STD | 0.011105554166
varray: [2.181, 2.154, 2.165, 2.168]
inst: std_all
operation: std_all(residuals)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04775
varray: [-0.987, -0.94, -1.05, -1.214]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.119644960891
varray: [-0.987, -0.94, -1.05, -1.214]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.46875
varray: [67.673, 67.292, 67.465, 67.445]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.156563884724
varray: [67.673, 67.292, 67.465, 67.445]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.50225
varray: [28.741, 28.036, 28.495, 28.737]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.33143664553
varray: [28.741, 28.036, 28.495, 28.737]
inst: std_all
operation: std_all(slope_shift)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0727094675174
varray: [-4.058, -4.184, -4.156, -4.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.07375881836
varray: [-28.361, -30.226, -30.293, -30.812]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.24225
varray: [65.041, 64.402, 64.397, 63.129]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.801414322724
varray: [65.041, 64.402, 64.397, 63.129]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.167
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.46875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.50225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.044
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.044
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.99004975124
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.39775
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-4.058, -4.184, -4.156, -4.23]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-28.361, -30.226, -30.293, -30.812]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.157
varray: [-4.058, -4.184, -4.156, -4.23]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
dD_mean | -29.923
varray: [-28.361, -30.226, -30.293, -30.812]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0727094675174
varray: [-4.058, -4.184, -4.156, -4.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
dD_stdev | 1.07375881836
varray: [-28.361, -30.226, -30.293, -30.812]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.157
varray: [-4.058, -4.184, -4.156, -4.23]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
measured_dD_mean | -29.923
varray: [-28.361, -30.226, -30.293, -30.812]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0727094675174
varray: [-4.058, -4.184, -4.156, -4.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.07375881836
varray: [-28.361, -30.226, -30.293, -30.812]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.14609779455
actual_array: [ 1.93 -15.84]
varray: [-4.058, -4.184, -4.156, -4.23]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.157
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.2852820745
actual_array: [ 18.1 -119. ]
varray: [-28.361, -30.226, -30.293, -30.812]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -29.923
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [28265.0, 28801.0, 29337.0, 29874.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6828
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.38105
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.178
varray: [2.144, 2.19, 2.194, 2.184]
inst: mean_all
operation: mean_all(residuals)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: residuals
SAMP_residuals_STD | 0.0230362033909
varray: [2.144, 2.19, 2.194, 2.184]
inst: std_all
operation: std_all(residuals)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04575
varray: [-1.047, -0.914, -1.076, -1.146]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0971678101705
varray: [-1.047, -0.914, -1.076, -1.146]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.70675
varray: [67.389, 67.606, 67.832, 68.0]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.26633234251
varray: [67.389, 67.606, 67.832, 68.0]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.918
varray: [28.525, 29.021, 29.139, 28.987]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.269975307513
varray: [28.525, 29.021, 29.139, 28.987]
inst: std_all
operation: std_all(slope_shift)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0309771421105
varray: [-4.483, -4.416, -4.474, -4.44]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.160778056131
varray: [-32.608, -32.326, -32.684, -32.447]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.595
varray: [63.882, 65.142, 64.844, 64.512]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.540514569646
varray: [63.882, 65.142, 64.844, 64.512]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.178
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.70675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.918
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.033
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.033
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.49253731343
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-4.483, -4.416, -4.474, -4.44]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-32.608, -32.326, -32.684, -32.447]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.45325
varray: [-4.483, -4.416, -4.474, -4.44]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
dD_mean | -32.51625
varray: [-32.608, -32.326, -32.684, -32.447]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0309771421105
varray: [-4.483, -4.416, -4.474, -4.44]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
dD_stdev | 0.160778056131
varray: [-32.608, -32.326, -32.684, -32.447]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.45325
varray: [-4.483, -4.416, -4.474, -4.44]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.51625
varray: [-32.608, -32.326, -32.684, -32.447]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0309771421105
varray: [-4.483, -4.416, -4.474, -4.44]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.160778056131
varray: [-32.608, -32.326, -32.684, -32.447]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.44979407879
actual_array: [ 1.93 -15.84]
varray: [-4.483, -4.416, -4.474, -4.44]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.45325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.0875802479
actual_array: [ 18.1 -119. ]
varray: [-32.608, -32.326, -32.684, -32.447]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -32.51625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [30462.0, 30997.0, 31534.0, 32071.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6718
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.14305
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.194
varray: [2.193, 2.223, 2.191, 2.169]
inst: mean_all
operation: mean_all(residuals)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: residuals
SAMP_residuals_STD | 0.0221810730128
varray: [2.193, 2.223, 2.191, 2.169]
inst: std_all
operation: std_all(residuals)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0295
varray: [-1.039, -0.876, -1.153, -1.05]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114491629971
varray: [-1.039, -0.876, -1.153, -1.05]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.91475
varray: [67.97, 68.082, 67.981, 67.626]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.198989740104
varray: [67.97, 68.082, 67.981, 67.626]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.12125
varray: [29.139, 29.483, 29.291, 28.572]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.392287628321
varray: [29.139, 29.483, 29.291, 28.572]
inst: std_all
operation: std_all(slope_shift)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: slope_shift
SAMP_d18O_STD | 0.031336879232
varray: [-4.654, -4.729, -4.705, -4.692]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0499699909946
varray: [-33.081, -33.081, -33.187, -33.117]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.055
varray: [65.2, 66.125, 64.556, 64.339]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.801578026978
varray: [65.2, 66.125, 64.556, 64.339]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.194
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0295
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.91475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.12125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.017
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.017
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.768882858435
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-4.654, -4.729, -4.705, -4.692]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-33.081, -33.081, -33.187, -33.117]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.695
varray: [-4.654, -4.729, -4.705, -4.692]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
dD_mean | -33.1165
varray: [-33.081, -33.081, -33.187, -33.117]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
d18O_stdev | 0.031336879232
varray: [-4.654, -4.729, -4.705, -4.692]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
dD_stdev | 0.0499699909946
varray: [-33.081, -33.081, -33.187, -33.117]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.695
varray: [-4.654, -4.729, -4.705, -4.692]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
measured_dD_mean | -33.1165
varray: [-33.081, -33.081, -33.187, -33.117]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.031336879232
varray: [-4.654, -4.729, -4.705, -4.692]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0499699909946
varray: [-33.081, -33.081, -33.187, -33.117]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.69762049809
actual_array: [ 1.93 -15.84]
varray: [-4.654, -4.729, -4.705, -4.692]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.7362178565
actual_array: [ 18.1 -119. ]
varray: [-33.081, -33.081, -33.187, -33.117]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -33.1165
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [32660.0, 33195.0, 33732.0, 34268.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6558
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.93505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10175
varray: [2.125, 2.097, 2.126, 2.059]
inst: mean_all
operation: mean_all(residuals)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: residuals
SAMP_residuals_STD | 0.0315105802337
varray: [2.125, 2.097, 2.126, 2.059]
inst: std_all
operation: std_all(residuals)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03225
varray: [-1.035, -1.212, -0.935, -0.947]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.127857668262
varray: [-1.035, -1.212, -0.935, -0.947]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.092
varray: [68.276, 68.11, 68.35, 67.632]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.322668044074
varray: [68.276, 68.11, 68.35, 67.632]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.99425
varray: [29.168, 29.103, 29.148, 28.558]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.292100867282
varray: [29.168, 29.103, 29.148, 28.558]
inst: std_all
operation: std_all(slope_shift)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: slope_shift
SAMP_d18O_STD | 0.18784279775
varray: [-12.914, -13.251, -13.302, -13.306]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.17143952594
varray: [-94.713, -99.801, -101.26, -101.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.462
varray: [63.912, 62.895, 64.559, 62.482]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.946537197719
varray: [63.912, 62.895, 64.559, 62.482]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.092
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.99425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.94120307553
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.178
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-12.914, -13.251, -13.302, -13.306]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-94.713, -99.801, -101.26, -101.546]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.19325
varray: [-12.914, -13.251, -13.302, -13.306]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
dD_mean | -99.33
varray: [-94.713, -99.801, -101.26, -101.546]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
d18O_stdev | 0.18784279775
varray: [-12.914, -13.251, -13.302, -13.306]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
dD_stdev | 3.17143952594
varray: [-94.713, -99.801, -101.26, -101.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.19325
varray: [-12.914, -13.251, -13.302, -13.306]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
measured_dD_mean | -99.33
varray: [-94.713, -99.801, -101.26, -101.546]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.18784279775
varray: [-12.914, -13.251, -13.302, -13.306]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.17143952594
varray: [-94.713, -99.801, -101.26, -101.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -14.4094751735
actual_array: [ 1.93 -15.84]
varray: [-12.914, -13.251, -13.302, -13.306]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -13.19325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -108.287348716
actual_array: [ 18.1 -119. ]
varray: [-94.713, -99.801, -101.26, -101.546]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -99.33
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [34857.0, 35394.0, 35929.0, 36465.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.74805
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.7578
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.00625
varray: [2.047, 1.99, 2.012, 1.976]
inst: mean_all
operation: mean_all(residuals)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: residuals
SAMP_residuals_STD | 0.0309448434046
varray: [2.047, 1.99, 2.012, 1.976]
inst: std_all
operation: std_all(residuals)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.92675
varray: [-1.011, -0.982, -0.852, -0.862]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0815081795486
varray: [-1.011, -0.982, -0.852, -0.862]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.09775
varray: [68.517, 67.918, 68.144, 67.812]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.311914918955
varray: [68.517, 67.918, 68.144, 67.812]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.7875
varray: [29.58, 28.68, 28.539, 28.351]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.545251318201
varray: [29.58, 28.68, 28.539, 28.351]
inst: std_all
operation: std_all(slope_shift)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: slope_shift
SAMP_d18O_STD | 0.127766388381
varray: [-19.354, -19.532, -19.592, -19.649]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.83219520278
varray: [-155.421, -159.526, -161.0, -161.779]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.81775
varray: [62.774, 60.926, 62.306, 61.265]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.866708861152
varray: [62.774, 60.926, 62.306, 61.265]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.00625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.92675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.09775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.7875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.20475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.2605156038
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.82225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-19.354, -19.532, -19.592, -19.649]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-155.421, -159.526, -161.0, -161.779]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -19.53175
varray: [-19.354, -19.532, -19.592, -19.649]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
dD_mean | -159.4315
varray: [-155.421, -159.526, -161.0, -161.779]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
d18O_stdev | 0.127766388381
varray: [-19.354, -19.532, -19.592, -19.649]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
dD_stdev | 2.83219520278
varray: [-155.421, -159.526, -161.0, -161.779]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
measured_d18O_mean | -19.53175
varray: [-19.354, -19.532, -19.592, -19.649]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
measured_dD_mean | -159.4315
varray: [-155.421, -159.526, -161.0, -161.779]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.127766388381
varray: [-19.354, -19.532, -19.592, -19.649]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.83219520278
varray: [-155.421, -159.526, -161.0, -161.779]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -20.9072942374
actual_array: [ 1.93 -15.84]
varray: [-19.354, -19.532, -19.592, -19.649]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -19.53175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -173.233776429
actual_array: [ 18.1 -119. ]
varray: [-155.421, -159.526, -161.0, -161.779]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -159.4315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [37056.0, 37592.0, 38127.0, 38664.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.84355
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.75205
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13275
varray: [2.152, 2.137, 2.126, 2.116]
inst: mean_all
operation: mean_all(residuals)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: residuals
SAMP_residuals_STD | 0.0154353490404
varray: [2.152, 2.137, 2.126, 2.116]
inst: std_all
operation: std_all(residuals)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.80625
varray: [-0.65, -0.79, -0.818, -0.967]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129944539452
varray: [-0.65, -0.79, -0.818, -0.967]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.88525
varray: [68.073, 67.933, 67.685, 67.85]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.162144739867
varray: [68.073, 67.933, 67.685, 67.85]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.781
varray: [28.901, 28.642, 28.5, 29.081]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.259924347968
varray: [28.901, 28.642, 28.5, 29.081]
inst: std_all
operation: std_all(slope_shift)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: slope_shift
SAMP_d18O_STD | 0.215673170948
varray: [-9.608, -9.306, -9.169, -9.134]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.00856927926
varray: [-77.179, -70.983, -68.951, -68.461]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.99375
varray: [66.414, 65.25, 65.201, 63.11]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.37528672768
varray: [66.414, 65.25, 65.201, 63.11]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.80625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.88525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.781
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.53912256897
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.64625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-9.608, -9.306, -9.169, -9.134]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-77.179, -70.983, -68.951, -68.461]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.30425
varray: [-9.608, -9.306, -9.169, -9.134]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
dD_mean | -71.3935
varray: [-77.179, -70.983, -68.951, -68.461]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
d18O_stdev | 0.215673170948
varray: [-9.608, -9.306, -9.169, -9.134]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
dD_stdev | 4.00856927926
varray: [-77.179, -70.983, -68.951, -68.461]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.30425
varray: [-9.608, -9.306, -9.169, -9.134]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.3935
varray: [-77.179, -70.983, -68.951, -68.461]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.215673170948
varray: [-9.608, -9.306, -9.169, -9.134]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.00856927926
varray: [-77.179, -70.983, -68.951, -68.461]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4227246269
actual_array: [ 1.93 -15.84]
varray: [-9.608, -9.306, -9.169, -9.134]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.30425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.0988196812
actual_array: [ 18.1 -119. ]
varray: [-77.179, -70.983, -68.951, -68.461]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -71.3935
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [39253.0, 39790.0, 40327.0, 40864.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.71705
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.96455
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.04425
varray: [2.02, 2.065, 2.047, 2.045]
inst: mean_all
operation: mean_all(residuals)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: residuals
SAMP_residuals_STD | 0.0185
varray: [2.02, 2.065, 2.047, 2.045]
inst: std_all
operation: std_all(residuals)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9405
varray: [-0.915, -0.796, -1.023, -1.028]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109533860822
varray: [-0.915, -0.796, -1.023, -1.028]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.04025
varray: [67.771, 68.147, 68.15, 68.093]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.181400799337
varray: [67.771, 68.147, 68.15, 68.093]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.62575
varray: [28.035, 28.667, 28.821, 28.98]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.414046193075
varray: [28.035, 28.667, 28.821, 28.98]
inst: std_all
operation: std_all(slope_shift)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: slope_shift
SAMP_d18O_STD | 0.161473424026
varray: [-16.663, -16.953, -16.982, -17.012]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.00778870878
varray: [-129.461, -134.313, -135.525, -136.054]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.44925
varray: [61.775, 63.488, 62.304, 62.23]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.730931540342
varray: [61.775, 63.488, 62.304, 62.23]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.04425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9405
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.04025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.62575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.54183627318
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.19075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-16.663, -16.953, -16.982, -17.012]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-129.461, -134.313, -135.525, -136.054]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -16.9025
varray: [-16.663, -16.953, -16.982, -17.012]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
dD_mean | -133.83825
varray: [-129.461, -134.313, -135.525, -136.054]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
d18O_stdev | 0.161473424026
varray: [-16.663, -16.953, -16.982, -17.012]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
dD_stdev | 3.00778870878
varray: [-129.461, -134.313, -135.525, -136.054]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
measured_d18O_mean | -16.9025
varray: [-16.663, -16.953, -16.982, -17.012]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
measured_dD_mean | -133.83825
varray: [-129.461, -134.313, -135.525, -136.054]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.161473424026
varray: [-16.663, -16.953, -16.982, -17.012]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.00778870878
varray: [-129.461, -134.313, -135.525, -136.054]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -18.2119576793
actual_array: [ 1.93 -15.84]
varray: [-16.663, -16.953, -16.982, -17.012]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -16.9025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -145.577392461
actual_array: [ 18.1 -119. ]
varray: [-129.461, -134.313, -135.525, -136.054]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -133.83825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [41455.0, 41990.0, 42527.0, 43063.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.80555
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.80955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.9665
varray: [1.982, 1.955, 1.959, 1.97]
inst: mean_all
operation: mean_all(residuals)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: residuals
SAMP_residuals_STD | 0.012124355653
varray: [1.982, 1.955, 1.959, 1.97]
inst: std_all
operation: std_all(residuals)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.05725
varray: [-1.281, -1.054, -0.987, -0.907]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.160815370327
varray: [-1.281, -1.054, -0.987, -0.907]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.77975
varray: [68.012, 67.86, 67.62, 67.627]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.190816447579
varray: [68.012, 67.86, 67.62, 67.627]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.35
varray: [28.839, 28.159, 28.3, 28.102]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.336455049004
varray: [28.839, 28.159, 28.3, 28.102]
inst: std_all
operation: std_all(slope_shift)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0110302614052
varray: [-18.877, -18.89, -18.891, -18.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.55190407242
varray: [-166.247, -168.692, -169.177, -169.787]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.98475
varray: [59.659, 60.266, 59.099, 60.915]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.782119502804
varray: [59.659, 60.266, 59.099, 60.915]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.9665
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.05725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.77975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.35
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2445
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2445
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.0583446404
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.65525
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-18.877, -18.89, -18.891, -18.904]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-166.247, -168.692, -169.177, -169.787]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -18.8905
varray: [-18.877, -18.89, -18.891, -18.904]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
dD_mean | -168.47575
varray: [-166.247, -168.692, -169.177, -169.787]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0110302614052
varray: [-18.877, -18.89, -18.891, -18.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
dD_stdev | 1.55190407242
varray: [-166.247, -168.692, -169.177, -169.787]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
measured_d18O_mean | -18.8905
varray: [-18.877, -18.89, -18.891, -18.904]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
measured_dD_mean | -168.47575
varray: [-166.247, -168.692, -169.177, -169.787]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0110302614052
varray: [-18.877, -18.89, -18.891, -18.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.55190407242
varray: [-166.247, -168.692, -169.177, -169.787]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -20.249926331
actual_array: [ 1.93 -15.84]
varray: [-18.877, -18.89, -18.891, -18.904]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -18.8905
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -183.007105362
actual_array: [ 18.1 -119. ]
varray: [-166.247, -168.692, -169.177, -169.787]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -168.47575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [43650.0, 44186.0, 44721.0, 45257.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.8833
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.07005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11575
varray: [2.147, 2.101, 2.107, 2.108]
inst: mean_all
operation: mean_all(residuals)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: residuals
SAMP_residuals_STD | 0.0210614181226
varray: [2.147, 2.101, 2.107, 2.108]
inst: std_all
operation: std_all(residuals)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0215
varray: [-0.946, -1.013, -1.127, -1.0]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0760810970829
varray: [-0.946, -1.013, -1.127, -1.0]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.954
varray: [68.277, 67.821, 67.897, 67.821]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.218293380568
varray: [68.277, 67.821, 67.897, 67.821]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.0305
varray: [29.472, 28.684, 28.909, 29.057]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.331888033329
varray: [29.472, 28.684, 28.909, 29.057]
inst: std_all
operation: std_all(slope_shift)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: slope_shift
SAMP_d18O_STD | 0.132979948363
varray: [-11.338, -11.136, -11.134, -11.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.8112905076
varray: [-89.277, -83.297, -81.888, -80.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.6865
varray: [64.292, 63.96, 62.518, 63.976]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.793858719588
varray: [64.292, 63.96, 62.518, 63.976]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0215
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.954
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.0305
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.30800542741
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.9535
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-11.338, -11.136, -11.134, -11.018]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-89.277, -83.297, -81.888, -80.69]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.1565
varray: [-11.338, -11.136, -11.134, -11.018]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
dD_mean | -83.788
varray: [-89.277, -83.297, -81.888, -80.69]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
d18O_stdev | 0.132979948363
varray: [-11.338, -11.136, -11.134, -11.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
dD_stdev | 3.8112905076
varray: [-89.277, -83.297, -81.888, -80.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.1565
varray: [-11.338, -11.136, -11.134, -11.018]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
measured_dD_mean | -83.788
varray: [-89.277, -83.297, -81.888, -80.69]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.132979948363
varray: [-11.338, -11.136, -11.134, -11.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.8112905076
varray: [-89.277, -83.297, -81.888, -80.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.3215311838
actual_array: [ 1.93 -15.84]
varray: [-11.338, -11.136, -11.134, -11.018]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -11.1565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.4924703935
actual_array: [ 18.1 -119. ]
varray: [-89.277, -83.297, -81.888, -80.69]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -83.788
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [45844.0, 46381.0, 46918.0, 47453.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.73405
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8958
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15925
varray: [2.171, 2.135, 2.162, 2.169]
inst: mean_all
operation: mean_all(residuals)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: residuals
SAMP_residuals_STD | 0.0166207701386
varray: [2.171, 2.135, 2.162, 2.169]
inst: std_all
operation: std_all(residuals)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0065
varray: [-1.065, -1.042, -0.975, -0.944]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0565125354826
varray: [-1.065, -1.042, -0.975, -0.944]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.10975
varray: [68.239, 67.829, 68.078, 68.293]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.20825845321
varray: [68.239, 67.829, 68.078, 68.293]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.7445
varray: [29.021, 28.435, 28.744, 28.778]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.240392040911
varray: [29.021, 28.435, 28.744, 28.778]
inst: std_all
operation: std_all(slope_shift)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0621952570539
varray: [-9.125, -9.011, -9.03, -8.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.16464683488
varray: [-65.359, -63.647, -63.11, -62.717]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.8495
varray: [64.849, 64.423, 64.824, 65.302]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.359309059168
varray: [64.849, 64.423, 64.824, 65.302]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0065
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.10975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.7445
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.34056987788
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-9.125, -9.011, -9.03, -8.981]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-65.359, -63.647, -63.11, -62.717]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.03675
varray: [-9.125, -9.011, -9.03, -8.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
dD_mean | -63.70825
varray: [-65.359, -63.647, -63.11, -62.717]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0621952570539
varray: [-9.125, -9.011, -9.03, -8.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
dD_stdev | 1.16464683488
varray: [-65.359, -63.647, -63.11, -62.717]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.03675
varray: [-9.125, -9.011, -9.03, -8.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.70825
varray: [-65.359, -63.647, -63.11, -62.717]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0621952570539
varray: [-9.125, -9.011, -9.03, -8.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.16464683488
varray: [-65.359, -63.647, -63.11, -62.717]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1485009778
actual_array: [ 1.93 -15.84]
varray: [-9.125, -9.011, -9.03, -8.981]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.03675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.7940430353
actual_array: [ 18.1 -119. ]
varray: [-65.359, -63.647, -63.11, -62.717]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.70825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [48041.0, 48577.0, 49114.0, 49650.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.69055
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.74005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.136
varray: [2.175, 2.122, 2.127, 2.12]
inst: mean_all
operation: mean_all(residuals)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: residuals
SAMP_residuals_STD | 0.0261661358757
varray: [2.175, 2.122, 2.127, 2.12]
inst: std_all
operation: std_all(residuals)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.80525
varray: [-0.695, -0.538, -0.842, -1.146]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.258868531627
varray: [-0.695, -0.538, -0.842, -1.146]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.80925
varray: [68.163, 67.454, 67.759, 67.861]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.292429336194
varray: [68.163, 67.454, 67.759, 67.861]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.7935
varray: [29.364, 28.269, 28.562, 28.979]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.479086978881
varray: [29.364, 28.269, 28.562, 28.979]
inst: std_all
operation: std_all(slope_shift)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0211187120819
varray: [-9.53, -9.58, -9.565, -9.553]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.420022519238
varray: [-68.95, -69.742, -69.84, -69.245]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.1385
varray: [66.983, 65.392, 64.946, 63.233]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.54212288745
varray: [66.983, 65.392, 64.946, 63.233]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.136
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.80525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.80925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.7935
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.3921302578
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.5015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-9.53, -9.58, -9.565, -9.553]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-68.95, -69.742, -69.84, -69.245]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.557
varray: [-9.53, -9.58, -9.565, -9.553]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
dD_mean | -69.44425
varray: [-68.95, -69.742, -69.84, -69.245]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0211187120819
varray: [-9.53, -9.58, -9.565, -9.553]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
dD_stdev | 0.420022519238
varray: [-68.95, -69.742, -69.84, -69.245]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.557
varray: [-9.53, -9.58, -9.565, -9.553]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.44425
varray: [-68.95, -69.742, -69.84, -69.245]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0211187120819
varray: [-9.53, -9.58, -9.565, -9.553]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.420022519238
varray: [-68.95, -69.742, -69.84, -69.245]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.6818275327
actual_array: [ 1.93 -15.84]
varray: [-9.53, -9.58, -9.565, -9.553]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.557
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.9924359101
actual_array: [ 18.1 -119. ]
varray: [-68.95, -69.742, -69.84, -69.245]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -69.44425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [50239.0, 50776.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.7138
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.04055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.154
varray: [2.162, 2.14, 2.16, 2.154]
inst: mean_all
operation: mean_all(residuals)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: residuals
SAMP_residuals_STD | 0.00993310961717
varray: [2.162, 2.14, 2.16, 2.154]
inst: std_all
operation: std_all(residuals)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0685
varray: [-1.208, -1.186, -0.773, -1.107]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.201716798177
varray: [-1.208, -1.186, -0.773, -1.107]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.1685
varray: [68.425, 68.244, 67.96, 68.045]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.208340266551
varray: [68.425, 68.244, 67.96, 68.045]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.9095
varray: [29.374, 28.931, 28.536, 28.797]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.350419272682
varray: [29.374, 28.931, 28.536, 28.797]
inst: std_all
operation: std_all(slope_shift)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0402367990775
varray: [-9.195, -9.156, -9.111, -9.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.357161494752
varray: [-66.475, -66.057, -65.788, -65.67]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.11975
varray: [64.134, 62.887, 65.459, 63.999]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.05325633949
varray: [64.134, 62.887, 65.459, 63.999]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.154
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0685
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.1685
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.9095
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.057
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.057
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.57801899593
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.52025
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-9.195, -9.156, -9.111, -9.112]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-66.475, -66.057, -65.788, -65.67]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1435
varray: [-9.195, -9.156, -9.111, -9.112]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
dD_mean | -65.9975
varray: [-66.475, -66.057, -65.788, -65.67]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0402367990775
varray: [-9.195, -9.156, -9.111, -9.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
dD_stdev | 0.357161494752
varray: [-66.475, -66.057, -65.788, -65.67]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1435
varray: [-9.195, -9.156, -9.111, -9.112]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.9975
varray: [-66.475, -66.057, -65.788, -65.67]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0402367990775
varray: [-9.195, -9.156, -9.111, -9.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.357161494752
varray: [-66.475, -66.057, -65.788, -65.67]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2579341537
actual_array: [ 1.93 -15.84]
varray: [-9.195, -9.156, -9.111, -9.112]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.1435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.2678350312
actual_array: [ 18.1 -119. ]
varray: [-66.475, -66.057, -65.788, -65.67]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.9975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [52436.0, 52972.0, 53509.0, 54045.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6958
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.6813
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15025
varray: [2.169, 2.146, 2.146, 2.14]
inst: mean_all
operation: mean_all(residuals)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: residuals
SAMP_residuals_STD | 0.012816005618
varray: [2.169, 2.146, 2.146, 2.14]
inst: std_all
operation: std_all(residuals)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.13575
varray: [-1.242, -1.285, -1.008, -1.008]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.148553862286
varray: [-1.242, -1.285, -1.008, -1.008]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.029
varray: [68.404, 68.046, 67.784, 67.882]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.272364951245
varray: [68.404, 68.046, 67.784, 67.882]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.3345
varray: [29.886, 29.151, 29.164, 29.137]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.367831936261
varray: [29.886, 29.151, 29.164, 29.137]
inst: std_all
operation: std_all(slope_shift)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0338673884438
varray: [-8.077, -8.158, -8.105, -8.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.222427516283
varray: [-60.396, -60.426, -59.957, -60.141]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.64225
varray: [63.758, 62.891, 63.465, 64.455]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.650577884141
varray: [63.758, 62.891, 63.465, 64.455]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.13575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.029
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.3345
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.74762550882
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.99775
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.077, -8.158, -8.105, -8.122]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-60.396, -60.426, -59.957, -60.141]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.1155
varray: [-8.077, -8.158, -8.105, -8.122]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
dD_mean | -60.23
varray: [-60.396, -60.426, -59.957, -60.141]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0338673884438
varray: [-8.077, -8.158, -8.105, -8.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
dD_stdev | 0.222427516283
varray: [-60.396, -60.426, -59.957, -60.141]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.1155
varray: [-8.077, -8.158, -8.105, -8.122]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.23
varray: [-60.396, -60.426, -59.957, -60.141]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0338673884438
varray: [-8.077, -8.158, -8.105, -8.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.222427516283
varray: [-60.396, -60.426, -59.957, -60.141]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.20409523315
actual_array: [ 1.93 -15.84]
varray: [-8.077, -8.158, -8.105, -8.122]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.1155
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.0354028651
actual_array: [ 18.1 -119. ]
varray: [-60.396, -60.426, -59.957, -60.141]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -60.23
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [54633.0, 55169.0, 55705.0, 56240.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.69955
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8208
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.129
varray: [2.141, 2.096, 2.111, 2.168]
inst: mean_all
operation: mean_all(residuals)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: residuals
SAMP_residuals_STD | 0.0320312347561
varray: [2.141, 2.096, 2.111, 2.168]
inst: std_all
operation: std_all(residuals)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98225
varray: [-0.922, -1.236, -0.881, -0.89]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.170079148242
varray: [-0.922, -1.236, -0.881, -0.89]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.65975
varray: [67.65, 67.534, 67.362, 68.093]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.312126443395
varray: [67.65, 67.534, 67.362, 68.093]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.64075
varray: [28.487, 28.379, 28.413, 29.284]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431197074047
varray: [28.487, 28.379, 28.413, 29.284]
inst: std_all
operation: std_all(slope_shift)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0356697817954
varray: [-8.508, -8.588, -8.519, -8.547]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.360509246298
varray: [-62.634, -63.492, -62.893, -62.948]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.16225
varray: [64.688, 62.332, 63.928, 65.701]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.41995924237
varray: [64.688, 62.332, 63.928, 65.701]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.129
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.65975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.64075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.082
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.082
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.70872908186
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.47775
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.508, -8.588, -8.519, -8.547]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-62.634, -63.492, -62.893, -62.948]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.5405
varray: [-8.508, -8.588, -8.519, -8.547]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
dD_mean | -62.99175
varray: [-62.634, -63.492, -62.893, -62.948]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0356697817954
varray: [-8.508, -8.588, -8.519, -8.547]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
dD_stdev | 0.360509246298
varray: [-62.634, -63.492, -62.893, -62.948]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.5405
varray: [-8.508, -8.588, -8.519, -8.547]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.99175
varray: [-62.634, -63.492, -62.893, -62.948]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0356697817954
varray: [-8.508, -8.588, -8.519, -8.547]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.360509246298
varray: [-62.634, -63.492, -62.893, -62.948]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.63977766625
actual_array: [ 1.93 -15.84]
varray: [-8.508, -8.588, -8.519, -8.547]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.5405
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.0197842322
actual_array: [ 18.1 -119. ]
varray: [-62.634, -63.492, -62.893, -62.948]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -62.99175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [56829.0, 57366.0, 57902.0, 58437.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.7208
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.19005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.118
varray: [2.113, 2.108, 2.15, 2.101]
inst: mean_all
operation: mean_all(residuals)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: residuals
SAMP_residuals_STD | 0.0218936824982
varray: [2.113, 2.108, 2.15, 2.101]
inst: std_all
operation: std_all(residuals)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00325
varray: [-1.108, -0.798, -0.931, -1.176]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.171428070436
varray: [-1.108, -0.798, -0.931, -1.176]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.66875
varray: [67.699, 67.404, 68.047, 67.525]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.279728886603
varray: [67.699, 67.404, 68.047, 67.525]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.665
varray: [28.887, 28.058, 29.085, 28.63]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.445480265182
varray: [28.887, 28.058, 29.085, 28.63]
inst: std_all
operation: std_all(slope_shift)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0352455670972
varray: [-9.187, -9.104, -9.137, -9.125]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.132066586741
varray: [-67.18, -67.008, -66.987, -67.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.4785
varray: [62.903, 64.457, 64.759, 61.795]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.38594312053
varray: [62.903, 64.457, 64.759, 61.795]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.118
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.66875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.665
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.093
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.093
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.20624151967
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.1615
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-9.187, -9.104, -9.137, -9.125]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-67.18, -67.008, -66.987, -67.258]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.13825
varray: [-9.187, -9.104, -9.137, -9.125]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
dD_mean | -67.10825
varray: [-67.18, -67.008, -66.987, -67.258]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0352455670972
varray: [-9.187, -9.104, -9.137, -9.125]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
dD_stdev | 0.132066586741
varray: [-67.18, -67.008, -66.987, -67.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.13825
varray: [-9.187, -9.104, -9.137, -9.125]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.10825
varray: [-67.18, -67.008, -66.987, -67.258]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0352455670972
varray: [-9.187, -9.104, -9.137, -9.125]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.132066586741
varray: [-67.18, -67.008, -66.987, -67.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2525521942
actual_array: [ 1.93 -15.84]
varray: [-9.187, -9.104, -9.137, -9.125]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.13825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.4681252833
actual_array: [ 18.1 -119. ]
varray: [-67.18, -67.008, -66.987, -67.258]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.10825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [59023.0, 59559.0, 60095.0, 60631.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.7318
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.18105
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14675
varray: [2.148, 2.148, 2.155, 2.136]
inst: mean_all
operation: mean_all(residuals)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: residuals
SAMP_residuals_STD | 0.00788986691903
varray: [2.148, 2.148, 2.155, 2.136]
inst: std_all
operation: std_all(residuals)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0015
varray: [-0.935, -0.919, -1.08, -1.072]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.086334620325
varray: [-0.935, -0.919, -1.08, -1.072]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.837
varray: [67.908, 67.806, 67.797, 67.837]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0503388517946
varray: [67.908, 67.806, 67.797, 67.837]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.85575
varray: [28.761, 28.643, 29.203, 28.816]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.242489003187
varray: [28.761, 28.643, 29.203, 28.816]
inst: std_all
operation: std_all(slope_shift)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00942956343988
varray: [-8.618, -8.596, -8.601, -8.606]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.358373524878
varray: [-63.912, -63.143, -63.245, -63.213]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.38925
varray: [64.71, 64.923, 64.396, 63.528]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.613614088821
varray: [64.71, 64.923, 64.396, 63.528]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0015
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.837
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.85575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.90592492085
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.25075
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.618, -8.596, -8.601, -8.606]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-63.912, -63.143, -63.245, -63.213]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.60525
varray: [-8.618, -8.596, -8.601, -8.606]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
dD_mean | -63.37825
varray: [-63.912, -63.143, -63.245, -63.213]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00942956343988
varray: [-8.618, -8.596, -8.601, -8.606]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
dD_stdev | 0.358373524878
varray: [-63.912, -63.143, -63.245, -63.213]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.60525
varray: [-8.618, -8.596, -8.601, -8.606]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.37825
varray: [-63.912, -63.143, -63.245, -63.213]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00942956343988
varray: [-8.618, -8.596, -8.601, -8.606]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.358373524878
varray: [-63.912, -63.143, -63.245, -63.213]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.70615516635
actual_array: [ 1.93 -15.84]
varray: [-8.618, -8.596, -8.601, -8.606]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.60525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.4374409348
actual_array: [ 18.1 -119. ]
varray: [-63.912, -63.143, -63.245, -63.213]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.37825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [61217.0, 61753.0, 62290.0, 62827.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.70305
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0128
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15275
varray: [2.164, 2.17, 2.155, 2.122]
inst: mean_all
operation: mean_all(residuals)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: residuals
SAMP_residuals_STD | 0.0214067746286
varray: [2.164, 2.17, 2.155, 2.122]
inst: std_all
operation: std_all(residuals)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.18825
varray: [-1.102, -1.065, -1.372, -1.214]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137909088412
varray: [-1.102, -1.065, -1.372, -1.214]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.09325
varray: [68.043, 68.249, 68.321, 67.76]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.251468188843
varray: [68.043, 68.249, 68.321, 67.76]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.564
varray: [28.632, 28.65, 28.662, 28.312]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.168451773514
varray: [28.632, 28.65, 28.662, 28.312]
inst: std_all
operation: std_all(slope_shift)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0223960562004
varray: [-8.597, -8.634, -8.631, -8.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.106868688897
varray: [-63.328, -63.529, -63.525, -63.563]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.9305
varray: [64.022, 65.107, 63.487, 63.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.869677526443
varray: [64.022, 65.107, 63.487, 63.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.18825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.09325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.564
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.63455450023
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.597, -8.634, -8.631, -8.591]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-63.328, -63.529, -63.525, -63.563]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.61325
varray: [-8.597, -8.634, -8.631, -8.591]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
dD_mean | -63.48625
varray: [-63.328, -63.529, -63.525, -63.563]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0223960562004
varray: [-8.597, -8.634, -8.631, -8.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
dD_stdev | 0.106868688897
varray: [-63.328, -63.529, -63.525, -63.563]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.61325
varray: [-8.597, -8.634, -8.631, -8.591]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.48625
varray: [-63.328, -63.529, -63.525, -63.563]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0223960562004
varray: [-8.597, -8.634, -8.631, -8.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.106868688897
varray: [-63.328, -63.529, -63.525, -63.563]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.71435624744
actual_array: [ 1.93 -15.84]
varray: [-8.597, -8.634, -8.631, -8.591]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.61325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.5541470768
actual_array: [ 18.1 -119. ]
varray: [-63.328, -63.529, -63.525, -63.563]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.48625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [63414.0, 63949.0, 64485.0, 65020.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.69705
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.75655
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.155
varray: [2.149, 2.162, 2.158, 2.151]
inst: mean_all
operation: mean_all(residuals)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: residuals
SAMP_residuals_STD | 0.00605530070819
varray: [2.149, 2.162, 2.158, 2.151]
inst: std_all
operation: std_all(residuals)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95225
varray: [-0.982, -0.898, -1.009, -0.92]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0519254272972
varray: [-0.982, -0.898, -1.009, -0.92]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.75175
varray: [67.608, 67.749, 67.841, 67.809]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.103141892556
varray: [67.608, 67.749, 67.841, 67.809]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.728
varray: [28.61, 28.978, 28.867, 28.457]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.237477016432
varray: [28.61, 28.978, 28.867, 28.457]
inst: std_all
operation: std_all(slope_shift)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0632745341086
varray: [-6.348, -6.258, -6.202, -6.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.724099670856
varray: [-51.147, -50.006, -49.631, -49.602]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.683
varray: [64.077, 64.918, 64.533, 65.204]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.488672351036
varray: [64.077, 64.918, 64.533, 65.204]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.155
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.75175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.728
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.056
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.056
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.53279059249
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.957
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-6.348, -6.258, -6.202, -6.23]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-51.147, -50.006, -49.631, -49.602]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.2595
varray: [-6.348, -6.258, -6.202, -6.23]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
dD_mean | -50.0965
varray: [-51.147, -50.006, -49.631, -49.602]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0632745341086
varray: [-6.348, -6.258, -6.202, -6.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
dD_stdev | 0.724099670856
varray: [-51.147, -50.006, -49.631, -49.602]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.2595
varray: [-6.348, -6.258, -6.202, -6.23]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.0965
varray: [-51.147, -50.006, -49.631, -49.602]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0632745341086
varray: [-6.348, -6.258, -6.202, -6.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.724099670856
varray: [-51.147, -50.006, -49.631, -49.602]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.30144441945
actual_array: [ 1.93 -15.84]
varray: [-6.348, -6.258, -6.202, -6.23]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.2595
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.0850168476
actual_array: [ 18.1 -119. ]
varray: [-51.147, -50.006, -49.631, -49.602]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.0965
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [65607.0, 66143.0, 66680.0, 67216.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6948
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.09805
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.174
varray: [2.216, 2.146, 2.173, 2.161]
inst: mean_all
operation: mean_all(residuals)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: residuals
SAMP_residuals_STD | 0.0300998338866
varray: [2.216, 2.146, 2.173, 2.161]
inst: std_all
operation: std_all(residuals)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.92825
varray: [-0.984, -0.873, -0.807, -1.049]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.108693376063
varray: [-0.984, -0.873, -0.807, -1.049]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.88975
varray: [68.288, 67.666, 67.762, 67.843]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.275180395377
varray: [68.288, 67.666, 67.762, 67.843]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.154
varray: [30.189, 28.851, 28.864, 28.712]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.693420988049
varray: [30.189, 28.851, 28.864, 28.712]
inst: std_all
operation: std_all(slope_shift)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0211876536376
varray: [-6.518, -6.52, -6.561, -6.548]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.185061025971
varray: [-50.459, -50.143, -50.522, -50.539]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.13675
varray: [64.971, 65.202, 65.33, 65.044]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.160912761044
varray: [64.971, 65.202, 65.33, 65.044]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.174
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.92825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.88975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.154
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.037
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.037
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.67345092718
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.50325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-6.518, -6.52, -6.561, -6.548]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-50.459, -50.143, -50.522, -50.539]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.53675
varray: [-6.518, -6.52, -6.561, -6.548]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
dD_mean | -50.41575
varray: [-50.459, -50.143, -50.522, -50.539]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0211876536376
varray: [-6.518, -6.52, -6.561, -6.548]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
dD_stdev | 0.185061025971
varray: [-50.459, -50.143, -50.522, -50.539]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.53675
varray: [-6.518, -6.52, -6.561, -6.548]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.41575
varray: [-50.459, -50.143, -50.522, -50.539]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0211876536376
varray: [-6.518, -6.52, -6.561, -6.548]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.185061025971
varray: [-50.459, -50.143, -50.522, -50.539]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.58566313609
actual_array: [ 1.93 -15.84]
varray: [-6.518, -6.52, -6.561, -6.548]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.53675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.4300023646
actual_array: [ 18.1 -119. ]
varray: [-50.459, -50.143, -50.522, -50.539]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.41575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [67803.0, 68339.0, 68876.0, 69411.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6758
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.96005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13925
varray: [2.126, 2.133, 2.116, 2.182]
inst: mean_all
operation: mean_all(residuals)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: residuals
SAMP_residuals_STD | 0.0293413814717
varray: [2.126, 2.133, 2.116, 2.182]
inst: std_all
operation: std_all(residuals)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.90475
varray: [-0.916, -0.85, -0.941, -0.912]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0386900073232
varray: [-0.916, -0.85, -0.941, -0.912]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.4055
varray: [67.268, 67.223, 67.199, 67.932]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.352163314387
varray: [67.268, 67.223, 67.199, 67.932]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.416
varray: [28.365, 28.407, 27.788, 29.104]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.538643357581
varray: [28.365, 28.407, 27.788, 29.104]
inst: std_all
operation: std_all(slope_shift)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: slope_shift
SAMP_d18O_STD | 0.046306946923
varray: [-6.558, -6.518, -6.461, -6.465]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.153703990406
varray: [-50.273, -50.246, -50.131, -49.935]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.411
varray: [63.371, 64.128, 64.254, 65.891]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.06093323698
varray: [63.371, 64.128, 64.254, 65.891]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.90475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.4055
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.416
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.24513794663
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.229
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-6.558, -6.518, -6.461, -6.465]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-50.273, -50.246, -50.131, -49.935]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.5005
varray: [-6.558, -6.518, -6.461, -6.465]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
dD_mean | -50.14625
varray: [-50.273, -50.246, -50.131, -49.935]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
d18O_stdev | 0.046306946923
varray: [-6.558, -6.518, -6.461, -6.465]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
dD_stdev | 0.153703990406
varray: [-50.273, -50.246, -50.131, -49.935]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.5005
varray: [-6.558, -6.518, -6.461, -6.465]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.14625
varray: [-50.273, -50.246, -50.131, -49.935]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.046306946923
varray: [-6.558, -6.518, -6.461, -6.465]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.153703990406
varray: [-50.273, -50.246, -50.131, -49.935]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.54850198739
actual_array: [ 1.93 -15.84]
varray: [-6.558, -6.518, -6.461, -6.465]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.5005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.1387773158
actual_array: [ 18.1 -119. ]
varray: [-50.273, -50.246, -50.131, -49.935]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.14625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [70000.0, 70536.0, 71072.0, 71607.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.71055
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4443
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14025
varray: [2.201, 2.142, 2.114, 2.104]
inst: mean_all
operation: mean_all(residuals)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: residuals
SAMP_residuals_STD | 0.043576560978
varray: [2.201, 2.142, 2.114, 2.104]
inst: std_all
operation: std_all(residuals)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.097
varray: [-1.019, -1.094, -1.024, -1.251]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.108225074113
varray: [-1.019, -1.094, -1.024, -1.251]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.9285
varray: [68.694, 67.967, 67.482, 67.571]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.552159699121
varray: [68.694, 67.967, 67.482, 67.571]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.21325
varray: [30.063, 29.087, 28.695, 29.008]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.591248608173
varray: [30.063, 29.087, 28.695, 29.008]
inst: std_all
operation: std_all(slope_shift)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0467430921813
varray: [-8.94, -8.99, -9.051, -9.016]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.749514954265
varray: [-64.819, -65.756, -66.462, -66.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.669
varray: [66.347, 63.767, 62.994, 61.568]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.0042549738
varray: [66.347, 63.767, 62.994, 61.568]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.097
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.9285
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.21325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.19990954319
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.971
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.94, -8.99, -9.051, -9.016]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-64.819, -65.756, -66.462, -66.339]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.99925
varray: [-8.94, -8.99, -9.051, -9.016]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
dD_mean | -65.844
varray: [-64.819, -65.756, -66.462, -66.339]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0467430921813
varray: [-8.94, -8.99, -9.051, -9.016]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
dD_stdev | 0.749514954265
varray: [-64.819, -65.756, -66.462, -66.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.99925
varray: [-8.94, -8.99, -9.051, -9.016]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.844
varray: [-64.819, -65.756, -66.462, -66.339]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0467430921813
varray: [-8.94, -8.99, -9.051, -9.016]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.749514954265
varray: [-64.819, -65.756, -66.462, -66.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1100584102
actual_array: [ 1.93 -15.84]
varray: [-8.94, -8.99, -9.051, -9.016]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.99925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.1019610239
actual_array: [ 18.1 -119. ]
varray: [-64.819, -65.756, -66.462, -66.339]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [72191.0, 72726.0, 73262.0, 73798.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.70955
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9213
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2115
varray: [2.227, 2.194, 2.222, 2.203]
inst: mean_all
operation: mean_all(residuals)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: residuals
SAMP_residuals_STD | 0.0155884572681
varray: [2.227, 2.194, 2.222, 2.203]
inst: std_all
operation: std_all(residuals)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.07775
varray: [-0.953, -1.184, -1.004, -1.17]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.116620109758
varray: [-0.953, -1.184, -1.004, -1.17]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.45025
varray: [67.493, 67.409, 67.494, 67.405]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0499691571539
varray: [67.493, 67.409, 67.494, 67.405]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.5745
varray: [28.57, 28.556, 28.6, 28.572]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.018430047929
varray: [28.57, 28.556, 28.6, 28.572]
inst: std_all
operation: std_all(slope_shift)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: slope_shift
SAMP_d18O_STD | 0.171099142799
varray: [-0.117, 0.223, 0.179, 0.256]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.89499964018
varray: [-6.387, -1.57, -0.729, 0.049]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.09325
varray: [66.043, 64.151, 65.52, 64.659]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.848635915259
varray: [66.043, 64.151, 65.52, 64.659]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2115
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.07775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.45025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.5745
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0005
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0005
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0226142017187
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.54675
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-0.117, 0.223, 0.179, 0.256]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-6.387, -1.57, -0.729, 0.049]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.13525
varray: [-0.117, 0.223, 0.179, 0.256]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
dD_mean | -2.15925
varray: [-6.387, -1.57, -0.729, 0.049]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
d18O_stdev | 0.171099142799
varray: [-0.117, 0.223, 0.179, 0.256]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
dD_stdev | 2.89499964018
varray: [-6.387, -1.57, -0.729, 0.049]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.13525
varray: [-0.117, 0.223, 0.179, 0.256]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.15925
varray: [-6.387, -1.57, -0.729, 0.049]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.171099142799
varray: [-0.117, 0.223, 0.179, 0.256]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.89499964018
varray: [-6.387, -1.57, -0.729, 0.049]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.74596150407
actual_array: [ 1.93 -15.84]
varray: [-0.117, 0.223, 0.179, 0.256]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.13525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.28342883604
actual_array: [ 18.1 -119. ]
varray: [-6.387, -1.57, -0.729, 0.049]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -2.15925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [74383.0, 74918.0, 75454.0, 75989.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6383
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.39955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.20675
varray: [2.248, 2.176, 2.201, 2.202]
inst: mean_all
operation: mean_all(residuals)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: residuals
SAMP_residuals_STD | 0.0300152738896
varray: [2.248, 2.176, 2.201, 2.202]
inst: std_all
operation: std_all(residuals)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.09625
varray: [-1.19, -1.071, -1.133, -0.991]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0853517232788
varray: [-1.19, -1.071, -1.133, -0.991]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.32275
varray: [67.929, 67.056, 67.219, 67.087]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.410299382565
varray: [67.929, 67.056, 67.219, 67.087]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.47
varray: [29.28, 27.909, 28.218, 28.473]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.587177996863
varray: [29.28, 27.909, 28.218, 28.473]
inst: std_all
operation: std_all(slope_shift)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0553074738771
varray: [0.36, 0.455, 0.448, 0.49]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.401027430483
varray: [0.92, 1.619, 1.754, 1.761]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.60875
varray: [65.184, 63.698, 64.873, 64.68]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.64168599538
varray: [65.184, 63.698, 64.873, 64.68]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.20675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.09625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.32275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.47
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.192220714609
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.03125
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [0.36, 0.455, 0.448, 0.49]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [0.92, 1.619, 1.754, 1.761]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.43825
varray: [0.36, 0.455, 0.448, 0.49]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
dD_mean | 1.5135
varray: [0.92, 1.619, 1.754, 1.761]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0553074738771
varray: [0.36, 0.455, 0.448, 0.49]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
dD_stdev | 0.401027430483
varray: [0.92, 1.619, 1.754, 1.761]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.43825
varray: [0.36, 0.455, 0.448, 0.49]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
measured_dD_mean | 1.5135
varray: [0.92, 1.619, 1.754, 1.761]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0553074738771
varray: [0.36, 0.455, 0.448, 0.49]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.401027430483
varray: [0.92, 1.619, 1.754, 1.761]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.435345557652
actual_array: [ 1.93 -15.84]
varray: [0.36, 0.455, 0.448, 0.49]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.43825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 0.685390451043
actual_array: [ 18.1 -119. ]
varray: [0.92, 1.619, 1.754, 1.761]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.5135
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [76575.0, 77110.0, 77646.0, 78182.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.64305
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.52705
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13925
varray: [2.152, 2.134, 2.162, 2.109]
inst: mean_all
operation: mean_all(residuals)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: residuals
SAMP_residuals_STD | 0.0232576152403
varray: [2.152, 2.134, 2.162, 2.109]
inst: std_all
operation: std_all(residuals)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.944
varray: [-1.034, -1.016, -0.736, -0.99]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.139838001511
varray: [-1.034, -1.016, -0.736, -0.99]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.63525
varray: [67.787, 67.531, 67.833, 67.39]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.210672534201
varray: [67.787, 67.531, 67.833, 67.39]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.62025
varray: [28.996, 28.535, 28.717, 28.233]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.320301290454
varray: [28.996, 28.535, 28.717, 28.233]
inst: std_all
operation: std_all(slope_shift)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: slope_shift
SAMP_d18O_STD | 0.174066079407
varray: [-8.651, -8.966, -8.995, -9.026]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.02169367519
varray: [-59.091, -63.889, -65.164, -65.744]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.388
varray: [64.489, 63.549, 65.698, 63.816]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.958715460047
varray: [64.489, 63.549, 65.698, 63.816]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.944
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.63525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.62025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.24513794663
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.252
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.651, -8.966, -8.995, -9.026]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-59.091, -63.889, -65.164, -65.744]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.9095
varray: [-8.651, -8.966, -8.995, -9.026]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
dD_mean | -63.472
varray: [-59.091, -63.889, -65.164, -65.744]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
d18O_stdev | 0.174066079407
varray: [-8.651, -8.966, -8.995, -9.026]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
dD_stdev | 3.02169367519
varray: [-59.091, -63.889, -65.164, -65.744]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.9095
varray: [-8.651, -8.966, -8.995, -9.026]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.472
varray: [-59.091, -63.889, -65.164, -65.744]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.174066079407
varray: [-8.651, -8.966, -8.995, -9.026]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.02169367519
varray: [-59.091, -63.889, -65.164, -65.744]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0180525317
actual_array: [ 1.93 -15.84]
varray: [-8.651, -8.966, -8.995, -9.026]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.9095
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.5387483497
actual_array: [ 18.1 -119. ]
varray: [-59.091, -63.889, -65.164, -65.744]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.472
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [78766.0, 79302.0, 79837.0, 80374.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.71055
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.21455
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2525
varray: [2.297, 2.246, 2.238, 2.229]
inst: mean_all
operation: mean_all(residuals)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: residuals
SAMP_residuals_STD | 0.0304685630336
varray: [2.297, 2.246, 2.238, 2.229]
inst: std_all
operation: std_all(residuals)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.134
varray: [-1.362, -0.97, -1.056, -1.148]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.168483431431
varray: [-1.362, -0.97, -1.056, -1.148]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.9485
varray: [68.723, 67.563, 67.783, 67.725]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.524659572167
varray: [68.723, 67.563, 67.783, 67.725]
inst: std_all
operation: std_all(baseline_shift)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.13925
varray: [30.144, 29.06, 28.581, 28.772]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.698168735574
varray: [30.144, 29.06, 28.581, 28.772]
inst: std_all
operation: std_all(slope_shift)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: slope_shift
SAMP_d18O_STD | 0.1831602213
varray: [0.327, 0.635, 0.679, 0.737]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.89890201916
varray: [-4.323, 0.38, 1.391, 2.119]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.10225
varray: [66.139, 66.784, 66.005, 65.481]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.535892635391
varray: [66.139, 66.784, 66.005, 65.481]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.134
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.9485
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.13925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0415
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0415
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.87697874265
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.46225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [0.327, 0.635, 0.679, 0.737]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-4.323, 0.38, 1.391, 2.119]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.5945
varray: [0.327, 0.635, 0.679, 0.737]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
dD_mean | -0.10825
varray: [-4.323, 0.38, 1.391, 2.119]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
d18O_stdev | 0.1831602213
varray: [0.327, 0.635, 0.679, 0.737]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
dD_stdev | 2.89890201916
varray: [-4.323, 0.38, 1.391, 2.119]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.5945
varray: [0.327, 0.635, 0.679, 0.737]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
measured_dD_mean | -0.10825
varray: [-4.323, 0.38, 1.391, 2.119]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.1831602213
varray: [0.327, 0.635, 0.679, 0.737]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.89890201916
varray: [-4.323, 0.38, 1.391, 2.119]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.275168192543
actual_array: [ 1.93 -15.84]
varray: [0.327, 0.635, 0.679, 0.737]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.5945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -1.0670927506
actual_array: [ 18.1 -119. ]
varray: [-4.323, 0.38, 1.391, 2.119]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -0.10825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [80960.0, 81496.0, 82032.0, 82567.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.5973
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9013
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.22325
varray: [2.244, 2.215, 2.211, 2.223]
inst: mean_all
operation: mean_all(residuals)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: residuals
SAMP_residuals_STD | 0.0147054411699
varray: [2.244, 2.215, 2.211, 2.223]
inst: std_all
operation: std_all(residuals)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.021
varray: [-1.142, -1.077, -0.875, -0.99]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115547969836
varray: [-1.142, -1.077, -0.875, -0.99]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.4075
varray: [67.739, 67.275, 67.297, 67.319]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.221728813343
varray: [67.739, 67.275, 67.297, 67.319]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.453
varray: [28.874, 28.394, 28.269, 28.275]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.2865088713
varray: [28.874, 28.394, 28.269, 28.275]
inst: std_all
operation: std_all(slope_shift)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0241177666185
varray: [0.448, 0.417, 0.449, 0.4]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0806158586549
varray: [1.894, 1.781, 1.722, 1.874]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.59225
varray: [65.563, 65.222, 66.234, 65.35]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.450358653372
varray: [65.563, 65.222, 66.234, 65.35]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.22325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.021
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.4075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.453
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.554047942108
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.04775
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [0.448, 0.417, 0.449, 0.4]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [1.894, 1.781, 1.722, 1.874]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.4285
varray: [0.448, 0.417, 0.449, 0.4]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
dD_mean | 1.81775
varray: [1.894, 1.781, 1.722, 1.874]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0241177666185
varray: [0.448, 0.417, 0.449, 0.4]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
dD_stdev | 0.0806158586549
varray: [1.894, 1.781, 1.722, 1.874]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.4285
varray: [0.448, 0.417, 0.449, 0.4]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
measured_dD_mean | 1.81775
varray: [1.894, 1.781, 1.722, 1.874]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0241177666185
varray: [0.448, 0.417, 0.449, 0.4]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0806158586549
varray: [1.894, 1.781, 1.722, 1.874]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.445340625234
actual_array: [ 1.93 -15.84]
varray: [0.448, 0.417, 0.449, 0.4]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.4285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 1.01416678161
actual_array: [ 18.1 -119. ]
varray: [1.894, 1.781, 1.722, 1.874]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.81775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [83154.0, 83690.0, 84225.0, 84761.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.62655
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4423
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.155
varray: [2.168, 2.16, 2.158, 2.134]
inst: mean_all
operation: mean_all(residuals)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: residuals
SAMP_residuals_STD | 0.0146515073172
varray: [2.168, 2.16, 2.158, 2.134]
inst: std_all
operation: std_all(residuals)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9395
varray: [-1.097, -0.763, -0.931, -0.967]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137582702401
varray: [-1.097, -0.763, -0.931, -0.967]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.849
varray: [68.112, 67.734, 67.855, 67.695]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.18810103668
varray: [68.112, 67.734, 67.855, 67.695]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.8725
varray: [29.186, 29.111, 28.704, 28.489]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.331978412953
varray: [29.186, 29.111, 28.704, 28.489]
inst: std_all
operation: std_all(slope_shift)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: slope_shift
SAMP_d18O_STD | 0.212805858315
varray: [-8.515, -8.901, -8.917, -8.985]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.0236974055
varray: [-58.58, -63.425, -64.575, -65.288]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.89875
varray: [63.922, 65.589, 65.33, 64.754]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.738786110229
varray: [63.922, 65.589, 65.33, 64.754]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.155
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9395
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.849
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.8725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.056
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.056
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.53279059249
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.74125
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.515, -8.901, -8.917, -8.985]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-58.58, -63.425, -64.575, -65.288]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8295
varray: [-8.515, -8.901, -8.917, -8.985]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
dD_mean | -62.967
varray: [-58.58, -63.425, -64.575, -65.288]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
d18O_stdev | 0.212805858315
varray: [-8.515, -8.901, -8.917, -8.985]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
dD_stdev | 3.0236974055
varray: [-58.58, -63.425, -64.575, -65.288]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8295
varray: [-8.515, -8.901, -8.917, -8.985]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.967
varray: [-58.58, -63.425, -64.575, -65.288]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.212805858315
varray: [-8.515, -8.901, -8.917, -8.985]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.0236974055
varray: [-58.58, -63.425, -64.575, -65.288]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.93604172075
actual_array: [ 1.93 -15.84]
varray: [-8.515, -8.901, -8.917, -8.985]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.8295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.9930390747
actual_array: [ 18.1 -119. ]
varray: [-58.58, -63.425, -64.575, -65.288]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -62.967
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [85347.0, 85883.0, 86418.0, 86954.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.6948
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0008
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.879
varray: [1.817, 1.823, 1.881, 1.995]
inst: mean_all
operation: mean_all(residuals)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: residuals
SAMP_residuals_STD | 0.0825429181295
varray: [1.817, 1.823, 1.881, 1.995]
inst: std_all
operation: std_all(residuals)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.37775
varray: [-1.498, -1.44, -1.24, -1.333]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114473796128
varray: [-1.498, -1.44, -1.24, -1.333]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 65.63975
varray: [64.944, 64.98, 65.669, 66.966]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.94501018513
varray: [64.944, 64.98, 65.669, 66.966]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 21.4835
varray: [20.781, 20.928, 21.285, 22.94]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.993794244298
varray: [20.781, 20.928, 21.285, 22.94]
inst: std_all
operation: std_all(slope_shift)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: slope_shift
SAMP_d18O_STD | 0.121365838137
varray: [-2.853, -2.642, -2.603, -2.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.317019846908
varray: [-54.364, -53.884, -53.685, -53.7]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 56.08025
varray: [53.687, 53.963, 56.881, 59.79]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.8643695496
varray: [53.687, 53.963, 56.881, 59.79]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.879
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.37775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 65.63975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 21.4835
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.332
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.332
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -15.0158299412
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.55975
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-2.853, -2.642, -2.603, -2.596]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-54.364, -53.884, -53.685, -53.7]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.6735
varray: [-2.853, -2.642, -2.603, -2.596]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
dD_mean | -53.90825
varray: [-54.364, -53.884, -53.685, -53.7]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
d18O_stdev | 0.121365838137
varray: [-2.853, -2.642, -2.603, -2.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
dD_stdev | 0.317019846908
varray: [-54.364, -53.884, -53.685, -53.7]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.6735
varray: [-2.853, -2.642, -2.603, -2.596]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.90825
varray: [-54.364, -53.884, -53.685, -53.7]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.121365838137
varray: [-2.853, -2.642, -2.603, -2.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.317019846908
varray: [-54.364, -53.884, -53.685, -53.7]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.62530981926
actual_array: [ 1.93 -15.84]
varray: [-2.853, -2.642, -2.603, -2.596]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -2.6735
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.2040412619
actual_array: [ 18.1 -119. ]
varray: [-54.364, -53.884, -53.685, -53.7]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -53.90825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [87543.0, 88077.0, 88613.0, 89149.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.9708
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.21005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.099
varray: [2.111, 2.1, 2.088, 2.097]
inst: mean_all
operation: mean_all(residuals)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: residuals
SAMP_residuals_STD | 0.00948683298051
varray: [2.111, 2.1, 2.088, 2.097]
inst: std_all
operation: std_all(residuals)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03
varray: [-0.983, -0.93, -1.124, -1.083]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0891702491492
varray: [-0.983, -0.93, -1.124, -1.083]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.29175
varray: [67.388, 67.137, 67.243, 67.399]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.125287336418
varray: [67.388, 67.137, 67.243, 67.399]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.26725
varray: [28.535, 28.075, 28.452, 28.007]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.26489793632
varray: [28.535, 28.075, 28.452, 28.007]
inst: std_all
operation: std_all(slope_shift)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: slope_shift
SAMP_d18O_STD | 0.123718430317
varray: [-8.793, -9.015, -9.029, -9.066]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.594845288009
varray: [-65.125, -66.288, -66.248, -66.39]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.60125
varray: [62.83, 63.221, 61.751, 62.603]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.621649619963
varray: [62.83, 63.221, 61.751, 62.603]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.099
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.29175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.26725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.06558118498
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.03875
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.5887
varray: [-8.793, -9.015, -9.029, -9.066]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.2435
varray: [-65.125, -66.288, -66.248, -66.39]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.97575
varray: [-8.793, -9.015, -9.029, -9.066]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
dD_mean | -66.01275
varray: [-65.125, -66.288, -66.248, -66.39]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
d18O_stdev | 0.123718430317
varray: [-8.793, -9.015, -9.029, -9.066]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
dD_stdev | 0.594845288009
varray: [-65.125, -66.288, -66.248, -66.39]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.97575
varray: [-8.793, -9.015, -9.029, -9.066]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.01275
varray: [-65.125, -66.288, -66.248, -66.39]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.123718430317
varray: [-8.793, -9.015, -9.029, -9.066]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.594845288009
varray: [-65.125, -66.288, -66.248, -66.39]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0859677345
actual_array: [ 1.93 -15.84]
varray: [-8.793, -9.015, -9.029, -9.066]
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.97575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.2843143707
actual_array: [ 18.1 -119. ]
varray: [-65.125, -66.288, -66.248, -66.39]
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.01275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: [89732.0, 90267.0, 90803.0, 91339.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 66.7508
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.55805
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.28
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.503
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3978650998
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.28
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0559218901
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -69.503
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.187
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.842
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3025275321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.187
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3416370766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.842
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.384
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2840750997
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8467165855
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.384
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.157
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.187
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.271773478
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.157
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6338359377
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.187
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.104
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.863
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2174413158
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.104
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2837175117
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.863
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.136
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.903
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2502456401
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.136
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3269420087
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.903
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.101
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.973
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2143659104
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.101
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4025848785
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.973
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.141
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.883
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2553713158
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.141
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3053297602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.883
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.646
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2215418563
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0492246153
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.646
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.753
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2153910455
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1648501448
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.753
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.251
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.338
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.42296816139
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.251
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 9.14064237719
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 9.338
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.707
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.82482113498
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 16.0230629175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 15.707
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.704
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.88735437831
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.704
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.9530367101
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 17.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.775
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.145
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.96013897302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.6575960117
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 18.145
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.835
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.667
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.02164708122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.835
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.221675698
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 18.667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.973
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.00729518931
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5523431004
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 18.973
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.86
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.152
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04727545964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.86
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.7457727246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 19.152
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.863
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.669
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05035086505
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.863
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3044493488
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 19.669
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.632
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04112464882
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.264466689
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 19.632
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.85
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.514
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7456
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.629
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.85
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.514
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03702410827
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 2.85
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1369544227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 19.514
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.817
operation: col(d(18_16)Mean)
inst: col
dD_actual | -97.112
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.048903215
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -13.817
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -105.890550356
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -97.112
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.452
operation: col(d(18_16)Mean)
inst: col
dD_actual | -106.547
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6998640268
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.452
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.086128594
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -106.547
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.558
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.938
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8085283513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.558
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.669872904
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -108.938
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.646
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.062
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8987402433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.646
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.884481271
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -110.062
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.684
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.899
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9376953785
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.684
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.788953871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -110.899
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.723
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.157
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9776756489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.723
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.067751877
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -111.157
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.755
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.676
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0104799732
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.755
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.628589726
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -111.676
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.729
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.8
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9838264597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.729
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.762585667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -111.8
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.783
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.1
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0391837571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.783
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.086769394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -112.1
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.74
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.144
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9951029462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -14.74
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.134316341
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -112.144
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.303
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.016
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.421443208
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.303
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.7715009163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -72.016
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.222
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.931
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3384072619
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.222
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4378115825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.931
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.21
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3261056403
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.21
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9990829376
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.717
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2297429374
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1259480975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.717
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.152
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.778
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2666478023
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1918654555
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.778
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.107
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.675
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2205167212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0805623756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.77
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2615221267
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1832205561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.77
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.564
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2246172617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9606143963
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.564
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.12
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.233843478
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.12
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8374245798
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.163
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.247
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2779242888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.163
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6180602573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.222
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.153
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.31327212521
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.222
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2746454117
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -64.153
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.161
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.977
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.25073888187
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.161
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0844576248
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.977
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.183
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.136
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.27329185488
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.183
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2562750005
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -64.136
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.12
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.757
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.20870834127
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.12
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8467228911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.757
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.797
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.386
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.70164373525
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -0.797
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.85090740704
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.386
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.49
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.38692724829
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -0.49
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.50619834874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -1.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.453
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.375
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.34899724823
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -0.453
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.35534611519
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -0.375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.437
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.045
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.33259508604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -0.437
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.99874401466
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -0.045
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.233
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.14860265485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -1.233
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.44243591007
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.378
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.616
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.29724724967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -1.378
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.09944826499
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.616
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.307
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.22446265497
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -1.307
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.924389052
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.344
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.18
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.26239265503
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -1.344
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.62830124732
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.18
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.058
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.361
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.04460941601
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.5973654653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -28.361
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.184
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.226
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17377644324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.184
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.6127076396
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -30.226
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.156
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.293
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.14507265941
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.156
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.6851086721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -30.293
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.23
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.812
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.22093265952
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.23
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.2459465211
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -30.812
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.483
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.608
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.48029184911
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.186726438
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -32.608
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.416
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.326
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.41160779495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.416
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.8819937339
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -32.326
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.474
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.684
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.47106563288
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.474
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.2688529823
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -32.684
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.447
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.43621103823
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.0127478374
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -32.447
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.654
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.65558995748
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.654
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.6978561154
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -33.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.729
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.73247509274
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.729
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.6978561154
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -33.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.705
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.187
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.70787184945
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.705
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.8124010325
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -33.187
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.692
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.117
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.69454509268
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -4.692
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.7367581627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -33.117
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.713
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.1232061866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -12.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.298161146
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -94.713
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.251
operation: col(d(18_16)Mean)
inst: col
dD_actual | -99.801
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.4686767276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -13.251
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -108.796317169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -99.801
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.302
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.26
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.5209586196
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -13.302
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.372930698
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -101.26
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.306
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.546
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.5250591602
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -13.306
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.681985852
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -101.546
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.354
operation: col(d(18_16)Mean)
inst: col
dD_actual | -155.421
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.7250764669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -19.354
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -168.899980295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -155.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -159.526
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.9075505212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -19.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -173.335894303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -159.526
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.592
operation: col(d(18_16)Mean)
inst: col
dD_actual | -161.0
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.9690586294
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -19.592
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -174.928717019
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -161.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.649
operation: col(d(18_16)Mean)
inst: col
dD_actual | -161.779
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -21.0274913322
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -19.649
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -175.770514099
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -161.779
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.608
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.179
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7341094247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.608
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.350702871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -77.179
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.306
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4245186134
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.306
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.6552282804
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -70.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.951
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2840750997
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.459423831
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.951
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.134
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.461
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2481953699
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9299237423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.461
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -16.663
operation: col(d(18_16)Mean)
inst: col
dD_actual | -129.461
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.966437814
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -16.663
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -140.84728172
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -129.461
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -16.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | -134.313
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -18.2637270037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -16.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -146.09041321
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -134.313
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -16.982
operation: col(d(18_16)Mean)
inst: col
dD_actual | -135.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -18.2934559227
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -16.982
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -147.40011547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -135.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -17.012
operation: col(d(18_16)Mean)
inst: col
dD_actual | -136.054
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -18.3242099768
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -17.012
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -147.971759444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -136.054
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -18.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | -166.247
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.2360870067
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -18.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -180.598690418
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -166.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -18.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -168.692
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.2494137635
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -18.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -183.240787799
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -168.692
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -18.891
operation: col(d(18_16)Mean)
inst: col
dD_actual | -169.177
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.2504388986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -18.891
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -183.764884825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -169.177
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -18.904
operation: col(d(18_16)Mean)
inst: col
dD_actual | -169.787
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -20.2637656554
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -18.904
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -184.424058405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -169.787
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.338
operation: col(d(18_16)Mean)
inst: col
dD_actual | -89.277
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.5075932111
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -11.338
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.4239519991
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -89.277
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.136
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.297
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3005159135
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -11.136
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9618896924
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -83.297
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.134
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.888
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2984656433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -11.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.4393067844
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -81.888
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.018
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.69
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.1795499674
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -11.018
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.1447330982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -80.69
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.359
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2389691536
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5778639973
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.359
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.011
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.647
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1221037481
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.011
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7278555242
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.647
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.03
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.11
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1415813157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.03
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1475666516
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.11
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.717
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.091349694
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.7228859682
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -62.717
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.53
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.95
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.654148884
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.53
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4583432186
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -68.95
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.742
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7054056408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.3141882599
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -69.742
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.84
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6900286138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.4200882776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -69.84
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.553
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.245
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6777269921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.553
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7771238842
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -69.245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.195
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.475
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3107286132
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.7838274646
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.156
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.057
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2707483429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.156
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3321314706
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.057
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.788
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2246172617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.041446728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.788
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.112
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2256423969
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.112
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9139344618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.077
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.16462753039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.077
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.2147845278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -60.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.158
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.426
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.24766347646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.158
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.2472029005
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -60.426
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.105
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.957
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.19333131422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.105
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.7403956728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -59.957
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.21075861154
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.9392283592
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -60.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.508
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.634
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.6064607743
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.508
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.6331951368
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -62.634
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.588
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.492
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.68847158524
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.5603605982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.492
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.519
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.893
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.61773726081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.519
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.9130737551
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -62.893
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.547
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.948
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.64644104463
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.547
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.9725074386
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -62.948
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.187
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.18
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3025275321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.187
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5456592248
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.18
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.104
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2174413158
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.104
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3597938876
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.137
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.987
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2512707753
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.137
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3371010266
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2389691536
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.629946994
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -67.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.912
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.71922563934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0142178171
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.912
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69667266633
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1832268616
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.601
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.245
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.70179834202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.601
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.2934493291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.606
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.213
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7069240177
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.606
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.2588697314
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.213
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69769780147
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3831401604
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.634
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.529
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73562780153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.634
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.600343258
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.529
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.631
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73255239612
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.631
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.5960208083
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.563
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69154699065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6370840805
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.563
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.348
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.39216887904
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.348
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.220200201
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -51.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.006
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.29990671674
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9872214231
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.006
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.202
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.631
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.24249914909
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.202
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.5819917634
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -49.631
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.23
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.602
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.27120293291
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.23
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.550654003
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -49.602
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.518
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.459
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56644185228
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.518
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.476738852
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.459
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.52
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56849212255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.52
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.1352653254
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.522
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.61052266316
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.5448174348
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.522
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.548
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.539
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.59719590638
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.548
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.5631878461
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.539
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.558
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.60744725775
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.558
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.2757449408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.518
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.246
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56644185228
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.518
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.2465684053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.246
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.131
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.50800914949
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.1222979763
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -50.131
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.465
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.935
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.51210969004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -6.465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9104979408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -49.935
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.94
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.819
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0493191534
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.94
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9943332874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -64.819
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.99
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.756
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1005759102
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.99
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0068671304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.756
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.051
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.462
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1631091535
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.051
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.769779503
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.462
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.016
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1272294237
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.016
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6368641747
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.387
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.0045518423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -0.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.85198801947
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -6.387
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.223
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.57
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.656005895825
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.6466779641
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -1.57
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.179
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.729
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.70111184184
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.179
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.73788291395
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -0.729
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.256
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.049
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.622176436314
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.256
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.89716644663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 0.049
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.36
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.92
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.515562382098
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.36
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.0440469762951
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 0.92
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.455
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.619
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.418174544112
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.799395061972
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.619
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.425350490069
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.945277739463
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.49
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.761
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.382294814328
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.49
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.952842026444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.761
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.651
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.091
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75305509885
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.651
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.804585312
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -59.091
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.966
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.889
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0759726669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.9893637313
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.889
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.995
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.164
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1057015859
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.995
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.3671445743
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.164
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.026
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.744
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1374807751
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.026
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9938997813
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.744
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.327
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.323
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.549391841609
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.327
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.62160397249
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -4.323
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.635
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.38
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.233650219507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.53948373367
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 0.38
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.679
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.391
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.188544273492
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.679
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.553015428875
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.391
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.737
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.119
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.129086435564
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.737
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 1.3397012749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 2.119
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.425350490069
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 1.09656347908
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.417
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.781
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.457129679306
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.417
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.974454274961
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.781
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.449
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.722
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.424325354932
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.449
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 0.910698141835
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.722
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.4
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.47455697663
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: 0.4
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 1.07495123057
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: 1.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.515
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.61363672026
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.515
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.2523923624
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -58.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.901
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.425
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.009338883
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.901
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4879595657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -63.425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.917
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.575
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0257410452
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.917
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.7306638554
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -64.575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.985
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.288
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0954502345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5011405151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.288
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.8093215763
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -2.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.696530375
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -54.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.642
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.884
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.59301806245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -2.642
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.1778364106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -53.884
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.603
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.55303779212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -2.603
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.9627945378
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -53.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.54586184617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -2.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.9790037242
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -53.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:41:42
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.793
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.125
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.89862428826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -8.793
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.3250006897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -65.125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.288
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1262042886
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5817529409
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.288
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.029
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.248
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1405561805
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.029
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5385284439
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.248
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7456
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.629
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.74
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.144
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.5887
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.2435
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.066
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.39
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1784861806
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.862921384356
measure_array: [ 2.7456 -14.5887]
inst: fit_type1
actual_mean: -9.066
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.975481148002
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6919754084
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.879238876732
measure_array: [ 17.629 -109.2435]
inst: fit_type1
actual_mean: -66.39
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.925401167031
tarray: []
source_col: d(D_H)Mean